## Readme
Created on Thu Oct 04 2012
author: Janwillem van Dijk
email: jwe.van.dijk@xs4all.nl
| 1 | Public Sub QuickSort(vArray As Variant, inLow As Integer, inHigh As Integer)
Calculate uncertainty in dose measurement using statistical sampling, | 2
the Monte Carlo method (JCGM 101:2008) | 3 | Dim vPivot As Variant
| 4 | Dim vSwap As Variant
The dose measurement is implemented as a two stage measurement, see e.g. | 5 | Dim lLow As Integer
Introduction to the GUM JCGM 101:2009, section 6.8: | 6 | Dim lHigh As Integer
a) The detector in the dosemeter is the primary measurement device | 7
that will generate a signal x_det wen read | 8 | lLow = inLow
b) The reader that will generate counts dependant on x_det, f_Rd and | 9 | lHigh = inHigh
the blank signal z. | 10
Steps: | 11 | vPivot = vArray((inLow + inHigh) / 2)
1) Define input quatities and uncertainties therein. | 12
2) Calculate M random samples of the signal a detector would generate | 13 | Do While lLow <= lHigh
given a set op input quantities and the uncertainties therein. | 14
3) Calculate the dose using the signal samples and expectation values | 15 | Do While vArray(lLow) < vPivot And lLow < inHigh
for the other input quantities. | 16 | lLow = lLow + 1
4) Perform statistical analysis. | 17 | Loop
18
The sheet needs the macro subroutine SortArray to sort the measurand | 19 | Do While (vPivot < vArray(lHigh)) And (lHigh > inLow)
Apart from sorting this macro also calculates the statistics for improved performance | 20 | lHigh = lHigh - 1
| 21 | Loop
22
| 23 | If (lLow <= lHigh) Then
| 24 | vSwap = vArray(lLow)
| 25 | vArray(lLow) = vArray(lHigh)
| 26 | vArray(lHigh) = vSwap
| 27 | lLow = lLow + 1
| 28 | lHigh = lHigh - 1
| 29 | End If
30
| 31 | Loop
32
| 33 | If (inLow < lHigh) Then QuickSort vArray, inLow, lHigh
| 34 | If (lLow < inHigh) Then QuickSort vArray, lLow, inHigh
35
| 36 | End Sub
37
| 38 | Function SortArray(a As Variant) As Variant
| 39 | Dim n0 As Integer, n As Integer
| 40 | Dim i As Integer
| 41 | Dim ai As Double
| 42 | Dim k As Integer, m As Integer
| 43 | Dim sum As Double, ssum As Double
| 44 | Dim sumB As Double, ssumB As Double
| 45 | Dim avg As Double, stdev As Double
| 46 | Dim avgB As Double, stdevB As Double
47
| 48 | If TypeName(a) = "Range" Then a = a.Value2
| 49 | n0 = LBound(a)
| 50 | n = UBound(a)
| 51 | Dim b() As Variant
| 52 | ReDim b(n0 To n + 8)
53
| 54 | '8 elements space for statistics at the bottom
| 55 | For i = n + 1 To n + 8
| 56 | b(i) = 1000000000#
| 57 | Next i
58
| 59 | k = 0
| 60 | For i = n0 To n
| 61 | ai = a(i, 1)
| 62 | b(i) = ai
| 63 | If ai >= 0 Then
| 64 | sumB = sumB + ai
| 65 | ssumB = ssumB + ai * ai
| 66 | k = k + 1
| 67 | Else
| 68 | sum = sum + ai
| 69 | ssum = ssum + ai * ai
| 70 | End If
| 71 | Next i
| 72 | sum = sum + sumB
| 73 | ssum = ssum + ssumB
74
| 75 | m = n - n0 + 1
| 76 | avg = sum / m
| 77 | avgB = sumB / k
| 78 | stdev = Sqr((ssum - sum * sum / m) / (m - 1))
| 79 | stdevB = Sqr((ssumB - sumB * sumB / k) / (k - 1))
80
| 81 | Call QuickSort(b, n0, n)
82
| 83 | 'store statistics in 8 extra array elements
| 84 | i = n0
| 85 | Do While (b(i) < 0) And (i < n)
| 86 | i = i + 1
| 87 | Loop
| 88 | m = n - i + 1
| 89 | b(n + 1) = avg
| 90 | b(n + 2) = stdev
| 91 | b(n + 3) = avgB
| 92 | b(n + 4) = stdevB
| 93 | 'JCGM 101 7.7
| 94 | b(n + 5) = b(Int(n * 0.025 + 0.5))
| 95 | b(n + 6) = b(Int(n * 0.975 + 0.5))
| 96 | b(n + 7) = b(i + Int(m * 0.025 + 0.5))
| 97 | b(n + 8) = b(i + Int(m * 0.975 + 0.5))
98
| 99 | SortArray = b
100
| 101 | End Function

## MonteCarlo
Monte Carlo uncertainty evaluation of TNO C-detector, LiF:Mg,Ti + 1mm Al
Bayesian
Htrue | 1000 | μSv |  |  | Hmcm | 1001.8633 | μSv | 1001.8633 | μSv
Hnat | 100 | μSv |  |  | stdev | 164.941 | 16.46% | 164.941 | 16.46% | [5 empty] | u=165 uSv
[6 empty] | Low 95% | 689.5205 | 690.6994 | μSv
N | 10000 |  |  |  |  | High 95% | 1335.7719 | 1336.2755 | μSv
fE,a | min |  | 0.39275
| max |  | 1.257
uBg,rel | 0.1 | [9 empty] | Count | 10000
[11 empty] | First >0 | 0
|  | Detector |  |  |  | Reader |  |  | Dose (μSv)
| Hnat | fBg | fE,a | fDet | xblank | freader | x | ux,rel |  | Hocc |  |  | 95% coverage
mean | 100 | 1 | 1 | 1 | 200 | 50 | μSv | 1 |  | 1001.8633 | 1001.8633 | Low | 689.5205 | 690.6994
stdev | 10 | 0.05 | 0.14404 | 0.05 | 10 | 2.5 | μSv | 0.02 |  | 164.941 | 164.941 | High | 1335.7719 | 1336.2755
Bayesian
| Natural | Bg resp, | Response | Detector | Blank | Sensitivity | Signal | Signal | Hgross | Hocc | Sorted |  | Frequency | Frequency
0 | 86.1852 | 1.0313 | 1.2051 | 1.0601 | 202.3986 | 44.6159 | 61166.031 | 58820.3144 | 1172.4063 | 1072.4063 | 430.1615 | 342.0992 | 0 | 0 | 0.56003
1 | 119.2969 | 1.0487 | 0.92057 | 0.95633 | 213.3268 | 52.1315 | 52630.1637 | 50717.6013 | 1010.352 | 910.352 | 435.56 | 408.0756 | 0 | 0 | 2.549
2 | 102.0861 | 0.91574 | 0.8744 | 1.0037 | 189.4888 | 49.6292 | 48384.0244 | 48418.2412 | 964.3648 | 864.3648 | 467.4089 | 474.052 | 3 | 3 | 9.8868
3 | 112.2597 | 0.98931 | 1.0337 | 1.0235 | 190.941 | 50.3633 | 59068.7983 | 57367.2893 | 1143.3458 | 1043.3458 | 497.8412 | 540.0284 | 10 | 10 | 32.6772
4 | 102.5806 | 0.96363 | 0.97081 | 1.0323 | 188.5124 | 53.2722 | 58841.7157 | 58984.2662 | 1175.6853 | 1075.6853 | 500.7585 | 606.0048 | 34 | 34 | 92.0341
5 | 72.619 | 1.1267 | 0.87851 | 1.0535 | 187.3196 | 52.8165 | 53390.4573 | 53534.557 | 1066.6911 | 966.6911 | 504.3176 | 671.9812 | 130 | 130 | 220.8847
6 | 88.705 | 0.99306 | 0.98163 | 1.0437 | 188.277 | 49.4934 | 55253.5469 | 54855.8633 | 1093.1173 | 993.1173 | 505.8591 | 737.9576 | 328 | 328 | 451.7469
7 | 121.5008 | 1.0837 | 1.1537 | 1.0402 | 197.1119 | 49.7355 | 66432.4692 | 68485.2102 | 1365.7042 | 1265.7042 | 509.171 | 803.9341 | 621 | 621 | 787.2952
8 | 102.1886 | 1.0017 | 0.843 | 1.0625 | 198.029 | 48.6915 | 48794.2785 | 47753.8358 | 951.0767 | 851.0767 | 514.0024 | 869.9105 | 1030 | 1030 | 1169.2113
9 | 120.0313 | 0.93275 | 1.2267 | 0.99958 | 191.1996 | 51.3423 | 68893.6262 | 69151.7237 | 1379.0345 | 1279.0345 | 515.2222 | 935.8869 | 1380 | 1380 | 1479.6576
10 | 105.3155 | 1.0079 | 1.1771 | 1.031 | 188.3211 | 48.4486 | 64124.7661 | 65702.7321 | 1310.0546 | 1210.0546 | 517.5887 | 1001.8633 | 1546 | 1546 | 1595.6673
11 | 86.9778 | 1.011 | 1.153 | 0.97009 | 196.2521 | 48.5381 | 58753.94 | 57692.3264 | 1149.8465 | 1049.8465 | 526.0478 | 1067.8397 | 1565 | 1565 | 1466.3456
12 | 100.121 | 0.99062 | 1.195 | 0.90836 | 186.3413 | 50.4025 | 59896.9795 | 57719.4984 | 1150.39 | 1050.39 | 538.4171 | 1133.8161 | 1270 | 1270 | 1148.2679
13 | 92.0032 | 1.0087 | 0.99598 | 1.0853 | 188.8509 | 50.5098 | 59477.0121 | 60205.5896 | 1200.1118 | 1100.1118 | 541.6248 | 1199.7925 | 925 | 925 | 766.2368
14 | 85.4113 | 1.0746 | 0.88953 | 1.0538 | 203.6371 | 45.8355 | 47375.2673 | 47999.2164 | 955.9843 | 855.9843 | 545.31 | 1265.7689 | 551 | 551 | 435.7081
15 | 76.1226 | 1.0591 | 0.86632 | 1.0041 | 201.1264 | 48.2619 | 46072.1756 | 44754.7538 | 891.0951 | 791.0951 | 546.4744 | 1331.7454 | 343 | 343 | 211.1257
16 | 109.2479 | 0.99729 | 0.72918 | 1.0106 | 215.8731 | 49.6181 | 42185.0236 | 41299.2841 | 821.9857 | 721.9857 | 552.2519 | 1397.7218 | 148 | 148 | 87.1765
17 | 92.3421 | 0.98062 | 0.94281 | 1.0399 | 197.7888 | 53.9165 | 57940.6274 | 57654.3442 | 1149.0869 | 1049.0869 | 555.23 | 1463.6982 | 76 | 76 | 30.674
18 | 86.9485 | 0.94395 | 0.85609 | 0.96476 | 195.9428 | 50.1185 | 45703.1249 | 44546.4723 | 886.9294 | 786.9294 | 562.9737 | 1529.6746 | 27 | 27 | 9.1972
19 | 103.4915 | 1.0538 | 0.99735 | 1.0407 | 201.7863 | 50.7283 | 58385.7555 | 56764.7852 | 1131.2957 | 1031.2957 | 564.2922 | 1595.651 | 10 | 10 | 2.3499
20 | 91.4346 | 1.0136 | 0.99757 | 1.0681 | 206.204 | 47.4676 | 55181.5151 | 56777.0415 | 1131.5408 | 1031.5408 | 567.0763 |  | 3 | 3
21 | 102.5353 | 1.0933 | 0.98277 | 1.0675 | 214.4029 | 50.3631 | 58695.588 | 58433.9471 | 1164.6789 | 1064.6789 | 569.1559
22 | 96.3747 | 1.0306 | 0.89849 | 1.0735 | 210.908 | 49.973 | 53375.676 | 54846.1123 | 1092.9222 | 992.9222 | 569.6083
23 | 88.9803 | 0.99534 | 0.76878 | 1.0528 | 203.2968 | 47.1651 | 42553.6853 | 41525.007 | 826.5001 | 726.5001 | 572.4994
24 | 85.3305 | 0.95326 | 1.1159 | 1.0789 | 215.6335 | 52.5209 | 67724.5602 | 68702.0423 | 1370.0408 | 1270.0408 | 574.416
25 | 90.9235 | 1.0092 | 0.83996 | 0.97302 | 190.93 | 49.3474 | 45050.7091 | 44432.2533 | 884.6451 | 784.6451 | 575.111
26 | 104.8281 | 1.0653 | 1.3846 | 1.0511 | 173.9334 | 47.1758 | 74101.7469 | 76579.0316 | 1527.5806 | 1427.5806 | 576.2151
27 | 98.2665 | 0.99006 | 1.0565 | 1.0559 | 207.5964 | 54.2803 | 66046.4178 | 63802.3121 | 1272.0462 | 1172.0462 | 578.9699
28 | 94.8046 | 0.97587 | 0.99786 | 0.97144 | 185.7164 | 45.0536 | 48026.9135 | 47093.7161 | 937.8743 | 837.8743 | 579.6389
29 | 106.988 | 1.0094 | 0.74562 | 0.99392 | 200.795 | 53.4227 | 45561.0437 | 46517.0615 | 926.3412 | 826.3412 | 581.2983
30 | 94.0874 | 0.97859 | 0.88655 | 0.94 | 197.9402 | 49.0536 | 45593.5626 | 44915.6594 | 894.3132 | 794.3132 | 581.7577
31 | 90.1824 | 0.92909 | 0.50893 | 0.94696 | 192.3905 | 51.792 | 29492.5591 | 30658.5518 | 609.171 | 509.171 | 582.9798
32 | 104.0136 | 1.0251 | 0.65446 | 1.0313 | 195.3367 | 48.0259 | 37731.3146 | 38112.4398 | 758.2488 | 658.2488 | 587.3558
33 | 98.8713 | 0.94358 | 0.91567 | 1.0938 | 211.9472 | 50.6121 | 55623.0145 | 54530.5831 | 1086.6117 | 986.6117 | 588.2002
34 | 101.1493 | 0.96395 | 0.80644 | 1.0633 | 207.4366 | 49.38 | 47365.623 | 46346.1422 | 922.9228 | 822.9228 | 589.7568
35 | 107.1511 | 1.0393 | 1.107 | 0.96204 | 205.1126 | 49.7354 | 58710.4607 | 57249.366 | 1140.9873 | 1040.9873 | 590.6925
36 | 104.26 | 1.0129 | 1.0734 | 0.96323 | 187.8292 | 46.0076 | 52612.9403 | 53285.7531 | 1061.7151 | 961.7151 | 592.5603
37 | 97.7802 | 0.95888 | 1.0698 | 0.98379 | 189.8423 | 50.4204 | 57983.51 | 59600.086 | 1188.0017 | 1088.0017 | 593.3567
38 | 89.3122 | 0.935 | 1.0943 | 1.0827 | 201.0666 | 51.0645 | 64966.7723 | 64712.589 | 1290.2518 | 1190.2518 | 594.7991
39 | 112.4259 | 1.1452 | 1.168 | 0.99777 | 200.8364 | 54.1988 | 70344.2622 | 69225.3374 | 1380.5067 | 1280.5067 | 595.2205
40 | 89.3474 | 1.1071 | 0.80732 | 1.0491 | 200.2741 | 45.7672 | 43491.7559 | 44775.9624 | 891.5192 | 791.5192 | 595.2976
41 | 99.9445 | 0.9959 | 1.0397 | 0.97932 | 186.9032 | 49.7927 | 55841.0494 | 57067.2385 | 1137.3448 | 1037.3448 | 596.5247
42 | 96.0697 | 1.0204 | 1.2194 | 0.93057 | 192.6639 | 50.5409 | 62495.575 | 61650.4432 | 1229.0089 | 1129.0089 | 600.9324
43 | 99.251 | 1.0298 | 1.3687 | 0.98991 | 194.2721 | 53.1692 | 77668.3669 | 76355.7204 | 1523.1144 | 1423.1144 | 601.7861
44 | 110.9705 | 1.0035 | 0.8522 | 1.0372 | 196.8761 | 52.6225 | 52570.8134 | 55326.0919 | 1102.5218 | 1002.5218 | 601.9589
45 | 95.4633 | 1.0138 | 1.3845 | 0.98539 | 201.3805 | 47.283 | 69284.7281 | 68467.5983 | 1365.352 | 1265.352 | 604.3777
46 | 86.6839 | 0.9705 | 0.94031 | 0.93347 | 202.5797 | 51.6575 | 49890.7277 | 49082.2393 | 977.6448 | 877.6448 | 604.4117
47 | 80.5193 | 1.0229 | 1.0102 | 1.0449 | 197.7617 | 52.7423 | 60211.7758 | 62062.616 | 1237.2523 | 1137.2523 | 606.2743
48 | 84.042 | 1.0076 | 0.80512 | 1.0724 | 190.9573 | 55.0535 | 52387.4562 | 51920.007 | 1034.4001 | 934.4001 | 606.7624
49 | 104.3375 | 0.94917 | 1.2331 | 0.94816 | 196.2395 | 43.9185 | 55894.5118 | 56639.4834 | 1128.7897 | 1028.7897 | 609.1144
50 | 100.4605 | 1.0675 | 0.96815 | 1.0006 | 185.5504 | 49.3945 | 53334.6457 | 53694.1317 | 1069.8826 | 969.8826 | 610.0768
51 | 88.9536 | 1.0438 | 1.1848 | 0.98547 | 202.5567 | 48.0443 | 60758.6834 | 59141.9606 | 1178.8392 | 1078.8392 | 610.6442
52 | 91.5716 | 0.92049 | 1.0268 | 0.97506 | 182.0511 | 55.4001 | 60317.6438 | 58961.6242 | 1175.2325 | 1075.2325 | 612.3108
53 | 94.1079 | 0.9742 | 1.1043 | 1.0793 | 193.1411 | 46.8942 | 60383.2868 | 59824.1562 | 1192.4831 | 1092.4831 | 612.8447
54 | 96.4374 | 1.0408 | 0.41646 | 1.0409 | 192.1801 | 48.7372 | 26211.1358 | 26708.0738 | 530.1615 | 430.1615 | 612.8779
55 | 107.5189 | 0.9898 | 1.0296 | 0.99464 | 203.5242 | 47.9992 | 54465.2306 | 54490.582 | 1085.8116 | 985.8116 | 613.0114
56 | 112.6679 | 1.0029 | 1.1916 | 1.012 | 180.5061 | 44.5841 | 58977.495 | 58083.1051 | 1157.6621 | 1057.6621 | 613.5786
57 | 96.1965 | 0.89439 | 0.92007 | 0.99682 | 198.7307 | 53.7061 | 54075.4925 | 53664.0484 | 1069.281 | 969.281 | 614.1915
58 | 102.8386 | 1.0102 | 1.0731 | 0.94694 | 216.0952 | 48.2042 | 54206.5047 | 55217.1541 | 1100.3431 | 1000.3431 | 614.8231
59 | 101.952 | 1.1305 | 0.85849 | 1.0195 | 202.4552 | 50.0323 | 49758.5095 | 49748.0318 | 990.9606 | 890.9606 | 616.4364
60 | 82.2293 | 1.0523 | 0.98389 | 1.0276 | 186.155 | 50.469 | 55577.5188 | 54800.399 | 1092.008 | 992.008 | 616.7194
61 | 79.0934 | 0.97111 | 1.0283 | 0.97024 | 197.0525 | 47.6439 | 51390.3426 | 52905.0558 | 1054.1011 | 954.1011 | 617.1605
62 | 104.4407 | 0.93258 | 0.79971 | 0.96753 | 192.3492 | 49.0433 | 42916.323 | 43176.6674 | 859.5333 | 759.5333 | 617.8914
63 | 87.5344 | 1.016 | 0.82274 | 0.91489 | 202.8341 | 52.5528 | 44433.4919 | 43215.738 | 860.3148 | 760.3148 | 617.9957
64 | 104.9441 | 0.89811 | 1.0905 | 1.04 | 195.2393 | 48.866 | 60221.5686 | 59249.7898 | 1180.9958 | 1080.9958 | 618.0499
65 | 123.6061 | 0.95325 | 0.82651 | 1.0383 | 180.7264 | 55.0702 | 53929.4237 | 54038.7583 | 1076.7752 | 976.7752 | 618.7235
66 | 117.6631 | 0.96872 | 1.1264 | 0.9915 | 207.0595 | 50.808 | 62742.7988 | 61839.3152 | 1232.7863 | 1132.7863 | 619.0481
67 | 106.1191 | 0.99687 | 1.0737 | 0.96151 | 193.9563 | 50.1382 | 57258.3784 | 55283.0487 | 1101.661 | 1001.661 | 619.6857
68 | 81.2 | 1.002 | 0.97848 | 0.94145 | 211.2568 | 48.7772 | 49113.1935 | 49233.3372 | 980.6667 | 880.6667 | 620.9499
69 | 94.7501 | 1.0445 | 1.0367 | 1.0858 | 198.9229 | 48.6784 | 59813.3823 | 57757.6132 | 1151.1523 | 1051.1523 | 621.7669
70 | 109.7831 | 1.0434 | 0.96901 | 1.0015 | 204.6592 | 46.3462 | 50492.2407 | 50233.0641 | 1000.6613 | 900.6613 | 623.0721
71 | 87.8664 | 0.97047 | 1.0419 | 0.96115 | 204.0941 | 48.055 | 52425.488 | 51373.9108 | 1023.4782 | 923.4782 | 624.8692
72 | 112.7726 | 0.95016 | 1.1182 | 0.99104 | 216.6509 | 48.8642 | 59604.5069 | 60101.1161 | 1198.0223 | 1098.0223 | 625.1397
73 | 88.4238 | 1.0727 | 1.0213 | 0.96371 | 216.6612 | 52.6131 | 56989.6004 | 58111.5068 | 1158.2301 | 1058.2301 | 625.9109
74 | 115.6973 | 0.97718 | 1.0199 | 1.0435 | 174.8586 | 45.3391 | 53556.1262 | 53310.5215 | 1062.2104 | 962.2104 | 626.2965
75 | 103.2224 | 0.98878 | 0.9767 | 1.0199 | 206.1868 | 50.9399 | 56148.2684 | 56906.3469 | 1134.1269 | 1034.1269 | 626.5088
76 | 87.3212 | 0.98135 | 1.0655 | 0.91538 | 204.7363 | 47.9948 | 51129.1219 | 50143.2868 | 998.8657 | 898.8657 | 626.7291
77 | 89.0137 | 0.95914 | 0.95197 | 1.0283 | 186.6972 | 48.7435 | 52062.7888 | 51885.7034 | 1033.7141 | 933.7141 | 626.9681
78 | 97.1911 | 0.90912 | 0.85157 | 1.0725 | 182.0594 | 53.0144 | 53286.6026 | 51821.6755 | 1032.4335 | 932.4335 | 628.0275
79 | 100.1639 | 0.89777 | 1.044 | 0.99712 | 206.0087 | 49.8627 | 56594.6809 | 54006.7648 | 1076.1353 | 976.1353 | 628.6908
80 | 112.4881 | 0.98106 | 0.79646 | 0.98532 | 197.5063 | 48.4094 | 43529.9631 | 43057.1686 | 857.1434 | 757.1434 | 628.7237
81 | 93.262 | 0.98509 | 1.0185 | 0.99305 | 201.3412 | 49.2257 | 54511.4321 | 54862.9557 | 1093.2591 | 993.2591 | 629.2725
82 | 106.4087 | 0.89943 | 1.0377 | 0.96877 | 194.6031 | 47.1695 | 52127.3359 | 52375.017 | 1043.5003 | 943.5003 | 630.5617
83 | 118.9518 | 0.94938 | 1.0081 | 1.0187 | 200.164 | 43.8332 | 50162.241 | 50194.5112 | 999.8902 | 899.8902 | 630.8488
84 | 83.7519 | 0.96349 | 0.88664 | 1.1661 | 211.8827 | 51.7387 | 57881.7758 | 58112.5179 | 1158.2504 | 1058.2504 | 631.7105
85 | 106.2955 | 1.0141 | 0.67622 | 1.042 | 209.5324 | 54.3386 | 44354.3065 | 43019.1033 | 856.3821 | 756.3821 | 632.1587
86 | 108.3237 | 1.0217 | 0.91834 | 1.005 | 202.8218 | 56.6812 | 58788.625 | 58704.0317 | 1170.0806 | 1070.0806 | 632.2983
87 | 117.3029 | 1.029 | 0.79137 | 1.0353 | 213.8365 | 49.199 | 46461.1492 | 46676.8666 | 929.5373 | 829.5373 | 633.3553
88 | 94.9638 | 0.94254 | 0.69708 | 1.0672 | 202.1809 | 47.7744 | 40018.2065 | 40092.0882 | 797.8418 | 697.8418 | 633.9189
89 | 96.3226 | 1.0833 | 0.88533 | 1.0935 | 186.2379 | 46.5121 | 50066.9576 | 50587.0538 | 1007.7411 | 907.7411 | 634.0475
90 | 98.9648 | 0.95826 | 1.1704 | 0.97207 | 195.034 | 52.0471 | 64343.6766 | 64055.8492 | 1277.117 | 1177.117 | 634.0866
91 | 92.7791 | 0.93746 | 1.0508 | 0.98813 | 214.4013 | 51.5125 | 58180.5425 | 60080.9983 | 1197.62 | 1097.62 | 634.7166
92 | 108.9062 | 1.0165 | 0.99609 | 0.85956 | 206.975 | 49.8196 | 48377.4243 | 48993.0696 | 975.8614 | 875.8614 | 636.6385
93 | 119.9416 | 1.038 | 1.2952 | 1.0022 | 215.6551 | 48.5791 | 69319.2427 | 69878.9687 | 1393.5794 | 1293.5794 | 637.39
94 | 94.8374 | 1.027 | 0.75688 | 0.98436 | 204.85 | 50.8508 | 43043.5678 | 42239.5784 | 840.7916 | 740.7916 | 638.4801
95 | 99.3907 | 0.98003 | 1.2562 | 1.0642 | 193.2891 | 48.0404 | 69094.5204 | 67172.0485 | 1339.441 | 1239.441 | 639.0773
96 | 104.1586 | 1.0318 | 0.86656 | 1.0311 | 195.2661 | 49.2963 | 49542.1498 | 48028.3809 | 956.5676 | 856.5676 | 639.3044
97 | 99.1828 | 0.94693 | 0.79939 | 0.97958 | 186.8821 | 46.1029 | 40618.6012 | 41077.9671 | 817.5593 | 717.5593 | 639.3774
98 | 120.6001 | 1.0251 | 1.1127 | 0.96896 | 207.0996 | 50.451 | 60840.8556 | 57659.8409 | 1149.1968 | 1049.1968 | 639.5619
99 | 103.225 | 0.98877 | 1.1137 | 0.93661 | 204.2955 | 52.8128 | 60682.2004 | 62479.5848 | 1245.5917 | 1145.5917 | 641.3016
100 | 102.4346 | 0.97275 | 0.95915 | 1.0369 | 191.8404 | 45.6572 | 50149.6716 | 49136.2705 | 978.7254 | 878.7254 | 641.3479
101 | 87.6366 | 0.98683 | 0.71337 | 1.0626 | 206.9586 | 48.7055 | 41340.2525 | 41346.8242 | 822.9365 | 722.9365 | 641.4186
102 | 97.096 | 0.97832 | 0.79953 | 0.9757 | 200.1645 | 54.0667 | 47513.4867 | 47572.5082 | 947.4502 | 847.4502 | 642.6884
103 | 98.7585 | 0.99888 | 0.94554 | 0.96287 | 199.3788 | 47.9818 | 48616.6867 | 47111.8292 | 938.2366 | 838.2366 | 642.7797
104 | 82.0134 | 0.9675 | 1.0142 | 0.95104 | 188.2946 | 51.1077 | 53541.0077 | 51948.8907 | 1034.9778 | 934.9778 | 643.4067
105 | 98.8805 | 1.0362 | 1.0183 | 0.97145 | 201.0099 | 51.034 | 55912.9724 | 55795.511 | 1111.9102 | 1011.9102 | 643.4249
106 | 99.9471 | 0.99539 | 0.90611 | 1.0093 | 208.5006 | 44.1963 | 45023.8112 | 45400.4836 | 904.0097 | 804.0097 | 643.6998
107 | 93.6167 | 1.0801 | 0.89999 | 0.95058 | 193.7772 | 50.1265 | 48145.8822 | 47937.2094 | 954.7442 | 854.7442 | 643.731
108 | 87.982 | 1.0458 | 0.93479 | 0.97055 | 196.298 | 49.6674 | 49827.5967 | 49275.655 | 981.5131 | 881.5131 | 643.83
109 | 105.2844 | 1.0458 | 0.97634 | 0.9576 | 187.1669 | 50.1351 | 52581.2335 | 50624.0569 | 1008.4811 | 908.4811 | 643.9136
110 | 105.3163 | 0.92233 | 1.1646 | 1.0032 | 195.6915 | 50.3308 | 63884.4888 | 64542.2492 | 1286.845 | 1186.845 | 644.0902
111 | 94.9913 | 0.91886 | 0.86712 | 0.97506 | 204.9085 | 46.2799 | 43373.8949 | 44013.398 | 876.268 | 776.268 | 644.4551
112 | 101.4185 | 0.90626 | 0.76441 | 1.071 | 207.7767 | 49.9477 | 45690.9485 | 45825.1027 | 912.5021 | 812.5021 | 645.3529
113 | 113.0712 | 1.0674 | 1.0288 | 1.0795 | 204.899 | 52.3367 | 64645.3593 | 64362.9658 | 1283.2593 | 1183.2593 | 645.596
114 | 103.0343 | 1.0882 | 0.72137 | 1.0417 | 175.5133 | 45.928 | 39838.8788 | 40631.5691 | 808.6314 | 708.6314 | 646.4826
115 | 101.249 | 0.93867 | 1.0521 | 0.93276 | 188.7694 | 49.0327 | 52967.7396 | 51808.23 | 1032.1646 | 932.1646 | 647.566
116 | 102.3673 | 1.0059 | 0.88112 | 0.96242 | 196.561 | 48.9137 | 46712.7585 | 44946.5287 | 894.9306 | 794.9306 | 648.2455
117 | 108.3993 | 1.0363 | 1.1691 | 1.0121 | 205.8028 | 51.2055 | 66549.1319 | 65758.8562 | 1311.1771 | 1211.1771 | 649.4316
118 | 104.9687 | 0.98666 | 1.1037 | 0.95188 | 178.4781 | 53.5801 | 62016.5529 | 64413.759 | 1284.2752 | 1184.2752 | 649.9504
119 | 74.7197 | 0.98334 | 0.84186 | 1.0098 | 210.9372 | 49.7404 | 46149.2928 | 46633.0069 | 928.6601 | 828.6601 | 650.2649
120 | 94.685 | 0.9911 | 0.83602 | 1.0596 | 192.263 | 51.0454 | 50200.6221 | 49822.3377 | 992.4468 | 892.4468 | 651.0096
121 | 90.5576 | 0.98193 | 1.2638 | 0.99658 | 214.1168 | 51.3307 | 69429.841 | 69081.1172 | 1377.6223 | 1277.6223 | 651.2271
122 | 117.4759 | 1.0122 | 0.90604 | 0.91762 | 210.3713 | 48.6557 | 46448.5038 | 45370.7734 | 903.4155 | 803.4155 | 651.3602
123 | 92.0985 | 0.95893 | 1.0523 | 0.9618 | 176.867 | 50.0395 | 55240.0243 | 55616.7268 | 1108.3345 | 1008.3345 | 651.3748
124 | 86.9311 | 1.0464 | 0.92192 | 1.0815 | 194.1866 | 51.3235 | 56034.3697 | 57717.1735 | 1150.3435 | 1050.3435 | 651.6583
125 | 101.3386 | 1.0317 | 0.78498 | 0.9638 | 192.2901 | 42.7978 | 37046.0848 | 37165.2187 | 739.3044 | 639.3044 | 652.7024
126 | 110.3229 | 0.94994 | 0.76451 | 0.8727 | 205.924 | 47.8417 | 37138.8304 | 36495.5455 | 725.9109 | 625.9109 | 652.7692
127 | 96.2867 | 1.0184 | 1.1266 | 0.98502 | 204.047 | 47.8654 | 58014.4037 | 58470.8386 | 1165.4168 | 1065.4168 | 653.1882
128 | 121.2067 | 0.99655 | 0.82908 | 1.0362 | 201.0221 | 48.6471 | 47870.7172 | 46663.3999 | 929.268 | 829.268 | 653.4876
129 | 99.1599 | 0.9861 | 1.0378 | 1.0585 | 201.9606 | 50.7537 | 60916.7721 | 61095.8094 | 1217.9162 | 1117.9162 | 653.5405
130 | 95.9778 | 0.99583 | 1.1112 | 1.0929 | 210.484 | 50.7286 | 66661.4814 | 66664.4087 | 1329.2882 | 1229.2882 | 654.415
131 | 92.9843 | 0.95667 | 0.76649 | 0.99895 | 195.6912 | 49.2745 | 42307.7811 | 41727.762 | 830.5552 | 730.5552 | 654.5124
132 | 102.5026 | 0.93984 | 1.1677 | 1.0238 | 194.3145 | 47.6669 | 61771.7739 | 64289.0527 | 1281.7811 | 1181.7811 | 654.9689
133 | 87.1309 | 0.96218 | 1.0323 | 0.94396 | 210.774 | 51.7301 | 54958.0656 | 56002.7524 | 1116.055 | 1016.055 | 656.0441
134 | 112.4802 | 0.9797 | 0.99461 | 1.0303 | 207.2606 | 53.7459 | 61203.31 | 62183.3763 | 1239.6675 | 1139.6675 | 656.2507
135 | 89.8082 | 0.95255 | 1.0266 | 0.96416 | 197.5944 | 50.9381 | 54975.5408 | 54282.7735 | 1081.6555 | 981.6555 | 656.5805
136 | 104.7733 | 0.91338 | 1.1369 | 1.0542 | 203.2713 | 45.1721 | 58664.9792 | 59388.7451 | 1183.7749 | 1083.7749 | 657.0579
137 | 94.3948 | 0.99067 | 0.96555 | 0.93437 | 200.626 | 47.1982 | 47195.5327 | 46420.6441 | 924.4129 | 824.4129 | 658.2488
138 | 85.4831 | 0.97061 | 1.203 | 1.0368 | 193.9169 | 47.262 | 63064.2289 | 60344.3454 | 1202.8869 | 1102.8869 | 659.1689
139 | 107.9012 | 1.01 | 1.0338 | 0.96825 | 211.939 | 50.3269 | 56070.7992 | 55643.1844 | 1108.8637 | 1008.8637 | 659.4009
140 | 100.853 | 1.0155 | 1.1119 | 0.9224 | 204.6331 | 48.4526 | 54862.6037 | 52201.0966 | 1040.0219 | 940.0219 | 660.168
141 | 84.739 | 1.0592 | 1.2284 | 1.1059 | 205.3426 | 49.7341 | 72229.2324 | 68837.0579 | 1372.7412 | 1272.7412 | 660.9137
142 | 106.102 | 0.90692 | 1.0488 | 0.95245 | 194.9721 | 50.3649 | 55354.3128 | 53450.8579 | 1065.0172 | 965.0172 | 661.0659
143 | 82.5574 | 0.95873 | 0.98943 | 1.047 | 191.8908 | 51.8417 | 57998.1612 | 57512.6832 | 1146.2537 | 1046.2537 | 661.0949
144 | 108.5969 | 1.0084 | 1.0881 | 0.95501 | 203.5347 | 52.0511 | 59993.9222 | 59015.9466 | 1176.3189 | 1076.3189 | 661.3979
145 | 102.3587 | 1.0151 | 1.1937 | 0.94894 | 203.7068 | 49.0464 | 60855.829 | 61639.5283 | 1228.7906 | 1128.7906 | 661.6049
146 | 97.8514 | 0.98034 | 0.94569 | 0.92589 | 193.7828 | 49.9797 | 48750.9711 | 48657.0233 | 969.1405 | 869.1405 | 662.5091
147 | 92.6898 | 0.93068 | 1.0259 | 1.0373 | 199.7571 | 51.741 | 59728.1715 | 58146.2763 | 1158.9255 | 1058.9255 | 663.0144
148 | 92.7077 | 0.98676 | 1.1419 | 1.0373 | 197.442 | 50.4704 | 64596.5624 | 65251.5852 | 1301.0317 | 1201.0317 | 663.2072
149 | 95.0112 | 0.95253 | 1.039 | 0.93247 | 197.8878 | 50.2059 | 53383.936 | 54141.3099 | 1078.8262 | 978.8262 | 663.5357
150 | 85.9171 | 0.97827 | 0.7149 | 0.90997 | 218.9298 | 49.1282 | 36307.7372 | 35850.5705 | 713.0114 | 613.0114 | 663.8409
151 | 109.057 | 1.0035 | 1.2421 | 1.0085 | 208.1843 | 50.7282 | 69306.3559 | 68664.064 | 1369.2813 | 1269.2813 | 664.2373
152 | 77.5773 | 1.0266 | 1.0555 | 0.99716 | 183.1338 | 50.4154 | 57263.0492 | 58829.6408 | 1172.5928 | 1072.5928 | 664.4686
153 | 121.1172 | 0.96965 | 1.0248 | 1.0736 | 198.1469 | 49.7505 | 60775.2801 | 61099.8792 | 1217.9976 | 1117.9976 | 664.798
154 | 95.5199 | 1.0305 | 1.2215 | 0.98426 | 193.8053 | 49.752 | 64907.8156 | 66391.2577 | 1323.8252 | 1223.8252 | 666.2338
155 | 111.7371 | 1.0257 | 1.0005 | 0.96712 | 203.1526 | 49.3723 | 53633.7271 | 53588.6313 | 1067.7726 | 967.7726 | 666.2809
156 | 100.9944 | 1.1139 | 0.78304 | 1.0465 | 209.5703 | 51.3363 | 48051.8564 | 48436.8616 | 964.7372 | 864.7372 | 666.3271
157 | 118.638 | 0.9968 | 1.1091 | 0.97436 | 202.7312 | 47.6588 | 57340.4215 | 56865.0533 | 1133.3011 | 1033.3011 | 666.4082
158 | 98.5276 | 0.9658 | 1.2189 | 1.0238 | 201.5059 | 50.5272 | 68057.9211 | 68139.4868 | 1358.7897 | 1258.7897 | 666.5622
159 | 93.8414 | 0.97737 | 1.0891 | 0.96947 | 210.5639 | 43.9812 | 50680.9904 | 51125.6693 | 1018.5134 | 918.5134 | 666.609
160 | 101.3458 | 1.0607 | 1.0561 | 1.0123 | 205.9841 | 52.3394 | 61791.147 | 61768.7486 | 1231.375 | 1131.375 | 667.229
161 | 91.7962 | 0.98655 | 0.94716 | 0.96395 | 196.3689 | 46.356 | 46717.9187 | 46329.4757 | 922.5895 | 822.5895 | 667.5072
162 | 104.626 | 1.0025 | 0.91202 | 1.005 | 189.9284 | 48.2443 | 49469.0642 | 47977.2238 | 955.5445 | 855.5445 | 667.6301
163 | 106.1426 | 1.0531 | 1.0178 | 1.0241 | 207.545 | 50.5349 | 58531.1312 | 58024.4324 | 1156.4886 | 1056.4886 | 667.6516
164 | 100.0438 | 1.0214 | 1.1373 | 1.021 | 182.6992 | 45.8546 | 58115.6375 | 57501.9645 | 1146.0393 | 1046.0393 | 667.9577
165 | 93.1758 | 0.95654 | 0.74345 | 1.0144 | 204.6861 | 52.7033 | 44648.1502 | 42774.6982 | 851.494 | 751.494 | 668.1811
166 | 111.4669 | 0.9743 | 0.9338 | 1.0224 | 205.1293 | 49.6259 | 52972.1892 | 53583.8415 | 1067.6768 | 967.6768 | 668.3232
167 | 98.498 | 1.0548 | 1.0084 | 1.0654 | 204.8394 | 52.1904 | 61698.736 | 63355.0579 | 1263.1012 | 1163.1012 | 668.4567
168 | 113.3179 | 1.0708 | 1.01 | 1.019 | 198.2515 | 47.1509 | 54442.6731 | 55922.5932 | 1114.4519 | 1014.4519 | 668.7289
169 | 75.9663 | 0.9939 | 1.001 | 1.0038 | 213.5302 | 50.1359 | 54374.2216 | 54282.7922 | 1081.6558 | 981.6558 | 669.0599
170 | 95.2805 | 1.0796 | 0.88912 | 1.1196 | 199.0661 | 50.9668 | 56176.86 | 56214.5333 | 1120.2907 | 1020.2907 | 669.2748
171 | 98.7593 | 0.92125 | 0.83177 | 0.97201 | 210.6428 | 47.6774 | 43095.0402 | 43768.8116 | 871.3762 | 771.3762 | 670.2244
172 | 100.4153 | 1.0288 | 0.93927 | 0.98258 | 204.8962 | 49.591 | 51096.4815 | 50281.2943 | 1001.6259 | 901.6259 | 670.2383
173 | 94.5208 | 0.99499 | 0.90087 | 1.0521 | 182.7582 | 51.34 | 53670.9921 | 52513.286 | 1046.2657 | 946.2657 | 671.2273
174 | 97.2153 | 1.0563 | 1.0593 | 0.98807 | 210.5211 | 47.6729 | 55003.9204 | 56163.6315 | 1119.2726 | 1019.2726 | 671.2568
175 | 84.6614 | 1.035 | 1.1122 | 1.0728 | 216.5629 | 54.362 | 69845.4514 | 69187.3998 | 1379.748 | 1279.748 | 671.4539
176 | 95.1928 | 1.0918 | 0.9208 | 0.95764 | 197.2891 | 54.0177 | 53444.0074 | 54262.2592 | 1081.2452 | 981.2452 | 671.896
177 | 100.531 | 1.0377 | 1.0317 | 1.0607 | 209.4624 | 51.1082 | 61470.0037 | 61487.1129 | 1225.7423 | 1125.7423 | 672.6712
178 | 92.4912 | 0.99766 | 0.76795 | 1.0544 | 199.8994 | 51.3378 | 46507.2614 | 46284.0868 | 921.6817 | 821.6817 | 672.7255
179 | 102.1105 | 0.95874 | 0.84197 | 0.93318 | 208.2749 | 49.6921 | 44116.7871 | 43744.107 | 870.8821 | 770.8821 | 673.1515
180 | 76.9925 | 0.98004 | 1.1573 | 1.008 | 224.1683 | 46.2209 | 57630.596 | 59374.1491 | 1183.483 | 1083.483 | 673.9416
181 | 107.3944 | 1.0166 | 1.0022 | 1.0219 | 190.0494 | 53.9869 | 61375.7631 | 62000.5604 | 1236.0112 | 1136.0112 | 674.0405
182 | 74.976 | 1.0089 | 1.1023 | 0.99725 | 193.2735 | 53.5719 | 63136.0802 | 62816.0222 | 1252.3204 | 1152.3204 | 674.2812
183 | 102.2804 | 0.96369 | 0.97952 | 1.0075 | 197.1526 | 50.6816 | 55210.8873 | 56257.7016 | 1121.154 | 1021.154 | 674.9036
184 | 98.2676 | 0.99676 | 0.98113 | 1.0366 | 199.1712 | 53.5096 | 59859.3044 | 60724.6618 | 1210.4932 | 1110.4932 | 674.9796
185 | 112.8419 | 0.926 | 0.95905 | 1.0472 | 205.3556 | 52.347 | 58246.4021 | 58987.0134 | 1175.7403 | 1075.7403 | 675.0469
186 | 100.729 | 1.004 | 1.0172 | 1.0464 | 195.714 | 50.1545 | 58655.6239 | 61134.1219 | 1218.6824 | 1118.6824 | 675.0947
187 | 106.0501 | 1.0168 | 0.90691 | 0.9582 | 215.1749 | 47.9676 | 47071.3049 | 47029.2285 | 936.5846 | 836.5846 | 675.9354
188 | 116.4389 | 0.95985 | 0.82195 | 0.97087 | 195.2829 | 48.0894 | 43945.5202 | 44304.5232 | 882.0905 | 782.0905 | 676.0909
189 | 98.9447 | 0.9402 | 0.79761 | 1.08 | 220.8074 | 49.5408 | 47503.6807 | 45487.0469 | 905.7409 | 805.7409 | 676.4096
190 | 111.3533 | 0.98165 | 0.99295 | 1.0882 | 197.7812 | 52.5066 | 62669.6337 | 63416.7709 | 1264.3354 | 1164.3354 | 676.6071
191 | 99.0484 | 0.95023 | 0.88906 | 0.91492 | 201.937 | 49.6403 | 45252.0769 | 45508.9342 | 906.1787 | 806.1787 | 676.7283
192 | 112.6139 | 1.0827 | 0.7642 | 0.97997 | 207.1601 | 47.6445 | 41696.7406 | 41751.9402 | 831.0388 | 731.0388 | 676.7339
193 | 77.7918 | 0.96894 | 1.0137 | 1.0404 | 193.0225 | 51.4787 | 58368.419 | 58623.2331 | 1168.4647 | 1068.4647 | 676.9671
194 | 84.8279 | 1.0141 | 1.1601 | 1.0093 | 180.6763 | 52.457 | 66112.4972 | 67214.3888 | 1340.2878 | 1240.2878 | 676.9767
195 | 103.4673 | 1.0357 | 1.0501 | 0.94638 | 196.5654 | 48.5267 | 53621.5064 | 53515.9077 | 1066.3182 | 966.3182 | 677.0913
196 | 99.2397 | 0.96048 | 0.83559 | 0.96931 | 182.7921 | 43.7443 | 39782.7183 | 39751.636 | 791.0327 | 691.0327 | 677.4744
197 | 105.0773 | 1.0264 | 1.06 | 1.0402 | 206.7981 | 46.5296 | 56529.1121 | 55417.3991 | 1104.348 | 1004.348 | 677.4812
198 | 102.4904 | 1.0769 | 1.0489 | 1.0131 | 221.4396 | 50.8882 | 59910.5855 | 60013.5439 | 1196.2709 | 1096.2709 | 677.6331
199 | 87.4735 | 0.91954 | 0.99118 | 0.98091 | 192.5648 | 53.1656 | 56160.0703 | 56660.4667 | 1129.2093 | 1029.2093 | 678.1294
200 | 88.3063 | 0.97968 | 0.9086 | 0.98385 | 200.8652 | 51.7533 | 50941.9212 | 51795.0666 | 1031.9013 | 931.9013 | 678.3428
201 | 97.5512 | 1.0144 | 0.95572 | 0.99328 | 202.4801 | 51.0224 | 53687.0527 | 53580.9221 | 1067.6184 | 967.6184 | 678.4345
202 | 89.6356 | 0.93469 | 1.2923 | 0.94058 | 198.0336 | 47.9571 | 62508.5315 | 63412.7034 | 1264.2541 | 1164.2541 | 678.5856
203 | 102.3328 | 1.0322 | 1.2604 | 1.0388 | 172.1699 | 48.0158 | 68111.9851 | 69988.1496 | 1395.763 | 1295.763 | 678.6946
204 | 87.8362 | 1.04 | 1.1746 | 0.97264 | 197.8376 | 50.2938 | 62252.695 | 61768.8807 | 1231.3776 | 1131.3776 | 678.769
205 | 93.9023 | 1.0839 | 1.1575 | 0.90829 | 201.3178 | 51.4144 | 59487.5201 | 61353.1241 | 1223.0625 | 1123.0625 | 679.0098
206 | 117.8522 | 1.0768 | 1.0957 | 1.0478 | 192.1645 | 51.1895 | 65458.4582 | 64644.4071 | 1288.8881 | 1188.8881 | 679.0435
207 | 114.848 | 1.0985 | 1.0225 | 0.986 | 180.3238 | 49.7187 | 56579.9632 | 57863.1603 | 1153.2632 | 1053.2632 | 679.3643
208 | 97.4954 | 0.96434 | 1.1726 | 1.0956 | 198.4074 | 49.9282 | 69037.8949 | 69074.2602 | 1377.4852 | 1277.4852 | 679.8354
209 | 107.8512 | 0.95742 | 0.85379 | 0.94805 | 202.6983 | 49.9547 | 45796.0147 | 47241.3814 | 940.8276 | 840.8276 | 680.0529
210 | 86.3935 | 1.0509 | 0.85237 | 1.0135 | 195.7525 | 49.8533 | 47789.8214 | 48749.4192 | 970.9884 | 870.9884 | 680.5943
211 | 101.7623 | 0.91485 | 0.98769 | 0.95568 | 187.9997 | 47.7733 | 49729.4 | 50651.288 | 1009.0258 | 909.0258 | 680.8787
212 | 99.4786 | 0.93363 | 0.86187 | 0.93905 | 207.6785 | 50.5242 | 45791.5872 | 45513.2557 | 906.2651 | 806.2651 | 681.1569
213 | 105.3806 | 1.0099 | 0.85737 | 0.96444 | 192.5384 | 49.1327 | 46048.1839 | 46693.3712 | 929.8674 | 829.8674 | 681.9234
214 | 87.5075 | 1.0716 | 0.93205 | 1.0385 | 206.051 | 49.4597 | 52716.1355 | 52709.1129 | 1050.1823 | 950.1823 | 682.4057
215 | 98.5188 | 0.97666 | 0.69701 | 1.0392 | 184.7252 | 50.257 | 41421.9446 | 40381.6522 | 803.633 | 703.633 | 682.4745
216 | 110.093 | 0.97516 | 1.0132 | 0.95229 | 211.8231 | 51.3911 | 55315.1541 | 55744.3251 | 1110.8865 | 1010.8865 | 682.4926
217 | 94.8315 | 1.0253 | 0.87063 | 0.94816 | 209.4156 | 51.8173 | 48022.3566 | 50518.4117 | 1006.3682 | 906.3682 | 682.5104
218 | 98.6772 | 1.0487 | 1.0011 | 0.9793 | 203.2091 | 51.1774 | 55673.3815 | 57001.4897 | 1136.0298 | 1036.0298 | 682.6007
219 | 119.6488 | 0.97202 | 0.96539 | 1.002 | 187.0881 | 51.3038 | 55779.1095 | 55606.2869 | 1108.1257 | 1008.1257 | 682.7302
220 | 92.5793 | 0.94944 | 1.0988 | 1.0824 | 202.3975 | 51.0433 | 65399.7271 | 65257.7758 | 1301.1555 | 1201.1555 | 682.8409
221 | 101.0911 | 0.8937 | 0.89768 | 0.9845 | 202.5284 | 53.5378 | 52354.4541 | 51579.4539 | 1027.5891 | 927.5891 | 683.1275
222 | 98.6504 | 0.93755 | 1.1104 | 1.0756 | 207.4748 | 49.3936 | 63767.0681 | 66449.5377 | 1324.9908 | 1224.9908 | 683.1509
223 | 114.58 | 0.99671 | 0.95299 | 1.0254 | 211.3126 | 53.1784 | 58247.9127 | 58959.5049 | 1175.1901 | 1075.1901 | 683.3764
224 | 88.7383 | 0.99759 | 1.1419 | 0.99239 | 189.6818 | 49.4634 | 60619.8007 | 59692.3847 | 1189.8477 | 1089.8477 | 683.5612
225 | 109.5357 | 0.98342 | 1.0626 | 1.0264 | 192.3402 | 52.6064 | 63235.8468 | 62937.14 | 1254.7428 | 1154.7428 | 683.5721
226 | 85.3075 | 1.0117 | 1.2669 | 1.023 | 208.8706 | 48.8276 | 67707.1319 | 69581.0571 | 1387.6211 | 1287.6211 | 683.7354
227 | 100.1855 | 0.97861 | 0.91985 | 1.0191 | 192.4877 | 51.9304 | 53964.1794 | 52665.012 | 1049.3002 | 949.3002 | 683.7375
228 | 89.2032 | 1.0247 | 0.94024 | 0.94804 | 203.3136 | 47.7126 | 47094.6958 | 48008.0236 | 956.1605 | 856.1605 | 684.2069
229 | 98.8854 | 1.0288 | 0.89676 | 0.95898 | 184.7864 | 47.1112 | 45492.0264 | 45364.1535 | 903.2831 | 803.2831 | 684.4193
230 | 104.1674 | 0.9069 | 1.0898 | 1.031 | 196.6374 | 48.1033 | 58790.2654 | 58730.0523 | 1170.601 | 1070.601 | 684.4562
231 | 102.797 | 1.0141 | 1.0757 | 1.0255 | 215.0858 | 47.4923 | 57553.7284 | 55415.2983 | 1104.306 | 1004.306 | 684.9766
232 | 99.5805 | 0.96536 | 0.69183 | 0.98638 | 199.3354 | 49.5875 | 38805.3806 | 39134.7289 | 778.6946 | 678.6946 | 685.04
233 | 78.8712 | 1.1457 | 1.0096 | 0.96397 | 181.234 | 48.598 | 51867.3933 | 52634.4618 | 1048.6892 | 948.6892 | 685.0801
234 | 115.3382 | 0.98258 | 0.68853 | 0.92608 | 201.8397 | 50.2945 | 37971.3501 | 37750.4816 | 751.0096 | 651.0096 | 685.5988
235 | 108.7173 | 0.96447 | 0.95081 | 0.98843 | 197.6987 | 53.1878 | 55760.8386 | 57221.3457 | 1140.4269 | 1040.4269 | 685.6224
236 | 100.1816 | 0.99473 | 0.94515 | 1.037 | 212.0899 | 52.0641 | 56431.9234 | 56245.4103 | 1120.9082 | 1020.9082 | 685.8443
237 | 102.6274 | 1.0208 | 0.97577 | 1.0595 | 214.3032 | 48.8473 | 55833.5825 | 55597.9773 | 1107.9595 | 1007.9595 | 686.1652
238 | 104.1719 | 0.94018 | 1.059 | 1.0205 | 211.8326 | 49.9985 | 59144.1841 | 57954.0959 | 1155.0819 | 1055.0819 | 686.8381
239 | 91.0098 | 1.0127 | 0.75932 | 1.0287 | 182.9463 | 46.627 | 40903.149 | 43549.1554 | 866.9831 | 766.9831 | 687.0217
240 | 91.5612 | 1.0142 | 0.73721 | 0.9858 | 192.4053 | 49.0086 | 40359.699 | 41704.144 | 830.0829 | 730.0829 | 687.6142
241 | 88.5741 | 0.97345 | 1.1984 | 1.1038 | 191.3772 | 52.7627 | 74536.2412 | 73995.527 | 1475.9105 | 1375.9105 | 687.8709
242 | 88.272 | 1.0515 | 0.85954 | 1.0129 | 190.3394 | 49.3102 | 47696.62 | 48510.2987 | 966.206 | 866.206 | 688.0778
243 | 100.342 | 0.88743 | 1.2393 | 0.94475 | 201.8106 | 49.9928 | 63185.9505 | 61800.9921 | 1232.0198 | 1132.0198 | 688.2612
244 | 93.2617 | 0.97241 | 1.2108 | 1.0537 | 190.5495 | 51.2357 | 70206.3786 | 68640.7045 | 1368.8141 | 1268.8141 | 688.4537
245 | 114.9423 | 1.0181 | 1.0505 | 1.0992 | 209.7431 | 47.0994 | 60106.4354 | 59198.8925 | 1179.9778 | 1079.9778 | 688.5293
246 | 110.9978 | 0.9986 | 1.1363 | 0.95977 | 203.1946 | 54.8687 | 66124.7831 | 64737.2955 | 1290.7459 | 1190.7459 | 688.7554
247 | 105.7559 | 0.94958 | 0.81631 | 1.0252 | 191.9371 | 52.3289 | 49238.8506 | 49842.6061 | 992.8521 | 892.8521 | 688.9303
248 | 94.8278 | 0.93763 | 0.98477 | 0.99696 | 202.2439 | 48.7577 | 52406.9985 | 53460.2298 | 1065.2046 | 965.2046 | 689.519
249 | 108.9 | 1.0405 | 1.1312 | 1.0106 | 202.8992 | 48.8845 | 61627.7324 | 59737.633 | 1190.7527 | 1090.7527 | 689.5205
250 | 96.1299 | 0.98944 | 0.83187 | 0.99577 | 205.1153 | 54.4559 | 50493.578 | 51207.8471 | 1020.1569 | 920.1569 | 690.6994
251 | 111.7668 | 1.0275 | 0.85471 | 1.0404 | 203.1305 | 43.7747 | 44157.0909 | 44120.2063 | 878.4041 | 778.4041 | 691.0327
252 | 96.5645 | 1.0602 | 0.83569 | 1.018 | 195.2983 | 49.0733 | 46967.907 | 46991.7517 | 935.835 | 835.835 | 691.1107
253 | 99.2646 | 1.0379 | 1.0334 | 1.0468 | 199.9522 | 53.3212 | 63374.0558 | 61648.3576 | 1228.9672 | 1128.9672 | 691.4439
254 | 103.9263 | 0.98822 | 0.97935 | 1.0269 | 200.3444 | 50.4572 | 56127.4751 | 57131.0038 | 1138.6201 | 1038.6201 | 691.5765
255 | 108.6388 | 1.0854 | 1.1283 | 1.0758 | 180.5567 | 47.3991 | 63305.9929 | 63191.6761 | 1259.8335 | 1159.8335 | 692.2731
256 | 85.7516 | 1.0419 | 1.0755 | 0.94191 | 190.2785 | 47.2788 | 52308.5518 | 52198.2829 | 1039.9657 | 939.9657 | 692.4741
257 | 97.2523 | 0.96487 | 1.0318 | 1.036 | 201.9374 | 48.9104 | 57073.6903 | 58666.7699 | 1169.3354 | 1069.3354 | 692.8435
258 | 88.6381 | 1.004 | 0.76015 | 1.0531 | 206.2695 | 52.6871 | 47071.1243 | 47670.5647 | 949.4113 | 849.4113 | 693.1071
259 | 107.1299 | 0.9315 | 0.97175 | 1.0131 | 185.7052 | 51.7826 | 56332.5678 | 56643.595 | 1128.8719 | 1028.8719 | 693.3628
260 | 94.3613 | 0.97989 | 1.0367 | 0.98436 | 178.2527 | 45.241 | 50527.1042 | 51654.5892 | 1029.0918 | 929.0918 | 693.5952
261 | 98.0635 | 1.0892 | 1.2392 | 1.0429 | 182.6781 | 49.63 | 69623.6412 | 68512.6549 | 1366.2531 | 1266.2531 | 693.7498
262 | 102.8008 | 1.0837 | 0.8168 | 0.98757 | 189.0734 | 52.7037 | 48574.0516 | 48415.5991 | 964.312 | 864.312 | 694.3357
263 | 105.0016 | 0.9347 | 0.93329 | 0.96286 | 205.9953 | 51.4154 | 51455.4085 | 51805.7542 | 1032.1151 | 932.1151 | 694.4437
264 | 91.1677 | 0.98648 | 1.112 | 0.97642 | 200.4356 | 46.2569 | 54584.0884 | 54508.0859 | 1086.1617 | 986.1617 | 694.4798
265 | 95.6134 | 1.0962 | 1.255 | 1.0668 | 176.038 | 48.3769 | 70016.5674 | 70410.0939 | 1404.2019 | 1304.2019 | 694.5245
266 | 90.8141 | 1.0081 | 1.2509 | 0.94759 | 198.3101 | 50.848 | 65125.6546 | 64307.1526 | 1282.1431 | 1182.1431 | 694.5314
267 | 97.3102 | 1.0125 | 1.0239 | 0.96971 | 215.1156 | 51.3395 | 56246.4315 | 56948.6744 | 1134.9735 | 1034.9735 | 694.5421
268 | 101.1731 | 0.98259 | 0.84389 | 0.98384 | 201.8513 | 49.9158 | 46607.0103 | 46255.7721 | 921.1154 | 821.1154 | 694.8143
269 | 121.2749 | 0.9209 | 1.1095 | 0.91827 | 211.0753 | 52.7462 | 59840.3838 | 59840.7712 | 1192.8154 | 1092.8154 | 695.9146
270 | 102.6539 | 1.0228 | 0.75511 | 1.0162 | 194.5165 | 52.4677 | 45963.5254 | 45689.3967 | 909.7879 | 809.7879 | 696.0047
271 | 92.1845 | 0.95351 | 0.98263 | 1.0365 | 192.1014 | 45.6508 | 50701.4467 | 50221.3239 | 1000.4265 | 900.4265 | 696.0985
272 | 124.5325 | 0.98129 | 0.88184 | 0.94823 | 208.0643 | 56.3195 | 54184.0878 | 54363.1113 | 1083.2622 | 983.2622 | 696.1975
273 | 105.9036 | 1.0032 | 0.69796 | 0.97034 | 194.1115 | 52.6786 | 41467.8512 | 42569.7542 | 847.3951 | 747.3951 | 697.3171
274 | 102.4164 | 1.0914 | 1.176 | 0.98183 | 203.7763 | 47.9127 | 60882.9952 | 60382.5106 | 1203.6502 | 1103.6502 | 697.6924
275 | 102.8628 | 0.89011 | 0.73107 | 0.96097 | 198.6216 | 46.6783 | 37265.3991 | 37270.93 | 741.4186 | 641.4186 | 697.808
276 | 95.9034 | 0.92589 | 1.0302 | 1.0913 | 211.8725 | 51.1028 | 62198.3862 | 61819.0227 | 1232.3805 | 1132.3805 | 697.8418
277 | 104.249 | 0.96048 | 1.0931 | 1.0436 | 205.3122 | 43.5249 | 54218.6823 | 55146.7325 | 1098.9346 | 998.9346 | 698.1335
278 | 100.0813 | 0.98879 | 1.0334 | 1.0913 | 207.0339 | 51.5649 | 63459.6569 | 61946.6579 | 1234.9332 | 1134.9332 | 698.2287
279 | 107.938 | 1.0107 | 1.2419 | 0.98024 | 213.3326 | 45.8625 | 61049.2401 | 60990.9017 | 1215.818 | 1115.818 | 698.5433
280 | 103.0294 | 0.99463 | 0.95462 | 1.0547 | 194.2105 | 51.5206 | 57344.7233 | 55819.647 | 1112.3929 | 1012.3929 | 699.1323
281 | 105.966 | 1.0977 | 0.92783 | 1.043 | 184.4496 | 48.7073 | 52984.7478 | 53565.6158 | 1067.3123 | 967.3123 | 699.1964
282 | 100.0006 | 0.9979 | 1.0306 | 0.94225 | 196.7466 | 47.4828 | 51046.9464 | 51162.034 | 1019.2407 | 919.2407 | 699.3693
283 | 87.3713 | 1.0634 | 0.87609 | 0.94486 | 202.9136 | 50.3064 | 46519.6414 | 45132.3513 | 898.647 | 798.647 | 699.5581
284 | 95.4899 | 1.0764 | 0.94839 | 1.0029 | 210.2891 | 51.8733 | 54879.8613 | 54718.1979 | 1090.364 | 990.364 | 699.8241
285 | 111.2518 | 0.94705 | 0.94725 | 0.98211 | 205.4402 | 51.6573 | 53704.7496 | 52902.7573 | 1054.0551 | 954.0551 | 700.1878
286 | 98.2981 | 0.98307 | 1.1519 | 1.0127 | 203.3578 | 54.3767 | 68888.237 | 69336.4917 | 1382.7298 | 1282.7298 | 700.8969
287 | 109.1968 | 0.9106 | 1.169 | 0.92624 | 208.4235 | 51.0793 | 60594.5357 | 60402.6158 | 1204.0523 | 1104.0523 | 700.9745
288 | 96.668 | 0.96044 | 1.0953 | 0.99915 | 203.4973 | 51.2246 | 61018.6057 | 61155.7751 | 1219.1155 | 1119.1155 | 701.1584
289 | 84.5254 | 0.94444 | 0.8038 | 1.0185 | 190.8021 | 49.6787 | 44827.2453 | 44843.7498 | 892.875 | 792.875 | 701.974
290 | 97.8185 | 0.9982 | 0.94584 | 1.0869 | 184.3614 | 51.1352 | 57747.9769 | 54944.9585 | 1094.8992 | 994.8992 | 702.632
291 | 104.228 | 1.1036 | 0.8748 | 0.90992 | 192.8965 | 50.4663 | 46169.2416 | 45333.3463 | 902.6669 | 802.6669 | 702.6341
292 | 79.8188 | 1.0319 | 0.90096 | 0.9902 | 183.0679 | 52.4802 | 51324.86 | 51653.1813 | 1029.0636 | 929.0636 | 702.841
293 | 95.867 | 0.95571 | 1.0011 | 1.0131 | 204.1731 | 49.6316 | 55089.853 | 55974.0019 | 1115.48 | 1015.48 | 703.1918
294 | 93.9234 | 1.0441 | 1.0434 | 1.0387 | 197.9214 | 52.5151 | 62266.9946 | 63322.3344 | 1262.4467 | 1162.4467 | 703.3747
295 | 99.0727 | 0.97314 | 0.87054 | 1.0199 | 195.9737 | 52.6589 | 52027.4067 | 53232.6388 | 1060.6528 | 960.6528 | 703.633
296 | 100.0942 | 1.0316 | 0.86552 | 1.0531 | 202.5146 | 47.6349 | 48537.6952 | 49260.2279 | 981.2046 | 881.2046 | 704.1015
297 | 94.843 | 0.98882 | 0.92858 | 0.90578 | 186.0676 | 49.7251 | 46672.8521 | 46469.7997 | 925.396 | 825.396 | 704.4458
298 | 112.2789 | 1.1387 | 1.1763 | 0.93864 | 218.0733 | 48.62 | 60117.606 | 58130.488 | 1158.6098 | 1058.6098 | 704.4982
299 | 109.6848 | 1.012 | 1.1299 | 0.98247 | 201.2712 | 46.7224 | 57253.083 | 57592.1158 | 1147.8423 | 1047.8423 | 704.5703
300 | 111.2131 | 0.9441 | 0.87799 | 0.92931 | 186.538 | 50.2532 | 46465.6615 | 46956.1317 | 935.1226 | 835.1226 | 704.6609
301 | 104.1347 | 1.0123 | 1.0908 | 1.0297 | 205.5134 | 51.6959 | 63722.5055 | 63002.237 | 1256.0447 | 1156.0447 | 704.6751
302 | 110.4106 | 0.93941 | 1.0363 | 1.1475 | 190.0868 | 49.5473 | 64247.8489 | 64183.7673 | 1279.6753 | 1179.6753 | 704.729
303 | 87.3696 | 0.99173 | 1.1869 | 1.0216 | 206.1128 | 52.4824 | 68389.5114 | 69930.6233 | 1394.6125 | 1294.6125 | 705.0816
304 | 105.5629 | 0.95302 | 1.0069 | 0.96366 | 194.1986 | 48.6677 | 52310.927 | 52360.3504 | 1043.207 | 943.207 | 705.2029
305 | 107.6808 | 0.96644 | 0.86736 | 0.93486 | 211.1957 | 48.4508 | 44539.9688 | 44749.3069 | 890.9861 | 790.9861 | 705.3166
306 | 100.2608 | 0.98805 | 1.0637 | 1.0078 | 191.2888 | 48.0658 | 56478.8741 | 55040.3815 | 1096.8076 | 996.8076 | 705.4649
307 | 89.3789 | 0.98805 | 1.0105 | 1.0171 | 200.9268 | 53.0533 | 59412.5287 | 59893.3119 | 1193.8662 | 1093.8662 | 706.1357
308 | 102.3509 | 0.97203 | 0.93455 | 1.0824 | 191.6886 | 50.3427 | 56124.1553 | 54809.7862 | 1092.1957 | 992.1957 | 706.1937
309 | 100.6498 | 0.93205 | 1.0641 | 1.0031 | 200.8309 | 50.4211 | 58751.2295 | 59041.7902 | 1176.8358 | 1076.8358 | 706.5889
310 | 74.3682 | 0.97259 | 1.0629 | 1.078 | 203.8504 | 54.0352 | 66022.6557 | 64042.2187 | 1276.8444 | 1176.8444 | 706.6857
311 | 101.7638 | 1.0289 | 1.0427 | 1.0816 | 215.4495 | 51.7006 | 63933.1968 | 62216.9797 | 1240.3396 | 1140.3396 | 706.7399
312 | 98.2205 | 0.94825 | 0.83711 | 0.98437 | 202.9591 | 50.8647 | 46854.2417 | 47475.399 | 945.508 | 845.508 | 706.9581
313 | 91.3087 | 0.96466 | 0.99842 | 1.0687 | 211.9689 | 50.1675 | 58158.6978 | 58915.7299 | 1174.3146 | 1074.3146 | 707.1642
314 | 108.6134 | 0.99049 | 1.1096 | 0.984 | 199.6905 | 50.8371 | 61175.0057 | 61559.8494 | 1227.197 | 1127.197 | 707.2722
315 | 106.8591 | 0.9605 | 1.014 | 0.97446 | 196.0915 | 54.8115 | 59979.5951 | 59927.1338 | 1194.5427 | 1094.5427 | 707.4445
316 | 100.9062 | 0.97508 | 0.92315 | 0.98252 | 209.4739 | 51.0573 | 51542.742 | 52233.2326 | 1040.6647 | 940.6647 | 707.5099
317 | 93.9129 | 0.93791 | 1.2434 | 1.043 | 202.1812 | 46.9527 | 65229.9395 | 64005.9377 | 1276.1188 | 1176.1188 | 707.7015
318 | 116.1227 | 0.98109 | 1.1686 | 1.0026 | 205.1552 | 51.1344 | 65943.1257 | 66482.1831 | 1325.6437 | 1225.6437 | 708.4146
319 | 113.7454 | 1.067 | 0.92502 | 1.0123 | 202.7142 | 46.4831 | 49372.2154 | 49598.5194 | 987.9704 | 887.9704 | 708.4325
320 | 84.105 | 1.0985 | 1.206 | 1.064 | 215.0329 | 51.7386 | 71386.6149 | 70515.8137 | 1406.3163 | 1306.3163 | 708.4793
321 | 98.5182 | 1.0244 | 1.1899 | 1.0342 | 196.1307 | 48.449 | 64708.0256 | 65928.8198 | 1314.5764 | 1214.5764 | 708.5602
322 | 94.6725 | 0.99876 | 0.9493 | 1.0263 | 204.7512 | 50.0898 | 53743.9029 | 54922.1274 | 1094.4425 | 994.4425 | 708.5851
323 | 108.5399 | 1.0868 | 0.94275 | 1.0109 | 187.7332 | 47.6969 | 51269.047 | 50629.1966 | 1008.5839 | 908.5839 | 708.6314
324 | 100.856 | 0.9586 | 0.83054 | 0.95351 | 197.8583 | 50.3327 | 44924.1917 | 44923.2609 | 894.4652 | 794.4652 | 708.6832
325 | 115.7279 | 1.0344 | 1.0014 | 0.96186 | 201.9301 | 48.0611 | 52248.5688 | 53166.27 | 1059.3254 | 959.3254 | 708.9971
326 | 103.0262 | 1.1059 | 1.1891 | 0.99409 | 165.3549 | 50.5031 | 65620.0897 | 63571.4712 | 1267.4294 | 1167.4294 | 709.1746
327 | 101.3471 | 1.0048 | 1.1971 | 1.0637 | 188.3166 | 51.5962 | 71138.7995 | 70479.9311 | 1405.5986 | 1305.5986 | 709.4744
328 | 103.8532 | 1.0063 | 0.96204 | 1.039 | 221.2914 | 49.2029 | 54543.516 | 53015.6478 | 1056.313 | 956.313 | 709.5154
329 | 103.1015 | 0.96586 | 0.99601 | 0.96521 | 187.8356 | 50.6721 | 53947.9633 | 54580.7347 | 1087.6147 | 987.6147 | 709.8373
330 | 108.7024 | 1.0825 | 0.88918 | 1.1204 | 195.4593 | 49.1199 | 54909.457 | 54474.532 | 1085.4906 | 985.4906 | 709.8741
331 | 102.4632 | 1.0523 | 1.0941 | 0.95678 | 209.3375 | 50.6912 | 58740.2525 | 55147.5514 | 1098.951 | 998.951 | 709.9091
332 | 86.6342 | 1.0182 | 0.9421 | 0.95472 | 193.3061 | 48.6263 | 48219.0973 | 48401.5416 | 964.0308 | 864.0308 | 710.3437
333 | 112.9567 | 1.0423 | 1.0421 | 0.98639 | 207.8129 | 52.1526 | 59957.1222 | 61536.957 | 1226.7391 | 1126.7391 | 710.3904
334 | 111.8198 | 0.96026 | 0.9795 | 0.9961 | 195.0326 | 50.4613 | 54847.4805 | 55246.5687 | 1100.9314 | 1000.9314 | 710.4149
335 | 98.4803 | 0.93149 | 1.0387 | 0.9698 | 203.1664 | 49.7759 | 54911.4556 | 55653.3689 | 1109.0674 | 1009.0674 | 710.5574
336 | 90.6403 | 0.95796 | 0.83327 | 1.0038 | 206.7188 | 53.7414 | 49822.5529 | 48871.9885 | 973.4398 | 873.4398 | 710.5775
337 | 118.5583 | 1.0146 | 0.92516 | 1.0873 | 187.1119 | 50.5935 | 57164.6626 | 58216.9434 | 1160.3389 | 1060.3389 | 710.8663
338 | 120.9548 | 1.0224 | 0.87921 | 0.95106 | 196.9832 | 48.9829 | 47212.9572 | 45569.3134 | 907.3863 | 807.3863 | 711.7209
339 | 101.884 | 0.96107 | 1.0181 | 1.0152 | 211.0456 | 48.3469 | 54915.5325 | 53781.7703 | 1071.6354 | 971.6354 | 712.1352
340 | 106.6873 | 0.99232 | 1.007 | 0.9565 | 199.3114 | 49.8597 | 53503.2863 | 53165.7787 | 1059.3156 | 959.3156 | 712.255
341 | 103.9509 | 0.9378 | 1.0554 | 0.98833 | 187.6916 | 54.2962 | 62117.3251 | 61136.6667 | 1218.7333 | 1118.7333 | 712.5974
342 | 95.1184 | 0.96046 | 0.93019 | 1.0133 | 212.7125 | 49.8163 | 51719.2277 | 52076.045 | 1037.5209 | 937.5209 | 713.1634
343 | 107.0094 | 1.089 | 0.90768 | 0.98598 | 185.0851 | 47.6341 | 48366.3918 | 50720.5842 | 1010.4117 | 910.4117 | 713.2952
344 | 85.7973 | 0.92684 | 1.0358 | 0.98718 | 195.8126 | 47.5318 | 52579.4662 | 51138.0216 | 1018.7604 | 918.7604 | 713.3771
345 | 123.8741 | 1.0834 | 1.0799 | 0.93338 | 186.5443 | 50.9403 | 58370.3159 | 60736.0278 | 1210.7206 | 1110.7206 | 713.5397
346 | 86.8008 | 1.099 | 1.0836 | 1.0499 | 183.4637 | 54.294 | 67132.2755 | 66967.8407 | 1335.3568 | 1235.3568 | 713.5764
347 | 100.3773 | 0.93187 | 1.0122 | 1.0471 | 200.9155 | 52.2916 | 60517.075 | 60387.0729 | 1203.7415 | 1103.7415 | 713.6577
348 | 112.6686 | 0.99102 | 0.73106 | 1.0677 | 199.1496 | 43.107 | 38660.2472 | 39622.6854 | 788.4537 | 688.4537 | 713.7454
349 | 98.1108 | 0.93724 | 1.1394 | 1.0794 | 190.4283 | 53.2152 | 70534.7423 | 70100.3654 | 1398.0073 | 1298.0073 | 713.7747
350 | 99.1278 | 1.0121 | 1.1782 | 1.015 | 200.2022 | 47.5533 | 61838.6495 | 61747.5682 | 1230.9514 | 1130.9514 | 713.9574
351 | 100.0199 | 1.0613 | 1.1925 | 0.94625 | 199.145 | 50.4501 | 62481.4364 | 62514.0602 | 1246.2812 | 1146.2812 | 713.9708
352 | 99.9366 | 0.98861 | 1.167 | 0.97188 | 203.6352 | 49.7829 | 61583.6943 | 60398.0218 | 1203.9604 | 1103.9604 | 714.0764
353 | 99.3758 | 1.091 | 0.78006 | 0.94479 | 198.8863 | 48.0507 | 40821.1876 | 39613.0621 | 788.2612 | 688.2612 | 714.3957
354 | 103.4422 | 0.92302 | 1.0966 | 1.01 | 213.8102 | 51.4284 | 62084.0272 | 60852.1084 | 1213.0422 | 1113.0422 | 714.4054
355 | 107.2744 | 0.98859 | 0.93685 | 0.89474 | 225.22 | 54.7951 | 51967.6321 | 51616.9647 | 1028.3393 | 928.3393 | 714.5263
356 | 92.9168 | 1.0537 | 0.90684 | 1.0127 | 208.9143 | 50.2988 | 51327.9077 | 51301.3084 | 1022.0262 | 922.0262 | 714.6334
357 | 93.4673 | 0.96866 | 0.93835 | 1.0968 | 202.2934 | 48.6283 | 54652.4465 | 57321.6232 | 1142.4325 | 1042.4325 | 715.1546
358 | 110.6891 | 0.98342 | 1.0365 | 0.98111 | 176.7896 | 53.1678 | 60030.8803 | 58783.9407 | 1171.6788 | 1071.6788 | 715.1941
359 | 90.4897 | 1.004 | 1.0617 | 1.042 | 190.0369 | 50.2445 | 60338.1952 | 59716.0185 | 1190.3204 | 1090.3204 | 715.2667
360 | 87.0474 | 1.0616 | 0.93733 | 0.99727 | 209.0258 | 47.4662 | 48965.6688 | 49366.7478 | 983.335 | 883.335 | 715.2895
361 | 116.4148 | 1.0644 | 1.1703 | 0.98598 | 214.6745 | 44.6055 | 57213.0796 | 57487.8844 | 1145.7577 | 1045.7577 | 715.4984
362 | 114.5698 | 0.95521 | 1.1564 | 0.99226 | 190.1601 | 51.6013 | 65047.9225 | 64464.6478 | 1285.293 | 1185.293 | 715.5313
363 | 94.263 | 0.92633 | 1.1283 | 1.1195 | 199.0085 | 51.3875 | 69594.0547 | 68781.3461 | 1371.6269 | 1271.6269 | 716.1399
364 | 107.2708 | 1.046 | 1.1199 | 0.98251 | 221.5858 | 48.0385 | 58470.2367 | 60027.9287 | 1196.5586 | 1096.5586 | 716.3491
365 | 105.1048 | 0.98512 | 0.82837 | 1.03 | 190.7206 | 52.9425 | 50841.739 | 49841.5719 | 992.8314 | 892.8314 | 716.4034
366 | 100.8398 | 1.05 | 0.82925 | 0.99894 | 206.5279 | 50.1369 | 47046.7593 | 46634.3458 | 928.6869 | 828.6869 | 717.3128
367 | 103.0077 | 0.95754 | 1.0886 | 0.92946 | 197.7392 | 50.0148 | 55736.1825 | 56785.0052 | 1131.7001 | 1031.7001 | 717.3601
368 | 96.0975 | 0.89783 | 0.95692 | 1.0643 | 203.5485 | 46.0273 | 51049.9141 | 50398.0146 | 1003.9603 | 903.9603 | 717.4965
369 | 99.5366 | 0.99732 | 1.1538 | 0.99105 | 202.1463 | 51.9876 | 64811.1513 | 63737.2676 | 1270.7454 | 1170.7454 | 717.5469
370 | 94.6008 | 0.99507 | 1.0935 | 0.98607 | 189.5561 | 51.7175 | 60825.2167 | 61079.2723 | 1217.5854 | 1117.5854 | 717.5593
371 | 101.4784 | 0.94533 | 1.1715 | 1.0111 | 188.9483 | 46.7707 | 60074.1614 | 59861.3071 | 1193.2261 | 1093.2261 | 718.0783
372 | 105.1794 | 1.0321 | 1.0702 | 0.96956 | 209.0477 | 50.3289 | 57892.8601 | 57341.3915 | 1142.8278 | 1042.8278 | 718.1358
373 | 103.5913 | 1.0495 | 1.1177 | 0.99595 | 212.1002 | 51.64 | 63308.5261 | 61751.3534 | 1231.0271 | 1131.0271 | 718.2823
374 | 101.3771 | 0.99721 | 0.94063 | 1.0261 | 190.5246 | 45.9377 | 49172.6333 | 48329.73 | 962.5946 | 862.5946 | 718.2969
375 | 105.498 | 1.0127 | 0.98373 | 1.0117 | 202.7606 | 51.8783 | 57374.3853 | 57324.8945 | 1142.4979 | 1042.4979 | 718.3682
376 | 110.9434 | 1.0243 | 1.1122 | 0.96467 | 205.0024 | 52.4473 | 62435.8104 | 61056.6714 | 1217.1334 | 1117.1334 | 718.4355
377 | 117.8495 | 1.0518 | 1.0451 | 1.0257 | 183.3835 | 47.5121 | 57001.7779 | 57927.0434 | 1154.5409 | 1054.5409 | 718.5695
378 | 106.8986 | 1.0221 | 0.84674 | 0.93583 | 184.4792 | 52.4678 | 47493.1998 | 48428.4965 | 964.5699 | 864.5699 | 718.7375
379 | 90.9399 | 1.0315 | 0.97931 | 1.1238 | 192.8123 | 52.3507 | 62716.9848 | 62609.0479 | 1248.181 | 1148.181 | 719.1311
380 | 97.4161 | 1.008 | 0.9984 | 0.99475 | 201.1319 | 50.5877 | 55410.786 | 54045.7406 | 1076.9148 | 976.9148 | 719.6061
381 | 106.8864 | 0.99857 | 0.96139 | 0.96817 | 198.2196 | 47.3112 | 49284.8292 | 48199.7997 | 959.996 | 859.996 | 719.7363
382 | 112.181 | 0.92706 | 1.0695 | 1.0422 | 203.0265 | 50.2449 | 61435.2573 | 61496.2016 | 1225.924 | 1125.924 | 719.7487
383 | 98.5442 | 0.98379 | 0.74446 | 1.0159 | 184.0151 | 50.4527 | 43233.8443 | 42662.5744 | 849.2515 | 749.2515 | 719.9917
384 | 101.7196 | 0.99883 | 1.2568 | 1.0485 | 198.3031 | 49.0364 | 69799.8096 | 70542.0133 | 1406.8403 | 1306.8403 | 720.062
385 | 105.4846 | 1.0338 | 1.0298 | 0.99927 | 199.6362 | 50.5234 | 57699.5018 | 58146.9645 | 1158.9393 | 1058.9393 | 720.1442
386 | 90.0697 | 0.94756 | 0.87504 | 1.0552 | 182.4852 | 53.423 | 54067.5835 | 54524.9903 | 1086.4998 | 986.4998 | 720.2733
387 | 111.6503 | 0.95731 | 0.80905 | 0.9953 | 194.4696 | 50.7458 | 46481.3025 | 46488.8526 | 925.7771 | 825.7771 | 720.2869
388 | 109.3237 | 1.0342 | 1.0635 | 0.87807 | 193.7658 | 53.0785 | 55763.1981 | 54620.5125 | 1088.4103 | 988.4103 | 720.5051
389 | 101.2821 | 1.068 | 1.0215 | 0.97688 | 213.6817 | 53.3378 | 59205.5612 | 60815.4395 | 1212.3088 | 1112.3088 | 720.6379
390 | 103.5233 | 0.93823 | 0.72886 | 0.96628 | 202.1001 | 46.995 | 37864.4609 | 38897.0779 | 773.9416 | 673.9416 | 720.8192
391 | 107.599 | 1.1014 | 1.0623 | 1.0301 | 211.4783 | 49.0874 | 59744.0691 | 60554.0109 | 1207.0802 | 1107.0802 | 720.9827
392 | 85.5655 | 0.98358 | 1.1463 | 1.0136 | 199.5395 | 48.1023 | 60135.8642 | 60765.2481 | 1211.305 | 1111.305 | 721.0749
393 | 100.9355 | 0.9592 | 0.85163 | 1.0344 | 210.3987 | 45.0826 | 44288.9612 | 45151.0815 | 899.0216 | 799.0216 | 721.1
394 | 108.1868 | 0.99501 | 0.94398 | 1.0144 | 213.9073 | 53.1901 | 56873.3087 | 56846.7172 | 1132.9343 | 1032.9343 | 721.6461
395 | 105.6273 | 0.9733 | 0.95997 | 1.0878 | 209.3634 | 50.7283 | 58397.3294 | 58240.6203 | 1160.8124 | 1060.8124 | 721.9027
396 | 95.5011 | 0.98847 | 0.98654 | 1.0976 | 210.7475 | 48.4951 | 57298.2338 | 56209.2816 | 1120.1856 | 1020.1856 | 721.9857
397 | 86.1696 | 0.98242 | 0.92405 | 1.0089 | 198.5139 | 44.1604 | 45105.5243 | 44920.6863 | 894.4137 | 794.4137 | 722.0529
398 | 94.288 | 0.99245 | 0.91374 | 1.0023 | 208.3049 | 46.8505 | 47501.3469 | 47363.7097 | 943.2742 | 843.2742 | 722.1891
399 | 108.7215 | 1.0067 | 0.8076 | 1.0266 | 209.1555 | 51.9583 | 48972.7884 | 49110.2869 | 978.2057 | 878.2057 | 722.205
400 | 91.7822 | 1.0087 | 1.1995 | 1.0235 | 195.1879 | 49.2509 | 65221.0275 | 66829.0268 | 1332.5805 | 1232.5805 | 722.3229
401 | 93.7924 | 0.99083 | 0.93027 | 0.95921 | 219.3484 | 50.2531 | 49731.7428 | 49174.8784 | 979.4976 | 879.4976 | 722.4291
402 | 91.2377 | 1.0412 | 1.0269 | 1.0118 | 191.5082 | 51.9833 | 59143.4409 | 59009.3738 | 1176.1875 | 1076.1875 | 722.608
403 | 109.5258 | 0.92999 | 1.0875 | 1.0138 | 194.689 | 50.5457 | 61068.9965 | 62109.6265 | 1238.1925 | 1138.1925 | 722.8463
404 | 107.7054 | 1.0269 | 0.89426 | 0.94009 | 181.225 | 56.8557 | 54267.3447 | 53490.9935 | 1065.8199 | 965.8199 | 722.8941
405 | 95.4156 | 0.97257 | 1.0113 | 0.97621 | 217.3015 | 49.6271 | 53816.1529 | 52837.5776 | 1052.7516 | 952.7516 | 722.9365
406 | 94.6521 | 0.99911 | 0.96483 | 0.96482 | 205.6779 | 53.7063 | 55278.8224 | 55592.3731 | 1107.8475 | 1007.8475 | 723.1757
407 | 82.2913 | 1.023 | 1.1887 | 1.0045 | 189.9023 | 49.0207 | 62846.9673 | 64127.804 | 1278.5561 | 1178.5561 | 723.3578
408 | 104.1687 | 1.0378 | 1.4077 | 0.97905 | 212.878 | 49.4898 | 73772.4129 | 73122.9671 | 1458.4593 | 1358.4593 | 723.4588
409 | 106.7428 | 1.0582 | 1.3147 | 1.0559 | 206.0342 | 50.097 | 75411.4331 | 74218.9195 | 1480.3784 | 1380.3784 | 723.5776
410 | 78.6692 | 1.073 | 1.0752 | 1.0108 | 209.1859 | 50.8441 | 59759.0729 | 59045.5399 | 1176.9108 | 1076.9108 | 723.8535
411 | 107.2057 | 0.95013 | 1.0582 | 0.98483 | 201.7456 | 46.0732 | 52911.8343 | 54021.4889 | 1076.4298 | 976.4298 | 724.0748
412 | 108.0956 | 0.89574 | 0.91501 | 1.0276 | 201.3553 | 48.4415 | 50439.0878 | 51017.9451 | 1016.3589 | 916.3589 | 724.2183
413 | 108.0187 | 1.0656 | 0.85355 | 1.0119 | 203.316 | 50.4081 | 49544.3402 | 50178.0951 | 999.5619 | 899.5619 | 724.2991
414 | 116.4554 | 0.99222 | 1.09 | 0.99464 | 199.0271 | 48.4387 | 58312.4068 | 56703.8142 | 1130.0763 | 1030.0763 | 724.4026
415 | 94.4457 | 1.036 | 1.0865 | 0.95472 | 209.4557 | 49.9872 | 56954.2976 | 58442.008 | 1164.8402 | 1064.8402 | 724.6217
416 | 109.9714 | 0.91338 | 1.0937 | 0.94136 | 194.3091 | 51.5951 | 58496.5933 | 55152.8528 | 1099.0571 | 999.0571 | 724.6696
417 | 87.5746 | 1.052 | 0.96216 | 1.0004 | 213.2661 | 51.4586 | 54486.0171 | 53125.9273 | 1058.5185 | 958.5185 | 724.915
418 | 89.3027 | 0.95998 | 0.9502 | 1.0436 | 195.8118 | 50.7545 | 54874.477 | 55935.7495 | 1114.715 | 1014.715 | 724.9383
419 | 88.0036 | 0.96795 | 1.0943 | 1.1044 | 212.0936 | 54.143 | 70262.8945 | 70601.9391 | 1408.0388 | 1308.0388 | 724.9526
420 | 92.9875 | 1.0518 | 1.0128 | 0.93663 | 192.9613 | 47.1426 | 49524.3496 | 48374.5878 | 963.4918 | 863.4918 | 724.9642
421 | 83.3962 | 1.0424 | 0.89238 | 1.0574 | 183.0839 | 48.6447 | 50313.6883 | 50728.9184 | 1010.5784 | 910.5784 | 725.2497
422 | 100.2616 | 1.0114 | 1.0908 | 0.95193 | 198.5563 | 54.3937 | 62194.5093 | 62527.0162 | 1246.5403 | 1146.5403 | 725.4078
423 | 99.064 | 0.98977 | 0.86288 | 0.96406 | 217.1538 | 50.5636 | 47236.9718 | 46732.7775 | 930.6555 | 830.6555 | 725.5936
424 | 98.2696 | 0.99152 | 0.77665 | 1.0013 | 200.4482 | 55.3707 | 48654.4372 | 50046.4066 | 996.9281 | 896.9281 | 725.9022
425 | 113.8122 | 1.0368 | 1.1192 | 1.0208 | 211.8005 | 47.2756 | 59800.8814 | 56805.6968 | 1132.1139 | 1032.1139 | 726.2417
426 | 118.0705 | 1.0295 | 1.0625 | 0.97033 | 222.2302 | 55.3725 | 64040.6181 | 62736.4422 | 1250.7288 | 1150.7288 | 726.2819
427 | 108.5959 | 0.96402 | 1.0285 | 1.0711 | 188.5719 | 53.6663 | 64925.2312 | 66119.6457 | 1318.3929 | 1218.3929 | 726.3791
428 | 109.1946 | 0.92266 | 1.0273 | 1.0341 | 187.3516 | 53.2509 | 62120.8637 | 60925.4937 | 1214.5099 | 1114.5099 | 726.464
429 | 93.7457 | 0.9983 | 0.8484 | 1.0191 | 197.9053 | 52.7049 | 50698.747 | 48431.8213 | 964.6364 | 864.6364 | 726.5001
430 | 99.8831 | 0.90983 | 0.93937 | 0.96463 | 202.9774 | 50.5313 | 50583.9952 | 52176.3108 | 1039.5262 | 939.5262 | 726.5372
431 | 92.5991 | 1.0237 | 1.1779 | 0.97494 | 216.7185 | 50.1154 | 62520.7654 | 64159.9606 | 1279.1992 | 1179.1992 | 726.5899
432 | 88.627 | 0.97859 | 0.67194 | 1.0282 | 201.4564 | 50.4048 | 39396.7649 | 39323.7248 | 782.4745 | 682.4745 | 726.7845
433 | 95.057 | 0.91346 | 1.1185 | 1.0831 | 192.7658 | 54.6022 | 71080.211 | 71916.6808 | 1434.3336 | 1334.3336 | 726.9484
434 | 101.0037 | 0.94652 | 1.0433 | 1.0921 | 201.3302 | 51.8082 | 64183.7438 | 64516.5503 | 1286.331 | 1186.331 | 726.9657
435 | 98.1672 | 0.90769 | 0.90528 | 0.97187 | 196.3827 | 51.7773 | 50364.1084 | 51651.7416 | 1029.0348 | 929.0348 | 727.1999
436 | 96.2239 | 1.0345 | 0.93483 | 0.91177 | 207.4087 | 49.8142 | 47624.7751 | 49865.1575 | 993.3032 | 893.3032 | 727.2443
437 | 105.1961 | 0.96331 | 0.8372 | 1.0227 | 213.4094 | 53.0007 | 50965.1272 | 50611.9159 | 1008.2383 | 908.2383 | 727.2549
438 | 103.7343 | 0.95868 | 1.1208 | 1.1539 | 200.1611 | 51.3158 | 71667.511 | 70651.7325 | 1409.0347 | 1309.0347 | 727.5
439 | 112.666 | 1.0166 | 0.84748 | 0.9997 | 179.4137 | 54.1732 | 52281.6502 | 52288.5989 | 1041.772 | 941.772 | 727.6489
440 | 92.1154 | 0.91697 | 0.99373 | 1.066 | 210.6112 | 46.6522 | 53568.415 | 51778.6184 | 1031.5724 | 931.5724 | 727.8026
441 | 103.9987 | 0.98967 | 0.94295 | 0.9862 | 196.8468 | 50.7565 | 52621.1731 | 51656.1161 | 1029.1223 | 929.1223 | 727.8758
442 | 106.6567 | 0.92946 | 0.86478 | 0.97606 | 193.9821 | 47.5353 | 45030.0331 | 44549.5986 | 886.992 | 786.992 | 728.1075
443 | 96.4425 | 0.95022 | 1.099 | 1.1299 | 192.2259 | 51.4672 | 68818.0521 | 65626.6936 | 1308.5339 | 1208.5339 | 728.4982
444 | 115.4598 | 0.98319 | 0.76274 | 0.96927 | 210.4091 | 49.2442 | 42206.8753 | 42388.1634 | 843.7633 | 743.7633 | 728.6937
445 | 85.7185 | 1.0255 | 1.0858 | 1.033 | 204.3152 | 49.6845 | 60300.3093 | 61459.5545 | 1225.1911 | 1125.1911 | 728.8559
446 | 98.6915 | 0.96451 | 0.81814 | 1.0327 | 193.1317 | 49.3079 | 46544.9697 | 48126.8958 | 958.5379 | 858.5379 | 729.1272
447 | 99.4879 | 1.0567 | 0.9143 | 1.0626 | 210.7509 | 53.974 | 58321.9528 | 60062.2537 | 1197.2451 | 1097.2451 | 729.2741
448 | 113.1331 | 0.92184 | 0.73885 | 1.0109 | 196.2676 | 49.1054 | 41994.7899 | 41747.2391 | 830.9448 | 730.9448 | 729.5365
449 | 100.7304 | 0.99451 | 1.0257 | 0.92762 | 196.2865 | 50.1327 | 52917.725 | 52296.931 | 1041.9386 | 941.9386 | 729.5978
450 | 80.1283 | 0.94111 | 0.9102 | 1.0677 | 202.5829 | 50.5248 | 53111.8885 | 53890.9528 | 1073.8191 | 973.8191 | 729.6316
451 | 114.5582 | 1.0248 | 0.92115 | 1.069 | 197.0874 | 50.4561 | 55805.2977 | 55701.2031 | 1110.0241 | 1010.0241 | 729.8421
452 | 109.6068 | 0.98272 | 0.94195 | 1.0061 | 209.051 | 52.2603 | 55367.141 | 57165.8652 | 1139.3173 | 1039.3173 | 730.0829
453 | 90.8438 | 1.0173 | 1.0643 | 0.99187 | 228.148 | 50.9444 | 58715.2011 | 60514.2148 | 1206.2843 | 1106.2843 | 730.2698
454 | 88.787 | 0.93104 | 1.1035 | 1.0031 | 208.1021 | 53.7518 | 64151.1892 | 65459.3425 | 1305.1868 | 1205.1868 | 730.2813
455 | 78.4704 | 0.96851 | 1.009 | 1.037 | 189.279 | 50.3012 | 56646.6596 | 54570.3376 | 1087.4068 | 987.4068 | 730.3865
456 | 104.313 | 0.98468 | 1.314 | 0.99407 | 206.0313 | 48.4332 | 68446.3651 | 67741.4398 | 1350.8288 | 1250.8288 | 730.4479
457 | 100.3136 | 0.97078 | 0.8834 | 0.98097 | 201.4494 | 50.1163 | 48511.6544 | 48381.0169 | 963.6203 | 863.6203 | 730.5552
458 | 99.7907 | 1.0006 | 1.0065 | 1.1483 | 188.1464 | 55.0386 | 69295.9504 | 67680.4971 | 1349.6099 | 1249.6099 | 730.6419
459 | 104.6 | 0.89754 | 0.76532 | 0.91037 | 203.4869 | 53.952 | 42858.4185 | 42859.8983 | 853.198 | 753.198 | 730.7277
460 | 96.3062 | 0.95132 | 1.0803 | 0.96352 | 194.3626 | 51.3 | 58292.4514 | 57143.4498 | 1138.869 | 1038.869 | 730.7451
461 | 86.0751 | 1.0327 | 1.0301 | 1.0597 | 205.2702 | 49.0466 | 58103.783 | 58223.477 | 1160.4695 | 1060.4695 | 730.9383
462 | 113.33 | 0.99066 | 1.0781 | 1.0443 | 189.2052 | 50.6418 | 62888.3501 | 62658.3841 | 1249.1677 | 1149.1677 | 730.9448
463 | 109.3248 | 0.95361 | 0.76157 | 0.97018 | 215.1912 | 48.5636 | 41160.068 | 40465.8282 | 805.3166 | 705.3166 | 731.0388
464 | 107.7885 | 0.97255 | 0.97765 | 1.0649 | 193.1115 | 48.0489 | 55253.2433 | 56063.5115 | 1117.2702 | 1017.2702 | 731.177
465 | 114.0057 | 0.94888 | 1.044 | 0.93965 | 196.6332 | 48.7154 | 53257.1196 | 54875.463 | 1093.5093 | 993.5093 | 731.2524
466 | 114.968 | 1.0127 | 0.81999 | 1.0267 | 206.9374 | 48.9971 | 47159.6956 | 46598.9434 | 927.9789 | 827.9789 | 731.2636
467 | 95.7628 | 0.99649 | 1.0764 | 1.0999 | 205.9552 | 49.8831 | 64021.2079 | 63032.2806 | 1256.6456 | 1156.6456 | 731.5594
468 | 95.6711 | 1.0283 | 0.87617 | 0.99929 | 204.7113 | 45.4288 | 44449.3683 | 43970.4279 | 875.4086 | 775.4086 | 731.6931
469 | 84.2528 | 0.98808 | 0.8405 | 0.9853 | 210.77 | 43.7595 | 40093.1569 | 38952.346 | 775.0469 | 675.0469 | 731.8065
470 | 108.7973 | 1.056 | 0.87216 | 1.0718 | 201.9426 | 46.4265 | 48933.5135 | 51051.1306 | 1017.0226 | 917.0226 | 732.0429
471 | 104.0298 | 0.90922 | 0.97253 | 0.98364 | 204.224 | 53.3829 | 56320.829 | 55589.4795 | 1107.7896 | 1007.7896 | 732.1599
472 | 102.3531 | 0.98996 | 1.094 | 0.99091 | 197.306 | 46.9921 | 55902.1735 | 56328.3559 | 1122.5671 | 1022.5671 | 732.5685
473 | 111.1261 | 0.95041 | 0.95987 | 0.96494 | 195.857 | 48.4933 | 50232.6065 | 49776.7762 | 991.5355 | 891.5355 | 733.1324
474 | 104.653 | 1.0625 | 0.93691 | 0.94833 | 190.7596 | 50.1248 | 50300.2709 | 49617.5792 | 988.3516 | 888.3516 | 733.2786
475 | 116.53 | 0.98182 | 1.0491 | 0.93903 | 179.0134 | 50.3142 | 55502.5024 | 55856.6068 | 1113.1321 | 1013.1321 | 733.3391
476 | 87.6079 | 0.95855 | 0.88534 | 0.99217 | 205.7061 | 51.9032 | 50156.3739 | 49485.9818 | 985.7196 | 885.7196 | 733.4976
477 | 101.7133 | 0.88987 | 0.79828 | 1.0267 | 188.4425 | 47.3851 | 43313.1575 | 43770.4988 | 871.41 | 771.41 | 733.8234
478 | 101.7381 | 0.99292 | 0.78144 | 0.94278 | 207.4104 | 52.1518 | 43897.4575 | 42799.5959 | 851.9919 | 751.9919 | 733.9608
479 | 116.0447 | 1.0134 | 1.0574 | 1.0479 | 196.4732 | 50.0782 | 61576.0732 | 61087.4765 | 1217.7495 | 1117.7495 | 734.2104
480 | 94.6931 | 1.1045 | 0.84908 | 1.0003 | 201.733 | 47.1955 | 45222.3213 | 45089.1252 | 897.7825 | 797.7825 | 734.3038
481 | 120.5432 | 1.0014 | 1.138 | 1.0385 | 202.3631 | 50.1721 | 65550.7027 | 65746.4781 | 1310.9296 | 1210.9296 | 734.3372
482 | 86.2723 | 0.89685 | 0.85711 | 0.99405 | 219.3826 | 54.8707 | 51215.6168 | 52585.5638 | 1047.7113 | 947.7113 | 734.3726
483 | 110.1487 | 1.011 | 0.93265 | 1.0453 | 200.7864 | 51.2388 | 55857.735 | 58255.1361 | 1161.1027 | 1061.1027 | 734.4474
484 | 98.0278 | 0.99235 | 1.0571 | 1.0084 | 200.9954 | 52.2869 | 61023.4707 | 61308.8484 | 1222.177 | 1122.177 | 734.4919
485 | 124.7017 | 0.99731 | 0.90557 | 1.0127 | 182.9479 | 47.5333 | 49686.906 | 50710.0833 | 1010.2017 | 910.2017 | 734.7281
486 | 101.7455 | 0.98636 | 1.0939 | 1.0142 | 199.2528 | 48.0659 | 58346.6603 | 60147.0693 | 1198.9414 | 1098.9414 | 734.7829
487 | 100.7393 | 0.97627 | 1.032 | 0.98605 | 195.0108 | 48.9682 | 54838.7694 | 53859.9117 | 1073.1982 | 973.1982 | 735.0159
488 | 97.7514 | 0.97891 | 1.0047 | 1.0367 | 190.707 | 54.5841 | 62263.5978 | 63419.8941 | 1264.3979 | 1164.3979 | 735.7065
489 | 115.9175 | 1.0008 | 0.86257 | 0.99888 | 184.9455 | 47.5856 | 46705.1757 | 45265.3595 | 901.3072 | 801.3072 | 735.8523
490 | 115.7391 | 1.0184 | 1.0406 | 0.96161 | 229.449 | 47.3654 | 53208.9795 | 52619.4872 | 1048.3897 | 948.3897 | 735.8967
491 | 85.5225 | 1.0695 | 0.98238 | 0.98547 | 206.1734 | 48.9124 | 52032.718 | 51953.3702 | 1035.0674 | 935.0674 | 735.9445
492 | 123.5185 | 1.0668 | 1.0779 | 1.027 | 186.935 | 51.1099 | 63500.5976 | 65397.5448 | 1303.9509 | 1203.9509 | 735.9849
493 | 110.1699 | 1.0412 | 1.0436 | 1.0167 | 191.53 | 52.6337 | 62073.2728 | 62430.0933 | 1244.6019 | 1144.6019 | 736.301
494 | 94.8302 | 0.98862 | 1.1235 | 0.98258 | 197.7783 | 50.5268 | 60714.6469 | 60904.2236 | 1214.0845 | 1114.0845 | 736.7679
495 | 105.9503 | 1.0714 | 1.1437 | 1.0339 | 191.0199 | 49.6388 | 64523.1153 | 63643.3387 | 1268.8668 | 1168.8668 | 736.8186
496 | 100.5833 | 1.0449 | 0.82173 | 0.89409 | 199.1832 | 51.4402 | 43398.4052 | 44380.4851 | 883.6097 | 783.6097 | 737.0079
497 | 105.5541 | 1.0365 | 1.0353 | 1.0466 | 220.3697 | 50.5737 | 60551.2389 | 61181.1101 | 1219.6222 | 1119.6222 | 737.0171
498 | 99.5863 | 0.97752 | 0.99025 | 1.0145 | 196.5233 | 49.7843 | 55056.9331 | 55134.7117 | 1098.6942 | 998.6942 | 737.297
499 | 110.0253 | 0.9833 | 0.87588 | 0.9852 | 213.5898 | 48.4883 | 47300.85 | 47940.2663 | 954.8053 | 854.8053 | 737.3679
500 | 120.5664 | 0.91606 | 1.1174 | 0.96855 | 207.5346 | 50.2584 | 60151.708 | 60400.6689 | 1204.0134 | 1104.0134 | 737.3806
501 | 108.0606 | 1.0154 | 1.0421 | 0.96734 | 198.5945 | 47.3506 | 53127.1418 | 52623.9352 | 1048.4787 | 948.4787 | 737.3894
502 | 90.5421 | 1.0147 | 0.92783 | 1.0209 | 195.4901 | 49.66 | 51797.7083 | 53810.6073 | 1072.2121 | 972.2121 | 737.5413
503 | 85.6083 | 0.96611 | 0.99146 | 0.98985 | 197.2058 | 48.8132 | 52139.2992 | 50720.5151 | 1010.4103 | 910.4103 | 737.6758
504 | 100.9005 | 1.0084 | 0.87723 | 1.0458 | 208.0778 | 46.7629 | 47866.71 | 48394.4407 | 963.8888 | 863.8888 | 737.7472
505 | 100.2714 | 1.0476 | 1.2774 | 1.0104 | 193.4336 | 51.0712 | 71479.0548 | 67235.5107 | 1340.7102 | 1240.7102 | 738.2029
506 | 102.5162 | 1.0268 | 1.0387 | 0.97644 | 208.2375 | 46.8533 | 52659.744 | 52644.0706 | 1048.8814 | 948.8814 | 738.2103
507 | 80.0591 | 0.88855 | 1.068 | 1.085 | 211.2601 | 49.1318 | 60639.6627 | 59754.2697 | 1191.0854 | 1091.0854 | 738.2992
508 | 92.3219 | 0.98066 | 0.69743 | 0.99399 | 205.0487 | 46.8631 | 36935.2454 | 37697.5211 | 749.9504 | 649.9504 | 738.4882
509 | 101.305 | 0.97394 | 1.2012 | 0.9889 | 214.2869 | 50.3216 | 64952.5909 | 67097.5563 | 1337.9511 | 1237.9511 | 738.6921
510 | 95.5934 | 1.0787 | 1.1377 | 1.0003 | 193.3691 | 54.1185 | 67366.0583 | 68176.9405 | 1359.5388 | 1259.5388 | 739.1098
511 | 110.9843 | 0.9582 | 0.83963 | 1.0382 | 202.9121 | 45.7412 | 44940.1727 | 45242.9801 | 900.8596 | 800.8596 | 739.4469
512 | 104.0387 | 1.0031 | 1.4777 | 1.1083 | 191.7244 | 50.8357 | 88755.652 | 88166.8994 | 1759.338 | 1659.338 | 739.5902
513 | 106.6297 | 0.91191 | 1.0637 | 1.034 | 213.4408 | 54.4849 | 65441.9151 | 66983.6923 | 1335.6738 | 1235.6738 | 739.6638
514 | 111.7307 | 1.0563 | 1.0059 | 0.94828 | 205.6925 | 53.2402 | 57273.7618 | 56356.1875 | 1123.1238 | 1023.1238 | 739.7004
515 | 93.3866 | 0.97157 | 0.9609 | 1.026 | 171.6278 | 52.2471 | 56421.2339 | 57717.215 | 1150.3443 | 1050.3443 | 739.7385
516 | 101.9202 | 1.0413 | 1.2121 | 0.89132 | 208.1568 | 47.6639 | 56759.9583 | 57811.9818 | 1152.2396 | 1052.2396 | 739.7539
517 | 107.7059 | 1.0384 | 1.1655 | 0.98264 | 199.6432 | 49.2301 | 62085.5579 | 61998.8423 | 1235.9768 | 1135.9768 | 739.7865
518 | 102.3203 | 0.98973 | 1.2975 | 0.93748 | 179.5716 | 48.2224 | 63720.6707 | 63001.408 | 1256.0282 | 1156.0282 | 739.8412
519 | 108.1084 | 1.0299 | 1.2059 | 0.95526 | 187.9047 | 47.9471 | 60757.3619 | 60998.2278 | 1215.9646 | 1115.9646 | 739.9589
520 | 95.7802 | 0.9236 | 1.0817 | 1.0666 | 217.8899 | 53.5919 | 66789.2293 | 67084.4535 | 1337.6891 | 1237.6891 | 739.9703
521 | 103.7208 | 1.0941 | 1.0565 | 1.0223 | 191.6207 | 50.358 | 60294.3622 | 59937.7514 | 1194.755 | 1094.755 | 740.1732
522 | 102.2733 | 0.93764 | 1.0784 | 1.0302 | 192.2931 | 48.4705 | 58691.7975 | 56285.0301 | 1121.7006 | 1021.7006 | 740.2934
523 | 94.5598 | 0.96917 | 1.1491 | 0.97126 | 198.4632 | 49.7002 | 60223.5417 | 60825.3439 | 1212.5069 | 1112.5069 | 740.3176
524 | 96.7839 | 1.0925 | 1.1605 | 0.92373 | 197.8462 | 47.3689 | 55983.2889 | 56944.3523 | 1134.887 | 1034.887 | 740.4826
525 | 111.1955 | 0.99607 | 1.2241 | 0.95325 | 195.0735 | 52.9778 | 67882.7011 | 63187.8414 | 1259.7568 | 1159.7568 | 740.5104
526 | 78.4252 | 0.98636 | 0.96039 | 0.94472 | 181.23 | 45.6097 | 45091.3256 | 43916.9982 | 874.34 | 774.34 | 740.5702
527 | 105.5795 | 0.99664 | 1.1316 | 0.99722 | 197.7394 | 46.621 | 57712.5008 | 59090.1739 | 1177.8035 | 1077.8035 | 740.5986
528 | 108.8553 | 1.0007 | 1.0234 | 0.9302 | 197.4389 | 50.9659 | 54265.7969 | 54808.0494 | 1092.161 | 992.161 | 740.7916
529 | 113.9479 | 1.0454 | 1.0357 | 1.104 | 216.8974 | 48.7279 | 61741.8917 | 63781.2825 | 1271.6257 | 1171.6257 | 740.8996
530 | 106.3504 | 1.0594 | 1.0809 | 0.96917 | 187.1058 | 52.9154 | 61583.1305 | 61832.668 | 1232.6534 | 1132.6534 | 740.908
531 | 91.5898 | 0.99163 | 1.2684 | 1.0135 | 198.8635 | 51.9551 | 71702.8007 | 72066.6267 | 1437.3325 | 1337.3325 | 741.2129
532 | 93.0225 | 0.97494 | 0.9434 | 1.0183 | 189.3829 | 53.6168 | 56560.6585 | 57742.318 | 1150.8464 | 1050.8464 | 741.4961
533 | 84.9775 | 1.0975 | 0.92361 | 1.0722 | 194.9101 | 54.2365 | 58963.6205 | 56938.885 | 1134.7777 | 1034.7777 | 741.7447
534 | 95.0556 | 0.95556 | 1.1178 | 1.0354 | 189.5569 | 44.9927 | 56348.4021 | 56043.7621 | 1116.8752 | 1016.8752 | 742.0729
535 | 95.0897 | 0.98184 | 1.1055 | 0.9775 | 199.0867 | 52.5501 | 61889.9433 | 62392.2801 | 1243.8456 | 1143.8456 | 742.5162
536 | 103.6796 | 1.023 | 1.1064 | 1.0342 | 190.2525 | 47.8568 | 60026.042 | 59370.6851 | 1183.4137 | 1083.4137 | 742.5698
537 | 101.1224 | 0.97148 | 0.81671 | 0.9625 | 197.4851 | 53.9988 | 47949.7593 | 49021.5399 | 976.4308 | 876.4308 | 742.6622
538 | 90.6144 | 1.0655 | 1.1754 | 0.93252 | 192.4427 | 46.2077 | 55301.0658 | 54851.6581 | 1093.0332 | 993.0332 | 742.9162
539 | 96.9756 | 1.018 | 0.79738 | 0.91649 | 215.6477 | 50.4731 | 42083.4711 | 42360.9611 | 843.2192 | 743.2192 | 743.0351
540 | 92.2371 | 1.0028 | 1.1759 | 1.0323 | 190.7162 | 51.2337 | 67123.8964 | 66038.0478 | 1316.761 | 1216.761 | 743.2192
541 | 87.9461 | 1.0628 | 1.0199 | 0.96842 | 203.2847 | 47.865 | 51954.7648 | 52126.199 | 1038.524 | 938.524 | 743.2956
542 | 107.5876 | 1.0112 | 0.98497 | 0.9868 | 220.3973 | 51.129 | 55479.1107 | 55569.592 | 1107.3918 | 1007.3918 | 743.3556
543 | 69.5379 | 0.97596 | 0.95504 | 1.0002 | 216.2104 | 49.3252 | 50681.3753 | 49766.1819 | 991.3236 | 891.3236 | 743.4377
544 | 101.8437 | 1.0485 | 0.92851 | 1.1038 | 208.8678 | 49.0722 | 55743.068 | 56894.1997 | 1133.884 | 1033.884 | 743.4675
545 | 97.3415 | 0.91821 | 0.82385 | 1.0507 | 182.1572 | 49.9114 | 47847.8785 | 47706.0117 | 950.1202 | 850.1202 | 743.5821
546 | 117.3251 | 1.014 | 1.1227 | 0.93155 | 206.3097 | 48.3588 | 56534.8032 | 56958.9573 | 1135.1791 | 1035.1791 | 743.6382
547 | 79.9277 | 0.90751 | 1.0811 | 1.0636 | 210.8358 | 51.1644 | 62752.1291 | 61988.2009 | 1235.764 | 1135.764 | 743.6788
548 | 103.8119 | 0.92049 | 1.0367 | 0.95471 | 198.9679 | 48.6892 | 53042.6508 | 53070.9199 | 1057.4184 | 957.4184 | 743.7161
549 | 90.5888 | 0.95916 | 1.0341 | 0.93082 | 200.3711 | 52.0038 | 54774.7083 | 55075.6855 | 1097.5137 | 997.5137 | 743.7633
550 | 106.7102 | 1.0435 | 1.0858 | 1.027 | 194.0418 | 48.6274 | 59835.4066 | 59137.8497 | 1178.757 | 1078.757 | 743.9691
551 | 83.6294 | 0.91961 | 1.0258 | 0.8901 | 199.9616 | 47.4035 | 47129.3695 | 47584.5533 | 947.6911 | 847.6911 | 744.0187
552 | 95.7172 | 1.0064 | 1.1401 | 0.98487 | 203.3391 | 48.3007 | 59088.7639 | 58120.4824 | 1158.4096 | 1058.4096 | 744.0717
553 | 115.2624 | 0.97213 | 0.71042 | 0.98082 | 205.1982 | 48.9189 | 39772.6431 | 39386.8729 | 783.7375 | 683.7375 | 744.2635
554 | 104.529 | 0.95968 | 1.0218 | 0.92602 | 203.8678 | 50.9699 | 53546.6471 | 52830.1314 | 1052.6026 | 952.6026 | 744.9017
555 | 102.7349 | 1.0303 | 1.1237 | 1.0219 | 215.5665 | 47.1095 | 59298.8381 | 60904.9381 | 1214.0988 | 1114.0988 | 745.0659
556 | 94.4399 | 1.0484 | 0.82367 | 0.97855 | 211.2434 | 51.5691 | 46882.0369 | 47987.1611 | 955.7432 | 855.7432 | 745.1263
557 | 108.1021 | 0.98906 | 0.99532 | 1.006 | 200.0096 | 46.6594 | 51910.2039 | 51531.7512 | 1026.635 | 926.635 | 745.236
558 | 106.9413 | 1.0703 | 1.0001 | 0.92698 | 200.7787 | 48.6194 | 50838.2139 | 51700.9396 | 1030.0188 | 930.0188 | 745.4341
559 | 90.8019 | 1.0323 | 0.97161 | 1.0205 | 182.4226 | 53.4687 | 58208.8511 | 58156.6512 | 1159.133 | 1059.133 | 745.6772
560 | 76.8583 | 1.0802 | 1.0464 | 1.0128 | 180.1432 | 48.4526 | 55554.5537 | 56415.5257 | 1124.3105 | 1024.3105 | 745.6983
561 | 113.3282 | 0.94311 | 0.93546 | 1.0457 | 197.7164 | 52.5152 | 57179.7599 | 56809.5526 | 1132.1911 | 1032.1911 | 745.9012
562 | 91.4426 | 1.0652 | 1.1869 | 1.01 | 208.1128 | 47.7108 | 62049.0024 | 61904.0412 | 1234.0808 | 1134.0808 | 746.0427
563 | 93.2138 | 1.0019 | 0.81882 | 0.93095 | 199.9811 | 49.1392 | 42246.7739 | 42214.6684 | 840.2934 | 740.2934 | 746.149
564 | 91.6675 | 1.0687 | 1.0327 | 1.0243 | 212.3737 | 47.9674 | 55647.8444 | 55600.2693 | 1108.0054 | 1008.0054 | 746.2127
565 | 101.4271 | 1.0169 | 1.0786 | 0.97263 | 209.0664 | 47.7082 | 55178.5794 | 54266.7753 | 1081.3355 | 981.3355 | 746.5662
566 | 96.5217 | 1.0353 | 1.0358 | 1.0303 | 221.9086 | 50.7363 | 59435.2494 | 58470.0092 | 1165.4002 | 1065.4002 | 746.5667
567 | 86.1866 | 1.0741 | 1.0413 | 0.93372 | 204.6621 | 53.5059 | 57178.4874 | 55243.515 | 1100.8703 | 1000.8703 | 746.697
568 | 113.4598 | 1.0129 | 0.9798 | 1.0042 | 185.2723 | 50.921 | 56140.7803 | 56787.7887 | 1131.7558 | 1031.7558 | 747.0042
569 | 100.434 | 0.97775 | 0.96853 | 0.91899 | 211.0016 | 48.2711 | 47915.6416 | 47703.7626 | 950.0753 | 850.0753 | 747.1578
570 | 91.3704 | 1.048 | 0.90736 | 1.0593 | 200.4534 | 47.5131 | 50417.6399 | 52210.5658 | 1040.2113 | 940.2113 | 747.1969
571 | 112.3485 | 0.9631 | 0.91558 | 1.0396 | 196.1467 | 50.7488 | 53989.872 | 54767.7406 | 1091.3548 | 991.3548 | 747.2882
572 | 81.6607 | 1.0103 | 0.9758 | 1.0023 | 209.0497 | 49.4288 | 52628.0561 | 51766.7723 | 1031.3354 | 931.3354 | 747.3951
573 | 105.4512 | 1.0146 | 0.95852 | 1.0084 | 196.9473 | 52.3671 | 56415.8468 | 56499.7308 | 1125.9946 | 1025.9946 | 747.5221
574 | 115.4851 | 0.93147 | 0.99798 | 1.043 | 183.6463 | 50.9187 | 58664.1813 | 59208.1434 | 1180.1629 | 1080.1629 | 747.8802
575 | 89.0774 | 0.94091 | 0.78896 | 1.0648 | 195.3309 | 51.1783 | 47477.7826 | 46591.5224 | 927.8304 | 827.8304 | 748.0164
576 | 95.4604 | 1.0218 | 0.72619 | 1.0394 | 198.537 | 48.3949 | 41446.3766 | 42328.4888 | 842.5698 | 742.5698 | 748.1381
577 | 128.4452 | 0.8924 | 1.0901 | 1.0073 | 197.0744 | 47.3425 | 57612.9305 | 58259.2569 | 1161.1851 | 1061.1851 | 748.3658
578 | 103.7948 | 0.94793 | 0.92256 | 0.96568 | 209.4912 | 43.6072 | 43349.6929 | 43418.7653 | 864.3753 | 764.3753 | 748.527
579 | 91.4359 | 0.99758 | 0.94555 | 0.96319 | 205.0381 | 46.2065 | 46501.9756 | 46488.6711 | 925.7734 | 825.7734 | 748.6021
580 | 81.3205 | 1.0262 | 1.0991 | 1.0028 | 186.3878 | 46.5861 | 55419.7619 | 54523.5734 | 1086.4715 | 986.4715 | 748.7628
581 | 112.5738 | 0.93018 | 1.3684 | 1.0223 | 202.1015 | 50.5229 | 76166.5823 | 77585.0865 | 1547.7017 | 1447.7017 | 748.8534
582 | 98.1937 | 1.0597 | 1.0902 | 1.0008 | 194.7963 | 50.4191 | 60449.4548 | 59670.8104 | 1189.4162 | 1089.4162 | 749.2515
583 | 114.9738 | 1.0129 | 1.049 | 0.93111 | 211.9726 | 49.724 | 54568.9405 | 53739.4889 | 1070.7898 | 970.7898 | 749.3244
584 | 88.2794 | 1.0065 | 0.80802 | 1.0233 | 233.4637 | 49.9295 | 45954.181 | 45328.9613 | 902.5792 | 802.5792 | 749.4605
585 | 113.5075 | 0.90161 | 1.011 | 0.94676 | 189.8054 | 48.5595 | 51637.7507 | 52480.904 | 1045.6181 | 945.6181 | 749.4939
586 | 105.9011 | 1.0831 | 0.86911 | 1.0945 | 202.1826 | 50.1944 | 53707.8986 | 53652.5837 | 1069.0517 | 969.0517 | 749.6181
587 | 99.7123 | 0.91712 | 0.97876 | 1.0116 | 210.2774 | 47.6696 | 51769.6911 | 51322.6422 | 1022.4528 | 922.4528 | 749.6219
588 | 105.4704 | 1.0062 | 0.91619 | 0.98772 | 174.8501 | 50.3294 | 51061.3261 | 50849.0717 | 1012.9814 | 912.9814 | 749.8103
589 | 112.6558 | 1.0097 | 0.9375 | 1.0431 | 206.5069 | 49.5542 | 54301.2086 | 54457.5141 | 1085.1503 | 985.1503 | 749.9095
590 | 107.3048 | 1.0501 | 1.1348 | 0.94905 | 187.9478 | 47.7016 | 56938.465 | 56035.6305 | 1116.7126 | 1016.7126 | 749.9738
591 | 107.195 | 1.0471 | 0.88834 | 1.0129 | 208.253 | 47.8572 | 48641.6163 | 49157.9984 | 979.16 | 879.16 | 749.9872
592 | 79.3251 | 0.99861 | 0.97237 | 1.0515 | 189.1612 | 47.3584 | 52361.3891 | 52240.945 | 1040.8189 | 940.8189 | 750.0565
593 | 103.7066 | 1.0138 | 1.0098 | 1.012 | 201.6151 | 51.0701 | 57762.0807 | 59819.4184 | 1192.3884 | 1092.3884 | 750.0769
594 | 99.7923 | 0.95627 | 1.1901 | 0.96769 | 211.771 | 56.068 | 70131.1404 | 70124.541 | 1398.4908 | 1298.4908 | 750.4366
595 | 90.4417 | 0.93591 | 1.1685 | 0.91983 | 217.9878 | 48.8854 | 56900.1019 | 55024.8659 | 1096.4973 | 996.4973 | 750.5931
596 | 120.2539 | 1.0031 | 1.1205 | 0.96318 | 193.302 | 52.4436 | 63117.1202 | 66512.9163 | 1326.2583 | 1226.2583 | 750.6821
597 | 104.5703 | 0.89922 | 0.863 | 0.9663 | 204.2473 | 50.2269 | 46812.3746 | 46750.9601 | 931.0192 | 831.0192 | 750.7422
598 | 85.0517 | 1.0476 | 0.90721 | 0.87112 | 219.1346 | 52.8083 | 46658.2148 | 47007.5266 | 936.1505 | 836.1505 | 750.8122
599 | 100.6285 | 1.0218 | 1.0945 | 0.99093 | 195.9935 | 54.0046 | 64323.1424 | 64325.0041 | 1282.5001 | 1182.5001 | 750.8411
600 | 95.5393 | 1.0142 | 0.71109 | 1.0464 | 193.0338 | 48.144 | 40682.0609 | 40717.1836 | 810.3437 | 710.3437 | 750.9992
601 | 94.661 | 0.93166 | 0.80426 | 0.94827 | 197.7425 | 51.0593 | 43641.3814 | 43407.9525 | 864.1591 | 764.1591 | 751.048
602 | 102.0619 | 0.92516 | 1.0436 | 1.0693 | 204.035 | 48.6701 | 59108.8062 | 58015.5987 | 1156.312 | 1056.312 | 751.2859
603 | 82.6895 | 0.93281 | 1.163 | 1.0773 | 204.4481 | 47.2145 | 63000.8359 | 64377.5856 | 1283.5517 | 1183.5517 | 751.3141
604 | 95.2471 | 1.0594 | 0.86315 | 1.0524 | 211.558 | 48.4784 | 49141.5794 | 48508.6654 | 966.1733 | 866.1733 | 751.413
605 | 92.3495 | 1.098 | 1.0123 | 1.0666 | 193.9918 | 51.5616 | 61093.1895 | 62277.2666 | 1241.5453 | 1141.5453 | 751.4867
606 | 78.4802 | 0.94886 | 0.89322 | 1.0264 | 188.1112 | 50.5704 | 50315.5616 | 50155.7758 | 999.1155 | 899.1155 | 751.494
607 | 95.5303 | 1.0016 | 1.2553 | 0.95012 | 206.7511 | 50.1305 | 64794.0223 | 64556.5642 | 1287.1313 | 1187.1313 | 751.9919
608 | 80.5818 | 1.0205 | 1.0197 | 0.99184 | 188.3849 | 48.1397 | 52835.5635 | 52218.7679 | 1040.3754 | 940.3754 | 752.0098
609 | 105.9519 | 0.90589 | 0.79635 | 0.98464 | 193.5833 | 48.7289 | 43079.8054 | 42737.1084 | 850.7422 | 750.7422 | 752.0722
610 | 107.6136 | 0.93985 | 1.2178 | 0.98185 | 199.6804 | 54.6969 | 71131.9951 | 72565.0157 | 1447.3003 | 1347.3003 | 752.1223
611 | 96.7358 | 1.0041 | 0.92682 | 1.0179 | 194.5082 | 52.6506 | 54979.9472 | 56130.7237 | 1118.6145 | 1018.6145 | 752.1564
612 | 96.6927 | 1.0348 | 1.0229 | 0.99565 | 198.0691 | 50.0203 | 56146.9568 | 57292.3513 | 1141.847 | 1041.847 | 752.5787
613 | 104.0191 | 0.93567 | 0.79713 | 1.0361 | 197.7777 | 51.2549 | 47518.9038 | 47174.0664 | 939.4813 | 839.4813 | 752.6125
614 | 116.76 | 1.0686 | 1.1404 | 1.0222 | 204.7196 | 48.3397 | 62584.5291 | 62124.2308 | 1238.4846 | 1138.4846 | 752.6574
615 | 114.2774 | 1.013 | 1.0124 | 0.96217 | 225.5473 | 47.1933 | 51659.8324 | 50379.3781 | 1003.5876 | 903.5876 | 752.8712
616 | 111.3228 | 0.89806 | 1.0707 | 0.9316 | 199.1286 | 50.9369 | 56098.8074 | 55301.9108 | 1102.0382 | 1002.0382 | 752.9331
617 | 106.9441 | 1.0176 | 1.0249 | 0.96954 | 194.2924 | 47.9892 | 53104.2058 | 53818.1186 | 1072.3624 | 972.3624 | 752.9855
618 | 93.2407 | 1.0363 | 0.72513 | 1.0593 | 182.7608 | 51.8345 | 45007.4874 | 45759.9369 | 911.1987 | 811.1987 | 753.0521
619 | 92.9845 | 1.0752 | 0.81036 | 1.0142 | 203.0025 | 52.0073 | 48144.4826 | 49236.8008 | 980.736 | 880.736 | 753.198
620 | 107.6782 | 1.0148 | 1.1528 | 0.967 | 205.9035 | 49.5694 | 60879.4185 | 60909.3506 | 1214.187 | 1114.187 | 753.7402
621 | 108.7421 | 1.0429 | 1.0446 | 1.0425 | 195.7906 | 55.2854 | 66671.5246 | 67770.8807 | 1351.4176 | 1251.4176 | 753.7779
622 | 105.3932 | 0.94851 | 1.1636 | 1.0445 | 218.1465 | 49.7171 | 65615.5843 | 65459.9901 | 1305.1998 | 1205.1998 | 754.268
623 | 94.6032 | 1.0424 | 1.0708 | 1.0487 | 197.2204 | 44.9907 | 55159.2583 | 54716.7566 | 1090.3351 | 990.3351 | 754.5966
624 | 99.8401 | 1.0979 | 1.256 | 1.0358 | 206.0022 | 52.0029 | 73565.5433 | 73048.2469 | 1456.9649 | 1356.9649 | 754.6013
625 | 97.6492 | 0.9966 | 0.98713 | 0.92665 | 207.8597 | 49.1312 | 49930.3446 | 51039.2606 | 1016.7852 | 916.7852 | 754.8114
626 | 117.147 | 0.88521 | 1.1482 | 0.97505 | 211.8885 | 46.6936 | 57328.6526 | 57491.7272 | 1145.8345 | 1045.8345 | 755.0785
627 | 80.4341 | 1.0704 | 0.69862 | 0.89981 | 196.4629 | 48.1595 | 34617.1373 | 36443.461 | 724.8692 | 624.8692 | 755.2655
628 | 97.4606 | 0.98908 | 1.1438 | 1.018 | 188.1362 | 49.6188 | 62747.5207 | 63553.7002 | 1267.074 | 1167.074 | 755.4123
629 | 84.1471 | 0.90536 | 1.0177 | 0.93121 | 217.9744 | 47.943 | 49306.5515 | 49996.9608 | 995.9392 | 895.9392 | 755.4575
630 | 111.7954 | 0.99636 | 1.115 | 0.85176 | 194.2315 | 47.4616 | 50556.0324 | 49223.3964 | 980.4679 | 880.4679 | 755.5894
631 | 95.9139 | 0.97932 | 0.97194 | 1.0421 | 204.021 | 49.8213 | 55344.383 | 56654.3178 | 1129.0864 | 1029.0864 | 755.6283
632 | 105.7666 | 1.0797 | 1.0099 | 1.0447 | 202.4987 | 47.5738 | 55825.9399 | 55168.5995 | 1099.372 | 999.372 | 755.6507
633 | 101.0932 | 1.0108 | 1.0998 | 0.98865 | 197.4072 | 46.4709 | 55472.5935 | 54992.3752 | 1095.8475 | 995.8475 | 755.8272
634 | 101.1671 | 1.033 | 1.0237 | 0.99501 | 190.0607 | 49.7979 | 56115.6431 | 55859.7643 | 1113.1953 | 1013.1953 | 755.9669
635 | 104.6008 | 1.0129 | 0.96764 | 1.0643 | 201.5809 | 49.671 | 56618.0953 | 54805.1967 | 1092.1039 | 992.1039 | 756.0742
636 | 80.5589 | 1.1157 | 0.97661 | 1.0425 | 207.442 | 45.5389 | 50662.8325 | 51086.1431 | 1017.7229 | 917.7229 | 756.1985
637 | 110.973 | 0.99726 | 1.1218 | 0.98192 | 204.9651 | 47.0872 | 57281.6013 | 55897.3436 | 1113.9469 | 1013.9469 | 756.3821
638 | 109.7063 | 1.0454 | 1.1292 | 1.0214 | 221.2763 | 50.7665 | 64595.8706 | 65580.2864 | 1307.6057 | 1207.6057 | 756.3979
639 | 95.8085 | 0.98587 | 1.1577 | 0.99088 | 202.8183 | 52.6501 | 65571.1229 | 69272.8869 | 1381.4577 | 1281.4577 | 756.4055
640 | 100.4347 | 0.93055 | 1.3561 | 1.0146 | 201.521 | 48.0425 | 70790.966 | 68342.7625 | 1362.8553 | 1262.8553 | 756.4639
641 | 102.2556 | 0.99505 | 1.2175 | 1.0031 | 196.8602 | 48.7328 | 64671.5053 | 62543.1072 | 1246.8621 | 1146.8621 | 756.4708
642 | 103.4542 | 1.0572 | 1.1716 | 0.9477 | 180.4562 | 53.6664 | 65634.7709 | 68799.2761 | 1371.9855 | 1271.9855 | 756.5112
643 | 85.0226 | 1.0705 | 1.1464 | 1.0139 | 212.4582 | 49.7771 | 62599.9279 | 62126.8428 | 1238.5369 | 1138.5369 | 756.6039
644 | 117.2586 | 0.9558 | 1.0808 | 1.056 | 199.342 | 51.7879 | 65109.6804 | 62460.5943 | 1245.2119 | 1145.2119 | 756.617
645 | 97.9297 | 1.0029 | 0.96592 | 0.95521 | 197.0929 | 51.9626 | 53244.0163 | 53283.9905 | 1061.6798 | 961.6798 | 756.7028
646 | 115.0026 | 1.0339 | 0.91199 | 1.0033 | 194.6651 | 51.9911 | 53950.1739 | 54743.458 | 1090.8692 | 990.8692 | 756.7273
647 | 95.8216 | 0.98872 | 1.035 | 1.0224 | 198.9671 | 50.8219 | 58795.5142 | 61241.2605 | 1220.8252 | 1120.8252 | 756.8445
648 | 100.7881 | 1.0079 | 0.90891 | 0.94363 | 195.2408 | 50.2612 | 48408.7941 | 48618.3035 | 968.3661 | 868.3661 | 757.0375
649 | 96.7118 | 1.0771 | 1.0279 | 0.93728 | 191.5419 | 51.9008 | 55599.9551 | 56550.3844 | 1127.0077 | 1027.0077 | 757.1319
650 | 82.9558 | 0.97823 | 1.1154 | 0.98343 | 199.6598 | 51.4363 | 60795.88 | 60824.084 | 1212.4817 | 1112.4817 | 757.1434
651 | 111.9732 | 0.92109 | 0.88067 | 1.0374 | 194.3379 | 49.5445 | 50570.0587 | 50978.223 | 1015.5645 | 915.5645 | 757.1848
652 | 91.4285 | 0.92859 | 1.1138 | 0.87935 | 203.9808 | 48.8592 | 52204.5194 | 53179.5035 | 1059.5901 | 959.5901 | 757.3266
653 | 89.9687 | 0.99887 | 0.78541 | 1.0636 | 197.3974 | 47.8227 | 44444.1001 | 44037.0072 | 876.7401 | 776.7401 | 757.327
654 | 77.6021 | 0.9811 | 0.92914 | 0.92045 | 200.2814 | 48.7823 | 45634.4462 | 46153.0725 | 919.0615 | 819.0615 | 757.559
655 | 109.7116 | 0.95853 | 1.0312 | 1.0111 | 186.8845 | 51.6588 | 59482.3079 | 57052.9911 | 1137.0598 | 1037.0598 | 757.6378
656 | 90.6134 | 0.98596 | 1.097 | 0.92617 | 194.143 | 51.2419 | 56832.9864 | 56209.2172 | 1120.1843 | 1020.1843 | 757.6836
657 | 94.6408 | 1.0303 | 1.2997 | 1.0132 | 190.4612 | 50.3082 | 71343.9065 | 71461.7785 | 1425.2356 | 1325.2356 | 757.7588
658 | 89.1198 | 0.9767 | 0.91556 | 1.0853 | 203.787 | 55.6472 | 60342.3136 | 58788.9943 | 1171.7799 | 1071.7799 | 757.9956
659 | 108.2684 | 0.92827 | 0.80736 | 1.0038 | 205.6252 | 49.7288 | 45504.1605 | 45112.883 | 898.2577 | 798.2577 | 758.0982
660 | 88.5701 | 1.006 | 0.92413 | 0.9592 | 178.6326 | 52.9073 | 51791.3659 | 51869.4529 | 1033.3891 | 933.3891 | 758.5326
661 | 97.8383 | 0.99124 | 1.2055 | 0.94777 | 203.1116 | 50.2919 | 62539.4114 | 60538.795 | 1206.7759 | 1106.7759 | 758.625
662 | 94.0504 | 0.96281 | 1.0736 | 1.0071 | 200.6479 | 46.6466 | 54860.0857 | 53376.2005 | 1063.524 | 963.524 | 758.6901
663 | 100.7452 | 0.97088 | 0.82936 | 0.99788 | 189.5406 | 50.5918 | 47007.6591 | 48655.5706 | 969.1114 | 869.1114 | 758.7931
664 | 95.2207 | 0.98675 | 0.9586 | 0.97664 | 205.7813 | 48.7859 | 50463.1777 | 50046.0765 | 996.9215 | 896.9215 | 758.8563
665 | 97.6338 | 1.0305 | 0.83449 | 1.0133 | 206.7006 | 53.8073 | 51116.9667 | 50354.4061 | 1003.0881 | 903.0881 | 758.9394
666 | 113.6276 | 0.99597 | 1.123 | 0.97914 | 191.6756 | 51.4663 | 62604.9447 | 61691.8756 | 1229.8375 | 1129.8375 | 758.9397
667 | 97.5985 | 1.0331 | 0.96108 | 0.92532 | 198.2574 | 47.0594 | 46793.4606 | 46674.8483 | 929.497 | 829.497 | 759.149
668 | 107.4546 | 1.1206 | 1.0912 | 1.0611 | 203.8115 | 51.7393 | 66340.1324 | 67972.335 | 1355.4467 | 1255.4467 | 759.339
669 | 102.4573 | 1.031 | 0.73731 | 0.99855 | 202.144 | 44.7662 | 37889.9305 | 38411.8632 | 764.2373 | 664.2373 | 759.3415
670 | 116.7748 | 1.0239 | 1.0292 | 1.0624 | 203.5123 | 52.8159 | 64272.2456 | 64562.3972 | 1287.2479 | 1187.2479 | 759.3442
671 | 113.6526 | 0.97677 | 0.90626 | 0.96272 | 190.3957 | 51.4585 | 50799.6288 | 50234.5625 | 1000.6912 | 900.6912 | 759.465
672 | 90.9825 | 0.99206 | 0.76833 | 1.0562 | 189.8157 | 48.6092 | 44024.0463 | 44017.7876 | 876.3558 | 776.3558 | 759.5333
673 | 107.2877 | 0.98319 | 0.85262 | 0.99782 | 211.1088 | 51.295 | 49262.0654 | 48668.5772 | 969.3715 | 869.3715 | 759.545
674 | 102.273 | 0.97672 | 0.96207 | 0.96782 | 203.6461 | 49.9693 | 51722.1799 | 51394.941 | 1023.8988 | 923.8988 | 759.7142
675 | 110.2606 | 1.0102 | 1.0194 | 1.0068 | 209.8212 | 47.4891 | 54239.3755 | 53755.3209 | 1071.1064 | 971.1064 | 759.7457
676 | 99.8747 | 0.96722 | 0.9616 | 1.0025 | 186.818 | 50.9556 | 54231.6228 | 54946.9862 | 1094.9397 | 994.9397 | 759.9216
677 | 88.3876 | 1.0253 | 0.92092 | 1.0629 | 212.5827 | 52.3335 | 56180.4314 | 56189.5591 | 1119.7912 | 1019.7912 | 759.9603
678 | 101.8241 | 1.007 | 1.0046 | 0.99865 | 182.9945 | 47.9752 | 53233.7158 | 53420.1945 | 1064.4039 | 964.4039 | 759.975
679 | 98.2687 | 1.0254 | 1.0792 | 1.0102 | 207.8293 | 49.6815 | 59377.0183 | 56547.1545 | 1126.9431 | 1026.9431 | 760.1004
680 | 109.2221 | 0.94925 | 0.99885 | 1.0262 | 203.5953 | 49.9717 | 56607.9495 | 54830.7088 | 1092.6142 | 992.6142 | 760.2747
681 | 92.589 | 0.99205 | 0.815 | 1.0141 | 189.3357 | 51.5327 | 47513.3842 | 49171.2257 | 979.4245 | 879.4245 | 760.3148
682 | 104.818 | 0.95321 | 1.0602 | 0.93167 | 209.312 | 51.2384 | 55941.3579 | 55745.4515 | 1110.909 | 1010.909 | 760.5743
683 | 99.8084 | 0.92271 | 1.1041 | 1.0349 | 194.7369 | 48.4409 | 60006.2135 | 60451.8822 | 1205.0376 | 1105.0376 | 760.6231
684 | 111.0018 | 0.98089 | 1.1952 | 0.99319 | 207.7228 | 47.3495 | 61567.8214 | 61278.7773 | 1221.5755 | 1121.5755 | 761.0096
685 | 93.7024 | 0.98286 | 1.1052 | 0.94312 | 211.9774 | 47.2533 | 53819.8107 | 52076.9593 | 1037.5392 | 937.5392 | 761.0573
686 | 97.1278 | 0.97599 | 1.1717 | 0.96379 | 204.0855 | 48.1358 | 59123.5077 | 60054.8678 | 1197.0974 | 1097.0974 | 761.2443
687 | 102.5088 | 0.96373 | 1.1753 | 1.1016 | 201.8197 | 52.9332 | 73967.8181 | 70811.1452 | 1412.2229 | 1312.2229 | 761.4081
688 | 108.1072 | 0.93144 | 0.88258 | 1.0285 | 194.6523 | 52.5923 | 53231.4592 | 50628.1911 | 1008.5638 | 908.5638 | 762.1323
689 | 129.4296 | 1.0667 | 1.0437 | 1.0374 | 195.5582 | 49.3127 | 60395.6134 | 61675.8287 | 1229.5166 | 1129.5166 | 762.2263
690 | 97.2261 | 0.90924 | 1.1055 | 1.0032 | 204.9156 | 45.0213 | 54116.0352 | 57710.4912 | 1150.2098 | 1050.2098 | 762.3146
691 | 100.1654 | 0.98164 | 0.87437 | 0.99979 | 203.4254 | 47.6271 | 46521.0445 | 44964.6267 | 895.2925 | 795.2925 | 762.4307
692 | 88.0064 | 1.017 | 0.84267 | 1.0474 | 215.5272 | 51.5518 | 50331.3447 | 48895.7303 | 973.9146 | 873.9146 | 762.5376
693 | 105.3554 | 0.92897 | 0.77069 | 1.0499 | 201.2324 | 53.7272 | 48933.7408 | 50456.0749 | 1005.1215 | 905.1215 | 762.5894
694 | 113.2487 | 1.0353 | 1.0286 | 0.983 | 210.4267 | 45.9376 | 52045.8426 | 51893.1185 | 1033.8624 | 933.8624 | 762.6728
695 | 86.2631 | 1.063 | 1.3135 | 0.97632 | 208.1165 | 45.8541 | 63214.5561 | 65160.9412 | 1299.2188 | 1199.2188 | 763.0118
696 | 98.3266 | 1.0531 | 0.95682 | 0.99427 | 186.6873 | 47.8549 | 50667.9783 | 49488.637 | 985.7727 | 885.7727 | 763.1389
697 | 96.5504 | 0.98711 | 0.9211 | 1.028 | 201.2853 | 48.1075 | 50338.1262 | 51087.1631 | 1017.7433 | 917.7433 | 763.4036
698 | 119.2436 | 1.0357 | 0.81411 | 0.91787 | 208.0905 | 54.882 | 47996.4262 | 47080.2803 | 937.6056 | 837.6056 | 763.5023
699 | 102.3619 | 0.93834 | 0.83798 | 1.0462 | 198.8244 | 52.2239 | 50999.3376 | 51687.5686 | 1029.7514 | 929.7514 | 763.6292
700 | 113.146 | 0.98495 | 1.004 | 1.054 | 192.1455 | 52.0171 | 61034.7166 | 60482.3731 | 1205.6475 | 1105.6475 | 763.6999
701 | 111.4139 | 1.0293 | 1.0605 | 1.0956 | 204.3761 | 51.5172 | 65973.8553 | 64546.9617 | 1286.9392 | 1186.9392 | 763.7184
702 | 104.6258 | 0.93369 | 0.90892 | 1.0141 | 205.8327 | 54.5653 | 55831.6321 | 55773.5226 | 1111.4705 | 1011.4705 | 763.8805
703 | 100.1644 | 0.90268 | 0.9404 | 0.86919 | 205.0855 | 51.3493 | 46820.351 | 48107.4848 | 958.1497 | 858.1497 | 764.1591
704 | 90.3589 | 0.9915 | 0.97718 | 1.023 | 204.5404 | 51.3892 | 56182.311 | 56424.1615 | 1124.4832 | 1024.4832 | 764.3753
705 | 96.2134 | 1.0358 | 0.97222 | 1.0283 | 208.9457 | 52.4798 | 57904.3447 | 58610.0913 | 1168.2018 | 1068.2018 | 764.3757
706 | 117.8576 | 0.96859 | 0.9599 | 1.0245 | 165.8052 | 55.8694 | 61488.9582 | 62893.7349 | 1253.8747 | 1153.8747 | 764.6269
707 | 100.7743 | 1.0661 | 1.1112 | 1.0177 | 191.8226 | 48.4482 | 60186.8582 | 58968.9655 | 1175.3793 | 1075.3793 | 764.7434
708 | 86.839 | 0.9656 | 0.80835 | 0.95235 | 188.5412 | 50.7094 | 43478.2081 | 44057.0483 | 877.141 | 777.141 | 765.3018
709 | 104.6145 | 0.99096 | 0.88183 | 0.97827 | 197.748 | 52.2054 | 50645.7169 | 49986.5616 | 995.7312 | 895.7312 | 765.3473
710 | 103.8943 | 0.96581 | 1.1274 | 0.86416 | 190.3815 | 47.5672 | 51305.5457 | 51462.982 | 1025.2596 | 925.2596 | 765.3618
711 | 100.6211 | 1.0296 | 1.1826 | 1.0181 | 197.8474 | 51.104 | 67025.5535 | 67812.0633 | 1352.2413 | 1252.2413 | 765.549
712 | 104.4906 | 0.97077 | 1.0023 | 1.0426 | 204.3157 | 50.6627 | 58285.8399 | 56784.5643 | 1131.6913 | 1031.6913 | 765.7828
713 | 101.8157 | 1.005 | 0.93599 | 1.0107 | 204.6531 | 49.2476 | 51831.8243 | 53403.1976 | 1064.064 | 964.064 | 765.9657
714 | 89.0701 | 1.0173 | 0.78432 | 1.0578 | 217.9696 | 43.9556 | 40668.148 | 41431.0849 | 824.6217 | 724.6217 | 766.0426
715 | 94.5484 | 1.0712 | 0.90125 | 0.90184 | 201.3391 | 50.6099 | 46461.7072 | 46989.7699 | 935.7954 | 835.7954 | 766.1173
716 | 101.9813 | 1.0683 | 1.2016 | 0.92623 | 191.4645 | 54.9082 | 67284.9565 | 67526.5657 | 1346.5313 | 1246.5313 | 766.4707
717 | 117.1451 | 0.9938 | 0.76207 | 0.95844 | 176.1565 | 50.7057 | 43114.4751 | 43477.4475 | 865.549 | 765.549 | 766.5564
718 | 107.2205 | 1.0013 | 0.84057 | 1.0463 | 190.3963 | 51.5449 | 51055.0532 | 51234.2563 | 1020.6851 | 920.6851 | 766.7781
719 | 103.8067 | 0.95331 | 1.1256 | 0.95983 | 197.5924 | 53.6219 | 63437.2609 | 63776.9328 | 1271.5387 | 1171.5387 | 766.7827
720 | 96.065 | 1.0279 | 0.90988 | 1.0151 | 196.6409 | 47.7278 | 48992.7029 | 50912.418 | 1014.2484 | 914.2484 | 766.7915
721 | 95.3181 | 1.0118 | 0.96905 | 1.017 | 194.5247 | 50.085 | 54383.8447 | 55347.6648 | 1102.9533 | 1002.9533 | 766.793
722 | 105.2539 | 1.0475 | 1.1627 | 0.93128 | 196.9372 | 52.6601 | 63023.5038 | 62684.4579 | 1249.6892 | 1149.6892 | 766.8489
723 | 95.2029 | 1.0308 | 0.98215 | 1.0395 | 182.2906 | 45.8225 | 51460.1688 | 49851.0341 | 993.0207 | 893.0207 | 766.9831
724 | 97.1838 | 1.0121 | 0.75902 | 1.048 | 208.8423 | 50.2691 | 45140.2887 | 45555.1735 | 907.1035 | 807.1035 | 767.2021
725 | 105.4876 | 1.0183 | 0.87999 | 1.0084 | 198.5854 | 47.0215 | 46975.3642 | 48424.8213 | 964.4964 | 864.4964 | 767.2953
726 | 100.2068 | 0.95981 | 0.9755 | 0.90804 | 204.5699 | 45.4099 | 44795.5278 | 44854.5501 | 893.091 | 793.091 | 767.3609
727 | 90.0545 | 0.97541 | 1.1695 | 1.0954 | 202.1776 | 51.0515 | 70088.0551 | 68308.9741 | 1362.1795 | 1262.1795 | 767.4916
728 | 110.8374 | 0.89607 | 0.88564 | 0.98526 | 192.3171 | 47.5065 | 46364.2517 | 47513.6239 | 946.2725 | 846.2725 | 767.5993
729 | 104.3412 | 1.0102 | 0.68902 | 1.0428 | 199.1881 | 46.4905 | 38504.1024 | 37524.129 | 746.4826 | 646.4826 | 767.6308
730 | 99.3106 | 1.0083 | 1.0443 | 0.9509 | 191.8325 | 51.6107 | 56611.406 | 56378.6593 | 1123.5732 | 1023.5732 | 767.6945
731 | 97.2744 | 0.91883 | 1.3404 | 1.0878 | 172.4174 | 51.4472 | 79787.2652 | 77146.1348 | 1538.9227 | 1438.9227 | 767.8974
732 | 101.419 | 0.98002 | 0.9938 | 0.99531 | 209.0075 | 50.5948 | 55282.7611 | 55236.0633 | 1100.7213 | 1000.7213 | 768.1201
733 | 98.216 | 1.0358 | 1.0694 | 0.96016 | 202.242 | 51.6603 | 58500.5627 | 58628.5619 | 1168.5712 | 1068.5712 | 768.7451
734 | 101.0493 | 0.98431 | 0.90741 | 0.98881 | 207.7126 | 48.1052 | 48155.227 | 48905.8782 | 974.1176 | 874.1176 | 768.8022
735 | 107.4646 | 0.99392 | 0.98147 | 0.96008 | 203.3395 | 44.225 | 46600.2153 | 45567.0868 | 907.3417 | 807.3417 | 768.9392
736 | 97.7229 | 1.0333 | 0.91767 | 0.92916 | 203.2337 | 56.3732 | 53962.7499 | 52744.6118 | 1050.8922 | 950.8922 | 769.0495
737 | 100.1597 | 0.92749 | 0.86417 | 1.0304 | 193.5046 | 49.4397 | 48811.3405 | 49464.5971 | 985.2919 | 885.2919 | 769.1111
738 | 93.4228 | 1.0172 | 0.82509 | 0.97442 | 193.5298 | 48.6997 | 43975.3345 | 45231.4655 | 900.6293 | 800.6293 | 769.2201
739 | 85.2564 | 0.98886 | 0.91147 | 0.99342 | 202.6131 | 51.0082 | 50689.3492 | 51040.2777 | 1016.8056 | 916.8056 | 769.2607
740 | 91.4583 | 1.0038 | 1.1388 | 0.96617 | 194.0991 | 52.0545 | 62247.6331 | 60321.733 | 1202.4347 | 1102.4347 | 769.3004
741 | 103.1919 | 1.0166 | 0.99882 | 1.0513 | 207.0577 | 50.9897 | 59100.5643 | 58870.8406 | 1173.4168 | 1073.4168 | 769.332
742 | 117.2275 | 1.0144 | 0.77875 | 1.0457 | 186.7368 | 52.7293 | 49395.1328 | 49693.0203 | 989.8604 | 889.8604 | 769.3908
743 | 104.0628 | 1.0592 | 0.95859 | 1.0036 | 196.3039 | 53.8386 | 57927.6833 | 58462.5431 | 1165.2509 | 1065.2509 | 769.4078
744 | 84.8514 | 1.0459 | 0.79586 | 0.95836 | 206.2824 | 49.4041 | 42272.2343 | 42274.8028 | 841.4961 | 741.4961 | 769.5147
745 | 122.5341 | 0.97646 | 0.90474 | 0.99613 | 212.1187 | 48.5091 | 49734.7089 | 49751.4545 | 991.0291 | 891.0291 | 769.5448
746 | 105.4983 | 1.0236 | 1.2826 | 1.0238 | 207.3783 | 51.6404 | 73591.956 | 73700.7418 | 1470.0148 | 1370.0148 | 769.6506
747 | 83.8589 | 0.91292 | 0.78309 | 0.95938 | 202.6146 | 48.0607 | 39988.9407 | 40257.9218 | 801.1584 | 701.1584 | 769.6975
748 | 92.823 | 1.0862 | 1.1077 | 1.0254 | 202.9871 | 51.437 | 63812.1163 | 62829.0072 | 1252.5801 | 1152.5801 | 769.7306
749 | 101.6951 | 0.93657 | 0.95458 | 0.93632 | 200.023 | 52.1686 | 51797.1727 | 51187.1373 | 1019.7427 | 919.7427 | 769.738
750 | 104.095 | 0.95412 | 1.3052 | 0.98177 | 191.5512 | 48.4834 | 67135.8887 | 66215.3413 | 1320.3068 | 1220.3068 | 769.8326
751 | 99.2745 | 1.0239 | 1.0083 | 0.93931 | 195.8703 | 49.9607 | 52590.3856 | 53247.2863 | 1060.9457 | 960.9457 | 769.8785
752 | 102.4449 | 1.0741 | 1.2112 | 1.0329 | 197.781 | 55.3647 | 75557.2486 | 75553.0377 | 1507.0608 | 1407.0608 | 769.9255
753 | 90.6711 | 1.0104 | 1.0269 | 1.0171 | 185.0366 | 48.802 | 55631.7309 | 54454.863 | 1085.0973 | 985.0973 | 770.0902
754 | 83.8382 | 0.91505 | 1.0443 | 0.89214 | 201.4663 | 55.0472 | 55711.2614 | 54622.9908 | 1088.4598 | 988.4598 | 770.0909
755 | 110.5817 | 0.95878 | 0.90004 | 1.0202 | 207.0777 | 46.8592 | 48202.6165 | 48223.3293 | 960.4666 | 860.4666 | 770.1387
756 | 104.17 | 1.0701 | 1.0309 | 0.97982 | 198.0573 | 46.9962 | 52907.9057 | 51794.9563 | 1031.8991 | 931.8991 | 770.1413
757 | 95.5589 | 1.0139 | 1.0116 | 0.96472 | 212.5478 | 50.0836 | 53943.035 | 53734.0627 | 1070.6813 | 970.6813 | 770.3464
758 | 102.1095 | 1.1262 | 0.80111 | 0.95606 | 197.6902 | 50.6462 | 44812.5662 | 43747.3401 | 870.9468 | 770.9468 | 770.3494
759 | 100.6359 | 0.94341 | 1.1162 | 0.93357 | 191.3503 | 50.1262 | 57185.3911 | 58344.6904 | 1162.8938 | 1062.8938 | 770.3804
760 | 90.4636 | 0.98151 | 0.83305 | 1.058 | 196.7268 | 49.2677 | 47995.0568 | 48440.4059 | 964.8081 | 864.8081 | 770.6183
761 | 113.7791 | 0.93347 | 1.2079 | 0.90424 | 199.7631 | 51.5714 | 62006.2863 | 61459.8248 | 1225.1965 | 1125.1965 | 770.6803
762 | 99.7181 | 1.1027 | 1.0854 | 0.97018 | 212.7916 | 54.1707 | 63214.3585 | 63805.3077 | 1272.1062 | 1172.1062 | 770.7601
763 | 104.4228 | 1.0093 | 0.97994 | 1.0066 | 206.2103 | 48.8342 | 53521.544 | 53036.5934 | 1056.7319 | 956.7319 | 770.7818
764 | 84.1914 | 0.91974 | 1.2297 | 0.97518 | 213.4777 | 49.7873 | 63770.9452 | 63498.8708 | 1265.9774 | 1165.9774 | 770.8821
765 | 109.7357 | 1.0161 | 0.83452 | 0.99556 | 210.3029 | 55.9468 | 52929.4664 | 51245.4719 | 1020.9094 | 920.9094 | 770.9468
766 | 91.103 | 1.0193 | 0.90712 | 0.89573 | 215.4861 | 50.5966 | 46025.0507 | 46730.9053 | 930.6181 | 830.6181 | 771.0754
767 | 99.2815 | 0.96198 | 1.0506 | 0.92219 | 209.4253 | 51.1562 | 54656.5535 | 53698.9601 | 1069.9792 | 969.9792 | 771.3476
768 | 93.5356 | 0.96772 | 0.70316 | 0.99031 | 216.8507 | 54.0798 | 42770.2501 | 42413.1729 | 844.2635 | 744.2635 | 771.3762
769 | 98.1701 | 1.0183 | 0.76304 | 0.9174 | 202.7316 | 48.0706 | 38658.5308 | 37467.6465 | 745.3529 | 645.3529 | 771.3854
770 | 94.2057 | 0.99615 | 1.189 | 0.92546 | 206.6653 | 52.0179 | 62329.0615 | 62564.7029 | 1247.2941 | 1147.2941 | 771.41
771 | 105.9332 | 1.0278 | 0.96507 | 0.98661 | 198.4333 | 49.1627 | 52360.9878 | 51461.5902 | 1025.2318 | 925.2318 | 771.4138
772 | 108.9734 | 1.0507 | 0.90372 | 0.97345 | 189.2673 | 53.3847 | 53265.6645 | 53574.1367 | 1067.4827 | 967.4827 | 771.5736
773 | 89.2295 | 1.0231 | 1.0074 | 0.99863 | 195.5224 | 53.4793 | 58876.8812 | 56285.938 | 1121.7188 | 1021.7188 | 771.5789
774 | 97.3156 | 0.93896 | 0.80672 | 1.1107 | 206.0101 | 55.3539 | 54863.7811 | 54058.1036 | 1077.1621 | 977.1621 | 771.7398
775 | 122.5897 | 1.0697 | 0.82439 | 0.99464 | 207.0409 | 48.5481 | 46381.629 | 46941.2445 | 934.8249 | 834.8249 | 771.8765
776 | 94.9533 | 0.9856 | 0.84971 | 1.0216 | 188.4976 | 46.8636 | 45255.4955 | 46217.0737 | 920.3415 | 820.3415 | 771.9921
777 | 102.5939 | 0.95179 | 1.0465 | 0.91688 | 204.5908 | 51.6795 | 54838.48 | 54377.7752 | 1083.5555 | 983.5555 | 772.0715
778 | 83.8881 | 0.9897 | 1.0376 | 0.99053 | 208.8632 | 50.1575 | 55923.5603 | 54511.9312 | 1086.2386 | 986.2386 | 772.1311
779 | 91.3182 | 0.99023 | 0.97893 | 0.9898 | 200.9021 | 44.4738 | 47314.9292 | 47473.1156 | 945.4623 | 845.4623 | 772.1377
780 | 105.7804 | 1.1142 | 1.0122 | 1.0579 | 196.4809 | 50.1684 | 59826.4634 | 60096.8053 | 1197.9361 | 1097.9361 | 772.1964
781 | 114.1611 | 1.0624 | 0.74772 | 0.9243 | 206.0229 | 48.8916 | 39925.3802 | 39879.7623 | 793.5952 | 693.5952 | 772.5879
782 | 103.1049 | 1.0231 | 0.9115 | 0.96182 | 206.5567 | 45.0149 | 44419.433 | 44027.828 | 876.5566 | 776.5566 | 772.5958
783 | 103.3359 | 0.95525 | 1.2719 | 0.92701 | 193.2483 | 45.4768 | 58303.4461 | 57562.8639 | 1147.2573 | 1047.2573 | 772.8035
784 | 91.866 | 1.0353 | 1.1113 | 1.0256 | 203.7818 | 48.828 | 60498.1887 | 61272.6828 | 1221.4537 | 1121.4537 | 773.2956
785 | 100.6082 | 1.0089 | 0.98679 | 1.0418 | 195.081 | 48.0005 | 54412.261 | 54786.4196 | 1091.7284 | 991.7284 | 773.3295
786 | 91.9905 | 1.0393 | 0.94783 | 1.0368 | 200.346 | 51.2539 | 55466.0281 | 55929.5825 | 1114.5916 | 1014.5916 | 773.5629
787 | 120.4199 | 0.97454 | 0.90501 | 0.94328 | 195.8089 | 52.2513 | 50933.5267 | 51626.1098 | 1028.5222 | 928.5222 | 773.6485
788 | 99.8064 | 0.97695 | 1.2069 | 0.91295 | 191.6305 | 49.9315 | 60076.8249 | 58965.9461 | 1175.3189 | 1075.3189 | 773.6707
789 | 107.4644 | 0.94858 | 1.1463 | 1.0057 | 172.8303 | 49.7238 | 62563.7551 | 63826.9187 | 1272.5384 | 1172.5384 | 773.7585
790 | 99.2297 | 1.0291 | 1.0074 | 1.113 | 196.1716 | 47.6459 | 58481.0571 | 60286.1698 | 1201.7234 | 1101.7234 | 773.8251
791 | 86.993 | 0.99856 | 1.0164 | 0.91631 | 197.2265 | 49.4715 | 50568.8142 | 50672.6324 | 1009.4526 | 909.4526 | 774.0857
792 | 97.7342 | 0.9805 | 0.95517 | 0.96439 | 201.0954 | 51.965 | 53048.7479 | 51885.1456 | 1033.7029 | 933.7029 | 774.2434
793 | 105.6909 | 1.0435 | 1.0916 | 0.96451 | 199.6875 | 53.1112 | 61977.5669 | 63041.3298 | 1256.8266 | 1156.8266 | 774.2741
794 | 88.2034 | 1.0074 | 0.98799 | 1.004 | 200.398 | 50.472 | 54749.2232 | 55324.6306 | 1102.4926 | 1002.4926 | 774.2984
795 | 93.8804 | 0.95984 | 1.0725 | 0.91917 | 205.3688 | 57.0536 | 61589.5614 | 61360.3213 | 1223.2064 | 1123.2064 | 774.3005
796 | 129.3025 | 0.99485 | 0.87196 | 1.0614 | 191.0958 | 49.7805 | 52665.3597 | 53269.5621 | 1061.3912 | 961.3912 | 774.34
797 | 108.1099 | 1.0097 | 1.1514 | 1.0054 | 199.3962 | 48.115 | 61151.8655 | 62714.9181 | 1250.2984 | 1150.2984 | 774.5199
798 | 96.5847 | 1.0129 | 1.097 | 0.95454 | 195.6627 | 51.1803 | 58796.0487 | 57646.9899 | 1148.9398 | 1048.9398 | 774.6798
799 | 111.583 | 0.96178 | 0.81446 | 0.96941 | 208.1837 | 50.2267 | 45254.7123 | 45924.0317 | 914.4806 | 814.4806 | 775.0657
800 | 93.9504 | 1.0125 | 0.94313 | 1.0061 | 208.4514 | 45.7186 | 47940.6439 | 47557.0534 | 947.1411 | 847.1411 | 775.3415
801 | 95.3053 | 1.0695 | 1.0322 | 0.97596 | 203.6362 | 50.4726 | 56192.7486 | 53384.4038 | 1063.6881 | 963.6881 | 775.4086
802 | 112.9986 | 1.0483 | 0.90223 | 0.9904 | 203.9058 | 49.1651 | 49960.3346 | 48942.338 | 974.8468 | 874.8468 | 775.4452
803 | 87.8221 | 0.9495 | 0.92995 | 1.0261 | 197.5894 | 47.0153 | 48980.5542 | 49473.558 | 985.4712 | 885.4712 | 775.5569
804 | 98.2006 | 1.0355 | 0.92378 | 0.96697 | 218.572 | 52.3691 | 52323.4398 | 53014.7897 | 1056.2958 | 956.2958 | 775.6791
805 | 111.7029 | 1.0282 | 1.0892 | 0.97163 | 185.8909 | 52.1543 | 61372.2263 | 62238.3921 | 1240.7678 | 1140.7678 | 775.9182
806 | 115.0159 | 1.0814 | 1.2039 | 1.0706 | 217.0555 | 48.4924 | 68752.7633 | 70141.1938 | 1398.8239 | 1298.8239 | 776.0547
807 | 102.546 | 1.0215 | 0.93548 | 1.0032 | 183.1645 | 48.1694 | 50435.5321 | 49864.4618 | 993.2892 | 893.2892 | 776.1596
808 | 100.9362 | 1.0169 | 0.90854 | 1.0622 | 197.0154 | 49.8815 | 53456.7651 | 52158.9713 | 1039.1794 | 939.1794 | 776.2368
809 | 103.1211 | 0.94015 | 0.88069 | 1.0573 | 192.1415 | 52.3178 | 53979.9832 | 54453.4449 | 1085.0689 | 985.0689 | 776.268
810 | 100.562 | 0.99125 | 1.021 | 1.007 | 191.2907 | 49.6556 | 56194.5217 | 57235.5786 | 1140.7116 | 1040.7116 | 776.3402
811 | 95.9202 | 0.88081 | 0.98736 | 0.9645 | 213.1326 | 53.1213 | 55289.2137 | 57769.539 | 1151.3908 | 1051.3908 | 776.3558
812 | 87.9975 | 1.0015 | 0.82332 | 0.82493 | 205.9835 | 50.3377 | 38830.453 | 37948.4446 | 754.9689 | 654.9689 | 776.4845
813 | 96.7113 | 0.96468 | 0.97897 | 1.0238 | 204.0505 | 47.0148 | 51711.6467 | 49884.3787 | 993.6876 | 893.6876 | 776.536
814 | 101.1847 | 1.0537 | 1.0338 | 1.0591 | 200.1875 | 48.1214 | 58020.3372 | 60773.9097 | 1211.4782 | 1111.4782 | 776.5566
815 | 100.448 | 0.97414 | 1.1769 | 1.0425 | 199.3474 | 46.5134 | 61818.0944 | 59867.576 | 1193.3515 | 1093.3515 | 776.6217
816 | 92.923 | 1.0551 | 1.0525 | 1.0139 | 199.097 | 49.9855 | 58438.7561 | 58782.1876 | 1171.6438 | 1071.6438 | 776.6295
817 | 107.6817 | 1.0132 | 0.94838 | 0.91873 | 208.5213 | 45.7052 | 45017.7745 | 45920.1408 | 914.4028 | 814.4028 | 776.7401
818 | 88.5766 | 0.97225 | 0.85917 | 0.95955 | 215.0934 | 51.1007 | 46744.1896 | 48263.4142 | 961.2683 | 861.2683 | 777.0285
819 | 96.8764 | 1.0923 | 1.1893 | 1.0934 | 197.0443 | 44.942 | 63396.8662 | 63929.7998 | 1274.596 | 1174.596 | 777.141
820 | 109.1524 | 1.0557 | 1.1436 | 1.0199 | 199.4375 | 48.1055 | 61853.1176 | 64149.0011 | 1278.98 | 1178.98 | 777.3326
821 | 108.4681 | 1.0133 | 1.0266 | 0.91703 | 211.7936 | 50.2006 | 52990.1627 | 51827.3921 | 1032.5478 | 932.5478 | 777.4963
822 | 93.96 | 0.95788 | 0.81986 | 1.0893 | 196.5736 | 49.5995 | 48954.6645 | 48270.8158 | 961.4163 | 861.4163 | 777.7434
823 | 99.2228 | 0.95171 | 1.2659 | 1.032 | 179.5284 | 50.6345 | 71108.8226 | 70616.0716 | 1408.3214 | 1308.3214 | 777.8119
824 | 90.7893 | 0.97887 | 1.2563 | 1.0226 | 194.6126 | 46.6574 | 64282.8401 | 63666.1831 | 1269.3237 | 1169.3237 | 777.9086
825 | 81.8041 | 1.0334 | 1.2836 | 0.95389 | 179.5322 | 46.0905 | 60511.1501 | 60538.9331 | 1206.7787 | 1106.7787 | 777.9171
826 | 113.2049 | 1.065 | 1.1532 | 0.98352 | 204.3746 | 53.0849 | 66815.5049 | 68821.4165 | 1372.4283 | 1272.4283 | 778.0727
827 | 105.5598 | 0.97542 | 1.3094 | 0.92642 | 192.2876 | 47.548 | 62765.2266 | 62699.305 | 1249.9861 | 1149.9861 | 778.3291
828 | 90.1436 | 0.97654 | 1.0574 | 0.95855 | 199.3674 | 44.9917 | 49760.6448 | 48087.6334 | 957.7527 | 857.7527 | 778.4041
829 | 97.9169 | 1.1045 | 1.2623 | 0.99043 | 208.2527 | 48.052 | 65480.4607 | 63738.7944 | 1270.7759 | 1170.7759 | 778.4679
830 | 106.813 | 1.07 | 1.0163 | 1.107 | 201.0518 | 51.1226 | 63559.921 | 63458.4613 | 1265.1692 | 1165.1692 | 778.8894
831 | 98.2408 | 1.0058 | 0.79614 | 1.0842 | 200.8775 | 49.9159 | 48217.3975 | 48661.8825 | 969.2376 | 869.2376 | 778.9579
832 | 85.2565 | 1.0349 | 0.98517 | 0.95059 | 195.2196 | 51.5932 | 53064.09 | 51391.638 | 1023.8328 | 923.8328 | 779.0821
833 | 102.358 | 1.0164 | 0.81445 | 1.1212 | 192.0502 | 48.9302 | 49961.9506 | 49120.8691 | 978.4174 | 878.4174 | 779.2275
834 | 88.3494 | 1.039 | 0.97498 | 0.97125 | 180.8426 | 48.8427 | 50915.8946 | 50713.5209 | 1010.2704 | 910.2704 | 779.2563
835 | 109.5953 | 1.0386 | 1.0664 | 0.94946 | 200.2297 | 48.7642 | 55124.7843 | 54346.7339 | 1082.9347 | 982.9347 | 779.465
836 | 90.392 | 0.9957 | 1.0998 | 1.0941 | 198.7155 | 47.3325 | 61412.7884 | 61199.5539 | 1219.9911 | 1119.9911 | 779.5217
837 | 90.2304 | 1.0154 | 0.96051 | 0.99615 | 186.6797 | 53.9235 | 56721.8344 | 57787.0013 | 1151.74 | 1051.74 | 779.5341
838 | 103.1986 | 0.97203 | 1.0966 | 1.0028 | 210.9208 | 51.7475 | 62306.043 | 63950.9828 | 1275.0197 | 1175.0197 | 779.535
839 | 102.7185 | 0.93507 | 0.98793 | 0.96722 | 200.1635 | 52.8796 | 55807.9594 | 57328.1653 | 1142.5633 | 1042.5633 | 779.5412
840 | 120.2554 | 1.053 | 0.73663 | 0.97186 | 201.3861 | 54.42 | 46051.8268 | 46596.5563 | 927.9311 | 827.9311 | 779.5533
841 | 104.7013 | 0.9568 | 1.098 | 0.93952 | 208.2149 | 52.643 | 59786.4345 | 60103.0432 | 1198.0609 | 1098.0609 | 779.5597
842 | 104.9351 | 1.0875 | 1.1337 | 0.97193 | 184.6402 | 48.5305 | 59196.2357 | 57394.6957 | 1143.8939 | 1043.8939 | 779.5653
843 | 94.2939 | 1.0082 | 0.87968 | 0.99318 | 191.7559 | 46.5786 | 45314.8331 | 46399.8966 | 923.9979 | 823.9979 | 779.6828
844 | 94.246 | 1.091 | 0.98588 | 0.98639 | 203.7784 | 53.7445 | 57994.3735 | 58418.8024 | 1164.376 | 1064.376 | 779.6971
845 | 112.6083 | 1.0367 | 1.126 | 0.98567 | 186.9727 | 47.0283 | 57870.761 | 58454.3092 | 1165.0862 | 1065.0862 | 779.8994
846 | 102.3881 | 0.94217 | 1.0275 | 1.029 | 191.0381 | 47.8961 | 55450.3843 | 56653.8665 | 1129.0773 | 1029.0773 | 780.0028
847 | 91.4715 | 0.97604 | 0.99451 | 1.021 | 198.2631 | 46.2659 | 51308.071 | 50202.8746 | 1000.0575 | 900.0575 | 780.0432
848 | 108.6719 | 0.94711 | 1.2356 | 0.99899 | 179.1941 | 50.6509 | 67915.1705 | 71136.0162 | 1418.7203 | 1318.7203 | 780.0819
849 | 98.3768 | 1.0376 | 1.3153 | 0.96029 | 195.1965 | 47.86 | 65532.8475 | 67544.2429 | 1346.8849 | 1246.8849 | 780.1376
850 | 90.3602 | 1.0202 | 1.3266 | 1.0515 | 196.8522 | 50.3041 | 75002.6427 | 77419.4125 | 1544.3883 | 1444.3883 | 780.1382
851 | 92.1469 | 1.0449 | 1.2522 | 0.94295 | 210.3015 | 51.819 | 66386.2427 | 65301.1963 | 1302.0239 | 1202.0239 | 780.172
852 | 108.4169 | 0.98672 | 1.0611 | 0.97012 | 199.9229 | 49.7485 | 56733.5483 | 57852.9273 | 1153.0585 | 1053.0585 | 780.1836
853 | 93.2959 | 0.97399 | 0.98238 | 0.9721 | 220.4768 | 47.2005 | 49584.5005 | 49708.9308 | 990.1786 | 890.1786 | 780.3243
854 | 83.5621 | 0.95832 | 0.80853 | 1.009 | 225.4469 | 51.0432 | 45954.2175 | 46346.3562 | 922.9271 | 822.9271 | 780.482
855 | 101.1207 | 0.99903 | 0.90986 | 1.048 | 198.6899 | 53.8334 | 56970.9074 | 58639.7734 | 1168.7955 | 1068.7955 | 780.5091
856 | 104.9608 | 0.94023 | 1.1263 | 1.0437 | 193.5278 | 54.0085 | 69010.0797 | 68587.5776 | 1367.7516 | 1267.7516 | 780.6181
857 | 104.2368 | 0.94284 | 0.7389 | 0.90868 | 202.8446 | 47.0556 | 36421.5313 | 37422.7565 | 744.4551 | 644.4551 | 780.8586
858 | 109.8367 | 0.90201 | 0.87203 | 0.87867 | 198.346 | 48.7993 | 42424.289 | 42069.0322 | 837.3806 | 737.3806 | 780.9236
859 | 102.1409 | 1.0211 | 1.0543 | 0.96961 | 190.6726 | 52.0186 | 58791.0754 | 57768.3745 | 1151.3675 | 1051.3675 | 780.9331
860 | 102.5555 | 1.0203 | 1.3965 | 1.0224 | 186.3298 | 49.0752 | 75391.0923 | 76735.7259 | 1530.7145 | 1430.7145 | 780.9989
861 | 97.7216 | 1.0544 | 0.83867 | 0.99336 | 206.0264 | 49.9899 | 47003.1 | 46621.6296 | 928.4326 | 828.4326 | 781.0265
862 | 110.5616 | 0.9204 | 1.3304 | 0.90868 | 201.9031 | 49.1118 | 64570.7987 | 66891.8334 | 1333.8367 | 1233.8367 | 781.131
863 | 96.8551 | 1.0881 | 0.79502 | 0.95262 | 206.3801 | 50.8422 | 44070.3571 | 44715.6038 | 890.3121 | 790.3121 | 781.3273
864 | 98.6483 | 0.96556 | 1.1765 | 1.0147 | 203.9655 | 50.4368 | 65216.5569 | 65854.3034 | 1313.0861 | 1213.0861 | 781.3563
865 | 90.024 | 0.99417 | 0.96781 | 0.96221 | 181.2482 | 51.8712 | 53128.0609 | 53272.0345 | 1061.4407 | 961.4407 | 781.52
866 | 98.6917 | 0.98025 | 0.90574 | 1.0173 | 216.1659 | 52.794 | 53969.7357 | 53208.2188 | 1060.1644 | 960.1644 | 781.6303
867 | 110.4247 | 1.101 | 1.2395 | 0.97761 | 195.8855 | 44.4522 | 59464.4274 | 59558.4435 | 1187.1689 | 1087.1689 | 781.6823
868 | 103.7205 | 0.99669 | 1.1187 | 0.9712 | 202.7191 | 48.9453 | 58439.8013 | 59355.182 | 1183.1036 | 1083.1036 | 781.691
869 | 97.7839 | 1.0091 | 1.2464 | 1.0435 | 210.3822 | 53.5077 | 75080.3815 | 74654.1368 | 1489.0827 | 1389.0827 | 781.8358
870 | 102.2223 | 0.94151 | 0.94483 | 0.95248 | 200.1498 | 47.743 | 47760.0763 | 47201.551 | 940.031 | 840.031 | 782.0058
871 | 106.5897 | 1.0484 | 0.99424 | 0.97154 | 205.0879 | 45.5768 | 49322.897 | 48941.5658 | 974.8313 | 874.8313 | 782.0905
872 | 105.9571 | 1.0471 | 1.0303 | 0.96883 | 184.1811 | 48.1064 | 53540.2884 | 52868.083 | 1053.3617 | 953.3617 | 782.1243
873 | 86.99 | 0.96659 | 1.0892 | 0.98144 | 194.0614 | 48.5456 | 56169.9381 | 56264.8494 | 1121.297 | 1021.297 | 782.3122
874 | 112.003 | 1.0346 | 0.70337 | 1.0169 | 204.0858 | 49.7361 | 41542.9667 | 41746.9126 | 830.9383 | 730.9383 | 782.415
875 | 117.0686 | 1.0506 | 0.92611 | 1.0337 | 198.3051 | 48.6369 | 52741.8606 | 52942.3584 | 1054.8472 | 954.8472 | 782.5185
876 | 89.8766 | 0.94432 | 0.75734 | 1.1175 | 195.641 | 52.9663 | 49519.6643 | 49799.7233 | 991.9945 | 891.9945 | 782.5434
877 | 99.4298 | 1.0887 | 1.1383 | 0.97809 | 214.6332 | 49.4091 | 60572.7911 | 61327.3348 | 1222.5467 | 1122.5467 | 782.7056
878 | 99.7468 | 1.0073 | 1.038 | 1.1274 | 206.5045 | 53.5156 | 68211.5259 | 71889.886 | 1433.7977 | 1333.7977 | 782.7262
879 | 87.0531 | 0.98513 | 0.90239 | 0.9856 | 193.4959 | 51.8588 | 50763.8944 | 51139.0106 | 1018.7802 | 918.7802 | 782.7764
880 | 103.6593 | 1.0239 | 0.96382 | 1.0256 | 210.2135 | 49.5656 | 54464.853 | 53215.3284 | 1060.3066 | 960.3066 | 782.8692
881 | 84.019 | 0.94877 | 1.2528 | 0.97213 | 221.2295 | 46.3725 | 60395.867 | 60717.1638 | 1210.3433 | 1110.3433 | 782.8796
882 | 104.9118 | 1.0326 | 1.0735 | 0.96789 | 211.0919 | 49.1801 | 56637.0775 | 58414.1893 | 1164.2838 | 1064.2838 | 782.9881
883 | 107.7469 | 0.997 | 1.1084 | 0.97409 | 194.0387 | 48.7401 | 58053.4374 | 58927.9851 | 1174.5597 | 1074.5597 | 783.0514
884 | 107.7344 | 0.9982 | 1.0856 | 0.9756 | 204.5906 | 51.2561 | 60003.7744 | 58716.9376 | 1170.3388 | 1070.3388 | 783.1037
885 | 94.9429 | 0.96713 | 0.79512 | 1.0603 | 202.2219 | 44.6067 | 41905.4759 | 42953.9227 | 855.0785 | 755.0785 | 783.174
886 | 106.1118 | 0.98896 | 1.003 | 1.029 | 201.5543 | 46.7997 | 53415.1532 | 54395.3318 | 1083.9066 | 983.9066 | 783.4144
887 | 93.1535 | 0.98526 | 1.0123 | 0.99362 | 200.9284 | 50.6723 | 55820.2889 | 55180.9422 | 1099.6188 | 999.6188 | 783.4684
888 | 92.4994 | 1.0789 | 1.0164 | 0.97616 | 212.7381 | 53.1825 | 58288.4678 | 58176.0898 | 1159.5218 | 1059.5218 | 783.6097
889 | 102.9901 | 0.97823 | 0.82999 | 1.0677 | 201.6675 | 49.745 | 49295.2954 | 47857.7739 | 953.1555 | 853.1555 | 783.6389
890 | 95.9789 | 1.0206 | 0.88541 | 1.0369 | 201.9741 | 51.0007 | 52018.5609 | 52376.8406 | 1043.5368 | 943.5368 | 783.8804
891 | 101.5639 | 0.98392 | 1.204 | 1.0909 | 193.8858 | 54.0324 | 76565.1518 | 75804.1815 | 1512.0836 | 1412.0836 | 783.9598
892 | 110.6744 | 0.93717 | 1.1986 | 1.013 | 207.7094 | 52.9668 | 70017.1414 | 70673.5965 | 1409.4719 | 1309.4719 | 784.069
893 | 98.6753 | 1.06 | 0.84213 | 1.0672 | 186.4944 | 45.0905 | 45427.2804 | 44284.1134 | 881.6823 | 781.6823 | 784.2199
894 | 86.8797 | 0.9848 | 1.1974 | 1.0144 | 216.6188 | 49.6879 | 64817.5768 | 67402.6503 | 1344.053 | 1244.053 | 784.2889
895 | 108.8376 | 1.0618 | 0.80047 | 1.0616 | 199.8793 | 48.5144 | 47034.8799 | 47413.7166 | 944.2743 | 844.2743 | 784.4112
896 | 116.6687 | 1.0389 | 1.2478 | 0.98191 | 187.3136 | 50.5211 | 68211.9021 | 68904.0401 | 1374.0808 | 1274.0808 | 784.5353
897 | 92.2725 | 0.92257 | 0.93429 | 0.93536 | 191.0575 | 51.0077 | 49108.7271 | 49587.6269 | 987.7525 | 887.7525 | 784.6446
898 | 102.4224 | 1.0026 | 1.1759 | 0.97423 | 198.8825 | 52.0184 | 65132.8754 | 66895.443 | 1333.9089 | 1233.9089 | 784.6451
899 | 92.4213 | 1.0969 | 1.1713 | 1.0413 | 235.3513 | 49.5655 | 65714.3091 | 64895.3144 | 1293.9063 | 1193.9063 | 784.6791
900 | 97.8199 | 1.0279 | 1.1504 | 0.93123 | 191.2137 | 49.6649 | 58390.5567 | 56952.2243 | 1135.0445 | 1035.0445 | 784.8924
901 | 93.4009 | 0.99184 | 1.0638 | 0.99203 | 203.5641 | 46.7781 | 53901.2274 | 52224.3296 | 1040.4866 | 940.4866 | 785.1069
902 | 109.1674 | 1.0285 | 1.0874 | 0.98943 | 175.4218 | 47.9708 | 57171.4559 | 57179.9232 | 1139.5985 | 1039.5985 | 785.3473
903 | 111.5846 | 0.9934 | 1.2025 | 1.018 | 209.9472 | 46.0797 | 61728.1251 | 63723.7631 | 1270.4753 | 1170.4753 | 785.4051
904 | 90.9662 | 0.99231 | 1.1391 | 0.95476 | 193.4422 | 54.1364 | 63954.7917 | 61622.7752 | 1228.4555 | 1128.4555 | 785.4501
905 | 87.9262 | 0.92857 | 0.93579 | 1.0422 | 198.3392 | 57.4587 | 60926.9443 | 61671.3348 | 1229.4267 | 1129.4267 | 785.5558
906 | 84.4521 | 0.98843 | 1.0353 | 1.0465 | 185.7641 | 50.7571 | 59413.8172 | 57813.0941 | 1152.2619 | 1052.2619 | 785.6136
907 | 101.3871 | 0.90756 | 0.90318 | 1.0192 | 200.7505 | 48.247 | 49053.4757 | 48543.8209 | 966.8764 | 866.8764 | 785.6239
908 | 93.2977 | 1.019 | 1.0028 | 1.0077 | 198.3273 | 47.5189 | 52735.9174 | 52496.5966 | 1045.9319 | 945.9319 | 785.624
909 | 94.0637 | 0.99213 | 0.92171 | 0.98532 | 214.958 | 58.5567 | 58859.6075 | 56035.856 | 1116.7171 | 1016.7171 | 785.6542
910 | 105.6094 | 1.1053 | 0.88951 | 0.95071 | 199.5533 | 48.8817 | 47243.0014 | 47074.0883 | 937.4818 | 837.4818 | 785.9089
911 | 108.0615 | 0.9011 | 1.2258 | 1.0754 | 202.4409 | 49.6033 | 70420.6202 | 71402.0663 | 1424.0413 | 1324.0413 | 785.9502
912 | 94.8316 | 0.96697 | 0.86846 | 1.0917 | 201.6792 | 48.6254 | 50764.5753 | 48691.5767 | 969.8315 | 869.8315 | 786.1853
913 | 95.463 | 0.99261 | 0.87412 | 1.0482 | 177.2798 | 46.7444 | 47436.2957 | 47019.7674 | 936.3953 | 836.3953 | 786.3083
914 | 118.1926 | 0.97468 | 0.82672 | 0.98008 | 187.4301 | 48.3298 | 44914.2403 | 43977.8436 | 875.5569 | 775.5569 | 786.6316
915 | 90.5995 | 1.0009 | 0.90347 | 0.99116 | 187.8319 | 51.5194 | 50994.6173 | 49647.4002 | 988.948 | 888.948 | 786.6932
916 | 101.7979 | 0.97988 | 0.89493 | 0.9628 | 198.5109 | 47.9103 | 46259.0284 | 45687.8576 | 909.7572 | 809.7572 | 786.6967
917 | 98.9558 | 0.98358 | 1.1143 | 0.96936 | 206.134 | 54.6342 | 64536.0497 | 62657.2714 | 1249.1454 | 1149.1454 | 786.9294
918 | 92.2253 | 0.90935 | 0.70617 | 1.0156 | 218.8287 | 49.8219 | 40128.2847 | 40084.6193 | 797.6924 | 697.6924 | 786.9347
919 | 107.8274 | 0.97235 | 1.047 | 0.93374 | 192.6008 | 50.7672 | 55148.6514 | 56054.3762 | 1117.0875 | 1017.0875 | 786.992
920 | 103.9636 | 1.0541 | 0.78334 | 0.94839 | 197.599 | 46.3137 | 39680.29 | 38514.0438 | 766.2809 | 666.2809 | 786.9992
921 | 108.8992 | 1.0512 | 1.0455 | 0.98794 | 185.1942 | 51.2163 | 58946.6495 | 57952.3521 | 1155.047 | 1055.047 | 787.0129
922 | 88.7104 | 1.0435 | 1.196 | 0.98444 | 189.525 | 43.9231 | 55970.272 | 55828.5941 | 1112.5719 | 1012.5719 | 787.0732
923 | 92.49 | 1.0563 | 1.2105 | 1.0799 | 207.0387 | 53.5629 | 75462.7365 | 75037.9408 | 1496.7588 | 1396.7588 | 787.1342
924 | 100.2809 | 1.0188 | 1.2858 | 0.9072 | 200.5577 | 47.8378 | 60889.9022 | 59377.339 | 1183.5468 | 1083.5468 | 787.4881
925 | 104.7608 | 1.031 | 1.0385 | 0.9723 | 214.5456 | 48.3665 | 54275.5817 | 54598.1667 | 1087.9633 | 987.9633 | 787.5649
926 | 91.3359 | 0.97082 | 1.0235 | 0.95533 | 213.8549 | 48.2371 | 51656.5463 | 51076.8092 | 1017.5362 | 917.5362 | 787.586
927 | 93.516 | 0.97509 | 0.81275 | 0.92968 | 182.6306 | 47.5692 | 40463.3674 | 40868.8565 | 813.3771 | 713.3771 | 787.6428
928 | 103.4123 | 1.0211 | 0.91719 | 0.99264 | 204.1852 | 51.5394 | 52569.8177 | 52420.9287 | 1044.4186 | 944.4186 | 787.8146
929 | 94.3578 | 1.0094 | 1.2257 | 0.93406 | 199.7185 | 50.0131 | 62224.095 | 59333.357 | 1182.6671 | 1082.6671 | 787.8338
930 | 114.5508 | 1.0527 | 0.99921 | 1.0544 | 188.9344 | 51.2827 | 60405.297 | 60972.6944 | 1215.4539 | 1115.4539 | 787.9244
931 | 80.5158 | 1.0213 | 0.97865 | 1.0559 | 198.0088 | 50.9305 | 57015.0353 | 58751.6624 | 1171.0332 | 1071.0332 | 787.9321
932 | 118.6304 | 1.0407 | 1.0116 | 0.94273 | 182.931 | 52.3651 | 56587.1341 | 58337.7222 | 1162.7544 | 1062.7544 | 788.0634
933 | 94.9689 | 1.0322 | 1.1027 | 0.96346 | 196.9382 | 47.5073 | 55325.1439 | 57668.2854 | 1149.3657 | 1049.3657 | 788.1108
934 | 91.6412 | 0.93259 | 1.1113 | 0.89797 | 193.2956 | 47.6175 | 51782.9905 | 50867.5902 | 1013.3518 | 913.3518 | 788.2942
935 | 113.3192 | 1.0363 | 0.89946 | 1.0293 | 193.4044 | 49.3329 | 51659.7965 | 50499.5344 | 1005.9907 | 905.9907 | 788.3558
936 | 99.1153 | 1.0318 | 0.93145 | 1.011 | 201.2827 | 54.038 | 56617.3387 | 57186.1664 | 1139.7233 | 1039.7233 | 788.612
937 | 88.7302 | 0.95252 | 1.1092 | 1.0338 | 198.8307 | 49.4886 | 61128.3882 | 60247.2614 | 1200.9452 | 1100.9452 | 788.6386
938 | 111.9038 | 1.0209 | 0.92678 | 0.89857 | 196.8736 | 51.0669 | 48558.2229 | 49574.9175 | 987.4983 | 887.4983 | 788.7343
939 | 98.8015 | 0.98328 | 1.0048 | 0.99583 | 204.6788 | 48.2191 | 53139.525 | 53309.0581 | 1062.1812 | 962.1812 | 788.9206
940 | 82.3828 | 1.0255 | 1.0314 | 1.0285 | 194.2869 | 50.7577 | 58325.696 | 58806.6669 | 1172.1333 | 1072.1333 | 788.9889
941 | 105.2728 | 0.98441 | 0.98177 | 1.015 | 191.1102 | 48.2681 | 53290.9428 | 52676.0366 | 1049.5207 | 949.5207 | 789.2227
942 | 119.3186 | 0.98835 | 1.1019 | 0.94963 | 213.386 | 55.8169 | 65200.7972 | 64960.185 | 1295.2037 | 1195.2037 | 789.268
943 | 111.9844 | 0.95144 | 0.9686 | 1.0168 | 189.5029 | 55.9877 | 61294.6083 | 60945.7834 | 1214.9157 | 1114.9157 | 789.3455
944 | 97.4346 | 0.95608 | 0.88575 | 1.0563 | 183.3612 | 48.7222 | 50306.1525 | 51178.8058 | 1019.5761 | 919.5761 | 789.3512
945 | 110.4437 | 1.0158 | 1.2382 | 0.9923 | 211.4718 | 51.1734 | 68828.3869 | 68116.7996 | 1358.336 | 1258.336 | 789.3693
946 | 105.5773 | 1.0764 | 0.95601 | 1.014 | 198.626 | 48.7468 | 52992.1047 | 53910.8852 | 1074.2177 | 974.2177 | 789.5578
947 | 110.4747 | 0.95804 | 1.0564 | 0.9526 | 190.3814 | 53.5343 | 59729.1889 | 58084.5796 | 1157.6916 | 1057.6916 | 789.5826
948 | 100.2539 | 1.0104 | 1.0258 | 1.0207 | 195.0215 | 49.641 | 57195.2876 | 57247.4461 | 1140.9489 | 1040.9489 | 789.6578
949 | 94.5707 | 1.0716 | 1.0266 | 0.97539 | 203.1413 | 53.2181 | 58884.3302 | 59048.8392 | 1176.9768 | 1076.9768 | 789.7359
950 | 104.9044 | 1.0035 | 0.88435 | 0.99936 | 190.4359 | 49.5163 | 49165.0377 | 47843.9528 | 952.8791 | 852.8791 | 789.745
951 | 101.9413 | 1.0705 | 1.1462 | 0.99078 | 179.4616 | 52.7597 | 65852.9905 | 66925.4167 | 1334.5083 | 1234.5083 | 789.764
952 | 98.0345 | 1.0203 | 1.0125 | 0.9941 | 212.4623 | 51.4567 | 57149.5833 | 54292.1402 | 1081.8428 | 981.8428 | 789.8482
953 | 96.1735 | 1.1224 | 0.87209 | 0.9448 | 194.4903 | 48.7903 | 45661.8953 | 46575.6369 | 927.5127 | 827.5127 | 789.9875
954 | 101.2698 | 1.0066 | 0.88031 | 1.0237 | 192.0113 | 54.1583 | 54517.2358 | 53375.8286 | 1063.5166 | 963.5166 | 790.0271
955 | 87.8163 | 1.0237 | 0.9631 | 1.0149 | 197.7993 | 50.2839 | 53870.3446 | 53369.2019 | 1063.384 | 963.384 | 790.0756
956 | 105.0032 | 1.0036 | 1.0936 | 0.98839 | 199.0654 | 47.6422 | 56714.2494 | 58087.2022 | 1157.744 | 1057.744 | 790.3121
957 | 92.5242 | 0.95246 | 0.81934 | 1.0045 | 180.1534 | 48.9457 | 44775.2953 | 45592.0322 | 907.8406 | 807.8406 | 790.377
958 | 81.4805 | 0.97746 | 0.96978 | 1.0934 | 200.2498 | 50.6687 | 57961.9782 | 57831.1321 | 1152.6226 | 1052.6226 | 790.5056
959 | 93.912 | 0.97288 | 0.98448 | 0.96166 | 193.9934 | 50.0987 | 52201.2563 | 50728.0828 | 1010.5617 | 910.5617 | 790.5158
960 | 93.266 | 0.99703 | 1.0195 | 1.0141 | 198.2265 | 51.4413 | 58164.7027 | 56843.6512 | 1132.873 | 1032.873 | 790.6457
961 | 102.5597 | 1.0435 | 0.821 | 0.92978 | 197.8941 | 52.173 | 45607.4918 | 46041.6303 | 916.8326 | 816.8326 | 790.6689
962 | 105.7377 | 0.90938 | 1.2239 | 0.9897 | 184.3001 | 51.1948 | 67118.4446 | 66868.2177 | 1333.3644 | 1233.3644 | 790.7309
963 | 89.1519 | 0.94103 | 0.87711 | 0.972 | 183.8874 | 49.1509 | 46211.1747 | 46196.6969 | 919.9339 | 819.9339 | 790.7479
964 | 96.4485 | 1.0423 | 0.89904 | 0.96763 | 209.076 | 50.5149 | 49232.5617 | 49421.8481 | 984.437 | 884.437 | 790.8373
965 | 79.8612 | 0.90107 | 1.0103 | 0.95812 | 190.6246 | 50.0674 | 52257.9335 | 52167.5324 | 1039.3506 | 939.3506 | 790.8486
966 | 118.2193 | 0.90089 | 0.82305 | 0.99344 | 190.7242 | 49.9588 | 46360.0057 | 46401.052 | 924.021 | 824.021 | 790.9861
967 | 107.3037 | 0.99489 | 0.94481 | 1.0815 | 188.4331 | 54.2648 | 61431.0122 | 62187.5439 | 1239.7509 | 1139.7509 | 791.0609
968 | 97.8783 | 1.0244 | 0.71621 | 1.014 | 201.7249 | 45.4533 | 37768.0297 | 37925.6209 | 754.5124 | 654.5124 | 791.0951
969 | 104.5945 | 0.92997 | 0.80154 | 0.99865 | 190.4189 | 52.9541 | 47728.5218 | 47996.905 | 955.9381 | 855.9381 | 791.336
970 | 104.2483 | 0.9015 | 0.99555 | 1.0527 | 207.0052 | 51.2209 | 58699.1807 | 57718.0094 | 1150.3602 | 1050.3602 | 791.396
971 | 105.3243 | 0.98988 | 1.0045 | 1.1012 | 218.4036 | 48.5826 | 59026.9616 | 58502.382 | 1166.0476 | 1066.0476 | 791.4093
972 | 105.4351 | 0.88364 | 0.99407 | 0.99986 | 206.0535 | 51.3231 | 55999.2859 | 54951.2047 | 1095.0241 | 995.0241 | 791.5192
973 | 92.2104 | 0.97297 | 0.94376 | 1.052 | 196.5927 | 50.9892 | 55393.2207 | 56067.5171 | 1117.3503 | 1017.3503 | 791.5743
974 | 110.7922 | 1.0202 | 1.1643 | 0.94517 | 209.2325 | 52.4159 | 63814.1292 | 64384.1432 | 1283.6829 | 1183.6829 | 791.6282
975 | 109.2221 | 0.94802 | 1.2015 | 1.0186 | 199.5468 | 45.5755 | 60697.1249 | 62101.9174 | 1238.0383 | 1138.0383 | 791.9179
976 | 92.0026 | 0.93224 | 0.98485 | 1.0063 | 214.0727 | 50.3096 | 54389.9659 | 54140.4587 | 1078.8092 | 978.8092 | 792.1802
977 | 122.1026 | 1.023 | 0.67784 | 1.0017 | 207.771 | 50.2814 | 40629.8738 | 39410.3475 | 784.2069 | 684.2069 | 792.2658
978 | 108.656 | 0.95598 | 1.1004 | 1.156 | 199.2375 | 47.4712 | 65520.3494 | 65704.7229 | 1310.0945 | 1210.0945 | 792.3482
979 | 87.3327 | 0.9953 | 0.9772 | 0.96116 | 203.4795 | 48.4694 | 49941.6172 | 50456.1193 | 1005.1224 | 905.1224 | 792.3871
980 | 100.8119 | 0.93285 | 0.89998 | 1.048 | 218.7029 | 51.4693 | 53604.2356 | 55663.7914 | 1109.2758 | 1009.2758 | 792.4873
981 | 100.6199 | 0.91659 | 1.1325 | 0.94761 | 194.8187 | 49.5866 | 57982.6531 | 59980.8611 | 1195.6172 | 1095.6172 | 792.5825
982 | 91.2555 | 0.98221 | 1.2361 | 0.97011 | 210.3418 | 44.2325 | 57214.5736 | 55787.286 | 1111.7457 | 1011.7457 | 792.6394
983 | 96.9662 | 0.97169 | 0.92131 | 0.90993 | 186.7996 | 45.1238 | 42267.0817 | 42245.4011 | 840.908 | 740.908 | 792.8456
984 | 98.4532 | 1.0263 | 0.91153 | 0.96054 | 215.4149 | 50.3536 | 49390.8318 | 50229.3744 | 1000.5875 | 900.5875 | 792.875
985 | 93.4784 | 0.98962 | 0.97491 | 0.97681 | 193.4669 | 52.3463 | 54885.4481 | 53228.5159 | 1060.5703 | 960.5703 | 793.0193
986 | 117.1715 | 0.96968 | 0.99558 | 0.94758 | 191.6643 | 52.7086 | 55905.4266 | 56066.4939 | 1117.3299 | 1017.3299 | 793.0244
987 | 88.039 | 0.91461 | 0.933 | 1.0119 | 184.6715 | 54.0003 | 55515.1192 | 56844.3077 | 1132.8862 | 1032.8862 | 793.0371
988 | 96.0305 | 1.0088 | 1.0426 | 0.98213 | 210.5203 | 48.4057 | 54466.4526 | 52767.2614 | 1051.3452 | 951.3452 | 793.0663
989 | 98.877 | 1.1085 | 0.95522 | 1.0251 | 184.7439 | 51.3485 | 56091.5611 | 55680.2798 | 1109.6056 | 1009.6056 | 793.091
990 | 98.1439 | 1.0863 | 0.89252 | 1.0713 | 180.9286 | 49.1671 | 52434.043 | 52809.3377 | 1052.1868 | 952.1868 | 793.1149
991 | 127.3709 | 1.0633 | 0.87731 | 1.0214 | 184.9338 | 48.817 | 50541.0297 | 49674.7618 | 989.4952 | 889.4952 | 793.1363
992 | 102.137 | 0.99079 | 1.3678 | 0.97685 | 202.5012 | 52.4084 | 75528.554 | 75837.6732 | 1512.7535 | 1412.7535 | 793.1713
993 | 105.2643 | 1.0623 | 1.0592 | 0.98899 | 201.0375 | 46.7126 | 54358.04 | 54237.9193 | 1080.7584 | 980.7584 | 793.4679
994 | 118.0463 | 1.1284 | 1.0935 | 1.0186 | 199.5102 | 48.8506 | 61118.0818 | 60735.0541 | 1210.7011 | 1110.7011 | 793.521
995 | 100.0324 | 1.0415 | 1.0838 | 1.0203 | 201.1362 | 52.5925 | 63837.4252 | 63662.8164 | 1269.2563 | 1169.2563 | 793.9307
996 | 86.6252 | 0.99339 | 0.68947 | 1.0772 | 200.6748 | 43.6261 | 36357.192 | 36288.3461 | 721.7669 | 621.7669 | 793.9775
997 | 107.0945 | 0.98437 | 0.82482 | 0.96759 | 222.5598 | 46.4781 | 42216.1735 | 42050.3959 | 837.0079 | 737.0079 | 794.0271
998 | 94.185 | 0.9614 | 1.0406 | 1.0672 | 201.5167 | 51.5551 | 62122.6405 | 62589.3636 | 1247.7873 | 1147.7873 | 794.18
999 | 90.1216 | 1.049 | 0.80193 | 0.92455 | 194.9686 | 51.1323 | 42939.4184 | 42888.8926 | 853.7779 | 753.7779 | 794.3132
1000 | 96.3457 | 1.0787 | 0.96058 | 1.0594 | 193.7837 | 52.4717 | 59042.6408 | 56322.2586 | 1122.4452 | 1022.4452 | 794.4137
1001 | 95.6354 | 1.0393 | 1.0284 | 0.97971 | 218.7629 | 46.248 | 51413.1361 | 51518.1909 | 1026.3638 | 926.3638 | 794.4652
1002 | 101.4903 | 0.99543 | 0.93905 | 0.96346 | 199.8009 | 46.971 | 47441.406 | 46556.3823 | 927.1276 | 827.1276 | 794.4715
1003 | 99.5916 | 1.0626 | 1.1928 | 1.0503 | 204.6561 | 49.1476 | 66978.9339 | 64375.4866 | 1283.5097 | 1183.5097 | 794.7433
1004 | 92.0383 | 0.98435 | 0.83262 | 0.91016 | 191.4477 | 50.4474 | 42991.8 | 42972.8728 | 855.4575 | 755.4575 | 794.9306
1005 | 86.9495 | 1.0012 | 1.0328 | 1.0004 | 200.3944 | 51.3905 | 57772.3806 | 58302.6502 | 1162.053 | 1062.053 | 794.9742
1006 | 76.3446 | 0.96692 | 0.9231 | 1.0681 | 209.5024 | 46.4026 | 49387.6737 | 49012.9324 | 976.2586 | 876.2586 | 795.02
1007 | 106.4691 | 0.98226 | 1.1188 | 1.0363 | 194.2046 | 50.07 | 63483.6566 | 65817.5857 | 1312.3517 | 1212.3517 | 795.2319
1008 | 93.8376 | 1.1094 | 1.0579 | 1.0081 | 193.5492 | 50.892 | 59765.5671 | 59686.5102 | 1189.7302 | 1089.7302 | 795.2345
1009 | 103.8672 | 0.89597 | 0.93989 | 1.008 | 195.9159 | 50.146 | 52372.8731 | 51289.0368 | 1021.7807 | 921.7807 | 795.2925
1010 | 92.1059 | 1.0151 | 1.0376 | 0.96448 | 210.3679 | 50.2767 | 55225.7257 | 56037.7033 | 1116.7541 | 1016.7541 | 795.3127
1011 | 108.9674 | 1.0093 | 1.0354 | 1.0027 | 204.2717 | 50.7154 | 58433.4504 | 58067.2874 | 1157.3457 | 1057.3457 | 795.3543
1012 | 96.5149 | 1.0134 | 0.96795 | 0.98913 | 214.8689 | 52.6626 | 55786.134 | 54947.3181 | 1094.9464 | 994.9464 | 795.3649
1013 | 81.9608 | 1.0235 | 1.0852 | 0.89436 | 163.9884 | 47.8413 | 50609.7104 | 51658.9097 | 1029.1782 | 929.1782 | 795.5572
1014 | 97.7492 | 1.0321 | 1.0829 | 0.93767 | 199.2099 | 49.2423 | 55169.4137 | 53254.1895 | 1061.0838 | 961.0838 | 795.5818
1015 | 118.1307 | 0.99284 | 1 | 0.99387 | 214.8747 | 53.3154 | 59457.5664 | 58584.9461 | 1167.6989 | 1067.6989 | 795.616
1016 | 116.4617 | 1.0841 | 0.99067 | 1.0132 | 196.8324 | 51.8339 | 58767.2827 | 58652.5895 | 1169.0518 | 1069.0518 | 795.7453
1017 | 105.1527 | 0.95701 | 0.94304 | 1.053 | 197.1225 | 51.9516 | 57013.4448 | 55085.3645 | 1097.7073 | 997.7073 | 795.844
1018 | 85.6698 | 1.0451 | 1.1208 | 1.0197 | 192.2773 | 52.6079 | 65025.217 | 64132.9307 | 1278.6586 | 1178.6586 | 795.9162
1019 | 106.4286 | 0.9829 | 1.1741 | 0.89546 | 188.6919 | 54.8032 | 63539.0167 | 65558.0653 | 1307.1613 | 1207.1613 | 796.0849
1020 | 95.795 | 1.026 | 0.79401 | 0.99583 | 195.8253 | 50.2144 | 44835.5429 | 44923.5761 | 894.4715 | 794.4715 | 796.0955
1021 | 105.4907 | 1.0146 | 1.2157 | 1.0572 | 192.0805 | 53.8091 | 75107.3577 | 76797.3573 | 1531.9471 | 1431.9471 | 796.2202
1022 | 86.7477 | 1.051 | 1.0406 | 1.0037 | 186.4179 | 48.9909 | 55822.4567 | 56576.3921 | 1127.5278 | 1027.5278 | 796.4111
1023 | 93.431 | 1.0425 | 0.78294 | 0.93056 | 186.9654 | 49.03 | 40684.2619 | 38636.4467 | 768.7289 | 668.7289 | 796.515
1024 | 110.4499 | 0.90169 | 1.3309 | 1.0252 | 199.481 | 51.4421 | 75510.7136 | 76945.0284 | 1534.9006 | 1434.9006 | 796.5775
1025 | 94.3522 | 0.9773 | 1.1095 | 1.0454 | 193.0001 | 55.5116 | 69700.2964 | 68772.13 | 1371.4426 | 1271.4426 | 796.6204
1026 | 90.3904 | 1.0071 | 0.98655 | 1.0303 | 198.5931 | 48.6537 | 54080.2893 | 56801.091 | 1132.0218 | 1032.0218 | 796.6528
1027 | 88.2693 | 1.0315 | 1.0131 | 1.0142 | 185.896 | 48.5974 | 54545.8181 | 54278.4012 | 1081.568 | 981.568 | 796.6871
1028 | 107.8078 | 1.0215 | 0.99808 | 0.99455 | 191.8705 | 50.818 | 56232.2532 | 58409.2099 | 1164.1842 | 1064.1842 | 796.7677
1029 | 97.4434 | 0.9776 | 1.1739 | 1.0255 | 178.5292 | 52.649 | 68570.5364 | 66103.4174 | 1318.0683 | 1218.0683 | 796.775
1030 | 116.6026 | 1.0235 | 1.0941 | 1.0187 | 187.0297 | 46.7583 | 57882.9496 | 56693.3127 | 1129.8663 | 1029.8663 | 796.925
1031 | 108.9072 | 1 | 0.90227 | 1.0122 | 209.4299 | 48.9297 | 50226.2231 | 51308.7261 | 1022.1745 | 922.1745 | 797.005
1032 | 92.4462 | 0.8771 | 1.2729 | 0.91993 | 210.1681 | 50.8017 | 63816.7131 | 64614.2576 | 1288.2852 | 1188.2852 | 797.0749
1033 | 126.0983 | 1.0434 | 1.3281 | 1.024 | 186.8158 | 54.3702 | 81282.1752 | 80698.7238 | 1609.9745 | 1509.9745 | 797.2705
1034 | 93.3628 | 0.99777 | 1.0157 | 0.98218 | 184.2009 | 49.8599 | 54569.2533 | 55622.5595 | 1108.4512 | 1008.4512 | 797.4056
1035 | 88.3981 | 0.95395 | 1.0269 | 1.0498 | 175.9352 | 51.8994 | 60498.6288 | 58965.2898 | 1175.3058 | 1075.3058 | 797.4488
1036 | 92.3917 | 1.0215 | 1.1676 | 0.99003 | 209.6654 | 51.5206 | 64627.6867 | 65503.0778 | 1306.0616 | 1206.0616 | 797.5092
1037 | 96.9821 | 1.0777 | 1.1619 | 0.98875 | 193.6354 | 50.5439 | 63544.0881 | 66019.9175 | 1316.3984 | 1216.3984 | 797.5234
1038 | 90.5358 | 0.9178 | 1.2621 | 1.0418 | 177.6091 | 46.6193 | 65349.1197 | 63923.3879 | 1274.4678 | 1174.4678 | 797.54
1039 | 92.5323 | 0.93014 | 1.0866 | 1.0295 | 201.6524 | 48.5776 | 58725.7787 | 57651.7522 | 1149.035 | 1049.035 | 797.5745
1040 | 97.1937 | 0.95095 | 1.0927 | 1.0229 | 204.5046 | 47.2404 | 57373.7265 | 57445.5421 | 1144.9108 | 1044.9108 | 797.6402
1041 | 98.7035 | 0.96285 | 1.0799 | 1.0078 | 196.0248 | 50.1063 | 59488.4336 | 60106.3524 | 1198.127 | 1098.127 | 797.6453
1042 | 96.0143 | 1.0297 | 1.1153 | 1.0022 | 188.8428 | 48.7075 | 59445.5941 | 57309.7846 | 1142.1957 | 1042.1957 | 797.6483
1043 | 99.2978 | 0.96263 | 0.70409 | 0.93972 | 194.5144 | 49.3504 | 37564.0412 | 36742.4412 | 730.8488 | 630.8488 | 797.7441
1044 | 94.603 | 1.0595 | 0.8268 | 1.0256 | 201.6643 | 48.9581 | 46625.0292 | 45063.5249 | 897.2705 | 797.2705 | 797.7749
1045 | 104.15 | 1.0145 | 1.1673 | 0.9702 | 194.9783 | 48.2443 | 59928.3577 | 57956.2354 | 1155.1247 | 1055.1247 | 797.7825
1046 | 108.1545 | 0.97311 | 1.0011 | 0.96958 | 217.4935 | 49.271 | 53229.6399 | 52191.5979 | 1039.832 | 939.832 | 797.8692
1047 | 97.7054 | 1.0287 | 0.99712 | 0.96173 | 200.8073 | 52.2768 | 55586.9025 | 56750.8571 | 1131.0171 | 1031.0171 | 797.8775
1048 | 103.9134 | 1.0073 | 0.97656 | 1.0137 | 201.2201 | 53.2127 | 58449.1308 | 58703.4179 | 1170.0684 | 1070.0684 | 797.9799
1049 | 100.1138 | 1.0393 | 0.88863 | 0.92723 | 221.7162 | 50.1143 | 46728.3044 | 46414.8509 | 924.297 | 824.297 | 798.0131
1050 | 92.228 | 0.97533 | 1.1211 | 0.96839 | 189.8908 | 53.6809 | 63297.2882 | 63421.2404 | 1264.4248 | 1164.4248 | 798.0156
1051 | 104.5448 | 1.0173 | 0.79253 | 0.91451 | 183.9798 | 49.6192 | 41424.2232 | 42325.8108 | 842.5162 | 742.5162 | 798.0728
1052 | 117.586 | 0.88213 | 0.99501 | 1.03 | 197.0985 | 52.8685 | 59864.786 | 61908.0135 | 1234.1603 | 1134.1603 | 798.2049
1053 | 95.4305 | 1.0581 | 0.79116 | 0.98735 | 209.0271 | 42.7617 | 37930.437 | 36663.6246 | 729.2725 | 629.2725 | 798.2474
1054 | 94.5106 | 0.99556 | 1.1137 | 1.0329 | 203.9848 | 49.0527 | 61248.0709 | 62098.4895 | 1237.9698 | 1137.9698 | 798.2577
1055 | 101.0406 | 1.0335 | 0.84751 | 0.98761 | 184.8832 | 44.8671 | 42424.1903 | 42453.2954 | 845.0659 | 745.0659 | 798.2705
1056 | 114.0763 | 1.0315 | 1.161 | 0.92659 | 211.1106 | 49.9476 | 59821.5351 | 60134.1049 | 1198.6821 | 1098.6821 | 798.2951
1057 | 92.3749 | 1.013 | 0.83511 | 0.92743 | 194.4053 | 47.0574 | 41044.2899 | 39927.1031 | 794.5421 | 694.5421 | 798.3621
1058 | 107.2758 | 0.97516 | 1.1017 | 1.0801 | 205.4764 | 49.2001 | 63897.0542 | 62773.4853 | 1251.4697 | 1151.4697 | 798.3835
1059 | 121.3667 | 0.93185 | 0.89629 | 0.9743 | 203.1234 | 49.238 | 48768.8819 | 49483.6435 | 985.6729 | 885.6729 | 798.4687
1060 | 97.755 | 1.0531 | 0.89494 | 0.99021 | 207.8733 | 47.5518 | 47242.6707 | 46685.1802 | 929.7036 | 829.7036 | 798.5074
1061 | 101.043 | 0.90452 | 1.0565 | 1.0638 | 203.9789 | 48.4624 | 59099.8864 | 61233.0996 | 1220.662 | 1120.662 | 798.5957
1062 | 81.1546 | 0.97541 | 1.2167 | 1.0683 | 208.6944 | 45.8896 | 63490.3647 | 60296.3048 | 1201.9261 | 1101.9261 | 798.6337
1063 | 110.3845 | 1.0036 | 1.026 | 1.0444 | 184.424 | 49.6535 | 58889.5249 | 57549.2757 | 1146.9855 | 1046.9855 | 798.6391
1064 | 104.1631 | 0.99488 | 0.91395 | 1.0112 | 193.6203 | 51.6245 | 53251.8925 | 53050.9926 | 1057.0199 | 957.0199 | 798.647
1065 | 68.2617 | 0.92139 | 1.1571 | 1.0316 | 216.1353 | 48.7461 | 61472.9326 | 62053.9036 | 1237.0781 | 1137.0781 | 798.7506
1066 | 90.6271 | 0.96976 | 1.0499 | 0.9959 | 197.7439 | 49.5537 | 56365.4355 | 56287.4813 | 1121.7496 | 1021.7496 | 798.8594
1067 | 107.3727 | 1.0472 | 1.0041 | 0.95985 | 184.8639 | 52.5802 | 56775.5884 | 57065.6151 | 1137.3123 | 1037.3123 | 799.0216
1068 | 95.0046 | 0.94379 | 1.2662 | 1.0545 | 191.2138 | 51.2391 | 73201.5051 | 74198.7724 | 1479.9754 | 1379.9754 | 799.0825
1069 | 85.2415 | 1.0071 | 0.83372 | 0.98824 | 187.0386 | 52.9126 | 48324.9999 | 48213.3367 | 960.2667 | 860.2667 | 799.0924
1070 | 93.8104 | 1.0804 | 1.1761 | 0.99402 | 198.343 | 53.2979 | 67908.6927 | 67280.4416 | 1341.6088 | 1241.6088 | 799.2209
1071 | 100.5605 | 1.0725 | 0.93424 | 1.0008 | 196.7605 | 53.8353 | 56339.0493 | 56697.8042 | 1129.9561 | 1029.9561 | 799.272
1072 | 97.1154 | 1.0304 | 1.2492 | 0.95947 | 198.5566 | 51.4985 | 67078.4469 | 66961.4043 | 1335.2281 | 1235.2281 | 799.3893
1073 | 90.8115 | 0.99704 | 0.79528 | 0.98569 | 213.3077 | 51.2174 | 44999.8493 | 45427.4127 | 904.5483 | 804.5483 | 799.4327
1074 | 114.5778 | 1.0537 | 0.90586 | 1.06 | 208.438 | 52.0865 | 56512.8506 | 56559.4796 | 1127.1896 | 1027.1896 | 799.5309
1075 | 93.2885 | 1.0005 | 0.91236 | 1.0277 | 195.2161 | 46.1114 | 47734.7317 | 47675.3295 | 949.5066 | 849.5066 | 799.7401
1076 | 104.6688 | 0.98022 | 1.1026 | 0.98932 | 186.4993 | 50.9811 | 61029.2015 | 58922.3903 | 1174.4478 | 1074.4478 | 799.8352
1077 | 118.1326 | 1.0145 | 0.92798 | 0.95373 | 203.0065 | 45.8239 | 46250.7589 | 45972.6978 | 915.454 | 815.454 | 799.861
1078 | 106.8133 | 0.96791 | 0.988 | 1.049 | 194.3259 | 46.7617 | 53493.0631 | 53919.2347 | 1074.3847 | 974.3847 | 800.0895
1079 | 103.4043 | 1.0137 | 1.0028 | 1.0517 | 210.8547 | 46.5531 | 54189.2937 | 53134.5991 | 1058.692 | 958.692 | 800.228
1080 | 89.9175 | 1.0253 | 1.0187 | 1.022 | 195.0793 | 53.1482 | 60423.7726 | 63439.515 | 1264.7903 | 1164.7903 | 800.2646
1081 | 79.908 | 0.99723 | 0.82522 | 0.93701 | 203.9678 | 49.7718 | 42655.9848 | 42774.3371 | 851.4867 | 751.4867 | 800.3006
1082 | 92.8774 | 1.024 | 1.0257 | 0.9659 | 195.4936 | 46.5002 | 50686.7777 | 51158.9667 | 1019.1793 | 919.1793 | 800.3311
1083 | 105.1977 | 0.99873 | 0.97105 | 0.99122 | 197.8389 | 50.8775 | 54514.1824 | 54207.5188 | 1080.1504 | 980.1504 | 800.3573
1084 | 111.5028 | 0.95463 | 0.71115 | 0.95767 | 204.8683 | 51.0128 | 40377.0037 | 41448.2119 | 824.9642 | 724.9642 | 800.364
1085 | 97.7812 | 1.0261 | 1.2453 | 1.0247 | 211.2282 | 54.1616 | 74761.4723 | 72799.991 | 1451.9998 | 1351.9998 | 800.4018
1086 | 86.4473 | 1.0239 | 0.77614 | 0.95173 | 210.2648 | 52.0008 | 43224.7666 | 43739.0881 | 870.7818 | 770.7818 | 800.4525
1087 | 107.7716 | 0.8574 | 1.0373 | 0.95137 | 181.6528 | 50.5195 | 54704.2277 | 54246.8734 | 1080.9375 | 980.9375 | 800.4929
1088 | 119.2917 | 0.96769 | 1.0892 | 0.9397 | 202.4857 | 48.3347 | 55253.7037 | 54379.1414 | 1083.5828 | 983.5828 | 800.6293
1089 | 108.6092 | 1.0009 | 0.89006 | 1.0056 | 201.3395 | 51.7446 | 52139.5648 | 50683.8208 | 1009.6764 | 909.6764 | 800.6917
1090 | 114.3578 | 0.91637 | 0.76894 | 1.0482 | 203.2067 | 49.1317 | 44950.695 | 44349.404 | 882.9881 | 782.9881 | 800.7256
1091 | 100.1605 | 1.0012 | 0.84337 | 1.0433 | 209.8524 | 49.8663 | 49086.9571 | 50860.7286 | 1013.2146 | 913.2146 | 800.8315
1092 | 104.1892 | 0.96687 | 1.0338 | 1.0061 | 217.3738 | 48.146 | 55143.8966 | 55210.2046 | 1100.2041 | 1000.2041 | 800.8596
1093 | 111.8265 | 1.0855 | 1.0017 | 0.97804 | 199.5662 | 47.8287 | 52864.8576 | 51846.2499 | 1032.925 | 932.925 | 800.8811
1094 | 120.8467 | 0.9829 | 1.0743 | 1.0055 | 198.1916 | 59.649 | 71714.7137 | 71395.0721 | 1423.9014 | 1323.9014 | 800.9103
1095 | 97.5478 | 0.93103 | 0.81972 | 1.062 | 188.9253 | 50.1921 | 48441.1856 | 46311.4746 | 922.2295 | 822.2295 | 800.9886
1096 | 90.6235 | 0.94418 | 0.99073 | 1.04 | 220.0159 | 53.0076 | 59372.2683 | 58391.0015 | 1163.82 | 1063.82 | 801.3072
1097 | 94.6337 | 0.95819 | 0.95103 | 0.94617 | 202.8498 | 50.314 | 50039.3723 | 51064.0646 | 1017.2813 | 917.2813 | 801.3225
1098 | 85.0855 | 1.0604 | 0.88785 | 0.95954 | 203.4252 | 53.0709 | 50204.4805 | 50342.1087 | 1002.8422 | 902.8422 | 801.4596
1099 | 96.7043 | 0.99936 | 1.0752 | 0.92865 | 202.6011 | 45.4323 | 49955.4838 | 48598.4371 | 967.9687 | 867.9687 | 801.6773
1100 | 109.1209 | 0.97873 | 0.98122 | 0.95637 | 196.2691 | 51.0914 | 53597.6613 | 54070.804 | 1077.4161 | 977.4161 | 801.6949
1101 | 99.2587 | 0.99876 | 1.0519 | 1.0251 | 181.0402 | 50.5746 | 59728.9143 | 61478.0526 | 1225.5611 | 1125.5611 | 801.7097
1102 | 87.0543 | 1.0516 | 0.94037 | 0.92665 | 201.3855 | 49.7019 | 48061.358 | 46737.29 | 930.7458 | 830.7458 | 801.9453
1103 | 109.882 | 0.94234 | 0.96917 | 1.097 | 212.4499 | 47.0111 | 55059.4887 | 54035.6839 | 1076.7137 | 976.7137 | 801.989
1104 | 105.4063 | 0.94123 | 1.1262 | 0.98487 | 194.8911 | 48.8371 | 59206.6015 | 61510.5543 | 1226.2111 | 1126.2111 | 802.1012
1105 | 97.4532 | 1.0361 | 1.0095 | 0.96995 | 201.8525 | 51.0807 | 55374.818 | 55401.8707 | 1104.0374 | 1004.0374 | 802.2219
1106 | 101.4443 | 1.0334 | 0.95915 | 1.0709 | 189.2472 | 49.4831 | 56205.5889 | 57010.7679 | 1136.2154 | 1036.2154 | 802.2397
1107 | 96.4223 | 0.98372 | 0.91284 | 0.98772 | 202.3857 | 51.8097 | 51830.0613 | 53446.3512 | 1064.927 | 964.927 | 802.2398
1108 | 105.6224 | 0.94831 | 1.111 | 0.96985 | 199.8703 | 50.5651 | 59749.7477 | 59478.6854 | 1185.5737 | 1085.5737 | 802.2654
1109 | 89.1879 | 0.92339 | 0.60104 | 0.97654 | 176.3531 | 51.8226 | 34860.5928 | 35026.2365 | 696.5247 | 596.5247 | 802.3112
1110 | 104.2458 | 0.98093 | 0.98043 | 1.0139 | 191.0333 | 49.5032 | 54461.6253 | 53567.7726 | 1067.3555 | 967.3555 | 802.3614
1111 | 109.7381 | 1.0262 | 1.1118 | 0.97103 | 197.3636 | 46.7673 | 55952.0781 | 55581.1251 | 1107.6225 | 1007.6225 | 802.5792
1112 | 104.7467 | 0.99155 | 0.82645 | 1.0405 | 194.4367 | 42.0026 | 40675.1771 | 40974.9217 | 815.4984 | 715.4984 | 802.6509
1113 | 94.0701 | 0.89238 | 0.92329 | 1.0411 | 202.3236 | 48.8743 | 51286.4377 | 51200.8169 | 1020.0163 | 920.0163 | 802.6669
1114 | 100.0001 | 1.0322 | 1.0112 | 1.0284 | 177.6016 | 55.1732 | 63244.2822 | 60520.1403 | 1206.4028 | 1106.4028 | 802.711
1115 | 95.4601 | 0.97787 | 1.0607 | 0.98858 | 182.1608 | 50.6564 | 58030.1458 | 58193.2539 | 1159.8651 | 1059.8651 | 802.9184
1116 | 94.408 | 0.97302 | 0.64379 | 1.0162 | 203.5001 | 51.0466 | 38290.0298 | 37370.335 | 743.4067 | 643.4067 | 803.2831
1117 | 95.7712 | 1.0331 | 0.91163 | 1.0886 | 204.7972 | 51.7229 | 56654.7934 | 56489.8713 | 1125.7974 | 1025.7974 | 803.3651
1118 | 97.3515 | 0.91979 | 1.0116 | 0.96776 | 205.3797 | 49.4827 | 53078.7768 | 52800.8189 | 1052.0164 | 952.0164 | 803.4146
1119 | 93.1876 | 1.0405 | 0.91607 | 1.0681 | 197.5088 | 50.5203 | 54529.2266 | 54582.8015 | 1087.656 | 987.656 | 803.4155
1120 | 122.5301 | 0.97382 | 0.87485 | 1.075 | 194.745 | 52.8442 | 56198.9363 | 55613.239 | 1108.2648 | 1008.2648 | 803.4619
1121 | 93.7778 | 1.0209 | 0.94012 | 1.0138 | 202.707 | 47.7815 | 50318.1477 | 50288.8023 | 1001.776 | 901.776 | 803.5504
1122 | 93.8742 | 0.97264 | 0.94004 | 1.0019 | 202.1991 | 46.7792 | 48530.5267 | 48532.3576 | 966.6472 | 866.6472 | 803.6447
1123 | 103.2048 | 1.1056 | 0.74632 | 1.0483 | 205.0284 | 46.8289 | 42185.3255 | 42630.1044 | 848.6021 | 748.6021 | 803.7027
1124 | 107.071 | 1.0568 | 1.0034 | 1.0243 | 199.4067 | 49.9184 | 57154.2714 | 57682.6207 | 1149.6524 | 1049.6524 | 803.7309
1125 | 86.0309 | 1.0003 | 1.2248 | 1.0692 | 186.0819 | 48.3669 | 67686.989 | 70205.6424 | 1400.1128 | 1300.1128 | 803.9108
1126 | 102.2192 | 1.0488 | 0.83833 | 0.94629 | 193.2495 | 50.308 | 45496.2027 | 44770.4635 | 891.4093 | 791.4093 | 804.0097
1127 | 110.3095 | 1.0381 | 0.80886 | 0.98552 | 183.2968 | 51.6784 | 47296.7581 | 47397.2643 | 943.9453 | 843.9453 | 804.0722
1128 | 112.3758 | 0.90939 | 0.68007 | 0.98697 | 204.2001 | 47.6216 | 37034.6522 | 38914.0585 | 774.2812 | 674.2812 | 804.2266
1129 | 90.6208 | 0.99288 | 1.0458 | 1.1003 | 217.3299 | 50.3992 | 62747.2964 | 64257.5396 | 1281.1508 | 1181.1508 | 804.3992
1130 | 116.5003 | 1.018 | 0.87492 | 1.011 | 193.3225 | 50.4635 | 50816.6048 | 51628.3688 | 1028.5674 | 928.5674 | 804.4943
1131 | 95.4307 | 0.95788 | 0.96347 | 1.0415 | 202.2351 | 52.8081 | 58021.1188 | 59227.4765 | 1180.5495 | 1080.5495 | 804.5483
1132 | 107.374 | 1.0157 | 1.1064 | 1.0217 | 194.2896 | 47.7311 | 59352.6485 | 60422.6425 | 1204.4528 | 1104.4528 | 804.5786
1133 | 114.142 | 1.0834 | 1.1587 | 0.88207 | 183.8326 | 49.8965 | 57349.1386 | 57058.1965 | 1137.1639 | 1037.1639 | 804.677
1134 | 130.1865 | 1.0763 | 0.70768 | 1.0154 | 198.0779 | 51.8856 | 44752.3964 | 45100.6547 | 898.0131 | 798.0131 | 804.7129
1135 | 80.7582 | 0.97763 | 0.94093 | 1.0131 | 215.7258 | 51.8323 | 53717.742 | 52748.2829 | 1050.9657 | 950.9657 | 804.8223
1136 | 101.3248 | 1.0235 | 0.98717 | 1.0135 | 189.3804 | 49.0458 | 54346.2265 | 54412.3488 | 1084.247 | 984.247 | 804.8539
1137 | 106.6447 | 0.99937 | 1.1426 | 0.95152 | 204.1622 | 52.2797 | 62613.6975 | 63854.0582 | 1273.0812 | 1173.0812 | 804.9139
1138 | 98.1921 | 0.96062 | 0.91592 | 0.97564 | 194.3532 | 52.107 | 51672.3559 | 53201.9392 | 1060.0388 | 960.0388 | 805.2106
1139 | 111.1977 | 0.95174 | 0.91542 | 1.0106 | 192.7398 | 51.8083 | 53603.2058 | 52426.8172 | 1044.5363 | 944.5363 | 805.2131
1140 | 93.5606 | 0.94818 | 1.1347 | 0.91286 | 220.8852 | 50.6928 | 57228.666 | 57167.5512 | 1139.351 | 1039.351 | 805.2553
1141 | 94.1321 | 0.97451 | 1.222 | 1.0111 | 201.1176 | 54.5035 | 72544.2276 | 71646.593 | 1428.9319 | 1328.9319 | 805.3503
1142 | 89.4009 | 0.90869 | 0.86121 | 0.99794 | 193.3691 | 48.6259 | 45934.5572 | 44090.5957 | 877.8119 | 777.8119 | 805.5165
1143 | 101.3654 | 1.0237 | 1.1911 | 1.0627 | 198.44 | 48.8397 | 67087.1187 | 69143.1538 | 1378.8631 | 1278.8631 | 805.5266
1144 | 114.758 | 1.0265 | 1.142 | 0.93639 | 217.1029 | 47.4873 | 56592.2425 | 55493.0773 | 1105.8615 | 1005.8615 | 805.6552
1145 | 100.5562 | 1.0361 | 0.84258 | 0.99552 | 199.188 | 51.3152 | 48589.3346 | 48932.6334 | 974.6527 | 874.6527 | 805.7269
1146 | 111.9559 | 1.0109 | 0.87627 | 0.89984 | 200.867 | 50.8321 | 46035.0337 | 48153.9725 | 959.0795 | 859.0795 | 805.7409
1147 | 111.2655 | 0.90489 | 0.95324 | 0.96279 | 200.0161 | 48.4956 | 49590.5895 | 50531.7858 | 1006.6357 | 906.6357 | 805.819
1148 | 92.3102 | 0.9141 | 0.87679 | 1.0649 | 202.8749 | 54.377 | 55564.3693 | 54096.9442 | 1077.9389 | 977.9389 | 806.0257
1149 | 106.327 | 0.9138 | 0.92777 | 0.9714 | 183.8258 | 48.1771 | 48283.8003 | 49202.8078 | 980.0562 | 880.0562 | 806.0576
1150 | 112.8585 | 1.0617 | 1.062 | 1.1088 | 209.3642 | 50.9601 | 66319.3371 | 66293.2717 | 1321.8654 | 1221.8654 | 806.1455
1151 | 94.9853 | 0.95873 | 1.0978 | 1.0296 | 184.6999 | 51.2705 | 62805.9389 | 61293.4982 | 1221.87 | 1121.87 | 806.1787
1152 | 97.9985 | 0.96118 | 1.1596 | 1.0006 | 191.4894 | 48.1714 | 60620.8285 | 61780.2428 | 1231.6049 | 1131.6049 | 806.206
1153 | 101.2732 | 0.97289 | 0.96338 | 1.0306 | 213.3189 | 47.7297 | 52305.7562 | 52036.6523 | 1036.733 | 936.733 | 806.2651
1154 | 104.2671 | 0.91419 | 0.97622 | 1.0204 | 190.2057 | 52.7405 | 57752.4472 | 56880.687 | 1133.6137 | 1033.6137 | 806.3033
1155 | 97.7148 | 0.97594 | 0.98656 | 0.97076 | 217.5074 | 51.392 | 54337.3704 | 53894.6453 | 1073.8929 | 973.8929 | 806.4202
1156 | 107.8475 | 0.87818 | 0.89175 | 0.88879 | 200.9343 | 50.721 | 45204.988 | 45435.6437 | 904.7129 | 804.7129 | 806.5992
1157 | 84.2596 | 1.0055 | 1.1083 | 1.0339 | 183.114 | 41.6965 | 51494.7408 | 52057.4108 | 1037.1482 | 937.1482 | 806.8545
1158 | 90.21 | 1.0897 | 0.89963 | 1.0153 | 209.8715 | 51.3216 | 52129.9827 | 49567.1091 | 987.3422 | 887.3422 | 806.9241
1159 | 95.1546 | 0.99667 | 0.91172 | 1.119 | 207.1669 | 47.7614 | 53464.3594 | 53591.0051 | 1067.8201 | 967.8201 | 807.0516
1160 | 84.4717 | 1.049 | 0.96092 | 0.98855 | 208.6201 | 53.5026 | 55772.292 | 55383.1556 | 1103.6631 | 1003.6631 | 807.1035
1161 | 102.4152 | 0.98137 | 0.8998 | 0.90362 | 214.7712 | 52.8674 | 48513.2644 | 48189.6474 | 959.7929 | 859.7929 | 807.1093
1162 | 102.4326 | 0.98263 | 1.0789 | 1.104 | 195.6391 | 51.1895 | 66319.6743 | 66929.448 | 1334.589 | 1234.589 | 807.1512
1163 | 102.8777 | 0.97622 | 0.91948 | 1.0165 | 204.7907 | 49.8577 | 51811.5249 | 51833.0475 | 1032.6609 | 932.6609 | 807.2008
1164 | 81.1255 | 0.98378 | 1.2412 | 0.95688 | 209.8259 | 47.5298 | 60454.7948 | 59769.1633 | 1191.3833 | 1091.3833 | 807.3105
1165 | 105.8926 | 0.92839 | 0.87305 | 1.0019 | 187.616 | 53.4139 | 52162.483 | 50941.4354 | 1014.8287 | 914.8287 | 807.3417
1166 | 109.2455 | 0.88783 | 1.0011 | 1.0667 | 206.6989 | 48.2929 | 56458.7084 | 56279.356 | 1121.5871 | 1021.5871 | 807.37
1167 | 94.0787 | 0.99586 | 1.0959 | 1.0437 | 204.0372 | 51.0998 | 63442.4908 | 65075.651 | 1297.513 | 1197.513 | 807.3863
1168 | 92.8761 | 1.107 | 1.006 | 0.99318 | 191.8225 | 49.2792 | 54496.2985 | 54365.8768 | 1083.3175 | 983.3175 | 807.4082
1169 | 109.1085 | 1.0057 | 0.89085 | 1.0967 | 209.7572 | 44.8792 | 48982.9081 | 48660.5071 | 969.2101 | 869.2101 | 807.4573
1170 | 90.5293 | 1.0245 | 1.1302 | 0.98594 | 211.3926 | 49.5756 | 60050.9509 | 58513.1016 | 1166.262 | 1066.262 | 807.4869
1171 | 91.8393 | 0.96232 | 0.80474 | 0.95219 | 201.3242 | 49.011 | 42088.6314 | 42124.4105 | 838.4882 | 738.4882 | 807.4908
1172 | 100.3024 | 1.0565 | 0.78606 | 0.9881 | 212.7502 | 50.4629 | 44754.9663 | 46338.9883 | 922.7798 | 822.7798 | 807.5626
1173 | 102.6588 | 0.95881 | 0.84702 | 1.0593 | 198.8541 | 50.1467 | 50129.4442 | 50476.5043 | 1005.5301 | 905.5301 | 807.6962
1174 | 106.079 | 1.0098 | 1.253 | 0.9426 | 180.0793 | 48.5427 | 62714.5295 | 63152.4701 | 1259.0494 | 1159.0494 | 807.8406
1175 | 96.7758 | 0.98483 | 0.98776 | 0.97935 | 216.0723 | 44.405 | 47403.9266 | 48248.5533 | 960.9711 | 860.9711 | 807.8601
1176 | 93.1523 | 0.98837 | 1.175 | 1.01 | 192.1541 | 49.6982 | 63750.3406 | 64441.9718 | 1284.8394 | 1184.8394 | 807.9808
1177 | 85.3948 | 1.0047 | 1.1261 | 0.96184 | 201.3623 | 46.784 | 54890.6035 | 56112.007 | 1118.2401 | 1018.2401 | 808.0809
1178 | 94.3361 | 1.0084 | 0.97551 | 0.94213 | 209.7652 | 49.5104 | 50422.4314 | 51770.4864 | 1031.4097 | 931.4097 | 808.1445
1179 | 116.0441 | 1.0508 | 1.1039 | 1.0009 | 202.5604 | 50.1881 | 61778.5706 | 63977.7363 | 1275.5547 | 1175.5547 | 808.1651
1180 | 123.9894 | 0.9593 | 0.91801 | 1.0177 | 180.9094 | 53.3795 | 56398.7488 | 57585.6329 | 1147.7127 | 1047.7127 | 808.1935
1181 | 93.3472 | 1.0252 | 1.1312 | 1.0189 | 199.2827 | 43.9286 | 55032.5804 | 56691.6158 | 1129.8323 | 1029.8323 | 808.254
1182 | 105.0386 | 1.001 | 1.1087 | 1.0312 | 202.9355 | 49.8658 | 62459.3576 | 64429.8221 | 1284.5964 | 1184.5964 | 808.2804
1183 | 102.2276 | 1.0073 | 0.89645 | 0.99306 | 214.4015 | 47.7531 | 47642.8657 | 47954.302 | 955.086 | 855.086 | 808.3863
1184 | 89.1457 | 0.95953 | 1.2192 | 0.97091 | 195.4039 | 49.8274 | 63442.0515 | 61628.7377 | 1228.5748 | 1128.5748 | 808.5386
1185 | 124.9398 | 0.91778 | 1.0054 | 0.9557 | 207.3703 | 50.4828 | 54504.7284 | 54645.1054 | 1088.9021 | 988.9021 | 808.6007
1186 | 103.4857 | 0.96679 | 0.82767 | 1.0102 | 186.7667 | 49.4742 | 46503.7469 | 45032.6406 | 896.6528 | 796.6528 | 808.7467
1187 | 103.2615 | 1.0675 | 1.1847 | 1.0552 | 216.5962 | 49.0829 | 66980.6938 | 64776.8844 | 1291.5377 | 1191.5377 | 808.7937
1188 | 105.2081 | 0.98848 | 0.98423 | 0.9856 | 208.5254 | 49.4815 | 53354.123 | 52166.7044 | 1039.3341 | 939.3341 | 808.8829
1189 | 101.3131 | 1.0546 | 1.3946 | 0.95676 | 210.2157 | 49.9511 | 72198.8907 | 70537.0915 | 1406.7418 | 1306.7418 | 808.8869
1190 | 106.5495 | 1.0117 | 0.9431 | 1.0202 | 186.2745 | 50.2266 | 53925.7246 | 55227.4262 | 1100.5485 | 1000.5485 | 809.109
1191 | 76.2238 | 1.0631 | 1.0878 | 1.0435 | 209.7595 | 50.547 | 61683.1475 | 60567.8371 | 1207.3567 | 1107.3567 | 809.2098
1192 | 103.9842 | 1.0326 | 1.3065 | 1.0232 | 198.6397 | 50.7611 | 73507.307 | 74531.9339 | 1486.6387 | 1386.6387 | 809.3586
1193 | 114.5011 | 1.0256 | 1.038 | 0.97173 | 205.7542 | 47.4364 | 53622.2234 | 53295.0901 | 1061.9018 | 961.9018 | 809.4338
1194 | 94.9951 | 0.9734 | 1.1083 | 0.99687 | 200.4456 | 51.4111 | 61754.9816 | 59814.2164 | 1192.2843 | 1092.2843 | 809.5936
1195 | 116.6378 | 1.0191 | 1.0384 | 0.94006 | 207.6368 | 46.3727 | 50988.1349 | 51711.2665 | 1030.2253 | 930.2253 | 809.7572
1196 | 97.8848 | 0.99978 | 1.0628 | 1.0319 | 212.9602 | 46.9048 | 56242.886 | 56879.4229 | 1133.5885 | 1033.5885 | 809.7879
1197 | 100.591 | 1.0045 | 1.0681 | 0.98136 | 192.0756 | 47.196 | 54431.2956 | 54311.2529 | 1082.2251 | 982.2251 | 809.8267
1198 | 95.5038 | 0.97077 | 0.97765 | 0.95935 | 204.0566 | 48.9158 | 50617.3839 | 51197.897 | 1019.9579 | 919.9579 | 809.8555
1199 | 96.7408 | 0.99616 | 1.1462 | 0.96575 | 200.9425 | 47.6967 | 57596.3417 | 57652.0238 | 1149.0405 | 1049.0405 | 809.8631
1200 | 98.31 | 1.0287 | 0.79755 | 0.96367 | 202.0318 | 53.1319 | 46411.3961 | 46232.3014 | 920.646 | 820.646 | 809.8658
1201 | 92.9248 | 0.92225 | 0.86687 | 0.99578 | 184.1345 | 49.9684 | 47599.7853 | 46606.3552 | 928.1271 | 828.1271 | 809.9567
1202 | 103.1513 | 1.0081 | 1.0311 | 1.038 | 206.8477 | 53.5261 | 63062.9129 | 62971.1216 | 1255.4224 | 1155.4224 | 810.0837
1203 | 100.8586 | 0.97436 | 1.0946 | 1.0261 | 191.3612 | 55.4972 | 67977.6159 | 67701.135 | 1350.0227 | 1250.0227 | 810.0937
1204 | 103.675 | 1.0374 | 1.1136 | 1.0898 | 189.4455 | 50.3405 | 66696.5223 | 66263.2433 | 1321.2649 | 1221.2649 | 810.123
1205 | 94.7205 | 1.007 | 0.95656 | 1.0417 | 183.4654 | 43.038 | 47174.829 | 47621.2797 | 948.4256 | 848.4256 | 810.1494
1206 | 102.999 | 0.97176 | 0.99672 | 1.0096 | 221.8542 | 49.4225 | 54901.9227 | 53311.7946 | 1062.2359 | 962.2359 | 810.3115
1207 | 97.777 | 1.0068 | 1.0465 | 1.0085 | 201.884 | 55.1854 | 63873.894 | 63146.5262 | 1258.9305 | 1158.9305 | 810.3192
1208 | 89.3057 | 0.87081 | 0.66962 | 0.98364 | 206.9764 | 55.3906 | 40998.6142 | 40575.4953 | 807.5099 | 707.5099 | 810.4172
1209 | 98.3847 | 0.99325 | 1.0387 | 0.87715 | 223.1084 | 52.5318 | 53216.024 | 53940.0023 | 1074.8 | 974.8 | 810.4601
1210 | 113.1767 | 1.0559 | 0.92067 | 1.0241 | 204.6582 | 51.8998 | 55343.1708 | 54634.0054 | 1088.6801 | 988.6801 | 810.4624
1211 | 101.213 | 1.0243 | 1.0486 | 0.8474 | 202.4551 | 50.0331 | 49846.3533 | 50238.2152 | 1000.7643 | 900.7643 | 810.4765
1212 | 108.3538 | 0.99319 | 1.0982 | 0.98259 | 208.0935 | 52.4856 | 62492.5397 | 66151.7677 | 1319.0354 | 1219.0354 | 810.5346
1213 | 88.4162 | 0.96655 | 0.88767 | 0.8867 | 196.7377 | 52.8099 | 46276.7692 | 48105.9496 | 958.119 | 858.119 | 810.5628
1214 | 100.0473 | 0.94973 | 1.0907 | 1.0443 | 192.3885 | 53.6627 | 66412.3941 | 66882.0243 | 1333.6405 | 1233.6405 | 810.5796
1215 | 108.5369 | 0.95136 | 1.1201 | 0.94437 | 190.4299 | 45.433 | 52941.0623 | 54752.7816 | 1091.0556 | 991.0556 | 810.7287
1216 | 94.7062 | 1.0729 | 1.1433 | 1.0231 | 200.1771 | 48.0108 | 61237.5614 | 58841.3718 | 1172.8274 | 1072.8274 | 810.8325
1217 | 91.749 | 1.0166 | 0.95249 | 1.0775 | 202.6645 | 51.8013 | 58196.4061 | 59364.3951 | 1183.2879 | 1083.2879 | 810.8662
1218 | 95.83 | 0.95005 | 0.94271 | 0.95191 | 212.2361 | 53.6619 | 53252.745 | 53270.8779 | 1061.4176 | 961.4176 | 810.8918
1219 | 100.7617 | 0.93557 | 1.1005 | 1.0366 | 199.5994 | 50.8235 | 62970.5138 | 63844.7774 | 1272.8955 | 1172.8955 | 810.971
1220 | 92.5961 | 1.0226 | 0.80046 | 0.94369 | 205.5026 | 47.3857 | 40486.771 | 41656.3593 | 829.1272 | 729.1272 | 810.9789
1221 | 103.5873 | 0.97242 | 0.91006 | 1.0428 | 207.0865 | 48.1189 | 50720.0228 | 51139.6832 | 1018.7937 | 918.7937 | 810.9957
1222 | 72.8197 | 1.0419 | 1.1662 | 1.066 | 221.3601 | 51.6477 | 68347.7776 | 68778.2483 | 1371.565 | 1271.565 | 811.0092
1223 | 79.5952 | 1.0898 | 0.97962 | 1.0328 | 211.7926 | 48.4273 | 53409.3833 | 51893.8893 | 1033.8778 | 933.8778 | 811.1645
1224 | 108.4216 | 1.0379 | 1.022 | 1.0304 | 181.7915 | 49.9798 | 58436.1485 | 59173.3878 | 1179.4678 | 1079.4678 | 811.1987
1225 | 105.7203 | 0.94039 | 1.1088 | 0.97117 | 181.1276 | 47.2856 | 55802.4218 | 57042.3943 | 1136.8479 | 1036.8479 | 811.2949
1226 | 105.6224 | 1.0583 | 1.217 | 0.94501 | 194.8673 | 43.7918 | 55453.4406 | 56369.0552 | 1123.3811 | 1023.3811 | 811.3024
1227 | 105.171 | 0.97447 | 1.0364 | 0.97172 | 204.2303 | 51.4453 | 57288.9386 | 57415.212 | 1144.3042 | 1044.3042 | 811.3227
1228 | 105.3713 | 1.0142 | 0.96758 | 0.9845 | 192.5612 | 51.9012 | 55179.5655 | 54809.4935 | 1092.1899 | 992.1899 | 811.3661
1229 | 88.3115 | 1.0252 | 0.81771 | 0.9872 | 190.6194 | 46.8669 | 42266.7877 | 41392.6757 | 823.8535 | 723.8535 | 811.6936
1230 | 89.2527 | 0.99663 | 0.84746 | 1.0778 | 190.705 | 49.3677 | 49674.7743 | 48581.0829 | 967.6217 | 867.6217 | 811.787
1231 | 110.4081 | 0.99489 | 1.054 | 1.035 | 204.2827 | 49.7403 | 59930.2962 | 61976.5278 | 1235.5306 | 1135.5306 | 811.839
1232 | 99.2316 | 0.95941 | 1.1104 | 1.0605 | 211.4121 | 48.7811 | 62296.0629 | 62338.8115 | 1242.7762 | 1142.7762 | 811.8637
1233 | 99.0923 | 1.016 | 0.94552 | 1.064 | 200.2849 | 52.5557 | 58366.6859 | 58906.7171 | 1174.1343 | 1074.1343 | 811.8657
1234 | 87.0409 | 1.0099 | 0.87514 | 0.93253 | 193.7461 | 50.2969 | 45662.102 | 45249.4295 | 900.9886 | 800.9886 | 811.8749
1235 | 105.2302 | 0.95777 | 1.4195 | 1.011 | 221.8615 | 50.3219 | 77507.6785 | 78051.1567 | 1557.0231 | 1457.0231 | 811.9203
1236 | 103.1749 | 0.98714 | 1.0764 | 0.9767 | 201.0786 | 51.1435 | 59175.7132 | 60414.4625 | 1204.2892 | 1104.2892 | 811.9947
1237 | 95.9277 | 1.0269 | 1.2587 | 0.97203 | 183.4098 | 49.9344 | 66197.6381 | 68136.3599 | 1358.7272 | 1258.7272 | 812.3037
1238 | 100.847 | 0.90793 | 1.1509 | 1.055 | 200.6688 | 49.4913 | 64821.7398 | 65615.0972 | 1308.3019 | 1208.3019 | 812.3485
1239 | 111.8205 | 0.97642 | 0.75479 | 1.0143 | 190.2003 | 47.4353 | 41684.9705 | 41922.3697 | 834.4474 | 734.4474 | 812.5021
1240 | 114.1727 | 0.9704 | 1.1536 | 0.9505 | 198.8491 | 46.6575 | 56525.7398 | 54695.3511 | 1089.907 | 989.907 | 812.5855
1241 | 120.2585 | 0.94141 | 1.1117 | 0.99782 | 224.2007 | 49.4117 | 60630.7796 | 62015.8339 | 1236.3167 | 1136.3167 | 812.6384
1242 | 77.0883 | 1.0595 | 1.0846 | 0.92599 | 206.5399 | 51.8899 | 56556.7854 | 57781.136 | 1151.6227 | 1051.6227 | 812.6727
1243 | 106.699 | 1.0366 | 0.82811 | 0.99577 | 203.1009 | 47.657 | 44772.606 | 43704.5091 | 870.0902 | 770.0902 | 812.7668
1244 | 84.3918 | 0.93842 | 1.1041 | 1.0443 | 189.786 | 47.5091 | 58730.7126 | 58201.5732 | 1160.0315 | 1060.0315 | 812.7972
1245 | 104.4352 | 1.0535 | 1.2198 | 0.95697 | 189.5853 | 56.8902 | 72857.311 | 73694.8809 | 1469.8976 | 1369.8976 | 812.9164
1246 | 92.8302 | 0.9421 | 0.86436 | 1.0721 | 201.8623 | 50.132 | 51043.0573 | 51009.6071 | 1016.1921 | 916.1921 | 812.9644
1247 | 95.6499 | 1.0536 | 1.0202 | 1.1057 | 200.4563 | 52.2423 | 64395.6121 | 63711.0843 | 1270.2217 | 1170.2217 | 813.0489
1248 | 106.8606 | 1.0074 | 1.0105 | 1.0137 | 209.3197 | 51.6575 | 58686.2589 | 59565.2206 | 1187.3044 | 1087.3044 | 813.0703
1249 | 114.2383 | 0.9498 | 0.83209 | 0.84303 | 206.3332 | 47.3909 | 38591.9693 | 39580.7096 | 787.6142 | 687.6142 | 813.0741
1250 | 108.7166 | 1.0885 | 1.0725 | 0.95663 | 209.9103 | 49.4223 | 56765.1701 | 57002.8988 | 1136.058 | 1036.058 | 813.1179
1251 | 98.9456 | 1.0442 | 0.75331 | 1.0227 | 198.3912 | 48.0265 | 42160.3371 | 40728.8747 | 810.5775 | 710.5775 | 813.2427
1252 | 105.5619 | 1.0222 | 1.1538 | 1.0028 | 196.6096 | 47.8127 | 60676.975 | 61526.7778 | 1226.5356 | 1126.5356 | 813.2998
1253 | 102.4218 | 0.98738 | 1.3128 | 1.0337 | 201.6142 | 48.5929 | 71057.4463 | 66189.4138 | 1319.7883 | 1219.7883 | 813.7289
1254 | 90.4409 | 1.0247 | 0.94918 | 1.0126 | 207.4506 | 53.1665 | 56236.1898 | 58482.2438 | 1165.6449 | 1065.6449 | 813.7579
1255 | 109.3677 | 0.87087 | 1.0583 | 0.97673 | 195.1822 | 47.7354 | 54082.5171 | 53820.3823 | 1072.4076 | 972.4076 | 813.836
1256 | 84.3659 | 1.0298 | 1.1233 | 0.98792 | 199.7856 | 48.1322 | 57793.0237 | 58641.9193 | 1168.8384 | 1068.8384 | 813.9577
1257 | 110.1823 | 0.96626 | 0.87993 | 1.014 | 207.1507 | 55.4513 | 55585.7523 | 54172.2684 | 1079.4454 | 979.4454 | 814.0381
1258 | 119.6117 | 1.002 | 0.97446 | 0.93233 | 197.0139 | 46.5451 | 48062.5579 | 49477.1963 | 985.5439 | 885.5439 | 814.192
1259 | 87.6386 | 0.98459 | 1.0281 | 0.97132 | 195.8005 | 50.7864 | 55292.7425 | 54731.6148 | 1090.6323 | 990.6323 | 814.3004
1260 | 99.1097 | 1.0332 | 1.126 | 1.0994 | 204.7596 | 48.0308 | 64578.0858 | 64183.3993 | 1279.668 | 1179.668 | 814.4028
1261 | 104.9188 | 1.006 | 1.0303 | 0.89122 | 206.0418 | 55.0495 | 56565.7906 | 57295.9003 | 1141.918 | 1041.918 | 814.4806
1262 | 107.1932 | 1.1254 | 0.80333 | 1.014 | 202.221 | 47.736 | 44844.6593 | 44603.1724 | 888.0634 | 788.0634 | 814.4907
1263 | 100.7681 | 1.0604 | 0.9783 | 0.95982 | 203.1368 | 51.7228 | 54297.0973 | 52802.72 | 1052.0544 | 952.0544 | 814.4973
1264 | 101.9855 | 1.0338 | 1.1797 | 0.93234 | 206.0973 | 52.8091 | 63858.6597 | 64359.002 | 1283.18 | 1183.18 | 814.5193
1265 | 98.3859 | 1.0609 | 0.84731 | 1.0007 | 189.0808 | 48.5797 | 46452.5445 | 45031.0209 | 896.6204 | 796.6204 | 814.5722
1266 | 114.263 | 1.0056 | 1.2211 | 1.1281 | 212.788 | 55.8992 | 83639.4178 | 83643.7156 | 1668.8743 | 1568.8743 | 814.7109
1267 | 110.2694 | 1.0045 | 0.92223 | 1.0212 | 195.6385 | 52.0875 | 55022.3372 | 55151.4932 | 1099.0299 | 999.0299 | 814.7723
1268 | 103.9051 | 0.98372 | 1.0588 | 1.008 | 186.8071 | 50.4728 | 59212.7397 | 56824.5552 | 1132.4911 | 1032.4911 | 814.8003
1269 | 89.8623 | 1.0159 | 0.61513 | 0.9861 | 193.4668 | 48.7391 | 34207.2559 | 34148.494 | 678.9699 | 578.9699 | 814.8803
1270 | 127.4085 | 0.95632 | 1.0861 | 0.94352 | 191.5466 | 50.476 | 58067.2935 | 55836.1489 | 1112.723 | 1012.723 | 814.9407
1271 | 92.8129 | 1.091 | 1.1695 | 0.93928 | 204.2702 | 51.149 | 61567.816 | 62986.4829 | 1255.7297 | 1155.7297 | 815.1017
1272 | 97.3159 | 1.0033 | 1.0141 | 0.98331 | 220.132 | 48.5706 | 53396.9017 | 50064.489 | 997.2898 | 897.2898 | 815.3513
1273 | 94.305 | 0.97404 | 0.85595 | 0.95737 | 205.1309 | 48.9921 | 44852.5022 | 46128.7327 | 918.5747 | 818.5747 | 815.3716
1274 | 96.5066 | 0.98653 | 1.0046 | 0.98181 | 209.2415 | 50.6574 | 54995.3084 | 54210.2504 | 1080.205 | 980.205 | 815.454
1275 | 97.345 | 0.94352 | 1.1597 | 0.99455 | 196.7784 | 53.0416 | 66243.4436 | 65146.1555 | 1298.9231 | 1198.9231 | 815.4762
1276 | 102.9251 | 0.99454 | 0.99242 | 0.92098 | 216.3505 | 46.8122 | 47794.3454 | 46380.8098 | 923.6162 | 823.6162 | 815.5618
1277 | 94.2391 | 0.98796 | 0.93184 | 1.0689 | 194.6136 | 46.5557 | 50900.2423 | 51758.4226 | 1031.1685 | 931.1685 | 815.5716
1278 | 94.0976 | 0.94621 | 1.1673 | 1.0063 | 194.6976 | 49.7331 | 63041.7853 | 62159.2229 | 1239.1845 | 1139.1845 | 815.662
1279 | 104.4354 | 0.95552 | 1.0497 | 0.99275 | 203.956 | 47.4936 | 54436.5528 | 55411.5848 | 1104.2317 | 1004.2317 | 815.6746
1280 | 109.9385 | 1.0712 | 0.97004 | 1.0126 | 190.2752 | 50.5206 | 55763.4868 | 55206.8456 | 1100.1369 | 1000.1369 | 815.7046
1281 | 95.2759 | 0.96804 | 1.1158 | 0.99707 | 208.7469 | 45.3828 | 54885.7051 | 54605.3072 | 1088.1061 | 988.1061 | 815.7049
1282 | 108.5863 | 0.97329 | 1.1449 | 1.056 | 189.6136 | 47.0158 | 62004.4926 | 62684.9491 | 1249.699 | 1149.699 | 815.9554
1283 | 107.5265 | 0.8935 | 0.99647 | 0.98913 | 200.3629 | 42.879 | 46583.1975 | 44995.8112 | 895.9162 | 795.9162 | 816.0446
1284 | 108.8686 | 1.0442 | 0.82319 | 1.0578 | 188.4699 | 51.9673 | 51346.3803 | 51301.5073 | 1022.0301 | 922.0301 | 816.0529
1285 | 94.8807 | 1.0491 | 0.8428 | 0.95387 | 198.5382 | 52.608 | 47727.5995 | 48700.066 | 970.0013 | 870.0013 | 816.2395
1286 | 99.4136 | 0.98828 | 1.0219 | 1.0253 | 200.7545 | 48.0844 | 55305.7682 | 54230.1894 | 1080.6038 | 980.6038 | 816.2401
1287 | 98.9753 | 0.94016 | 0.96667 | 1.0106 | 209.6818 | 50.7227 | 54480.7365 | 54544.9099 | 1086.8982 | 986.8982 | 816.6848
1288 | 100.2372 | 0.94954 | 0.86955 | 1.0467 | 188.7212 | 48.2532 | 48698.4183 | 49265.5676 | 981.3114 | 881.3114 | 816.7292
1289 | 90.9716 | 1.04 | 1.0029 | 0.99742 | 195.3034 | 48.8986 | 53733.515 | 53013.8141 | 1056.2763 | 956.2763 | 816.8326
1290 | 79.3134 | 0.97255 | 1.2123 | 1.0716 | 200.2805 | 48.5194 | 66973.6296 | 68498.5059 | 1365.9701 | 1265.9701 | 816.9926
1291 | 109.6594 | 0.89416 | 1.2307 | 1.1114 | 199.1697 | 51.3396 | 75449.8865 | 74952.0896 | 1495.0418 | 1395.0418 | 817.113
1292 | 94.0248 | 0.94515 | 0.95451 | 1.0027 | 196.0211 | 51.5496 | 54113.8839 | 52846.8855 | 1052.9377 | 952.9377 | 817.1132
1293 | 111.2143 | 0.9092 | 1.2847 | 0.91469 | 214.1395 | 49.0969 | 62874.4681 | 61037.2864 | 1216.7457 | 1116.7457 | 817.1168
1294 | 100.9142 | 0.99961 | 1.1431 | 0.94653 | 202.9228 | 49.5677 | 58833.3226 | 60050.7823 | 1197.0156 | 1097.0156 | 817.1269
1295 | 106.6659 | 1.0057 | 1.2228 | 0.9312 | 207.4257 | 53.6065 | 66999.2699 | 66624.1016 | 1328.482 | 1228.482 | 817.2586
1296 | 108.6023 | 0.987 | 0.90103 | 1.0969 | 192.8591 | 52.3856 | 57583.5285 | 58101.0422 | 1158.0208 | 1058.0208 | 817.3071
1297 | 95.7748 | 0.99923 | 1.0455 | 0.91837 | 201.7399 | 50.5366 | 53561.3182 | 52953.9871 | 1055.0797 | 955.0797 | 817.3662
1298 | 88.021 | 0.96463 | 0.94897 | 0.97423 | 213.8935 | 51.7282 | 52429.4115 | 54338.9658 | 1082.7793 | 982.7793 | 817.4394
1299 | 112.2843 | 1.0494 | 0.80132 | 0.99703 | 196.66 | 49.8928 | 45936.8477 | 45799.7372 | 911.9947 | 811.9947 | 817.4525
1300 | 91.3162 | 1.0658 | 0.98758 | 0.94839 | 199.6375 | 48.6309 | 50481.0418 | 51369.4268 | 1023.3885 | 923.3885 | 817.4759
1301 | 96.5218 | 1.0133 | 0.92676 | 1.069 | 178.2847 | 53.2296 | 58117.0735 | 58011.5023 | 1156.23 | 1056.23 | 817.5063
1302 | 93.4995 | 0.97236 | 1.0608 | 1.0596 | 192.2193 | 49.7015 | 60578.6503 | 59706.0956 | 1190.1219 | 1090.1219 | 817.6444
1303 | 90.4071 | 1.1177 | 0.99645 | 0.93617 | 208.34 | 50.9885 | 52924.9549 | 53218.0551 | 1060.3611 | 960.3611 | 817.6785
1304 | 100.1025 | 0.95979 | 0.81566 | 1.0217 | 178.5025 | 49.5828 | 46260.9002 | 44856.8152 | 893.1363 | 793.1363 | 817.6809
1305 | 101.1654 | 0.98768 | 0.9929 | 0.9144 | 217.16 | 49.1046 | 49705.811 | 48624.661 | 968.4932 | 868.4932 | 817.7507
1306 | 114.9517 | 0.98574 | 0.96526 | 1.0383 | 198.9907 | 51.9589 | 58163.4155 | 60261.2366 | 1201.2247 | 1101.2247 | 817.9045
1307 | 95.3897 | 0.97708 | 1.0796 | 0.95281 | 214.1476 | 50.2108 | 56544.641 | 56063.2198 | 1117.2644 | 1017.2644 | 817.9158
1308 | 110.1125 | 1.0515 | 1.0101 | 1.0528 | 186.6035 | 54.0405 | 63908.2639 | 63350.8895 | 1263.0178 | 1163.0178 | 817.9513
1309 | 102.0262 | 1.0861 | 1.1042 | 0.89867 | 192.0946 | 50.6264 | 56039.9639 | 58819.3543 | 1172.3871 | 1072.3871 | 817.9804
1310 | 108.7262 | 0.9119 | 0.87483 | 1.0296 | 211.4227 | 54.1765 | 54380.3674 | 53652.5647 | 1069.0513 | 969.0513 | 818.0235
1311 | 118.3877 | 0.98329 | 1.1611 | 1.1469 | 222.8877 | 49.1614 | 71410.4732 | 71254.0091 | 1421.0802 | 1321.0802 | 818.0946
1312 | 81.2718 | 0.98929 | 0.85821 | 1.0509 | 196.1552 | 47.5866 | 46938.6416 | 44992.2 | 895.844 | 795.844 | 818.1484
1313 | 86.8069 | 0.933 | 1.1568 | 0.92239 | 195.3519 | 45.277 | 52172.3954 | 51422.6071 | 1024.4521 | 924.4521 | 818.1531
1314 | 80.5386 | 0.88688 | 1.0966 | 1.0329 | 203.0899 | 48.7952 | 58958.8856 | 58899.3916 | 1173.9878 | 1073.9878 | 818.363
1315 | 101.2765 | 1.0135 | 0.96487 | 1.0512 | 203.2844 | 48.007 | 53822.6299 | 54318.1588 | 1082.3632 | 982.3632 | 818.4032
1316 | 102.9026 | 0.98283 | 0.76719 | 0.96785 | 202.6352 | 48.1406 | 40817.0312 | 41446.9169 | 824.9383 | 724.9383 | 818.5747
1317 | 99.1182 | 1.0211 | 0.88867 | 1.02 | 208.9302 | 49.4541 | 50043.4636 | 49899.8383 | 993.9968 | 893.9968 | 818.6674
1318 | 96.0887 | 1.0161 | 1.033 | 1.0597 | 205.4067 | 49.5272 | 59258.6774 | 60247.3347 | 1200.9467 | 1100.9467 | 818.6705
1319 | 99.4914 | 1.0245 | 1.011 | 0.98096 | 198.63 | 48.3612 | 53087.9314 | 51538.5762 | 1026.7715 | 926.7715 | 818.7248
1320 | 86.9658 | 0.96678 | 0.8777 | 1.0152 | 212.8091 | 47.3419 | 46377.2303 | 44682.8909 | 889.6578 | 789.6578 | 818.8867
1321 | 97.2367 | 0.99043 | 1.1011 | 0.99781 | 195.965 | 49.1638 | 58945.1269 | 59457.019 | 1185.1404 | 1085.1404 | 818.9637
1322 | 98.2625 | 0.99047 | 1.2092 | 0.98067 | 192.639 | 45.5801 | 58677.8558 | 59274.0251 | 1181.4805 | 1081.4805 | 818.9695
1323 | 118.3752 | 0.95512 | 0.75167 | 1.0007 | 198.498 | 49.4785 | 43010.1561 | 42991.3607 | 855.8272 | 755.8272 | 819.0202
1324 | 102.4531 | 1.1019 | 1.1794 | 1.0749 | 219.4347 | 49.9565 | 69193.6656 | 67674.4043 | 1349.4881 | 1249.4881 | 819.0608
1325 | 99.2435 | 1.0132 | 0.8602 | 0.94945 | 216.8439 | 47.761 | 44026.4036 | 45082.2674 | 897.6453 | 797.6453 | 819.0615
1326 | 104.3236 | 1.0015 | 0.87524 | 0.9995 | 201.7402 | 49.2549 | 48436.1175 | 47431.1473 | 944.6229 | 844.6229 | 819.0867
1327 | 112.8324 | 1.0342 | 0.98077 | 0.99186 | 201.9496 | 47.0039 | 51411.5752 | 50160.1541 | 999.2031 | 899.2031 | 819.1003
1328 | 99.5027 | 0.9837 | 0.93548 | 1.0529 | 183.1082 | 49.4451 | 53723.5463 | 54265.8654 | 1081.3173 | 981.3173 | 819.1056
1329 | 97.001 | 1.0408 | 0.65084 | 0.99349 | 196.8013 | 46.7963 | 35180.0007 | 34964.8788 | 695.2976 | 595.2976 | 819.1546
1330 | 86.6066 | 0.95382 | 0.82063 | 1.0832 | 205.6084 | 53.7117 | 52389.3713 | 49674.5618 | 989.4912 | 889.4912 | 819.2528
1331 | 90.3769 | 0.98295 | 1.3132 | 0.90996 | 204.9587 | 50.0083 | 64406.1446 | 63370.3986 | 1263.408 | 1163.408 | 819.4896
1332 | 86.7928 | 0.94077 | 0.91326 | 0.97788 | 202.2938 | 53.2665 | 52121.7053 | 53137.2951 | 1058.7459 | 958.7459 | 819.539
1333 | 103.5867 | 1.0277 | 0.90785 | 0.99667 | 203.1095 | 47.1866 | 47922.1958 | 48885.8865 | 973.7177 | 873.7177 | 819.6185
1334 | 115.8121 | 1.0438 | 1.1407 | 0.96089 | 208.1671 | 51.6251 | 63032.8099 | 63251.1528 | 1261.0231 | 1161.0231 | 819.6578
1335 | 105.9265 | 0.90851 | 1.2325 | 0.94887 | 217.5122 | 50.6098 | 64273.3611 | 65086.2494 | 1297.725 | 1197.725 | 819.6601
1336 | 101.9271 | 0.94689 | 0.99986 | 0.98748 | 204.091 | 48.6377 | 52920.2245 | 54345.5756 | 1082.9115 | 982.9115 | 819.6812
1337 | 105.2934 | 0.88007 | 1.0411 | 1.0222 | 203.2804 | 55.6938 | 64631.7267 | 65387.7997 | 1303.756 | 1203.756 | 819.9339
1338 | 80.0997 | 1.0338 | 1.2044 | 1.0048 | 202.5587 | 48.2727 | 62618.627 | 63162.0752 | 1259.2415 | 1159.2415 | 819.9512
1339 | 94.2283 | 1.1049 | 1.1399 | 1.0219 | 216.4351 | 50.3818 | 64151.1034 | 65125.158 | 1298.5032 | 1198.5032 | 819.9592
1340 | 102.0785 | 0.98284 | 1.1485 | 0.94392 | 172.7975 | 48.7492 | 57912.9905 | 56013.6865 | 1116.2737 | 1016.2737 | 820.0055
1341 | 101.8833 | 1.0175 | 0.90144 | 0.99608 | 197.2304 | 47.6821 | 47954.3826 | 48365.9248 | 963.3185 | 863.3185 | 820.0822
1342 | 103.9909 | 0.93291 | 1.0216 | 1.0299 | 220.4772 | 53.3756 | 61554.7137 | 60910.9967 | 1214.2199 | 1114.2199 | 820.1307
1343 | 104.3265 | 0.98808 | 0.9061 | 0.98538 | 204.633 | 46.8714 | 46885.5111 | 45985.2302 | 915.7046 | 815.7046 | 820.2827
1344 | 86.7763 | 1.0172 | 1.3608 | 1.0227 | 195.5824 | 49.7969 | 73894.4535 | 74221.0304 | 1480.4206 | 1380.4206 | 820.2912
1345 | 108.4007 | 0.96333 | 1.1996 | 1.0353 | 188.746 | 53.5449 | 72279.2015 | 74101.0217 | 1478.0204 | 1378.0204 | 820.3324
1346 | 95.7902 | 1.067 | 0.94308 | 1.0351 | 215.8947 | 51.3661 | 55606.9474 | 54572.7558 | 1087.4551 | 987.4551 | 820.3415
1347 | 82.7482 | 0.97084 | 0.94134 | 1.0176 | 196.4839 | 48.5924 | 50647.0081 | 50825.8878 | 1012.5178 | 912.5178 | 820.3519
1348 | 98.4427 | 1.0109 | 0.99917 | 1.0341 | 205.7316 | 43.3394 | 49298.5841 | 49018.9228 | 976.3785 | 876.3785 | 820.4294
1349 | 94.9523 | 1.0064 | 0.95936 | 0.97765 | 191.3479 | 49.6107 | 51463.0165 | 51868.2742 | 1033.3655 | 933.3655 | 820.4895
1350 | 107.5009 | 0.99955 | 0.96789 | 1.0261 | 207.1792 | 51.3462 | 56716.7192 | 56422.6842 | 1124.4537 | 1024.4537 | 820.6448
1351 | 98.369 | 1.0635 | 1.2668 | 1.021 | 213.0651 | 48.048 | 67387.6766 | 69753.0063 | 1391.0601 | 1291.0601 | 820.646
1352 | 92.2374 | 1.0749 | 0.7026 | 0.95822 | 198.1874 | 45.192 | 35103.9776 | 34828.0126 | 692.5603 | 592.5603 | 820.6479
1353 | 110.127 | 0.99438 | 1.0685 | 1.005 | 215.5685 | 48.7275 | 57877.6911 | 58114.4561 | 1158.2891 | 1058.2891 | 820.7292
1354 | 104.2827 | 0.97455 | 1.0758 | 0.94286 | 206.9305 | 49.1821 | 55092.1059 | 55088.6905 | 1097.7738 | 997.7738 | 820.8494
1355 | 106.3969 | 1.0115 | 0.81295 | 0.9592 | 198.1621 | 50.0921 | 44649.701 | 44154.1052 | 879.0821 | 779.0821 | 820.8964
1356 | 109.2933 | 0.99317 | 1.0765 | 1.0019 | 199.7192 | 47.1189 | 56133.3361 | 56941.2448 | 1134.8249 | 1034.8249 | 820.91
1357 | 86.8539 | 1.0833 | 1.2696 | 1.0044 | 224.7748 | 51.0015 | 70056.0852 | 67255.6937 | 1341.1139 | 1241.1139 | 820.9548
1358 | 91.5162 | 1.0204 | 0.82821 | 1.0061 | 208.385 | 47.8308 | 44529.4881 | 44225.4529 | 880.5091 | 780.5091 | 820.9585
1359 | 95.3666 | 1.0858 | 1.3212 | 0.94783 | 192.3556 | 52.5749 | 71472.69 | 70345.5106 | 1402.9102 | 1302.9102 | 821.0223
1360 | 115.5459 | 0.98728 | 1.276 | 0.92977 | 204.8506 | 52.8623 | 68949.0584 | 70274.817 | 1401.4963 | 1301.4963 | 821.092
1361 | 105.8487 | 1.031 | 1.083 | 1.0252 | 205.6972 | 46.7111 | 57170.0365 | 56650.8718 | 1129.0174 | 1029.0174 | 821.1154
1362 | 109.1596 | 1.0328 | 0.88818 | 1.068 | 207.9676 | 47.0254 | 50114.7996 | 50217.0525 | 1000.3411 | 900.3411 | 821.1206
1363 | 85.1079 | 1.0434 | 1.178 | 0.8987 | 198.1168 | 50.1357 | 57727.4317 | 59176.445 | 1179.5289 | 1079.5289 | 821.1671
1364 | 108.4567 | 0.99019 | 0.98559 | 1.0297 | 201.065 | 55.1639 | 62108.7318 | 61926.801 | 1234.536 | 1134.536 | 821.1851
1365 | 97.8968 | 0.99665 | 1.0435 | 0.91718 | 190.5254 | 49.3085 | 52191.632 | 52011.0349 | 1036.2207 | 936.2207 | 821.1893
1366 | 98.1711 | 0.95541 | 1.0737 | 1.1141 | 209.4198 | 53.698 | 69482.6691 | 69975.8615 | 1395.5172 | 1295.5172 | 821.2039
1367 | 101.2124 | 1.0401 | 0.91417 | 0.96266 | 204.6552 | 50.5749 | 50036.1836 | 50372.4372 | 1003.4487 | 903.4487 | 821.3739
1368 | 87.5381 | 1.0258 | 0.83886 | 1.0016 | 207.8222 | 52.3249 | 48868.5681 | 49114.0802 | 978.2816 | 878.2816 | 821.388
1369 | 101.4313 | 0.98559 | 1.0647 | 0.91735 | 211.8188 | 47.6823 | 51548.0298 | 51182.0728 | 1019.6415 | 919.6415 | 821.4497
1370 | 108.8885 | 1.1134 | 1.0519 | 0.99759 | 195.52 | 49.949 | 58668.2753 | 60067.649 | 1197.353 | 1097.353 | 821.4998
1371 | 124.813 | 0.98285 | 0.90922 | 1.0095 | 208.9817 | 47.0415 | 49156.473 | 50199.6548 | 999.9931 | 899.9931 | 821.511
1372 | 84.6123 | 1.0072 | 1.0467 | 0.93721 | 207.3601 | 55.222 | 59082.4373 | 59095.994 | 1177.9199 | 1077.9199 | 821.5936
1373 | 93.0593 | 0.96658 | 1.1502 | 1.0444 | 184.8754 | 51.8303 | 67108.9346 | 67559.8417 | 1347.1968 | 1247.1968 | 821.6817
1374 | 112.4486 | 1.0396 | 0.97687 | 1.0041 | 190.9526 | 48.9996 | 53982.6945 | 55525.2377 | 1106.5048 | 1006.5048 | 821.7119
1375 | 100.1217 | 1.0143 | 0.90492 | 1.0394 | 203.3441 | 54.0392 | 56517.2455 | 58039.0662 | 1156.7813 | 1056.7813 | 821.7836
1376 | 102.8051 | 1.0144 | 1.1336 | 1.0102 | 188.2777 | 49.6663 | 62243.9623 | 61934.3365 | 1234.6867 | 1134.6867 | 821.8545
1377 | 106.22 | 0.87886 | 0.92338 | 1.0224 | 187.8733 | 54.6503 | 56885.1346 | 54860.256 | 1093.2051 | 993.2051 | 821.9645
1378 | 91.0425 | 0.89175 | 1.2178 | 1.0329 | 208.3627 | 49.7634 | 66846.3908 | 65938.8731 | 1314.7775 | 1214.7775 | 821.993
1379 | 113.3002 | 1.0547 | 0.9844 | 1.0892 | 188.6341 | 53.0768 | 63439.8239 | 63595.036 | 1267.9007 | 1167.9007 | 821.9942
1380 | 91.0096 | 0.96935 | 0.74765 | 1.0486 | 201.6928 | 45.8508 | 40193.2875 | 39138.4523 | 778.769 | 678.769 | 822.003
1381 | 104.9719 | 0.94687 | 0.99551 | 0.95371 | 195.086 | 51.2887 | 53987.9046 | 54765.7004 | 1091.314 | 991.314 | 822.024
1382 | 81.0778 | 0.96629 | 0.90778 | 0.94771 | 197.8607 | 46.9143 | 44234.1768 | 43912.171 | 874.2434 | 774.2434 | 822.2295
1383 | 107.7078 | 1.0768 | 0.87012 | 1.012 | 203.6164 | 53.3559 | 53373.7602 | 52652.967 | 1049.0593 | 949.0593 | 822.3217
1384 | 116.6911 | 0.99002 | 1.3343 | 0.87452 | 209.8179 | 50.1063 | 64465.8377 | 66016.8302 | 1316.3366 | 1216.3366 | 822.4207
1385 | 101.3402 | 1.0045 | 1.1294 | 1.0665 | 197.7613 | 44.5086 | 58341.1603 | 58190.9813 | 1159.8196 | 1059.8196 | 822.4248
1386 | 99.2422 | 1.0347 | 0.8699 | 0.96686 | 196.9506 | 49.3534 | 46774.5269 | 46546.9578 | 926.9392 | 826.9392 | 822.4766
1387 | 84.9285 | 0.97329 | 0.84177 | 0.99502 | 198.0309 | 50.8029 | 46948.9609 | 47282.1364 | 941.6427 | 841.6427 | 822.5697
1388 | 101.2966 | 1.062 | 1.0441 | 0.95289 | 172.2595 | 49.9773 | 55272.028 | 55495.6407 | 1105.9128 | 1005.9128 | 822.5895
1389 | 93.2673 | 1.0223 | 0.99546 | 0.91689 | 207.7939 | 50.1081 | 50720.1244 | 51577.1863 | 1027.5437 | 927.5437 | 822.7798
1390 | 108.3869 | 1.007 | 0.8806 | 0.98235 | 196.1964 | 52.4716 | 51314.3692 | 51788.4817 | 1031.7696 | 931.7696 | 822.8119
1391 | 134.972 | 0.97381 | 0.95268 | 0.93353 | 186.623 | 48.4433 | 49637.351 | 49783.2698 | 991.6654 | 891.6654 | 822.825
1392 | 91.816 | 0.97548 | 0.89202 | 0.92224 | 202.8726 | 52.164 | 47787.9269 | 48538.6514 | 966.773 | 866.773 | 822.8351
1393 | 83.4006 | 1.002 | 0.90369 | 1.0366 | 206.1016 | 49.8737 | 51094.8644 | 51697.27 | 1029.9454 | 929.9454 | 822.8705
1394 | 108.1657 | 0.98884 | 0.96405 | 1.0468 | 199.0973 | 49.5217 | 55471.054 | 55851.6689 | 1113.0334 | 1013.0334 | 822.9228
1395 | 95.2044 | 1.0037 | 0.77006 | 1.0866 | 211.2355 | 50.3529 | 47155.4764 | 48315.9654 | 962.3193 | 862.3193 | 822.9271
1396 | 98.025 | 1.0033 | 1.0701 | 1.017 | 209.8659 | 53.1482 | 63280.1485 | 65803.7929 | 1312.0759 | 1212.0759 | 822.9332
1397 | 104.474 | 0.90034 | 1.2668 | 0.95121 | 207.0033 | 47.6294 | 62080.5035 | 61215.4919 | 1220.3098 | 1120.3098 | 823.0882
1398 | 99.2017 | 0.97703 | 0.38848 | 1.073 | 198.9942 | 52.6896 | 27268.4993 | 26978.0003 | 535.56 | 435.56 | 823.0942
1399 | 84.9752 | 1.0708 | 0.91414 | 0.9589 | 203.783 | 46.8355 | 45519.8536 | 46260.1962 | 921.2039 | 821.2039 | 823.1851
1400 | 92.1795 | 1.0134 | 0.76047 | 0.96205 | 204.7358 | 49.4717 | 41020.0789 | 40159.8176 | 799.1964 | 699.1964 | 823.2252
1401 | 112.2448 | 0.95761 | 0.78903 | 1.0977 | 203.1296 | 55.9045 | 54631.294 | 55521.0562 | 1106.4211 | 1006.4211 | 823.4413
1402 | 100.4995 | 1.0759 | 0.87856 | 1.033 | 202.9365 | 47.8583 | 48812.5095 | 47219.8827 | 940.3977 | 840.3977 | 823.4963
1403 | 95.8218 | 0.97609 | 0.81845 | 1.0422 | 203.189 | 45.6224 | 43385.8305 | 44534.6614 | 886.6932 | 786.6932 | 823.6162
1404 | 110.2035 | 0.94111 | 1.0402 | 1.0079 | 198.0579 | 50.2224 | 58063.6604 | 58387.3694 | 1163.7474 | 1063.7474 | 823.7075
1405 | 111.4453 | 1.0121 | 0.99829 | 0.97109 | 180.0317 | 45.6607 | 49595.3668 | 47863.3304 | 953.2666 | 853.2666 | 823.7183
1406 | 115.2848 | 1.0908 | 0.82579 | 1.0908 | 195.1279 | 46.885 | 48323.138 | 48236.8562 | 960.7371 | 860.7371 | 823.7623
1407 | 103.2764 | 0.95355 | 0.98826 | 1.0261 | 203.5727 | 51.469 | 57465.8412 | 57665.119 | 1149.3024 | 1049.3024 | 823.9213
1408 | 107.87 | 1.0429 | 1.1023 | 0.98534 | 186.7806 | 46.1852 | 55547.0046 | 56041.0005 | 1116.82 | 1016.82 | 823.9979
1409 | 91.4279 | 0.9886 | 0.7951 | 1.1072 | 190.7672 | 52.5473 | 51200.9562 | 50689.4916 | 1009.7898 | 909.7898 | 824.021
1410 | 102.1022 | 0.94087 | 0.78328 | 0.99105 | 198.847 | 45.3623 | 39770.2378 | 39454.005 | 785.0801 | 685.0801 | 824.099
1411 | 75.1227 | 0.97404 | 1.0974 | 0.99885 | 215.3775 | 50.7591 | 59566.9653 | 57463.5558 | 1145.2711 | 1045.2711 | 824.2207
1412 | 93.8424 | 0.91308 | 0.90071 | 0.92902 | 198.9144 | 49.4674 | 45830.7076 | 45568.4983 | 907.37 | 807.37 | 824.297
1413 | 109.8897 | 0.92974 | 0.90592 | 0.91371 | 212.7648 | 51.452 | 48058.9133 | 48543.7027 | 966.8741 | 866.8741 | 824.3666
1414 | 79.9631 | 0.99992 | 0.95286 | 1.0695 | 215.9744 | 49.4074 | 54517.2932 | 55475.2256 | 1105.5045 | 1005.5045 | 824.4129
1415 | 103.501 | 1.0225 | 0.63389 | 1.025 | 202.2577 | 50.005 | 37984.0554 | 37386.5495 | 743.731 | 643.731 | 824.561
1416 | 90.7849 | 0.93658 | 1.1609 | 1.068 | 196.6125 | 50.5155 | 67121.4879 | 67649.1409 | 1348.9828 | 1248.9828 | 824.6055
1417 | 92.7122 | 1.0306 | 0.8263 | 1.0146 | 189.7526 | 50.7477 | 47583.2425 | 48595.7939 | 967.9159 | 867.9159 | 824.6454
1418 | 109.7185 | 1.0083 | 1.129 | 1.0484 | 210.0605 | 50.2429 | 65239.645 | 63071.4958 | 1257.4299 | 1157.4299 | 824.6857
1419 | 108.3909 | 0.93453 | 0.95963 | 1.0319 | 197.9051 | 49.768 | 54519.238 | 53862.7009 | 1073.254 | 973.254 | 824.7096
1420 | 108.4168 | 1.0878 | 0.94387 | 0.96734 | 192.9102 | 49.6773 | 51409.2977 | 52327.273 | 1042.5455 | 942.5455 | 824.7421
1421 | 99.142 | 1.0028 | 0.88227 | 0.89593 | 208.8898 | 49.9812 | 44686.118 | 45191.7611 | 899.8352 | 799.8352 | 824.79
1422 | 112.5816 | 0.99707 | 0.88845 | 1.0049 | 184.7304 | 52.0525 | 52502.4752 | 53174.8972 | 1059.4979 | 959.4979 | 824.8742
1423 | 105.6364 | 1.0425 | 0.87361 | 0.98576 | 204.0038 | 47.54 | 46379.6122 | 45311.9841 | 902.2397 | 802.2397 | 825.0403
1424 | 93.0191 | 0.97692 | 0.94293 | 1.0509 | 198.4826 | 51.2156 | 55601.4974 | 57346.1642 | 1142.9233 | 1042.9233 | 825.0704
1425 | 111.1122 | 0.97268 | 0.93707 | 0.9797 | 194.4562 | 49.2053 | 50685.1457 | 50917.24 | 1014.3448 | 914.3448 | 825.309
1426 | 93.7409 | 0.95114 | 0.91251 | 0.89483 | 198.2391 | 50.1701 | 45637.8413 | 44147.8932 | 878.9579 | 778.9579 | 825.396
1427 | 91.2755 | 0.95983 | 1.1316 | 0.98783 | 199.5944 | 52.5842 | 63585.8993 | 63231.0139 | 1260.6203 | 1160.6203 | 825.512
1428 | 86.4235 | 1.015 | 1.0471 | 0.95188 | 205.1352 | 46.1984 | 50303.8432 | 48454.6746 | 965.0935 | 865.0935 | 825.5481
1429 | 97.5342 | 0.98013 | 1.03 | 1.067 | 191.7294 | 49.2278 | 59001.2335 | 59637.3828 | 1188.7477 | 1088.7477 | 825.5504
1430 | 103.4728 | 1.0539 | 1.1246 | 1.0694 | 224.7041 | 52.0233 | 68467.3424 | 70122.3287 | 1398.4466 | 1298.4466 | 825.554
1431 | 112.2389 | 0.964 | 1.144 | 1.0081 | 187.2773 | 50.5406 | 63942.2352 | 63638.9132 | 1268.7783 | 1168.7783 | 825.6006
1432 | 74.4701 | 1.0023 | 1.0569 | 0.98856 | 204.4296 | 50.1634 | 56360.9323 | 57174.5582 | 1139.4912 | 1039.4912 | 825.625
1433 | 102.7198 | 1.0939 | 1.1135 | 0.99565 | 183.5507 | 49.0017 | 60016.4849 | 61978.3447 | 1235.5669 | 1135.5669 | 825.6523
1434 | 104.3571 | 1.0569 | 0.94673 | 1.0137 | 197.9922 | 54.0881 | 58070.9231 | 57573.2021 | 1147.464 | 1047.464 | 825.7079
1435 | 89.0183 | 1.0778 | 0.92012 | 1.0196 | 181.0261 | 49.3915 | 51254.8285 | 52240.4689 | 1040.8094 | 940.8094 | 825.7734
1436 | 99.1485 | 1.0495 | 1.1322 | 0.89142 | 191.65 | 45.0463 | 50343.141 | 50089.4507 | 997.789 | 897.789 | 825.7771
1437 | 103.7494 | 1.0017 | 1.0398 | 0.96969 | 199.6798 | 52.1015 | 58148.365 | 57695.069 | 1149.9014 | 1049.9014 | 825.7791
1438 | 90.5208 | 0.99832 | 0.89068 | 0.9971 | 198.7741 | 51.9862 | 51065.2268 | 49808.4873 | 992.1697 | 892.1697 | 825.9039
1439 | 108.5857 | 0.94392 | 1.1627 | 0.92202 | 198.8969 | 49.4803 | 58316.3132 | 59825.4242 | 1192.5085 | 1092.5085 | 825.9273
1440 | 101.4447 | 0.99619 | 0.76869 | 1.084 | 205.4517 | 49.4554 | 46413.1691 | 48461.0657 | 965.2213 | 865.2213 | 826.1438
1441 | 102.6668 | 1.053 | 1.1118 | 1.0419 | 191.7124 | 53.097 | 67438.8278 | 68877.3362 | 1373.5467 | 1273.5467 | 826.1732
1442 | 107.9372 | 1.0218 | 0.99129 | 0.9695 | 200.9468 | 57.1898 | 61470.9464 | 63334.7206 | 1262.6944 | 1162.6944 | 826.2987
1443 | 113.3752 | 1.0125 | 0.9766 | 0.96204 | 197.8729 | 47.5544 | 50334.9737 | 48997.2668 | 975.9453 | 875.9453 | 826.3247
1444 | 106.3215 | 1.0405 | 1.0612 | 0.95247 | 206.7968 | 52.7242 | 59331.8925 | 61426.2802 | 1224.5256 | 1124.5256 | 826.3412
1445 | 106.6274 | 1.0147 | 0.97076 | 0.92992 | 192.4117 | 51.1185 | 51869.4196 | 53315.549 | 1062.311 | 962.311 | 826.4988
1446 | 107.7915 | 1.0969 | 1.1648 | 0.97209 | 217.2005 | 48.4638 | 60822.0193 | 60671.5411 | 1209.4308 | 1109.4308 | 826.5015
1447 | 102.6204 | 1.0536 | 1.0021 | 0.94731 | 218.3896 | 50.9561 | 54099.4428 | 53407.4955 | 1064.1499 | 964.1499 | 826.5688
1448 | 89.6331 | 0.97476 | 0.85528 | 0.98092 | 193.075 | 47.6543 | 44336.8557 | 43704.5436 | 870.0909 | 770.0909 | 826.616
1449 | 105.009 | 0.92418 | 1.1091 | 0.98393 | 210.5837 | 49.7889 | 59376.827 | 59534.3991 | 1186.688 | 1086.688 | 826.7269
1450 | 102.621 | 0.94014 | 1.0747 | 1.0038 | 201.0539 | 50.3115 | 59330.1096 | 57318.3745 | 1142.3675 | 1042.3675 | 826.7425
1451 | 100.6586 | 0.97537 | 0.87803 | 0.96802 | 211.9472 | 49.2907 | 46946.071 | 45853.5169 | 913.0703 | 813.0703 | 826.7501
1452 | 84.3977 | 1.0286 | 1.0287 | 0.95666 | 218.9008 | 52.1185 | 56033.1105 | 53959.5715 | 1075.1914 | 975.1914 | 826.9392
1453 | 109.4207 | 1.0004 | 0.97614 | 1.0882 | 191.7277 | 49.2364 | 57880.8032 | 58661.1969 | 1169.2239 | 1069.2239 | 827.0046
1454 | 97.3613 | 0.95337 | 1.1625 | 0.96743 | 185.7646 | 48.5083 | 59240.7353 | 59871.9665 | 1193.4393 | 1093.4393 | 827.0875
1455 | 101.9422 | 0.96396 | 0.72117 | 0.96347 | 199.6973 | 50.6928 | 40403.695 | 39922.183 | 794.4437 | 694.4437 | 827.1276
1456 | 100.3945 | 0.96963 | 0.92531 | 1.0324 | 206.3454 | 46.5255 | 49179.3934 | 47962.3106 | 955.2462 | 855.2462 | 827.1276
1457 | 108.0549 | 0.95827 | 0.95149 | 1.018 | 195.605 | 47.5284 | 51156.2423 | 52860.7547 | 1053.2151 | 953.2151 | 827.1789
1458 | 89.1624 | 0.84958 | 0.89409 | 1.0223 | 205.0625 | 45.8668 | 45603.2262 | 44769.8012 | 891.396 | 791.396 | 827.2111
1459 | 94.1232 | 0.99811 | 0.80324 | 1.0394 | 218.4218 | 51.57 | 48118.2468 | 49178.8307 | 979.5766 | 879.5766 | 827.23
1460 | 86.7236 | 0.94021 | 0.95335 | 1.0329 | 192.2068 | 57.2866 | 61276.9237 | 62072.2367 | 1237.4447 | 1137.4447 | 827.259
1461 | 124.7315 | 0.98842 | 0.9971 | 0.91514 | 194.2764 | 51.0728 | 53093.99 | 53527.2071 | 1066.5441 | 966.5441 | 827.2647
1462 | 107.2945 | 0.94384 | 1.0514 | 0.97781 | 215.2685 | 51.8906 | 58816.8146 | 58053.1776 | 1157.0636 | 1057.0636 | 827.2929
1463 | 101.8326 | 1.0507 | 1.0395 | 1.0224 | 196.3435 | 54.9965 | 64530.0898 | 63303.904 | 1262.0781 | 1162.0781 | 827.4161
1464 | 115.5481 | 0.93858 | 0.72927 | 1.027 | 204.2967 | 53.3158 | 45917.6886 | 46374.8155 | 923.4963 | 823.4963 | 827.4416
1465 | 111.3932 | 1.0549 | 1.0927 | 1.0322 | 198.0076 | 48.7548 | 60915.8371 | 59077.3379 | 1177.5468 | 1077.5468 | 827.4622
1466 | 116.9719 | 1.0316 | 1.0967 | 1.0555 | 193.5828 | 50.3438 | 64543.6874 | 63924.848 | 1274.497 | 1174.497 | 827.5127
1467 | 106.3829 | 0.9373 | 1.13 | 1.0137 | 196.4474 | 50.5493 | 63140.3341 | 64694.8065 | 1289.8961 | 1189.8961 | 827.5236
1468 | 94.9302 | 0.9308 | 1.0828 | 0.93418 | 182.156 | 53.1322 | 58621.3283 | 57403.2366 | 1144.0647 | 1044.0647 | 827.5389
1469 | 98.0176 | 1.0683 | 1.036 | 0.97804 | 207.1497 | 50.9267 | 57139.5451 | 56748.0904 | 1130.9618 | 1030.9618 | 827.5758
1470 | 114.7678 | 0.98551 | 1.1472 | 0.92165 | 201.7081 | 50.2817 | 59052.036 | 59215.0766 | 1180.3015 | 1080.3015 | 827.6316
1471 | 87.2716 | 0.96056 | 1.0095 | 1.0239 | 193.2087 | 50.1058 | 56181.8728 | 58052.416 | 1157.0483 | 1057.0483 | 827.7679
1472 | 98.5663 | 1.0057 | 0.89987 | 1.0118 | 192.1157 | 51.7503 | 52442.5075 | 52264.3399 | 1041.2868 | 941.2868 | 827.8304
1473 | 97.7424 | 1.0912 | 1.0206 | 0.99003 | 186.8064 | 50.3214 | 56397.9938 | 56184.2484 | 1119.685 | 1019.685 | 827.8431
1474 | 111.9222 | 1.0715 | 1.064 | 0.95674 | 206.5277 | 50.0684 | 57178.1478 | 56168.9553 | 1119.3791 | 1019.3791 | 827.8454
1475 | 96.9859 | 1.0301 | 1.126 | 1.0424 | 214.8665 | 51.3598 | 65630.7933 | 64358.1153 | 1283.1623 | 1183.1623 | 827.8942
1476 | 101.296 | 0.9049 | 1.5301 | 1.0253 | 212.1661 | 47.2065 | 78593.9325 | 81442.1776 | 1624.8436 | 1524.8436 | 827.9311
1477 | 94.5546 | 0.98982 | 0.79687 | 0.96108 | 199.8323 | 46.6264 | 40272.9422 | 39330.0326 | 782.6007 | 682.6007 | 827.9774
1478 | 76.7235 | 0.95576 | 0.85999 | 1.012 | 217.3851 | 49.0244 | 46480.2801 | 45028.8757 | 896.5775 | 796.5775 | 827.9789
1479 | 108.8913 | 0.96795 | 1.0237 | 0.95688 | 193.7291 | 49.8365 | 54264.2763 | 53492.9359 | 1065.8587 | 965.8587 | 828.0806
1480 | 73.1906 | 1.049 | 0.91759 | 1.0172 | 209.4357 | 52.6787 | 53424.0912 | 53216.356 | 1060.3271 | 960.3271 | 828.1271
1481 | 96.3571 | 1.0396 | 0.68889 | 0.98632 | 187.844 | 51.7117 | 40504.3983 | 40454.0819 | 805.0816 | 705.0816 | 828.2145
1482 | 103.8142 | 0.98767 | 1.1684 | 0.9511 | 197.5446 | 49.9258 | 60796.7006 | 61486.6887 | 1225.7338 | 1125.7338 | 828.3345
1483 | 91.7862 | 0.99213 | 0.85629 | 1.0339 | 192.8503 | 48.9954 | 48030.601 | 48775.1787 | 971.5036 | 871.5036 | 828.3808
1484 | 108.2923 | 0.90791 | 1.0999 | 1.0285 | 194.1028 | 47.2589 | 58297.81 | 58204.6493 | 1160.093 | 1060.093 | 828.4214
1485 | 103.7096 | 0.94491 | 1.1178 | 0.99532 | 199.6082 | 51.1439 | 62112.1147 | 61876.965 | 1233.5393 | 1133.5393 | 828.4326
1486 | 81.8948 | 0.964 | 0.78256 | 1.0191 | 202.9068 | 51.4462 | 45294.4708 | 45546.203 | 906.9241 | 806.9241 | 828.4592
1487 | 104.6998 | 0.91948 | 1.0027 | 0.93173 | 204.0437 | 49.3915 | 51103.9319 | 50749.2087 | 1010.9842 | 910.9842 | 828.5776
1488 | 83.243 | 0.97683 | 1.0728 | 1.01 | 212.718 | 54.0391 | 63160.589 | 64049.3924 | 1276.9878 | 1176.9878 | 828.6601
1489 | 79.8917 | 0.99079 | 0.9464 | 0.99849 | 207.9377 | 48.0986 | 49466.8033 | 51525.7262 | 1026.5145 | 926.5145 | 828.6756
1490 | 99.8911 | 0.94482 | 0.88164 | 0.95592 | 185.59 | 50.5564 | 47564.3399 | 46653.0256 | 929.0605 | 829.0605 | 828.6768
1491 | 94.3653 | 1.0009 | 0.8448 | 0.98578 | 186.9305 | 50.0316 | 46577.9671 | 46945.9151 | 934.9183 | 834.9183 | 828.6869
1492 | 104.2014 | 1.0037 | 0.89743 | 1.0329 | 197.1488 | 51.7108 | 53537.16 | 53129.638 | 1058.5928 | 958.5928 | 828.8324
1493 | 89.6172 | 1.1053 | 0.88571 | 1.0096 | 199.9646 | 50.8794 | 50735.0253 | 50370.9764 | 1003.4195 | 903.4195 | 828.8626
1494 | 118.4167 | 1.0799 | 1.1708 | 1.0151 | 200.4299 | 50.9558 | 67275.1327 | 67316.9729 | 1342.3395 | 1242.3395 | 828.8703
1495 | 87.7103 | 0.98455 | 0.90152 | 0.99927 | 208.1944 | 51.0432 | 50599.1129 | 52087.0736 | 1037.7415 | 937.7415 | 828.9882
1496 | 95.546 | 0.96543 | 1.3563 | 1.03 | 207.3972 | 50.2919 | 75105.2026 | 77733.8771 | 1550.6775 | 1450.6775 | 829.028
1497 | 96.3276 | 1.0797 | 1.0376 | 0.90325 | 189.7739 | 50.155 | 52410.1037 | 51705.7049 | 1030.1141 | 930.1141 | 829.0605
1498 | 75.3892 | 0.90587 | 1.011 | 1.0263 | 197.779 | 52.2688 | 58000.4378 | 60623.9212 | 1208.4784 | 1108.4784 | 829.0985
1499 | 85.0512 | 0.94748 | 0.85323 | 0.99501 | 208.5166 | 44.623 | 41688.2902 | 41898.0402 | 833.9608 | 733.9608 | 829.268
1500 | 110.1715 | 0.99864 | 1.0087 | 0.99273 | 201.8958 | 52.6479 | 58712.3381 | 59571.5728 | 1187.4315 | 1087.4315 | 829.4291
1501 | 108.7103 | 0.99356 | 1.0502 | 0.96429 | 203.8504 | 53.0519 | 59660.2478 | 61007.1369 | 1216.1427 | 1116.1427 | 829.4466
1502 | 91.4154 | 1.0903 | 0.94312 | 0.99974 | 211.2759 | 48.0153 | 50269.7733 | 51979.483 | 1035.5897 | 935.5897 | 829.4725
1503 | 98.6134 | 0.98773 | 0.99431 | 0.9863 | 190.7942 | 49.8858 | 53972.2788 | 53480.4322 | 1065.6086 | 965.6086 | 829.497
1504 | 110.2883 | 0.98442 | 0.98443 | 0.93142 | 204.7207 | 51.6537 | 53174.833 | 52462.2249 | 1045.2445 | 945.2445 | 829.5341
1505 | 77.4671 | 1.0068 | 1.0507 | 1.0523 | 206.7185 | 48.4484 | 57554.8223 | 57616.5787 | 1148.3316 | 1048.3316 | 829.5373
1506 | 96.4341 | 1.0024 | 1.2206 | 1.1075 | 194.4302 | 51.3757 | 74613.8291 | 73870.2679 | 1473.4054 | 1373.4054 | 829.6288
1507 | 119.5803 | 1.1172 | 0.86505 | 0.99703 | 196.5662 | 49.7902 | 49791.0991 | 50129.2033 | 998.5841 | 898.5841 | 829.7036
1508 | 103.5555 | 0.95943 | 0.93782 | 0.96853 | 203.1273 | 51.6696 | 52268.9242 | 53269.824 | 1061.3965 | 961.3965 | 829.8674
1509 | 107.7397 | 1.0285 | 1.2078 | 0.92893 | 185.6637 | 51.7436 | 63972.9226 | 64754.0937 | 1291.0819 | 1191.0819 | 829.8737
1510 | 99.9088 | 1.1059 | 1.0319 | 0.99285 | 200.1211 | 52.4475 | 59730.029 | 58837.148 | 1172.743 | 1072.743 | 829.8764
1511 | 106.8791 | 0.95526 | 1.1198 | 0.90749 | 193.268 | 50.8486 | 57055.2492 | 58680.9028 | 1169.6181 | 1069.6181 | 830.0076
1512 | 101.6632 | 0.96857 | 1.0861 | 1.036 | 195.2648 | 54.3238 | 66666.0713 | 66599.8484 | 1327.997 | 1227.997 | 830.0846
1513 | 102.2916 | 1.0267 | 1.2009 | 1.1028 | 192.8328 | 47.9949 | 68799.4475 | 70656.9182 | 1409.1384 | 1309.1384 | 830.1439
1514 | 97.9486 | 1.0808 | 0.81074 | 1.0274 | 199.6538 | 46.6348 | 43979.8959 | 43637.2539 | 868.7451 | 768.7451 | 830.2095
1515 | 97.9621 | 1.0231 | 0.83414 | 1.0357 | 196.6845 | 49.6951 | 48108.783 | 47511.2867 | 946.2257 | 846.2257 | 830.3291
1516 | 110.2304 | 0.98799 | 1.1466 | 0.96554 | 201.211 | 54.0934 | 65979.2574 | 64952.8159 | 1295.0563 | 1195.0563 | 830.3876
1517 | 114.044 | 0.95849 | 1.1014 | 0.97854 | 200.3563 | 49.9321 | 59472.1422 | 58622.5832 | 1168.4517 | 1068.4517 | 830.4544
1518 | 87.7486 | 1.0592 | 0.89978 | 1.0031 | 186.4819 | 49.6932 | 49655.1439 | 50132.5001 | 998.65 | 898.65 | 830.4954
1519 | 90.8069 | 0.93027 | 1.3237 | 1.0448 | 211.5351 | 48.9653 | 72071.1208 | 73763.284 | 1471.2657 | 1371.2657 | 830.5413
1520 | 105.5308 | 0.89047 | 0.99084 | 1.0276 | 191.9893 | 49.1072 | 54809.5179 | 55101.713 | 1098.0343 | 998.0343 | 830.6181
1521 | 98.5829 | 1.0509 | 0.97092 | 0.98081 | 215.7389 | 46.1997 | 48997.6517 | 49302.364 | 982.0473 | 882.0473 | 830.6281
1522 | 98.5863 | 0.96458 | 1.2119 | 0.98249 | 215.3411 | 52.464 | 67672.111 | 66916.4292 | 1334.3286 | 1234.3286 | 830.6555
1523 | 116.2576 | 1.0017 | 1.0403 | 1.0898 | 196.4281 | 49.5514 | 62143.1303 | 61122.0625 | 1218.4412 | 1118.4412 | 830.6904
1524 | 111.4039 | 0.97063 | 1.2012 | 1.0067 | 192.4461 | 51.5006 | 68038.2851 | 69351.592 | 1383.0318 | 1283.0318 | 830.7458
1525 | 105.3175 | 0.92253 | 0.88117 | 1.0286 | 192.1833 | 44.7713 | 45119.9507 | 45938.6133 | 914.7723 | 814.7723 | 830.8458
1526 | 109.336 | 1.0665 | 0.88843 | 0.91963 | 206.2822 | 50.5371 | 47389.2483 | 48376.5285 | 963.5306 | 863.5306 | 830.9127
1527 | 97.4746 | 0.99427 | 0.97621 | 1.0488 | 197.1177 | 48.7739 | 54862.1934 | 56195.1524 | 1119.903 | 1019.903 | 831.0161
1528 | 101.0458 | 1.0261 | 1.1039 | 1.0192 | 192.4164 | 49.6493 | 61204.9757 | 59518.254 | 1186.3651 | 1086.3651 | 831.0192
1529 | 86.5441 | 1.0684 | 1.1377 | 0.99614 | 192.4552 | 52.2466 | 64232.5862 | 64629.3939 | 1288.5879 | 1188.5879 | 831.336
1530 | 104.2571 | 0.97286 | 1.0427 | 0.847 | 212.56 | 52.8018 | 52202.0292 | 52888.4463 | 1053.7689 | 953.7689 | 831.3431
1531 | 99.0815 | 1.0245 | 0.90142 | 0.94975 | 193.2331 | 48.0334 | 46191.5735 | 47000.9718 | 936.0194 | 836.0194 | 831.3704
1532 | 96.3078 | 0.94342 | 1.237 | 1.0645 | 208.9859 | 50.568 | 71388.2313 | 70743.2804 | 1410.8656 | 1310.8656 | 831.4179
1533 | 103.0018 | 1.06 | 1.0611 | 1.0273 | 211.6711 | 49.4044 | 59458.5868 | 60150.8273 | 1199.0165 | 1099.0165 | 831.4496
1534 | 93.4683 | 0.95547 | 1.0705 | 0.9147 | 212.9779 | 48.2287 | 51746.0328 | 49914.5222 | 994.2904 | 894.2904 | 831.4618
1535 | 111.5026 | 1.0517 | 1.1196 | 0.94751 | 211.5865 | 53.0962 | 62761.9416 | 64259.3322 | 1281.1866 | 1181.1866 | 831.4987
1536 | 105.4697 | 0.96559 | 1.2179 | 1.0732 | 205.2075 | 46.0353 | 65067.288 | 63241.0022 | 1260.82 | 1160.82 | 831.6074
1537 | 101.8071 | 0.99834 | 1.1171 | 0.98878 | 209.457 | 52.9016 | 64022.1369 | 68476.3479 | 1365.527 | 1265.527 | 831.6632
1538 | 92.8801 | 0.95242 | 1.0391 | 1.0155 | 191.5512 | 48.4969 | 55656.8208 | 56465.6539 | 1125.3131 | 1025.3131 | 831.6761
1539 | 86.624 | 0.95732 | 0.94828 | 1.032 | 202.4201 | 46.6324 | 49706.4456 | 50008.8489 | 996.177 | 896.177 | 831.758
1540 | 95.66 | 1.0007 | 0.92926 | 1.0746 | 190.7489 | 48.0841 | 52809.3363 | 52596.1057 | 1047.9221 | 947.9221 | 831.785
1541 | 102.1821 | 1.022 | 1.0986 | 0.97897 | 208.2209 | 55.1122 | 65235.7766 | 64137.4167 | 1278.7483 | 1178.7483 | 831.8702
1542 | 103.3091 | 1.0856 | 0.98794 | 0.99101 | 216.8514 | 48.6086 | 53259.5238 | 55212.6808 | 1100.2536 | 1000.2536 | 831.9345
1543 | 96.1009 | 0.93768 | 1.0297 | 1.0776 | 229.3169 | 47.9265 | 57726.831 | 57505.0294 | 1146.1006 | 1046.1006 | 832.0189
1544 | 94.3017 | 0.97601 | 1.1674 | 0.98369 | 197.0222 | 45.7792 | 56979.7066 | 56751.9955 | 1131.0399 | 1031.0399 | 832.0536
1545 | 98.866 | 0.99749 | 1.1281 | 1.0796 | 208.1725 | 49.6024 | 65507.5603 | 66790.5274 | 1331.8105 | 1231.8105 | 832.0834
1546 | 101.1525 | 0.99344 | 0.97111 | 1.0058 | 189.0535 | 44.1385 | 47735.9571 | 48407.1192 | 964.1424 | 864.1424 | 832.1726
1547 | 98.2171 | 1.0308 | 0.86936 | 0.97355 | 208.5278 | 49.8609 | 47457.1366 | 46676.7068 | 929.5341 | 829.5341 | 832.4135
1548 | 92.6352 | 0.9852 | 0.97115 | 1.0201 | 200.7383 | 46.795 | 50829.6054 | 51237.2993 | 1020.746 | 920.746 | 832.4266
1549 | 93.3518 | 1.0186 | 1.0259 | 1.013 | 182.2467 | 50.5481 | 57521.456 | 54245.0345 | 1080.9007 | 980.9007 | 832.5361
1550 | 104.7361 | 0.97654 | 1.1777 | 0.96781 | 210.9671 | 51.6173 | 64321.5022 | 63445.9712 | 1264.9194 | 1164.9194 | 832.5623
1551 | 107.3932 | 1.0682 | 0.80794 | 1.0139 | 186.8458 | 47.2433 | 44308.5341 | 43146.9695 | 858.9394 | 758.9394 | 832.6049
1552 | 89.4998 | 0.93951 | 1.1782 | 0.9924 | 195.6243 | 49.3708 | 62072.4568 | 62204.1165 | 1240.0823 | 1140.0823 | 832.6113
1553 | 101.3968 | 0.96792 | 0.95187 | 1.0217 | 210.6815 | 46.8832 | 50406.2102 | 52092.6239 | 1037.8525 | 937.8525 | 832.6459
1554 | 96.8443 | 1.1348 | 0.92458 | 1.0242 | 200.993 | 48.3291 | 51276.4373 | 50711.7212 | 1010.2344 | 910.2344 | 832.7004
1555 | 95.6629 | 1.0577 | 1.0325 | 0.99799 | 195.6823 | 50.6834 | 57548.2115 | 56766.0397 | 1131.3208 | 1031.3208 | 832.7315
1556 | 96.5714 | 1.039 | 0.83515 | 0.96789 | 192.8715 | 48.8319 | 44564.8917 | 46120.1597 | 918.4032 | 818.4032 | 832.732
1557 | 108.1272 | 0.95973 | 0.99454 | 0.88277 | 221.9634 | 48.4538 | 47790.1653 | 47948.9479 | 954.979 | 854.979 | 832.8162
1558 | 111.8567 | 1.0526 | 0.88528 | 1.0228 | 214.7214 | 49.9152 | 51289.5166 | 51769.9383 | 1031.3988 | 931.3988 | 832.848
1559 | 102.0208 | 0.96104 | 0.99456 | 1.032 | 195.0485 | 51.8135 | 58455.3291 | 60914.6486 | 1214.293 | 1114.293 | 832.918
1560 | 102.0535 | 1.0115 | 1.0769 | 0.93068 | 201.335 | 48.3236 | 53619.7053 | 51811.4791 | 1032.2296 | 932.2296 | 833.0779
1561 | 90.6063 | 0.99514 | 0.91221 | 1.0112 | 194.1385 | 48.9539 | 49762.7015 | 51235.8287 | 1020.7166 | 920.7166 | 833.2418
1562 | 101.9111 | 0.98939 | 0.87735 | 0.9719 | 180.2445 | 49.3598 | 47245.727 | 45887.8966 | 913.7579 | 813.7579 | 833.4702
1563 | 95.0592 | 0.9513 | 1.0095 | 1.0208 | 193.9232 | 53.3121 | 59954.0378 | 58755.8272 | 1171.1165 | 1071.1165 | 833.5037
1564 | 99.5514 | 1.0016 | 1.1185 | 1.0426 | 187.8287 | 48.9489 | 62150.3676 | 61379.7684 | 1223.5954 | 1123.5954 | 833.5521
1565 | 84.897 | 0.99198 | 0.73553 | 1.0858 | 227.9334 | 52.7686 | 46815.8114 | 47252.5638 | 941.0513 | 841.0513 | 833.5939
1566 | 108.9188 | 0.96883 | 0.88499 | 1.0301 | 188.7246 | 52.2598 | 53343.1003 | 51361.2678 | 1023.2254 | 923.2254 | 833.679
1567 | 103.1004 | 0.97399 | 0.94099 | 1.0156 | 206.9189 | 49.5959 | 52586.3751 | 51708.8956 | 1030.1779 | 930.1779 | 833.7224
1568 | 93.0664 | 0.94795 | 1.2022 | 1.0071 | 217.2387 | 50.6144 | 65962.4608 | 64819.5665 | 1292.3913 | 1192.3913 | 833.7225
1569 | 110.955 | 0.95016 | 0.85838 | 0.93721 | 173.7075 | 49.8478 | 45530.2211 | 45691.3333 | 909.8267 | 809.8267 | 833.873
1570 | 79.4301 | 1.0529 | 1.1918 | 1.0594 | 211.3234 | 53.1382 | 71750.9052 | 71928.7557 | 1434.5751 | 1334.5751 | 833.9952
1571 | 94.1038 | 1.0271 | 1.0327 | 0.9873 | 203.3687 | 50.9188 | 57038.6459 | 56460.5023 | 1125.21 | 1025.21 | 834.0258
1572 | 105.2897 | 0.92207 | 1.1101 | 1.0132 | 210.8941 | 51.5201 | 63163.3439 | 62761.3541 | 1251.2271 | 1151.2271 | 834.0709
1573 | 94.8156 | 1.0233 | 1.227 | 1.0891 | 192.0996 | 46.7449 | 67195.5635 | 66676.807 | 1329.5361 | 1229.5361 | 834.0823
1574 | 93.8468 | 0.93948 | 1.0381 | 0.95897 | 190.9749 | 48.2891 | 52521.9068 | 52254.9727 | 1041.0995 | 941.0995 | 834.2018
1575 | 103.3731 | 0.96469 | 0.79954 | 0.9548 | 211.5774 | 49.3808 | 42833.3247 | 42087.3598 | 837.7472 | 737.7472 | 834.2765
1576 | 98.6024 | 1.0241 | 0.86608 | 1.0184 | 206.7649 | 53.552 | 52850.3003 | 53184.0513 | 1059.681 | 959.681 | 834.6298
1577 | 98.0536 | 0.97485 | 1.0606 | 1.0777 | 217.2197 | 45.9692 | 57154.0934 | 57835.6053 | 1152.7121 | 1052.7121 | 834.7062
1578 | 95.286 | 0.96044 | 0.95277 | 1.0783 | 225.8856 | 46.8535 | 52649.4315 | 52194.7526 | 1039.8951 | 939.8951 | 834.7147
1579 | 108.6745 | 0.92049 | 1.0659 | 0.99987 | 206.7995 | 51.9327 | 60748.7787 | 60766.9402 | 1211.3388 | 1111.3388 | 834.7754
1580 | 105.014 | 0.96784 | 1.1191 | 1.032 | 217.4507 | 52.2082 | 65817.142 | 66079.8578 | 1317.5972 | 1217.5972 | 834.8249
1581 | 92.4563 | 0.91848 | 0.92126 | 1.0238 | 206.9465 | 49.9437 | 51556.1821 | 53248.9161 | 1060.9783 | 960.9783 | 834.8558
1582 | 104.7917 | 1.0405 | 1.2965 | 1.0158 | 176.4167 | 47.6124 | 68070.1 | 66697.1884 | 1329.9438 | 1229.9438 | 834.8851
1583 | 108.7109 | 1.022 | 1.1161 | 0.95504 | 195.6949 | 46.0509 | 54399.3703 | 53251.1286 | 1061.0226 | 961.0226 | 834.9104
1584 | 79.181 | 1.0185 | 0.97497 | 0.98727 | 203.8629 | 50.2816 | 52658.4647 | 53105.9019 | 1058.118 | 958.118 | 834.9169
1585 | 110.3039 | 0.93382 | 0.73796 | 1.0778 | 184.0344 | 49.6012 | 44743.9886 | 45486.3425 | 905.7269 | 805.7269 | 834.9183
1586 | 102.4552 | 1.027 | 1.021 | 0.98783 | 195.4043 | 53.8384 | 60161.4258 | 58915.9003 | 1174.318 | 1074.318 | 834.9715
1587 | 92.7952 | 0.87327 | 0.97872 | 1.022 | 199.7805 | 50.7091 | 55031.0752 | 55158.5759 | 1099.1715 | 999.1715 | 835.0012
1588 | 86.8255 | 0.92893 | 1.0014 | 0.93071 | 204.1663 | 46.1271 | 46914.5941 | 46083.9259 | 917.6785 | 817.6785 | 835.0725
1589 | 97.8815 | 0.95324 | 0.81421 | 0.90855 | 207.0618 | 47.5139 | 39788.4715 | 39420.9639 | 784.4193 | 684.4193 | 835.1226
1590 | 99.7373 | 1.0956 | 0.75018 | 0.98204 | 203.4095 | 51.5299 | 43796.5445 | 43465.0895 | 865.3018 | 765.3018 | 835.1496
1591 | 97.3406 | 1.0076 | 1.0897 | 0.97746 | 174.1741 | 52.7784 | 61567.8003 | 62150.8576 | 1239.0172 | 1139.0172 | 835.3002
1592 | 108.5027 | 0.96415 | 0.87157 | 0.94154 | 211.6545 | 52.8121 | 49075.0815 | 49970.0643 | 995.4013 | 895.4013 | 835.3612
1593 | 95.0133 | 1.0079 | 1.2264 | 0.93855 | 190.1296 | 46.6963 | 58411.929 | 57489.3522 | 1145.787 | 1045.787 | 835.5132
1594 | 101.5154 | 1.0675 | 1.1903 | 1.0678 | 180.2392 | 53.1418 | 73484.5888 | 73455.609 | 1465.1122 | 1365.1122 | 835.5758
1595 | 101.2382 | 0.91848 | 1.0837 | 0.95288 | 196.2511 | 52.2278 | 58982.9305 | 59739.0871 | 1190.7817 | 1090.7817 | 835.7635
1596 | 123.407 | 1.0218 | 1.0029 | 0.94045 | 191.068 | 47.4567 | 50935.6308 | 50143.3863 | 998.8677 | 898.8677 | 835.7954
1597 | 102.4236 | 0.90605 | 0.94329 | 0.99181 | 194.337 | 51.6855 | 53345.5349 | 54064.7436 | 1077.2949 | 977.2949 | 835.835
1598 | 104.9207 | 1.0545 | 1.0456 | 0.99983 | 194.5858 | 48.1864 | 55902.2256 | 56763.7783 | 1131.2756 | 1031.2756 | 835.9197
1599 | 106.8087 | 0.95739 | 0.91876 | 1.018 | 185.8763 | 53.1036 | 55283.557 | 57693.4044 | 1149.8681 | 1049.8681 | 836.0056
1600 | 93.6497 | 1.0144 | 1.1133 | 1.0817 | 205.9191 | 47.4525 | 61857.5244 | 62697.0206 | 1249.9404 | 1149.9404 | 836.0194
1601 | 97.9365 | 1.0015 | 1.2128 | 1.0053 | 209.8274 | 52.1584 | 68917.4833 | 68601.0796 | 1368.0216 | 1268.0216 | 836.0292
1602 | 93.0467 | 0.93639 | 0.98514 | 0.94558 | 193.9952 | 50.6632 | 51802.8055 | 50875.3121 | 1013.5062 | 913.5062 | 836.1315
1603 | 90.0248 | 1.0468 | 1.1683 | 1.0798 | 218.388 | 49.8326 | 67781.5411 | 68149.584 | 1358.9917 | 1258.9917 | 836.1505
1604 | 114.2945 | 0.99459 | 1.1866 | 0.92336 | 199.861 | 49.8115 | 60439.2309 | 62708.2436 | 1250.1649 | 1150.1649 | 836.2049
1605 | 88.8417 | 0.95684 | 0.99904 | 1.0051 | 197.9516 | 51.4878 | 56272.7904 | 55604.4213 | 1108.0884 | 1008.0884 | 836.2409
1606 | 83.7783 | 1.0273 | 0.88722 | 0.96077 | 198.9085 | 48.5378 | 45750.7812 | 43953.2864 | 875.0657 | 775.0657 | 836.2649
1607 | 108.0217 | 0.98187 | 0.82234 | 1.064 | 206.6941 | 48.3999 | 47687.8033 | 47417.1972 | 944.3439 | 844.3439 | 836.2681
1608 | 89.7807 | 0.98739 | 1.1821 | 0.99451 | 203.4702 | 50.1235 | 63571.9992 | 62870.351 | 1253.407 | 1153.407 | 836.3048
1609 | 106.897 | 1.044 | 0.96371 | 1.0551 | 200.2502 | 49.547 | 56111.6555 | 55256.2413 | 1101.1248 | 1001.1248 | 836.3344
1610 | 88.6883 | 1.0387 | 1.1282 | 0.9069 | 213.1484 | 53.0226 | 59347.693 | 58559.1267 | 1167.1825 | 1067.1825 | 836.3953
1611 | 102.7281 | 0.94387 | 0.63403 | 0.95105 | 197.0648 | 48.2042 | 33937.8541 | 35297.9443 | 701.9589 | 601.9589 | 836.4059
1612 | 114.1817 | 0.95744 | 1.0895 | 1.0049 | 190.535 | 48.8688 | 59035.7769 | 58667.3704 | 1169.3474 | 1069.3474 | 836.4283
1613 | 94.6849 | 1.0847 | 1.0927 | 1.0199 | 201.0334 | 48.6196 | 59381.4729 | 59611.6277 | 1188.2326 | 1088.2326 | 836.4671
1614 | 97.6333 | 1.0247 | 1.1007 | 0.96546 | 203.5638 | 52.5675 | 61322.2831 | 59502.306 | 1186.0461 | 1086.0461 | 836.5042
1615 | 108.3118 | 1.0432 | 1.1704 | 0.97349 | 193.6854 | 56.6625 | 71157.1622 | 71370.8561 | 1423.4171 | 1323.4171 | 836.5846
1616 | 103.3013 | 0.96213 | 0.90007 | 0.9388 | 184.742 | 51.4471 | 48770.4492 | 47591.3368 | 947.8267 | 847.8267 | 836.6167
1617 | 111.8348 | 1.0287 | 1.2006 | 1.0731 | 203.9215 | 48.6004 | 68410.0713 | 68968.9302 | 1375.3786 | 1275.3786 | 836.6362
1618 | 99.3494 | 0.90371 | 1.2258 | 0.99139 | 206.7029 | 50.6447 | 66296.8839 | 64460.2311 | 1285.2046 | 1185.2046 | 836.9098
1619 | 92.028 | 0.95678 | 1.2285 | 0.92823 | 200.7069 | 51.381 | 63315.0983 | 61595.9094 | 1227.9182 | 1127.9182 | 836.9903
1620 | 103.5664 | 1.0008 | 1.069 | 1.0526 | 185.4965 | 46.5018 | 57330.9254 | 58563.9662 | 1167.2793 | 1067.2793 | 836.9976
1621 | 104.4899 | 1.0352 | 0.78493 | 1.0974 | 197.7576 | 47.4339 | 46188.4937 | 45154.1259 | 899.0825 | 799.0825 | 837.0352
1622 | 89.9125 | 0.93597 | 0.93162 | 1.0322 | 203.6377 | 52.8592 | 55484.6089 | 54958.0188 | 1095.1604 | 995.1604 | 837.0775
1623 | 107.0576 | 0.88534 | 0.96056 | 1.0654 | 205.6908 | 49.5504 | 55612.8107 | 58451.2337 | 1165.0247 | 1065.0247 | 837.1522
1624 | 105.694 | 0.97192 | 1.0537 | 0.89576 | 200.4983 | 49.2703 | 51764.674 | 52585.1746 | 1047.7035 | 947.7035 | 837.199
1625 | 110.9372 | 0.9431 | 0.99777 | 1.0631 | 197.8204 | 50.8124 | 59412.7174 | 58441.1053 | 1164.8221 | 1064.8221 | 837.2217
1626 | 115.2954 | 0.9947 | 1.0595 | 1.0441 | 182.9955 | 51.7547 | 63370.8876 | 64131.5904 | 1278.6318 | 1178.6318 | 837.2432
1627 | 99.7417 | 0.93356 | 0.97283 | 1.0346 | 189.6115 | 50.448 | 55662.8876 | 57924.9717 | 1154.4994 | 1054.4994 | 837.2682
1628 | 106.6305 | 1.0281 | 0.78461 | 0.97049 | 193.92 | 46.5819 | 40770.662 | 40342.0518 | 802.841 | 702.841 | 837.2756
1629 | 103.1247 | 0.95797 | 1.024 | 1.0252 | 197.7408 | 51.0343 | 58812.3561 | 59945.2197 | 1194.9044 | 1094.9044 | 837.3522
1630 | 93.8518 | 0.98923 | 0.74552 | 0.96862 | 195.6533 | 50.567 | 41405.9359 | 42576.1073 | 847.5221 | 747.5221 | 837.3546
1631 | 80.9202 | 0.96418 | 0.8994 | 0.99383 | 189.6416 | 53.4498 | 52135.8462 | 53750.4965 | 1071.0099 | 971.0099 | 837.3588
1632 | 89.4224 | 0.98729 | 0.90073 | 0.99182 | 210.9878 | 51.1824 | 50454.0533 | 49843.8562 | 992.8771 | 892.8771 | 837.4262
1633 | 88.8706 | 0.94561 | 0.92209 | 0.97484 | 199.4332 | 52.6756 | 51975.7268 | 51810.7226 | 1032.2145 | 932.2145 | 837.4778
1634 | 111.8531 | 0.93425 | 0.97137 | 1.0288 | 185.2711 | 49.4362 | 54754.5964 | 54166.8141 | 1079.3363 | 979.3363 | 837.4818
1635 | 113.1164 | 1.0374 | 0.96027 | 1.053 | 201.1474 | 54.0533 | 61200.9649 | 62048.806 | 1236.9761 | 1136.9761 | 837.5826
1636 | 127.8261 | 0.91458 | 0.91625 | 1.0217 | 210.3944 | 49.8124 | 52666.0459 | 54198.7258 | 1079.9745 | 979.9745 | 837.6056
1637 | 123.6625 | 0.99341 | 1.1193 | 1.0432 | 189.1454 | 45.2727 | 58610.0329 | 57752.3045 | 1151.0461 | 1051.0461 | 837.6275
1638 | 79.2453 | 0.92594 | 1.049 | 0.90347 | 206.6266 | 49.5872 | 50840.3988 | 52013.7602 | 1036.2752 | 936.2752 | 837.8743
1639 | 87.5362 | 1.1085 | 1.2129 | 0.97121 | 192.168 | 48.3296 | 61814.8935 | 63967.9134 | 1275.3583 | 1175.3583 | 837.9051
1640 | 98.0908 | 1.0596 | 1.1192 | 1.0126 | 201.0487 | 50.773 | 63015.57 | 63024.3758 | 1256.4875 | 1156.4875 | 837.9157
1641 | 88.7326 | 1.0332 | 1.1082 | 1.0016 | 174.4055 | 44.7276 | 53922.3356 | 54908.1093 | 1094.1622 | 994.1622 | 837.939
1642 | 100.422 | 1.039 | 1.0439 | 0.97897 | 199.6492 | 53.1432 | 60055.0263 | 60516.5633 | 1206.3313 | 1106.3313 | 837.949
1643 | 88.7067 | 1.0655 | 1.0085 | 1.0265 | 209.0226 | 49.3779 | 55993.2711 | 54850.9908 | 1093.0198 | 993.0198 | 838.0841
1644 | 89.5867 | 0.97335 | 1.106 | 0.94024 | 191.3012 | 48.9659 | 55381.1358 | 56899.8439 | 1133.9969 | 1033.9969 | 838.1236
1645 | 98.8746 | 1.0188 | 1.0094 | 0.9864 | 188.9092 | 48.3406 | 53188.4661 | 53068.729 | 1057.3746 | 957.3746 | 838.1716
1646 | 105.2993 | 0.99178 | 0.7231 | 0.98185 | 190.8625 | 49.2624 | 40310.6182 | 42189.3228 | 839.7865 | 739.7865 | 838.2165
1647 | 95.5532 | 0.86465 | 0.96018 | 1.0401 | 198.5928 | 50.3597 | 54653.3101 | 55760.3698 | 1111.2074 | 1011.2074 | 838.2366
1648 | 89.2724 | 0.91694 | 0.95478 | 1.0178 | 209.4419 | 52.5991 | 55631.0941 | 54575.4113 | 1087.5082 | 987.5082 | 838.3651
1649 | 118.1348 | 0.96866 | 1.097 | 1.0499 | 203.0197 | 47.6662 | 60556.9151 | 60559.1463 | 1207.1829 | 1107.1829 | 838.4182
1650 | 101.2429 | 1.0283 | 0.68755 | 1.0462 | 197.189 | 49.2891 | 40783.3909 | 40919.7855 | 814.3957 | 714.3957 | 838.4662
1651 | 91.6452 | 0.98334 | 1.1297 | 0.97609 | 182.2377 | 49.0059 | 58635.4206 | 57261.6886 | 1141.2338 | 1041.2338 | 838.55
1652 | 99.2924 | 0.97292 | 1.0815 | 0.9884 | 193.4255 | 50.7237 | 59316.2837 | 59500.8094 | 1186.0162 | 1086.0162 | 838.6469
1653 | 94.9123 | 1.055 | 0.90543 | 0.98572 | 204.6069 | 51.3755 | 51201.5079 | 50256.8549 | 1001.1371 | 901.1371 | 838.7071
1654 | 98.0692 | 0.97957 | 0.94846 | 0.95868 | 209.9909 | 49.8238 | 50299.497 | 50210.4335 | 1000.2087 | 900.2087 | 838.7772
1655 | 98.0216 | 1.0322 | 1.0653 | 0.99441 | 205.1836 | 48.6306 | 56640.5662 | 55879.4143 | 1113.5883 | 1013.5883 | 838.8771
1656 | 100.4507 | 0.8763 | 0.88117 | 1.0243 | 191.9127 | 45.4879 | 45251.3507 | 45441.1153 | 904.8223 | 804.8223 | 838.8895
1657 | 112.775 | 0.93344 | 1.0675 | 0.93519 | 186.9945 | 46.4484 | 51448.4743 | 50300.5793 | 1002.0116 | 902.0116 | 838.9228
1658 | 93.5576 | 0.96043 | 1.045 | 0.94763 | 198.9507 | 48.6077 | 52702.1227 | 54005.5702 | 1076.1114 | 976.1114 | 838.9263
1659 | 96.2884 | 1.104 | 1.1214 | 0.90763 | 195.4013 | 52.9331 | 59700.3016 | 56976.6951 | 1135.5339 | 1035.5339 | 838.9728
1660 | 93.3669 | 0.95613 | 1.0523 | 0.95776 | 206.6468 | 56.1489 | 61810.5809 | 61604.2384 | 1228.0848 | 1128.0848 | 838.9938
1661 | 88.9252 | 0.99461 | 0.77757 | 0.90452 | 201.1954 | 51.9725 | 41351.6187 | 40534.2863 | 806.6857 | 706.6857 | 839.1166
1662 | 107.187 | 1.0403 | 0.72101 | 1.0389 | 210.177 | 52.6014 | 45476.6846 | 46055.6505 | 917.113 | 817.113 | 839.144
1663 | 96.5552 | 0.92089 | 1.1168 | 1.0262 | 185.5836 | 51.7825 | 64137.5254 | 66028.6625 | 1316.5732 | 1216.5732 | 839.1509
1664 | 114.5585 | 1.0611 | 0.91885 | 1.041 | 209.3075 | 51.6607 | 55901.5538 | 57677.4884 | 1149.5498 | 1049.5498 | 839.2092
1665 | 106.2831 | 0.97297 | 1.3203 | 0.96273 | 189.7057 | 47.5658 | 65570.4289 | 66859.8035 | 1333.1961 | 1233.1961 | 839.3177
1666 | 107.1526 | 0.95822 | 0.94698 | 0.96838 | 200.8027 | 51.1797 | 52389.1274 | 52292.6502 | 1041.853 | 941.853 | 839.3288
1667 | 98.3688 | 1.0366 | 0.92528 | 1.0251 | 173.6557 | 47.2055 | 49761.3689 | 50959.6313 | 1015.1926 | 915.1926 | 839.332
1668 | 113.3785 | 1.0695 | 1.0692 | 0.94473 | 189.6269 | 52.5514 | 59645.9111 | 59455.8242 | 1185.1165 | 1085.1165 | 839.4813
1669 | 95.3872 | 1.0129 | 1.0225 | 0.9471 | 188.8836 | 52.5523 | 56159.6892 | 55497.1413 | 1105.9428 | 1005.9428 | 839.5106
1670 | 99.2044 | 1.0006 | 1.0931 | 0.9563 | 200.3081 | 49.2694 | 56594.0606 | 57056.4502 | 1137.129 | 1037.129 | 839.5165
1671 | 106.628 | 0.94533 | 1.0226 | 1.0646 | 195.0587 | 48.2051 | 57531.8034 | 55182.9161 | 1099.6583 | 999.6583 | 839.6039
1672 | 73.2386 | 0.92489 | 0.9603 | 1.0025 | 192.166 | 48.2265 | 49886.6335 | 50363.8215 | 1003.2764 | 903.2764 | 839.639
1673 | 83.7877 | 0.95767 | 1.0336 | 0.9961 | 205.9576 | 53.0568 | 59087.0977 | 58507.634 | 1166.1527 | 1066.1527 | 839.7186
1674 | 115.1741 | 1.0085 | 1.0975 | 1.0407 | 187.7638 | 47.4327 | 59871.1963 | 57791.3556 | 1151.8271 | 1051.8271 | 839.7191
1675 | 99.0797 | 0.91617 | 1.0213 | 0.93645 | 201.7923 | 53.0602 | 55763.697 | 56100.8653 | 1118.0173 | 1018.0173 | 839.9363
1676 | 102.76 | 1.003 | 0.91518 | 1.0602 | 182.6604 | 50.8464 | 54759.1574 | 56405.5063 | 1124.1101 | 1024.1101 | 839.9671
1677 | 103.0942 | 1.0383 | 0.84815 | 1.104 | 205.8446 | 53.5254 | 56054.2611 | 56446.1557 | 1124.9231 | 1024.9231 | 840.031
1678 | 104.3559 | 1.0019 | 1.1329 | 0.94954 | 209.1657 | 47.3995 | 56153.8871 | 53469.9509 | 1065.399 | 965.399 | 840.1195
1679 | 100.0744 | 0.9546 | 1.0037 | 0.99335 | 206.3693 | 48.1922 | 52859.034 | 53586.719 | 1067.7344 | 967.7344 | 840.1351
1680 | 105.4014 | 1.1404 | 0.97881 | 1.0103 | 205.3552 | 48.158 | 53619.6357 | 52980.6694 | 1055.6134 | 955.6134 | 840.2479
1681 | 112.2203 | 1.1031 | 0.85245 | 1.06 | 209.8193 | 44.7792 | 46215.6069 | 45370.7278 | 903.4146 | 803.4146 | 840.3977
1682 | 106.6299 | 1.0203 | 0.86096 | 0.99712 | 212.182 | 50.9187 | 49465.0429 | 47428.5532 | 944.5711 | 844.5711 | 840.5461
1683 | 109.1944 | 0.9541 | 1.2985 | 1.0477 | 193.168 | 46.5096 | 68308.7395 | 67969.9123 | 1355.3982 | 1255.3982 | 840.5545
1684 | 109.4644 | 1.0049 | 0.71067 | 1.0124 | 207.8769 | 51.2129 | 42688.318 | 44251.3247 | 881.0265 | 781.0265 | 840.5984
1685 | 75.2606 | 0.98797 | 0.92223 | 0.98689 | 213.9842 | 50.619 | 50048.0056 | 49809.9797 | 992.1996 | 892.1996 | 840.7327
1686 | 90.2214 | 1.0192 | 1.1422 | 1.0711 | 199.5128 | 52.1688 | 68820.2359 | 69965.4497 | 1395.309 | 1295.309 | 840.7396
1687 | 87.4221 | 1.0309 | 1.1284 | 1.0177 | 199.6235 | 52.9916 | 65830.1855 | 66428.4877 | 1324.5698 | 1224.5698 | 840.8233
1688 | 80.8053 | 0.94561 | 0.89308 | 0.99544 | 199.8325 | 50.8895 | 49329.3529 | 49059.1289 | 977.1826 | 877.1826 | 840.8276
1689 | 96.3674 | 0.95 | 0.99144 | 0.96509 | 201.595 | 49.8829 | 52497.6953 | 53108.926 | 1058.1785 | 958.1785 | 841.0513
1690 | 88.5506 | 1.0099 | 0.88551 | 1.0232 | 212.2659 | 49.681 | 49669.1374 | 48957.8278 | 975.1566 | 875.1566 | 841.053
1691 | 94.3452 | 1.0462 | 1.3405 | 0.96235 | 195.2665 | 49.8992 | 69493.0024 | 71545.9871 | 1426.9197 | 1326.9197 | 841.1128
1692 | 108.9172 | 1.0628 | 1.1243 | 0.94833 | 198.3416 | 48.2891 | 57272.8347 | 59052.0431 | 1177.0409 | 1077.0409 | 841.2115
1693 | 112.1606 | 0.97216 | 0.94168 | 0.99028 | 196.3761 | 51.1184 | 53439.4555 | 52578.2546 | 1047.5651 | 947.5651 | 841.2517
1694 | 85.5569 | 1.0485 | 0.71797 | 1.012 | 183.4767 | 45.663 | 37457.78 | 36099.7874 | 717.9957 | 617.9957 | 841.3188
1695 | 101.9882 | 1.0011 | 0.88219 | 0.95747 | 182.3367 | 51.9823 | 49397.5239 | 48281.6865 | 961.6337 | 861.6337 | 841.4007
1696 | 116.3618 | 1.0844 | 0.8895 | 1.0153 | 199.5872 | 51.3695 | 53072.8109 | 52140.9779 | 1038.8196 | 938.8196 | 841.411
1697 | 106.5764 | 1.0353 | 0.86197 | 1.0801 | 199.354 | 49.7824 | 52041.1438 | 51726.6534 | 1030.5331 | 930.5331 | 841.4471
1698 | 91.2469 | 0.99942 | 1.0122 | 1.0897 | 215.5839 | 49.924 | 59836.5515 | 58914.8442 | 1174.2969 | 1074.2969 | 841.6427
1699 | 94.6298 | 0.99735 | 0.77755 | 0.96348 | 196.5933 | 47.2006 | 40011.9783 | 39755.5369 | 791.1107 | 691.1107 | 841.7271
1700 | 116.3518 | 0.9659 | 1.1134 | 1.1085 | 194.5975 | 47.4354 | 64068.4422 | 66037.4259 | 1316.7485 | 1216.7485 | 841.8053
1701 | 102.0649 | 1.0309 | 0.84952 | 1.0131 | 199.7974 | 52.5577 | 50964.4928 | 51725.5248 | 1030.5105 | 930.5105 | 841.8464
1702 | 103.0485 | 1.0114 | 0.7005 | 1.0312 | 206.2345 | 47.8637 | 39768.7558 | 39048.3559 | 776.9671 | 676.9671 | 842.1127
1703 | 101.5034 | 0.96384 | 1.1893 | 0.97631 | 195.1382 | 51.595 | 65150.0156 | 65014.1469 | 1296.2829 | 1196.2829 | 842.1241
1704 | 93.6544 | 0.97646 | 0.9024 | 0.92846 | 191.2474 | 50.1462 | 46791.8779 | 46681.44 | 929.6288 | 829.6288 | 842.1563
1705 | 105.7449 | 1.028 | 1.0163 | 1.0036 | 208.5371 | 51.4069 | 58227.1742 | 57140.7394 | 1138.8148 | 1038.8148 | 842.1646
1706 | 106.3291 | 1.024 | 0.67147 | 1.0184 | 193.7449 | 48.0963 | 38319.1092 | 37859.411 | 753.1882 | 653.1882 | 842.3627
1707 | 110.384 | 0.92493 | 0.928 | 1.0018 | 202.66 | 50.5256 | 52332.9353 | 51786.2737 | 1031.7255 | 931.7255 | 842.4057
1708 | 101.9496 | 1.0234 | 1.113 | 1.0508 | 203.7206 | 52.7122 | 67351.4279 | 68114.5589 | 1358.2912 | 1258.2912 | 842.4355
1709 | 99.3186 | 1.0788 | 0.91963 | 1.0005 | 197.3958 | 50.1534 | 51716.2636 | 51300.1073 | 1022.0021 | 922.0021 | 842.6449
1710 | 102.7661 | 0.92441 | 1.1429 | 1.0132 | 203.833 | 51.6827 | 64960.5338 | 66371.7875 | 1323.4357 | 1223.4357 | 842.7552
1711 | 102.9214 | 0.92809 | 0.94632 | 0.99937 | 210.153 | 48.5618 | 50775.114 | 50462.4216 | 1005.2484 | 905.2484 | 842.7698
1712 | 103.2468 | 0.99085 | 0.89943 | 1.0144 | 194.6646 | 50.6483 | 51588.8418 | 54226.0386 | 1080.5208 | 980.5208 | 842.7907
1713 | 95.6808 | 0.93527 | 0.80729 | 0.98702 | 198.3225 | 47.5832 | 42371.1523 | 42367.7784 | 843.3556 | 743.3556 | 842.8137
1714 | 95.8926 | 0.99299 | 1.253 | 1.0693 | 203.9577 | 49.5858 | 71363.3407 | 71487.9214 | 1425.7584 | 1325.7584 | 842.9411
1715 | 91.6817 | 0.94096 | 1.0427 | 1.0509 | 208.3501 | 45.6898 | 54214.5405 | 53397.9289 | 1063.9586 | 963.9586 | 842.9555
1716 | 88.4976 | 1.067 | 0.89053 | 0.98249 | 202.5898 | 51.1377 | 49773.655 | 47926.6838 | 954.5337 | 854.5337 | 843.2301
1717 | 88.2977 | 0.97971 | 0.99808 | 0.92449 | 216.7866 | 51.9514 | 52647.1996 | 54439.8397 | 1084.7968 | 984.7968 | 843.2368
1718 | 105.6819 | 0.94157 | 0.93711 | 1.0159 | 184.9281 | 48.7582 | 51454.591 | 52263.7685 | 1041.2754 | 941.2754 | 843.2742
1719 | 102.6828 | 0.96243 | 1.1665 | 1.0957 | 211.6914 | 50.6647 | 69970.8594 | 69408.0415 | 1384.1608 | 1284.1608 | 843.2958
1720 | 87.5831 | 0.9616 | 0.92823 | 0.96423 | 185.0815 | 53.9815 | 53046.4625 | 54222.055 | 1080.4411 | 980.4411 | 843.3723
1721 | 89.8221 | 1.0226 | 0.90934 | 1.0221 | 194.4238 | 44.631 | 45773.8447 | 47012.0472 | 936.2409 | 836.2409 | 843.4217
1722 | 127.9436 | 1.0183 | 1.0092 | 0.94081 | 211.2307 | 51.2836 | 55586.3946 | 56469.2426 | 1125.3849 | 1025.3849 | 843.4397
1723 | 107.8821 | 0.96991 | 0.68778 | 0.97908 | 209.9294 | 52.5588 | 41101.8918 | 41302.6443 | 822.0529 | 722.0529 | 843.4401
1724 | 82.3248 | 0.94249 | 1.2312 | 0.94837 | 194.9386 | 47.8491 | 59778.1948 | 63824.465 | 1272.4893 | 1172.4893 | 843.5082
1725 | 112.1079 | 1.0481 | 1.0781 | 0.95847 | 190.4151 | 46.8401 | 54097.0991 | 54175.9343 | 1079.5187 | 979.5187 | 843.5375
1726 | 98.5985 | 0.97723 | 0.81438 | 0.98686 | 183.9828 | 53.7986 | 48604.1245 | 49177.3987 | 979.548 | 879.548 | 843.7796
1727 | 113.8013 | 0.95913 | 1.0582 | 0.96662 | 192.6487 | 48.3877 | 54968.4278 | 55537.2025 | 1106.744 | 1006.744 | 843.9165
1728 | 92.7395 | 1.0013 | 1.0717 | 0.99054 | 187.8158 | 52.1894 | 60434.2896 | 61299.7519 | 1221.995 | 1121.995 | 843.9255
1729 | 100.8546 | 0.97403 | 0.97837 | 1.0561 | 179.943 | 53.614 | 60841.803 | 60463.7487 | 1205.275 | 1105.275 | 843.9453
1730 | 109.1991 | 1.0103 | 1.2916 | 1.0595 | 197.1274 | 48.5471 | 71988.6349 | 71316.3199 | 1422.3264 | 1322.3264 | 843.9743
1731 | 89.0147 | 0.99905 | 0.72313 | 1.0372 | 213.6223 | 47.6776 | 40214.6707 | 40888.7368 | 813.7747 | 713.7747 | 844.022
1732 | 113.5175 | 0.97952 | 1.1373 | 0.90697 | 202.3752 | 48.1458 | 55220.4753 | 54540.8725 | 1086.8174 | 986.8174 | 844.0654
1733 | 86.1978 | 0.92413 | 1.0513 | 0.92405 | 189.4186 | 48.4824 | 51149.0958 | 52726.3538 | 1050.5271 | 950.5271 | 844.1169
1734 | 125.6772 | 0.91103 | 0.73781 | 0.94509 | 221.2056 | 48.168 | 39323.5739 | 38245.6844 | 760.9137 | 660.9137 | 844.1442
1735 | 98.2823 | 1.0396 | 0.87388 | 0.94258 | 212.7985 | 49.5496 | 46089.3571 | 45315.5616 | 902.3112 | 802.3112 | 844.2391
1736 | 96.5952 | 1.0908 | 0.99406 | 1.0953 | 181.0796 | 52.174 | 62486.502 | 60833.8966 | 1212.6779 | 1112.6779 | 844.2743
1737 | 86.2881 | 0.90808 | 0.88563 | 0.99772 | 201.7357 | 48.2664 | 46632.7199 | 46935.7334 | 934.7147 | 834.7147 | 844.3439
1738 | 93.9481 | 0.95994 | 1.154 | 1.0565 | 196.386 | 47.6301 | 62561.5381 | 63572.9608 | 1267.4592 | 1167.4592 | 844.5625
1739 | 107.0104 | 1.0123 | 0.94349 | 0.94249 | 182.4442 | 46.1151 | 46184.7753 | 47146.1411 | 938.9228 | 838.9228 | 844.5711
1740 | 81.7889 | 1.0577 | 0.95686 | 1.0277 | 193.3152 | 48.1682 | 51725.8508 | 51950.8381 | 1035.0168 | 935.0168 | 844.591
1741 | 105.986 | 0.95995 | 0.92296 | 1.0725 | 197.0798 | 47.4714 | 52016.2999 | 52722.3338 | 1050.4467 | 950.4467 | 844.6229
1742 | 93.8745 | 0.94094 | 1.1071 | 0.97859 | 209.2628 | 48.9652 | 57582.0153 | 58269.6987 | 1161.394 | 1061.394 | 844.6614
1743 | 106.3747 | 1.0159 | 0.95453 | 0.98463 | 214.3309 | 50.474 | 53106.8675 | 52975.1697 | 1055.5034 | 955.5034 | 844.7647
1744 | 102.3989 | 1.05 | 0.85187 | 1.076 | 209.0135 | 50.1842 | 51602.271 | 53097.7518 | 1057.955 | 957.955 | 844.8724
1745 | 112.8991 | 0.94186 | 0.95206 | 1.0415 | 182.6108 | 50.515 | 55641.0955 | 54845.8133 | 1092.9163 | 992.9163 | 845.2989
1746 | 104.2206 | 0.99944 | 0.95454 | 1.0144 | 198.0894 | 48.4637 | 52173.9124 | 52255.592 | 1041.1118 | 941.1118 | 845.3362
1747 | 104.972 | 1.0194 | 1.0025 | 1.0179 | 194.0911 | 48.2526 | 54593.7436 | 52574.2766 | 1047.4855 | 947.4855 | 845.4344
1748 | 106.1464 | 0.98998 | 0.84804 | 0.98588 | 196.5654 | 49.2896 | 46585.3467 | 45925.9634 | 914.5193 | 814.5193 | 845.4383
1749 | 115.8967 | 0.97987 | 1.0322 | 1.0036 | 188.133 | 51.6565 | 59568.0952 | 56908.7061 | 1134.1741 | 1034.1741 | 845.4623
1750 | 115.3598 | 1.0802 | 0.79588 | 0.99912 | 197.0861 | 49.3395 | 45579.1562 | 46272.4857 | 921.4497 | 821.4497 | 845.4859
1751 | 100.0878 | 1.0491 | 0.91431 | 1.0742 | 200.869 | 48.9128 | 53377.8383 | 53046.7215 | 1056.9344 | 956.9344 | 845.508
1752 | 103.1351 | 1.0553 | 1.0808 | 0.97305 | 195.1269 | 50.3261 | 58601.3781 | 58683.0982 | 1169.662 | 1069.662 | 845.5335
1753 | 114.2178 | 0.96659 | 1.2529 | 1.0017 | 201.3272 | 51.1899 | 70094.8872 | 69735.7834 | 1390.7157 | 1290.7157 | 845.6142
1754 | 106.9533 | 1.0534 | 0.67963 | 1.0006 | 203.1294 | 48.3006 | 38489.896 | 38711.218 | 770.2244 | 670.2244 | 845.8198
1755 | 86.287 | 0.91096 | 0.57242 | 1.0803 | 186.1365 | 49.5347 | 34710.8722 | 35246.6196 | 700.9324 | 600.9324 | 845.8382
1756 | 83.4852 | 1.0732 | 0.92313 | 0.96187 | 191.7028 | 47.5252 | 46648.7196 | 47474.2971 | 945.4859 | 845.4859 | 845.8465
1757 | 93.271 | 1.0194 | 1.1614 | 1.0041 | 192.391 | 52.4753 | 66379.015 | 63273.1494 | 1261.463 | 1161.463 | 845.9094
1758 | 92.2331 | 0.96588 | 0.97035 | 0.93529 | 201.7549 | 51.7906 | 51818.6433 | 50103.2207 | 998.0644 | 898.0644 | 846.0048
1759 | 115.1722 | 0.97867 | 0.89234 | 0.99117 | 207.559 | 46.9703 | 47045.2382 | 47307.8131 | 942.1563 | 842.1563 | 846.1406
1760 | 111.5971 | 1.0597 | 1.0776 | 1.0347 | 192.6957 | 52.0159 | 64340.1952 | 63534.4122 | 1266.6882 | 1166.6882 | 846.1835
1761 | 87.348 | 0.93305 | 1.055 | 0.95637 | 207.4429 | 53.7599 | 58829.8709 | 58716.4547 | 1170.3291 | 1070.3291 | 846.2257
1762 | 77.2931 | 1.0168 | 1.0908 | 1.0842 | 209.976 | 50.9178 | 64430.7616 | 65013.4067 | 1296.2681 | 1196.2681 | 846.2519
1763 | 93.8669 | 0.99103 | 0.91304 | 0.92698 | 196.9832 | 49.0957 | 46317.0712 | 48332.7717 | 962.6554 | 862.6554 | 846.2725
1764 | 103.8772 | 0.99695 | 1.182 | 1.0902 | 199.7287 | 51.7551 | 72250.5412 | 72261.3019 | 1441.226 | 1341.226 | 846.284
1765 | 100.8455 | 1.0163 | 0.90571 | 1.04 | 205.0219 | 51.5353 | 54027.9165 | 52102.9585 | 1038.0592 | 938.0592 | 846.3173
1766 | 86.944 | 1.0769 | 0.93577 | 1.0325 | 203.0073 | 48.8739 | 52000.1835 | 53025.8454 | 1056.5169 | 956.5169 | 846.3187
1767 | 102.1272 | 0.98743 | 1.1434 | 1.0154 | 197.2833 | 51.9829 | 65791.5178 | 66459.0736 | 1325.1815 | 1225.1815 | 846.5227
1768 | 109.8601 | 1.0674 | 0.83183 | 0.97955 | 204.2646 | 53.985 | 50522.6801 | 50830.6262 | 1012.6125 | 912.6125 | 846.5864
1769 | 87.3185 | 1.0526 | 1.2493 | 0.94682 | 210.9225 | 52.7925 | 67510.5888 | 67531.8634 | 1346.6373 | 1246.6373 | 846.6066
1770 | 95.2623 | 1.0249 | 0.98003 | 1.0374 | 205.8026 | 45.0672 | 50426.6426 | 51336.4913 | 1022.7298 | 922.7298 | 846.8068
1771 | 107.1965 | 1.0661 | 0.98742 | 1.0269 | 207.9652 | 48.135 | 54515.2897 | 53895.9056 | 1073.9181 | 973.9181 | 846.905
1772 | 88.5341 | 0.95998 | 0.88143 | 1.0022 | 177.1642 | 47.0928 | 45778.4579 | 46289.1782 | 921.7836 | 821.7836 | 847.0249
1773 | 119.7283 | 1.065 | 0.9595 | 0.96162 | 216.0836 | 52.1338 | 54966.0956 | 54222.9627 | 1080.4593 | 980.4593 | 847.0894
1774 | 100.7812 | 0.95111 | 1.0049 | 0.99067 | 181.8314 | 50.4569 | 55247.8193 | 56136.6679 | 1118.7334 | 1018.7334 | 847.1411
1775 | 97.6655 | 1.004 | 0.87999 | 0.96048 | 202.9653 | 47.7135 | 45209.7979 | 46432.2709 | 924.6454 | 824.6454 | 847.2616
1776 | 113.7983 | 1.0672 | 1.2518 | 0.97305 | 195.4884 | 55.5677 | 74630.2548 | 75770.4473 | 1511.4089 | 1411.4089 | 847.3396
1777 | 100.3947 | 1.0314 | 0.98935 | 1.0849 | 192.5208 | 51.5376 | 60845.1558 | 59337.062 | 1182.7412 | 1082.7412 | 847.393
1778 | 115.5301 | 0.93358 | 1.1166 | 0.93484 | 221.616 | 48.995 | 56648.6196 | 57036.0026 | 1136.7201 | 1036.7201 | 847.4502
1779 | 113.2955 | 0.97076 | 0.82407 | 1.0925 | 216.5384 | 49.7014 | 50427.5736 | 51539.8641 | 1026.7973 | 926.7973 | 847.551
1780 | 81.6876 | 0.97496 | 1.0326 | 1.0154 | 197.2528 | 50.5378 | 57209.8911 | 58049.4124 | 1156.9882 | 1056.9882 | 847.6093
1781 | 96.5429 | 1.015 | 0.99378 | 1.0583 | 209.6688 | 48.0586 | 55463.2214 | 54983.8213 | 1095.6764 | 995.6764 | 847.6911
1782 | 90.7962 | 0.96295 | 1.1804 | 0.96934 | 169.6342 | 47.5886 | 58781.103 | 60294.7464 | 1201.8949 | 1101.8949 | 847.8267
1783 | 104.1287 | 1.0653 | 1.0141 | 0.93826 | 187.1609 | 48.3435 | 51548.2312 | 50524.6393 | 1006.4928 | 906.4928 | 847.8271
1784 | 76.1448 | 1.0059 | 1.123 | 0.98852 | 194.3688 | 44.5513 | 53064.9438 | 52444.9382 | 1044.8988 | 944.8988 | 847.9028
1785 | 97.5547 | 0.98882 | 1.176 | 1.0456 | 205.4215 | 47.699 | 63455.3862 | 63919.072 | 1274.3814 | 1174.3814 | 848.0425
1786 | 102.8411 | 1.022 | 0.85764 | 1.0129 | 211.4357 | 46.6942 | 45682.7367 | 45924.8634 | 914.4973 | 814.4973 | 848.1299
1787 | 91.4255 | 1.0227 | 0.87207 | 1.0626 | 194.9769 | 48.1043 | 49269.4788 | 48071.1415 | 957.4228 | 857.4228 | 848.4134
1788 | 104.891 | 0.96246 | 1.1206 | 0.99632 | 201.8576 | 49.4589 | 60413.8021 | 59791.3846 | 1191.8277 | 1091.8277 | 848.4256
1789 | 101.6946 | 1.0566 | 0.86914 | 0.97477 | 214.763 | 50.7424 | 48656.5793 | 48961.9259 | 975.2385 | 875.2385 | 848.5293
1790 | 82.2503 | 1.0198 | 0.78063 | 0.95579 | 201.7359 | 48.8914 | 40781.2271 | 41106.7897 | 818.1358 | 718.1358 | 848.8254
1791 | 80.4815 | 1.0504 | 0.92269 | 0.9794 | 185.6864 | 47.597 | 47222.0313 | 45864.9876 | 913.2998 | 813.2998 | 848.8832
1792 | 127.8865 | 0.96189 | 0.88225 | 1.025 | 197.4621 | 53.2104 | 54862.6545 | 55954.0666 | 1115.0813 | 1015.0813 | 848.9447
1793 | 96.2219 | 0.96676 | 0.93935 | 0.97526 | 182.9901 | 52.862 | 53527.7507 | 52248.5196 | 1040.9704 | 940.9704 | 849.0305
1794 | 89.3816 | 1.168 | 0.90314 | 0.9217 | 170.2476 | 49.9242 | 46940.2947 | 48360.9138 | 963.2183 | 863.2183 | 849.1119
1795 | 92.3763 | 1.0262 | 0.80961 | 1.0572 | 181.1023 | 52.8207 | 50400.5672 | 49178.75 | 979.575 | 879.575 | 849.1753
1796 | 94.9189 | 1.0391 | 1.0085 | 1.0052 | 195.2448 | 53.3854 | 59579.8744 | 60148.97 | 1198.9794 | 1098.9794 | 849.1975
1797 | 106.1482 | 1.0638 | 1.0944 | 0.98164 | 208.6539 | 47.0986 | 56125.5598 | 56969.2119 | 1135.3842 | 1035.3842 | 849.4113
1798 | 99.0292 | 1.1093 | 0.87693 | 1.0167 | 215.6614 | 49.8255 | 50113.5848 | 50854.4013 | 1013.088 | 913.088 | 849.5066
1799 | 104.8678 | 0.92875 | 0.97896 | 1.027 | 207.2483 | 50.344 | 55728.6138 | 54296.2172 | 1081.9243 | 981.9243 | 849.5225
1800 | 103.1934 | 1.0013 | 0.71064 | 0.98963 | 192.6683 | 49.7077 | 40286.4799 | 39448.8285 | 784.9766 | 684.9766 | 849.5266
1801 | 111.6374 | 0.91747 | 0.84116 | 1.0019 | 200.8629 | 48.9951 | 46508.8744 | 46572.0814 | 927.4416 | 827.4416 | 849.6243
1802 | 91.1352 | 1.0285 | 1.2557 | 0.98193 | 219.2577 | 46.509 | 61923.2 | 62980.1572 | 1255.6031 | 1155.6031 | 849.641
1803 | 96.8333 | 0.92437 | 0.81874 | 1.0426 | 205.7827 | 45.251 | 42883.9847 | 43584.7243 | 867.6945 | 767.6945 | 849.7155
1804 | 90.9571 | 0.98855 | 0.97915 | 1.0162 | 205.7439 | 49.7851 | 54217.0528 | 53816.1359 | 1072.3227 | 972.3227 | 849.7443
1805 | 110.7596 | 0.94148 | 0.86607 | 1.1057 | 204.9613 | 54.2022 | 57759.5038 | 57389.3447 | 1143.7869 | 1043.7869 | 849.7703
1806 | 91.9683 | 0.98593 | 1.0304 | 1.0459 | 199.5177 | 49.6445 | 58203.6418 | 57733.0523 | 1150.661 | 1050.661 | 849.8035
1807 | 80.868 | 1.0399 | 0.88627 | 1.0156 | 208.4673 | 50.4326 | 49842.6603 | 51749.3778 | 1030.9876 | 930.9876 | 849.8495
1808 | 90.3247 | 0.96999 | 1.0211 | 1.0121 | 210.7043 | 45.4847 | 51206.008 | 49656.6686 | 989.1334 | 889.1334 | 849.8872
1809 | 97.3593 | 0.93202 | 0.61965 | 0.86544 | 201.9033 | 50.0737 | 31598.4474 | 30237.9248 | 600.7585 | 500.7585 | 850.0257
1810 | 101.0035 | 1.0408 | 1.408 | 0.99992 | 203.2579 | 46.1971 | 70102.3053 | 70268.8013 | 1401.376 | 1301.376 | 850.0753
1811 | 104.4898 | 0.96526 | 1.2004 | 1.0259 | 196.7937 | 47.4895 | 63467.0106 | 64285.6452 | 1281.7129 | 1181.7129 | 850.1092
1812 | 110.5931 | 0.93684 | 0.95558 | 1.0118 | 208.3747 | 53.2067 | 57166.2503 | 56612.7742 | 1128.2555 | 1028.2555 | 850.1202
1813 | 112.1619 | 0.98049 | 1.0168 | 0.94175 | 200.7807 | 44.5026 | 47710.4221 | 48600.0165 | 968.0003 | 868.0003 | 850.1474
1814 | 93.406 | 0.9877 | 0.96234 | 1.0395 | 178.9897 | 48.9317 | 53640.3682 | 54026.7547 | 1076.5351 | 976.5351 | 850.2458
1815 | 108.9652 | 0.97726 | 1.0373 | 0.95976 | 189.9108 | 54.5441 | 60299.7461 | 59105.5123 | 1178.1102 | 1078.1102 | 850.3175
1816 | 95.2941 | 1.0024 | 0.935 | 0.99007 | 211.2977 | 48.9085 | 50158.478 | 50150.2776 | 999.0056 | 899.0056 | 850.334
1817 | 75.5005 | 0.96993 | 1.2844 | 1.029 | 203.0711 | 52.5368 | 73485.468 | 73190.9428 | 1459.8189 | 1359.8189 | 850.3418
1818 | 83.2367 | 1.0094 | 0.71078 | 1.0452 | 204.4733 | 53.7069 | 44616.4294 | 44176.7026 | 879.5341 | 779.5341 | 850.3771
1819 | 91.3745 | 1.0616 | 1.1792 | 1.0943 | 190.48 | 53.2894 | 74125.4685 | 75635.0341 | 1508.7007 | 1408.7007 | 850.519
1820 | 99.5709 | 1.0055 | 1.0431 | 1.0409 | 192.6264 | 46.4666 | 55295.9639 | 55979.8732 | 1115.5975 | 1015.5975 | 850.7101
1821 | 94.2831 | 1.088 | 1.2156 | 1.0427 | 203.2916 | 50.1615 | 68927.2697 | 69820.1387 | 1392.4028 | 1292.4028 | 850.7822
1822 | 80.2019 | 0.93904 | 1.2623 | 0.88478 | 204.9675 | 46.2574 | 55352.8671 | 56157.3444 | 1119.1469 | 1019.1469 | 850.8462
1823 | 104.0146 | 1.0038 | 1.1165 | 1.0336 | 209.7307 | 52.5549 | 66341.9642 | 64997.9944 | 1295.9599 | 1195.9599 | 850.8958
1824 | 98.6509 | 1.0208 | 1.0291 | 1.0239 | 176.9714 | 51.8429 | 60026.5265 | 61291.6826 | 1221.8337 | 1121.8337 | 850.9411
1825 | 90.3664 | 0.99604 | 0.92844 | 0.99595 | 203.7966 | 53.6714 | 54663.6928 | 56031.822 | 1116.6364 | 1016.6364 | 850.9784
1826 | 97.3995 | 1.0183 | 1.0466 | 0.94866 | 195.3987 | 49.6168 | 54378.5292 | 53576.707 | 1067.5341 | 967.5341 | 851.0032
1827 | 88.0957 | 1.028 | 1.2343 | 0.97071 | 216.5274 | 53.5147 | 69182.7729 | 69312.9595 | 1382.2592 | 1282.2592 | 851.0046
1828 | 99.9254 | 0.88994 | 1.1884 | 1.0237 | 191.4244 | 51.7877 | 67797.5123 | 69186.6362 | 1379.7327 | 1279.7327 | 851.006
1829 | 98.617 | 1.0151 | 0.96536 | 1.0176 | 218.2221 | 49.9074 | 54242.3557 | 55205.7512 | 1100.115 | 1000.115 | 851.0183
1830 | 77.9475 | 1.0663 | 0.98558 | 1.0203 | 177.5851 | 52.969 | 57846.9013 | 57812.376 | 1152.2475 | 1052.2475 | 851.0597
1831 | 71.8188 | 1.028 | 1.038 | 0.89971 | 190.6871 | 47.27 | 47827.893 | 49038.8365 | 976.7767 | 876.7767 | 851.0767
1832 | 113.8885 | 1.0216 | 1.036 | 1.0509 | 198.9429 | 50.062 | 60529.0652 | 61247.887 | 1220.9577 | 1120.9577 | 851.1209
1833 | 112.0641 | 1.0164 | 1.1011 | 1.0018 | 197.2963 | 48.0269 | 58646.4551 | 59602.6335 | 1188.0527 | 1088.0527 | 851.1272
1834 | 93.1525 | 0.93984 | 0.9308 | 0.96639 | 202.4058 | 49.4194 | 48982.2805 | 48423.2848 | 964.4657 | 864.4657 | 851.227
1835 | 104.5185 | 0.96279 | 1.0197 | 0.97362 | 194.7465 | 53.4423 | 58628.7362 | 60050.4557 | 1197.0091 | 1097.0091 | 851.3065
1836 | 103.6343 | 0.94686 | 1.1081 | 0.9788 | 192.7353 | 47.2387 | 56062.0637 | 56889.708 | 1133.7942 | 1033.7942 | 851.5149
1837 | 95.8014 | 0.97938 | 1.0086 | 0.98287 | 197.3059 | 46.3202 | 50463.802 | 49887.217 | 993.7443 | 893.7443 | 851.6758
1838 | 93.914 | 0.97949 | 1.0647 | 0.97367 | 207.7742 | 48.5213 | 54974.045 | 55430.4658 | 1104.6093 | 1004.6093 | 851.8361
1839 | 113.0485 | 0.97114 | 1.1585 | 0.98161 | 194.2631 | 49.5679 | 62003.0292 | 63502.4806 | 1266.0496 | 1166.0496 | 851.9387
1840 | 87.1185 | 0.96086 | 1.0216 | 1.0479 | 217.2448 | 51.2246 | 59343.7408 | 59782.068 | 1191.6414 | 1091.6414 | 852.0033
1841 | 92.7643 | 0.9378 | 1.0658 | 0.98217 | 190.0101 | 51.6108 | 58705.0345 | 57793.9629 | 1151.8793 | 1051.8793 | 852.0732
1842 | 90.1902 | 0.88182 | 1.1435 | 1.0266 | 208.7442 | 51.8357 | 65180.4091 | 65520.6153 | 1306.4123 | 1206.4123 | 852.0858
1843 | 107.2935 | 0.97598 | 0.97094 | 0.99697 | 198.2687 | 47.194 | 50823.9977 | 49503.411 | 986.0682 | 886.0682 | 852.3209
1844 | 92.5248 | 0.92088 | 1.0085 | 1.0568 | 191.0373 | 51.0783 | 58984.7909 | 57827.2267 | 1152.5445 | 1052.5445 | 852.3875
1845 | 105.2093 | 0.96299 | 0.7674 | 1.0159 | 204.2541 | 50.6382 | 44813.4256 | 43770.6913 | 871.4138 | 771.4138 | 852.712
1846 | 90.944 | 0.9403 | 1.3574 | 1.0513 | 208.5204 | 48.3406 | 73329.1358 | 77428.4832 | 1544.5697 | 1444.5697 | 852.72
1847 | 115.4129 | 0.86705 | 1.169 | 1.0139 | 210.6909 | 51.9225 | 66948.0291 | 66352.2954 | 1323.0459 | 1223.0459 | 852.8791
1848 | 85.978 | 1.0612 | 0.98255 | 1.0049 | 203.7257 | 48.1494 | 52135.4934 | 50529.4826 | 1006.5897 | 906.5897 | 852.9213
1849 | 98.8107 | 1.0575 | 0.93045 | 0.97866 | 192.1292 | 49.1143 | 50047.1203 | 49097.4439 | 977.9489 | 877.9489 | 852.9349
1850 | 94.7504 | 1.0504 | 0.57697 | 0.93227 | 198.5283 | 48.729 | 31259.4509 | 30092.0587 | 597.8412 | 497.8412 | 852.9901
1851 | 107.5033 | 1.0104 | 1.1391 | 1.0554 | 195.806 | 45.0495 | 59247.1486 | 58280.561 | 1161.6112 | 1061.6112 | 853.0036
1852 | 118.8976 | 1.0794 | 1.0622 | 0.99404 | 213.5768 | 51.3035 | 60965.8411 | 62214.6998 | 1240.294 | 1140.294 | 853.0809
1853 | 88.1407 | 1.0327 | 1.3134 | 1.0319 | 201.6194 | 45.3598 | 65805.965 | 66960.4733 | 1335.2095 | 1235.2095 | 853.1555
1854 | 99.1106 | 1.0237 | 0.89165 | 0.98455 | 185.4667 | 52.0745 | 51183.6978 | 50413.3209 | 1004.2664 | 904.2664 | 853.1725
1855 | 107.4609 | 1.0176 | 1.0234 | 1.0091 | 201.5747 | 45.9455 | 52675.9412 | 55696.9382 | 1109.9388 | 1009.9388 | 853.2666
1856 | 93.1975 | 0.97432 | 0.98401 | 1.0035 | 208.2524 | 50.3211 | 54466.0717 | 55945.4452 | 1114.9089 | 1014.9089 | 853.3094
1857 | 93.6339 | 0.95011 | 0.90959 | 1.004 | 177.1331 | 50.8217 | 51110.9518 | 51563.5369 | 1027.2707 | 927.2707 | 853.3418
1858 | 88.9639 | 1.0203 | 1.0295 | 0.8905 | 219.2365 | 48.6742 | 49262.4446 | 49391.4485 | 983.829 | 883.829 | 853.531
1859 | 107.2159 | 1.1044 | 0.90853 | 0.93334 | 204.3618 | 48.2291 | 46812.1561 | 46323.8294 | 922.4766 | 822.4766 | 853.5678
1860 | 92.5856 | 0.97975 | 0.74462 | 1.0249 | 184.0094 | 54.4855 | 46706.2596 | 46570.806 | 927.4161 | 827.4161 | 853.5757
1861 | 103.2605 | 0.9611 | 0.92853 | 1.0351 | 195.9212 | 48.8929 | 52041.9211 | 54333.4624 | 1082.6692 | 982.6692 | 853.5853
1862 | 95.4645 | 0.99926 | 0.7819 | 0.97106 | 209.1452 | 50.5206 | 43387.1287 | 42471.7046 | 845.4341 | 745.4341 | 853.5936
1863 | 84.7183 | 0.97284 | 1.1961 | 1.0429 | 197.8861 | 48.6761 | 64925.2787 | 63247.4981 | 1260.95 | 1160.95 | 853.615
1864 | 97.809 | 0.9886 | 1.0043 | 1.0426 | 186.1787 | 45.7455 | 52511.4592 | 52940.4135 | 1054.8083 | 954.8083 | 853.6196
1865 | 117.6032 | 1.0246 | 0.95897 | 1.0287 | 188.7583 | 48.8425 | 54256.0839 | 56352.6898 | 1123.0538 | 1023.0538 | 853.6237
1866 | 87.5956 | 0.99308 | 1.0445 | 0.98833 | 204.9873 | 50.835 | 57104.6476 | 56149.713 | 1118.9943 | 1018.9943 | 853.6416
1867 | 84.5446 | 0.95864 | 0.97415 | 0.98005 | 200.4716 | 47.887 | 49800.0831 | 48576.9602 | 967.5392 | 867.5392 | 853.661
1868 | 91.7655 | 0.92859 | 1.0353 | 1.0287 | 204.0212 | 57.0459 | 65816.2142 | 64665.7156 | 1289.3143 | 1189.3143 | 853.6928
1869 | 111.6709 | 1.0286 | 1.0867 | 0.98507 | 189.4917 | 47.7547 | 56795.6601 | 55486.603 | 1105.7321 | 1005.7321 | 853.7207
1870 | 86.5066 | 1.0633 | 0.9839 | 0.95452 | 198.3018 | 54.1176 | 56000.8351 | 56343.7555 | 1122.8751 | 1022.8751 | 853.7513
1871 | 88.7469 | 1.0201 | 0.94302 | 1.0769 | 187.7891 | 51.7715 | 57450.6744 | 58613.612 | 1168.2722 | 1068.2722 | 853.7702
1872 | 92.6264 | 1.0108 | 0.8907 | 0.87949 | 206.6265 | 49.8769 | 43948.0384 | 43139.6542 | 858.7931 | 758.7931 | 853.9618
1873 | 111.4156 | 1.0388 | 1.121 | 0.97377 | 216.9631 | 46.9945 | 56953.6172 | 56800.2669 | 1132.0053 | 1032.0053 | 853.9986
1874 | 113.006 | 0.92615 | 1.0487 | 0.92386 | 206.2592 | 49.2016 | 53024.0643 | 53336.6249 | 1062.7325 | 962.7325 | 854.0005
1875 | 102.9075 | 0.92293 | 1.0105 | 0.90669 | 199.1723 | 45.8811 | 46592.7658 | 45924.5373 | 914.4907 | 814.4907 | 854.058
1876 | 96.6329 | 0.96083 | 0.7915 | 0.90937 | 216.9213 | 50.8615 | 41548.0936 | 42673.0265 | 849.4605 | 749.4605 | 854.1155
1877 | 101.5808 | 0.98807 | 0.88685 | 0.89911 | 216.5608 | 52.648 | 47481.0376 | 47761.349 | 951.227 | 851.227 | 854.1429
1878 | 99.4151 | 0.8481 | 1.0944 | 1.0141 | 199.1345 | 49.03 | 58748.178 | 58867.53 | 1173.3506 | 1073.3506 | 854.169
1879 | 75.967 | 1.0143 | 1.0021 | 1.0629 | 199.2452 | 51.2261 | 58711.8002 | 61594.6523 | 1227.893 | 1127.893 | 854.2353
1880 | 107.284 | 1.0419 | 1.224 | 1.0708 | 200.5472 | 52.4115 | 74748.2153 | 74336.3439 | 1482.7269 | 1382.7269 | 854.2417
1881 | 92.2039 | 0.98101 | 1.2959 | 1.048 | 186.3372 | 45.2492 | 65731.4856 | 66995.9003 | 1335.918 | 1235.918 | 854.391
1882 | 107.8013 | 0.92106 | 0.80904 | 0.99835 | 199.8405 | 49.035 | 44674.3071 | 44824.3651 | 892.4873 | 792.4873 | 854.5337
1883 | 87.8479 | 1.0104 | 1.0595 | 0.96487 | 188.9181 | 51.5473 | 57458.5855 | 58640.734 | 1168.8147 | 1068.8147 | 854.5585
1884 | 92.4581 | 0.90652 | 1.1113 | 0.97297 | 201.0862 | 52.4849 | 61351.9168 | 62278.6809 | 1241.5736 | 1141.5736 | 854.7442
1885 | 90.0457 | 1.0792 | 0.89308 | 0.99078 | 197.2601 | 47.4625 | 46806.2884 | 45833.6362 | 912.6727 | 812.6727 | 854.7592
1886 | 101.469 | 1.0354 | 1.0591 | 0.9787 | 210.3944 | 48.5434 | 55628.992 | 56432.0423 | 1124.6408 | 1024.6408 | 854.8053
1887 | 121.9929 | 1.0716 | 1.2468 | 0.94704 | 206.817 | 50.563 | 66519.1769 | 65909.5469 | 1314.1909 | 1214.1909 | 854.8634
1888 | 92.2229 | 1.0051 | 1.1993 | 1.0595 | 193.8016 | 50.1108 | 68513.2569 | 69400.1042 | 1384.0021 | 1284.0021 | 854.979
1889 | 98.3883 | 0.97143 | 0.93174 | 0.94572 | 200.0053 | 55.1475 | 54064.5509 | 55599.281 | 1107.9856 | 1007.9856 | 855.0728
1890 | 81.6798 | 0.98381 | 0.9525 | 1.0158 | 195.7634 | 47.0353 | 49482.467 | 50244.9298 | 1000.8986 | 900.8986 | 855.086
1891 | 96.7043 | 1.049 | 1.3902 | 1.0993 | 200.1869 | 51.9015 | 84786.7614 | 81664.6329 | 1629.2927 | 1529.2927 | 855.1545
1892 | 120.2917 | 0.972 | 1.2772 | 1.0544 | 190.6596 | 51.6583 | 75800.0673 | 78070.1771 | 1557.4035 | 1457.4035 | 855.1572
1893 | 104.1642 | 1.0139 | 1.1772 | 0.98931 | 186.1328 | 50.6279 | 64497.3536 | 64792.8134 | 1291.8563 | 1191.8563 | 855.2324
1894 | 82.2916 | 1.0399 | 1.09 | 0.98265 | 188.2724 | 53.3747 | 61923.4703 | 64250.6828 | 1281.0137 | 1181.0137 | 855.2462
1895 | 89.7126 | 1.091 | 1.156 | 0.94113 | 211.0979 | 47.2276 | 56212.8213 | 55864.599 | 1113.292 | 1013.292 | 855.2781
1896 | 104.4765 | 1.0783 | 1.0395 | 1.0464 | 203.0275 | 48.4096 | 58310.4096 | 57828.5175 | 1152.5704 | 1052.5704 | 855.4006
1897 | 89.4924 | 1.021 | 0.56934 | 1.0017 | 196.5664 | 49.9486 | 33246.4149 | 33414.6098 | 664.2922 | 564.2922 | 855.5304
1898 | 110.0115 | 0.97712 | 1.1224 | 0.93123 | 174.3903 | 47.2675 | 54659.1119 | 53186.3848 | 1059.7277 | 959.7277 | 855.539
1899 | 88.3895 | 0.96738 | 1.3025 | 0.93345 | 203.3147 | 53.4231 | 69724.9631 | 66928.6458 | 1334.5729 | 1234.5729 | 855.5445
1900 | 99.4562 | 1.0079 | 1.0069 | 1.0972 | 212.5951 | 51.1445 | 61844.7935 | 61893.9902 | 1233.8798 | 1133.8798 | 855.6722
1901 | 98.5226 | 1.0195 | 0.99937 | 0.98053 | 181.126 | 51.5618 | 55886.5661 | 55629.5423 | 1108.5908 | 1008.5908 | 855.6872
1902 | 103.3037 | 0.98177 | 0.77364 | 0.92963 | 202.2397 | 48.477 | 39983.0495 | 39734.9677 | 790.6994 | 690.6994 | 855.7432
1903 | 113.0649 | 0.95614 | 0.98198 | 0.9556 | 184.1575 | 52.4797 | 55103.2356 | 56755.1279 | 1131.1026 | 1031.1026 | 855.7883
1904 | 90.2007 | 1.0439 | 0.82012 | 0.956 | 191.6197 | 48.9583 | 43186.6355 | 44074.817 | 877.4963 | 777.4963 | 855.7985
1905 | 89.8957 | 1.0597 | 1.0144 | 1.0251 | 200.921 | 51.2647 | 58391.4872 | 56996.9342 | 1135.9387 | 1035.9387 | 855.8673
1906 | 97.9082 | 0.9853 | 0.84604 | 1.0486 | 198.1045 | 48.3856 | 47789.5343 | 49961.1967 | 995.2239 | 895.2239 | 855.9159
1907 | 100.4025 | 1.0388 | 1.0863 | 1.0116 | 201.4998 | 45.3949 | 54818.3472 | 54198.4914 | 1079.9698 | 979.9698 | 855.9381
1908 | 103.4218 | 1.0582 | 1.0567 | 0.96153 | 202.3062 | 50.2708 | 56783.1589 | 56727.4959 | 1130.5499 | 1030.5499 | 855.9608
1909 | 99.4352 | 0.99576 | 0.9896 | 0.92765 | 194.6655 | 49.4118 | 50447.2873 | 51419.6435 | 1024.3929 | 924.3929 | 855.9843
1910 | 97.9729 | 1.0712 | 0.96591 | 1.004 | 193.9696 | 46.7273 | 50412.21 | 50233.823 | 1000.6765 | 900.6765 | 856.0151
1911 | 105.9982 | 0.99833 | 1.3297 | 1.0364 | 193.0722 | 55.6703 | 82801.7944 | 81323.0754 | 1622.4615 | 1522.4615 | 856.0256
1912 | 108.4326 | 0.9412 | 1.029 | 1.0025 | 201.6616 | 55.4496 | 63063.6131 | 61622.907 | 1228.4581 | 1128.4581 | 856.0715
1913 | 104.166 | 0.99449 | 1.0006 | 0.98612 | 198.229 | 54.6461 | 59781.8169 | 59233.2497 | 1180.665 | 1080.665 | 856.1605
1914 | 99.2257 | 0.92623 | 1.0699 | 0.91884 | 209.151 | 53.2621 | 57465.6509 | 57517.1851 | 1146.3437 | 1046.3437 | 856.3131
1915 | 91.1014 | 1.0928 | 1.1266 | 0.95293 | 200.5744 | 50.7541 | 59741.8092 | 58327.4145 | 1162.5483 | 1062.5483 | 856.4318
1916 | 100.853 | 1.007 | 1.0511 | 1.0073 | 188.2174 | 47.604 | 55427.0718 | 55788.8619 | 1111.7772 | 1011.7772 | 856.4535
1917 | 91.87 | 0.92066 | 1.0515 | 1.0789 | 193.8657 | 51.9507 | 63522.611 | 62756.9342 | 1251.1387 | 1151.1387 | 856.5676
1918 | 112.3336 | 0.95488 | 1.3402 | 0.94522 | 200.1335 | 47.3624 | 65280.3222 | 66982.0619 | 1335.6412 | 1235.6412 | 856.5769
1919 | 91.8412 | 1.0159 | 1.1491 | 1.0416 | 202.7574 | 51.3614 | 66465.758 | 64939.1439 | 1294.7829 | 1194.7829 | 856.5961
1920 | 94.855 | 0.94364 | 0.90157 | 0.98356 | 195.1093 | 47.0776 | 46154.9989 | 46292.7252 | 921.8545 | 821.8545 | 856.6266
1921 | 106.6699 | 0.99262 | 0.94815 | 1.0828 | 196.8553 | 53.2672 | 60526.1936 | 60754.8763 | 1211.0975 | 1111.0975 | 856.6437
1922 | 90.4283 | 0.98429 | 0.85568 | 1.0173 | 195.3494 | 49.4505 | 47643.6806 | 47651.5266 | 949.0305 | 849.0305 | 856.6873
1923 | 106.0965 | 1.0536 | 1.0636 | 0.99366 | 192.3649 | 49.8207 | 58414.0165 | 57802.0432 | 1152.0409 | 1052.0409 | 856.8361
1924 | 124.9744 | 1.0095 | 0.86062 | 0.97654 | 195.0451 | 49.2345 | 47784.545 | 48681.9933 | 969.6399 | 869.6399 | 856.9283
1925 | 90.9841 | 1.0054 | 1.0539 | 0.98681 | 199.9544 | 49.6036 | 56322.7078 | 57652.4365 | 1149.0487 | 1049.0487 | 856.948
1926 | 104.4787 | 1.097 | 1.014 | 1.0479 | 208.3731 | 47.3067 | 55893.6642 | 56133.086 | 1118.6617 | 1018.6617 | 856.9497
1927 | 98.537 | 0.99448 | 1.0451 | 0.97046 | 201.8987 | 57.4388 | 64087.7425 | 62912.4375 | 1254.2488 | 1154.2488 | 857.1467
1928 | 102.6002 | 0.95063 | 0.86122 | 1.0947 | 193.0875 | 50.3513 | 52572.0748 | 53674.0032 | 1069.4801 | 969.4801 | 857.1639
1929 | 108.0973 | 0.96993 | 0.99809 | 1.0074 | 210.9311 | 50.5962 | 56387.3493 | 55088.5678 | 1097.7714 | 997.7714 | 857.2512
1930 | 96.2617 | 1.0405 | 1.0608 | 0.98895 | 193.147 | 47.6557 | 54963.0249 | 55178.4961 | 1099.5699 | 999.5699 | 857.321
1931 | 116.7833 | 1.0568 | 1.0349 | 0.97222 | 196.8564 | 47.189 | 53500.1187 | 53284.5311 | 1061.6906 | 961.6906 | 857.3432
1932 | 101.3522 | 0.97181 | 0.79543 | 1.0016 | 203.7212 | 52.1473 | 46883.7899 | 45515.1629 | 906.3033 | 806.3033 | 857.3727
1933 | 88.3899 | 1.0156 | 0.78171 | 1.0395 | 220.7792 | 52.3066 | 47420.1558 | 48496.8492 | 965.937 | 865.937 | 857.3949
1934 | 93.7565 | 0.9554 | 1.1421 | 0.8775 | 206.9035 | 53.8563 | 59005.7674 | 57197.0332 | 1139.9407 | 1039.9407 | 857.4228
1935 | 105.842 | 0.98627 | 1.1496 | 1.111 | 204.4865 | 51.3223 | 71108.1302 | 71794.8724 | 1431.8974 | 1331.8974 | 857.4421
1936 | 83.5567 | 1.0922 | 1.0446 | 0.97625 | 193.2245 | 46.6612 | 52034.6505 | 51719.7801 | 1030.3956 | 930.3956 | 857.5286
1937 | 98.9724 | 0.98877 | 0.86107 | 1.0668 | 199.2427 | 49.7182 | 50733.4298 | 50498.991 | 1005.9798 | 905.9798 | 857.7527
1938 | 85.9757 | 0.98398 | 0.85782 | 0.993 | 206.1996 | 52.9381 | 49778.3585 | 49340.794 | 982.8159 | 882.8159 | 857.851
1939 | 95.7033 | 1.01 | 0.88505 | 0.88357 | 184.5469 | 50.2022 | 44295.5076 | 43498.2842 | 865.9657 | 765.9657 | 857.9082
1940 | 90.4933 | 1.1039 | 0.90527 | 0.96859 | 192.7953 | 52.1372 | 51117.1403 | 51757.8224 | 1031.1564 | 931.1564 | 858.0534
1941 | 112.9251 | 0.95557 | 0.88718 | 0.91966 | 188.9136 | 49.6621 | 46067.1993 | 46480.0313 | 925.6006 | 825.6006 | 858.0746
1942 | 114.4077 | 1.059 | 1.0488 | 0.993 | 197.6104 | 48.1747 | 56207.1142 | 55115.2644 | 1098.3053 | 998.3053 | 858.0995
1943 | 93.626 | 1.0867 | 1.0573 | 0.98127 | 190.974 | 47.6094 | 54430.4487 | 53698.7182 | 1069.9744 | 969.9744 | 858.119
1944 | 96.8701 | 0.98891 | 1.0893 | 1.041 | 195.4119 | 48.3876 | 59698.8454 | 59848.674 | 1192.9735 | 1092.9735 | 858.1497
1945 | 109.3859 | 1.0311 | 1.1866 | 0.98245 | 197.8729 | 55.891 | 71658.3607 | 69685.0473 | 1389.7009 | 1289.7009 | 858.1651
1946 | 110.1962 | 0.96295 | 1.0379 | 1.0512 | 215.787 | 49.4672 | 59437.6977 | 59865.7642 | 1193.3153 | 1093.3153 | 858.2418
1947 | 111.56 | 1.0046 | 0.93105 | 0.99129 | 199.7814 | 54.1662 | 56262.3767 | 57871.579 | 1153.4316 | 1053.4316 | 858.2774
1948 | 104.2363 | 0.9131 | 0.83642 | 0.98583 | 200.7149 | 44.4705 | 41102.3746 | 41807.9969 | 832.1599 | 732.1599 | 858.4793
1949 | 89.3589 | 1.0076 | 1.25 | 0.95751 | 209.3247 | 49.0747 | 63362.9365 | 63920.2609 | 1274.4052 | 1174.4052 | 858.5379
1950 | 109.4504 | 0.94666 | 0.96024 | 1.015 | 190.4701 | 48.3689 | 52345.2404 | 52868.7866 | 1053.3757 | 953.3757 | 858.6106
1951 | 109.7809 | 0.94815 | 1.0447 | 0.96774 | 194.3536 | 49.9231 | 55862.0744 | 56177.4817 | 1119.5496 | 1019.5496 | 858.6962
1952 | 130.5841 | 1.0635 | 1.0106 | 1.0328 | 193.3476 | 44.3333 | 52621.9437 | 53761.1229 | 1071.2225 | 971.2225 | 858.7799
1953 | 99.9234 | 1.0109 | 1.0486 | 1.0472 | 197.5857 | 42.7962 | 51517.8482 | 51330.4193 | 1022.6084 | 922.6084 | 858.7955
1954 | 92.2699 | 0.9868 | 1.0045 | 0.96063 | 200.1109 | 46.6864 | 49501.2239 | 50161.2793 | 999.2256 | 899.2256 | 859.0631
1955 | 91.0431 | 1.0667 | 0.88281 | 1.0159 | 206.5965 | 52.9385 | 52827.7161 | 52855.73 | 1053.1146 | 953.1146 | 859.0795
1956 | 94.5616 | 1.0094 | 1.0611 | 0.97974 | 200.6087 | 44.8223 | 51074.5144 | 51491.9814 | 1025.8396 | 925.8396 | 859.153
1957 | 130.2024 | 1.0065 | 1.0513 | 1.0046 | 194.1803 | 55.3366 | 65891.6037 | 65512.6345 | 1306.2527 | 1206.2527 | 859.2075
1958 | 98.8683 | 1.0411 | 1.0336 | 0.95317 | 210.4694 | 49.6481 | 54231.5062 | 52691.9427 | 1049.8389 | 949.8389 | 859.5021
1959 | 121.0318 | 1.1025 | 0.84102 | 1.0191 | 197.6174 | 50.8151 | 50530.6163 | 49935.3875 | 994.7077 | 894.7077 | 859.5414
1960 | 100.147 | 1.0356 | 1.0716 | 1.0284 | 207.0202 | 45.6958 | 55303.1952 | 55413.8554 | 1104.2771 | 1004.2771 | 859.6098
1961 | 89.198 | 0.92886 | 0.83847 | 1.056 | 203.5068 | 49.4304 | 48064.2802 | 47290.2665 | 941.8053 | 841.8053 | 859.6184
1962 | 88.2721 | 1.0772 | 1.1052 | 0.9878 | 201.3389 | 50.7296 | 60409.1371 | 60003.0919 | 1196.0618 | 1096.0618 | 859.7023
1963 | 96.957 | 0.96306 | 1.077 | 1.0147 | 196.0973 | 52.8103 | 62840.1238 | 62127.5456 | 1238.5509 | 1138.5509 | 859.7677
1964 | 101.1363 | 1.0793 | 0.9093 | 0.97547 | 216.0225 | 48.7067 | 48734.9948 | 48104.9744 | 958.0995 | 858.0995 | 859.7929
1965 | 89.9885 | 0.9921 | 0.72349 | 1.0014 | 202.2417 | 48.705 | 39837.2065 | 40864.7599 | 813.2952 | 713.2952 | 859.9316
1966 | 100.8595 | 0.91808 | 1.1887 | 0.96966 | 197.2005 | 48.0125 | 59985.9774 | 59460.9309 | 1185.2186 | 1085.2186 | 859.9402
1967 | 81.5565 | 0.89133 | 0.98552 | 0.94354 | 219.3158 | 47.0979 | 47438.2747 | 47563.0816 | 947.2616 | 847.2616 | 859.996
1968 | 96.0785 | 0.9598 | 1.0443 | 1.0391 | 210.3077 | 45.3142 | 53562.0214 | 53217.7071 | 1060.3541 | 960.3541 | 860.0238
1969 | 95.0828 | 1.0219 | 1.0175 | 0.99082 | 208.7057 | 48.9015 | 54261.2755 | 55064.6187 | 1097.2924 | 997.2924 | 860.0256
1970 | 93.6981 | 0.95815 | 1.0783 | 1.0431 | 201.2522 | 45.1346 | 55023.3129 | 56832.2385 | 1132.6448 | 1032.6448 | 860.1364
1971 | 88.578 | 0.95128 | 1.0755 | 0.97791 | 200.2503 | 49.4279 | 56353.071 | 53408.7931 | 1064.1759 | 964.1759 | 860.1912
1972 | 106.6934 | 1.0722 | 1.1709 | 1.0008 | 205.2696 | 53.0078 | 68385.3978 | 68784.9459 | 1371.6989 | 1271.6989 | 860.2553
1973 | 99.2004 | 1.0075 | 1.2859 | 1.0095 | 191.5684 | 48.322 | 67749.1099 | 66793.9806 | 1331.8796 | 1231.8796 | 860.2624
1974 | 107.4769 | 0.91969 | 0.90806 | 0.96306 | 196.4131 | 48.165 | 47078.2964 | 47515.863 | 946.3173 | 846.3173 | 860.2667
1975 | 100.8699 | 1.0879 | 0.72575 | 0.96087 | 184.3053 | 51.3944 | 41664.1747 | 42742.0568 | 850.8411 | 750.8411 | 860.3192
1976 | 102.5717 | 1.0154 | 1.0726 | 1.0079 | 188.764 | 47.0678 | 55974.2242 | 54640.3814 | 1088.8076 | 988.8076 | 860.3735
1977 | 99.3338 | 0.97584 | 1.0554 | 0.93097 | 208.263 | 50.5421 | 54766.516 | 52993.6339 | 1055.8727 | 955.8727 | 860.4013
1978 | 93.5621 | 0.91622 | 0.89535 | 0.92052 | 196.3907 | 52.9261 | 48354.2683 | 47405.8455 | 944.1169 | 844.1169 | 860.4666
1979 | 82.5709 | 0.98966 | 0.8212 | 0.93796 | 193.6487 | 52.0519 | 44540.417 | 42333.1109 | 842.6622 | 742.6622 | 860.4941
1980 | 84.1346 | 0.97834 | 0.88546 | 0.99956 | 192.9443 | 49.8007 | 48369.0725 | 49112.4329 | 978.2487 | 878.2487 | 860.6071
1981 | 109.4895 | 0.9735 | 0.91237 | 1.0931 | 198.4858 | 47.2125 | 52315.6657 | 52406.6667 | 1044.1333 | 944.1333 | 860.7371
1982 | 91.5556 | 1.0479 | 0.89374 | 0.97036 | 201.7219 | 49.8307 | 48198.1776 | 48267.8589 | 961.3572 | 861.3572 | 860.7651
1983 | 116.4584 | 0.99802 | 0.79404 | 1.0222 | 204.4053 | 50.1078 | 46697.9248 | 47395.8272 | 943.9165 | 843.9165 | 860.8193
1984 | 112.2031 | 1.0448 | 1.0351 | 1.0913 | 194.7027 | 52.2141 | 65299.1628 | 66189.8396 | 1319.7968 | 1219.7968 | 860.85
1985 | 103.9508 | 0.97632 | 0.96056 | 0.98077 | 181.9128 | 51.6993 | 54134.0108 | 54320.5299 | 1082.4106 | 982.4106 | 860.9711
1986 | 90.7287 | 0.8864 | 0.86777 | 0.93633 | 200.0533 | 52.3339 | 46931.2144 | 45838.3422 | 912.7668 | 812.7668 | 860.9822
1987 | 93.8807 | 0.99329 | 1.1604 | 1.0091 | 181.3731 | 50.6525 | 64219.3945 | 64633.0433 | 1288.6609 | 1188.6609 | 860.9844
1988 | 99.3645 | 0.9968 | 0.95044 | 0.92382 | 206.5375 | 49.4438 | 48517.1798 | 49349.5965 | 982.9919 | 882.9919 | 861.0439
1989 | 120.6892 | 1.0043 | 1.1905 | 1.0876 | 207.5249 | 47.4365 | 67381.8123 | 67617.6872 | 1348.3537 | 1248.3537 | 861.092
1990 | 96.575 | 0.94775 | 0.94296 | 0.94295 | 205.075 | 54.1769 | 53336.2581 | 51552.4622 | 1027.0492 | 927.0492 | 861.2424
1991 | 100.4242 | 1.0944 | 1.3446 | 1.0561 | 197.7326 | 49.3482 | 75695.9853 | 77337.9416 | 1542.7588 | 1442.7588 | 861.2638
1992 | 111.3027 | 1.0178 | 0.97593 | 1.0204 | 177.4242 | 45.8494 | 51028.6319 | 51466.9662 | 1025.3393 | 925.3393 | 861.2683
1993 | 90.047 | 0.98656 | 0.76921 | 0.92751 | 203.5036 | 49.2687 | 39731.0702 | 40629.2558 | 808.5851 | 708.5851 | 861.2704
1994 | 120.4137 | 1.033 | 0.91136 | 0.93143 | 187.2798 | 49.8181 | 48673.1068 | 49482.7611 | 985.6552 | 885.6552 | 861.3293
1995 | 98.3835 | 1.0381 | 0.78303 | 0.98359 | 185.3758 | 48.0841 | 42129.8058 | 43077.9479 | 857.559 | 757.559 | 861.343
1996 | 102.4139 | 0.94543 | 0.92471 | 1.0308 | 204.818 | 51.4764 | 54257.9204 | 56128.3652 | 1118.5673 | 1018.5673 | 861.3539
1997 | 95.7154 | 1.0009 | 1.0554 | 0.97921 | 188.3436 | 49.9264 | 56567.3623 | 54860.8836 | 1093.2177 | 993.2177 | 861.3572
1998 | 84.8669 | 0.9955 | 1.1087 | 0.97316 | 206.7342 | 52.3257 | 61083.3242 | 61017.4886 | 1216.3498 | 1116.3498 | 861.3784
1999 | 85.8929 | 1.0657 | 1.259 | 1.109 | 184.8704 | 52.1266 | 77732.5719 | 75680.1371 | 1509.6027 | 1409.6027 | 861.4163
2000 | 102.8419 | 0.97122 | 1.0278 | 1.0876 | 195.8381 | 47.0603 | 57503.2899 | 57725.5067 | 1150.5101 | 1050.5101 | 861.5675
2001 | 101.7891 | 0.8924 | 1.0427 | 0.97858 | 206.4881 | 49.751 | 55489.56 | 54178.674 | 1079.5735 | 979.5735 | 861.6337
2002 | 104.2795 | 0.95345 | 0.97297 | 1.0178 | 186.7457 | 46.1167 | 50441.1496 | 50618.8968 | 1008.3779 | 908.3779 | 861.6805
2003 | 89.9498 | 0.91636 | 0.95064 | 1.0555 | 187.8369 | 51.6799 | 56303.334 | 56884.4168 | 1133.6883 | 1033.6883 | 861.7586
2004 | 75.0178 | 0.95745 | 0.79085 | 0.99276 | 207.4132 | 51.0233 | 43932.1219 | 44204.0965 | 880.0819 | 780.0819 | 861.9104
2005 | 103.8819 | 1.0158 | 1.3302 | 0.94858 | 205.4818 | 50.0682 | 68664.4591 | 71891.4308 | 1433.8286 | 1333.8286 | 861.9875
2006 | 95.2911 | 0.96297 | 1.1555 | 1.0793 | 186.6763 | 50.0445 | 67192.8785 | 67383.4267 | 1343.6685 | 1243.6685 | 862.0155
2007 | 83.6049 | 1.0277 | 1.026 | 1.0009 | 213.9261 | 51.1746 | 57164.0568 | 57078.9039 | 1137.5781 | 1037.5781 | 862.1315
2008 | 89.4874 | 0.9645 | 0.91602 | 0.97654 | 219.5556 | 48.2523 | 47547.4088 | 47658.7659 | 949.1753 | 849.1753 | 862.3193
2009 | 96.4613 | 0.99108 | 1.1563 | 0.97974 | 196.3124 | 51.6528 | 63650.2834 | 63672.4492 | 1269.449 | 1169.449 | 862.4442
2010 | 86.8922 | 1.1301 | 0.80119 | 1.0961 | 203.9898 | 57.521 | 56364.8379 | 56840.4974 | 1132.8099 | 1032.8099 | 862.4703
2011 | 97.1313 | 0.94937 | 0.90795 | 1.0127 | 202.0329 | 54.5326 | 55371.3225 | 53780.9644 | 1071.6193 | 971.6193 | 862.4831
2012 | 91.9971 | 0.96821 | 1.071 | 0.9839 | 198.8474 | 53.9904 | 61898.8279 | 60788.5635 | 1211.7713 | 1111.7713 | 862.5766
2013 | 103.0841 | 1.0847 | 1.2198 | 1.121 | 210.6861 | 50.24 | 74528.1198 | 72954.4119 | 1455.0882 | 1355.0882 | 862.5946
2014 | 87.7156 | 1.0068 | 1.0963 | 0.87686 | 202.8096 | 48.4705 | 51076.7012 | 51516.7389 | 1026.3348 | 926.3348 | 862.6554
2015 | 115.8596 | 1.0375 | 0.98219 | 1.0762 | 194.4227 | 49.505 | 58472.2061 | 57419.7462 | 1144.3949 | 1044.3949 | 862.8425
2016 | 92.2508 | 1.0371 | 0.85701 | 0.91212 | 206.7429 | 45.7848 | 40377.0146 | 40959.7067 | 815.1941 | 715.1941 | 862.854
2017 | 104.6805 | 1.0283 | 0.72835 | 1.0353 | 216.8725 | 48.0354 | 41610.2101 | 41367.8895 | 823.3578 | 723.3578 | 862.8834
2018 | 95.495 | 0.93141 | 1.1735 | 0.9027 | 186.3439 | 56.9675 | 65597.6431 | 66246.5626 | 1320.9313 | 1220.9313 | 862.9802
2019 | 97.4066 | 0.99463 | 1.1201 | 1.0049 | 177.3357 | 48.1222 | 59003.6195 | 58240.6633 | 1160.8133 | 1060.8133 | 863.0166
2020 | 103.4466 | 1.0161 | 0.95987 | 0.98247 | 213.0611 | 51.7018 | 54404.919 | 54164.6603 | 1079.2932 | 979.2932 | 863.1714
2021 | 83.9784 | 1.0212 | 0.85236 | 0.94916 | 212.3205 | 51.9442 | 46690.8837 | 47580.4663 | 947.6093 | 847.6093 | 863.2183
2022 | 86.1739 | 0.912 | 0.69789 | 1.0736 | 203.6178 | 50.3498 | 41886.6137 | 42680.9026 | 849.6181 | 749.6181 | 863.239
2023 | 99.5969 | 0.94939 | 0.91569 | 1.0123 | 188.8927 | 49.8242 | 51085.2782 | 50447.3519 | 1004.947 | 904.947 | 863.2417
2024 | 87.1029 | 0.98458 | 1.2781 | 0.94631 | 189.8814 | 46.3924 | 60280.8416 | 62022.0516 | 1236.441 | 1136.441 | 863.2955
2025 | 85.9197 | 1.0261 | 0.72068 | 0.98368 | 213.197 | 49.7354 | 39856.3228 | 39995.7295 | 795.9146 | 695.9146 | 863.3185
2026 | 106.7646 | 0.92595 | 0.81343 | 0.96724 | 187.4947 | 47.6846 | 42419.1395 | 41910.5218 | 834.2104 | 734.2104 | 863.3543
2027 | 95.7672 | 0.95979 | 0.98895 | 0.96713 | 201.1351 | 51.6398 | 54337.9978 | 54989.876 | 1095.7975 | 995.7975 | 863.4024
2028 | 98.2449 | 0.95435 | 0.98767 | 0.96081 | 214.6337 | 48.9752 | 51282.4768 | 49660.4185 | 989.2084 | 889.2084 | 863.4493
2029 | 117.7647 | 1.0156 | 0.75207 | 1.0054 | 184.8754 | 50.2668 | 44203.7901 | 45311.0926 | 902.2219 | 802.2219 | 863.4859
2030 | 92.4377 | 0.93069 | 0.98864 | 0.97138 | 191.9686 | 47.5214 | 49917.3819 | 49829.3993 | 992.588 | 892.588 | 863.4918
2031 | 111.8442 | 0.96572 | 1.0149 | 0.99312 | 191.2783 | 48.4678 | 54276.85 | 53666.0443 | 1069.3209 | 969.3209 | 863.4932
2032 | 96.4749 | 1.0044 | 0.92753 | 1.0345 | 193.7957 | 50.9209 | 53986.2883 | 53890.9176 | 1073.8184 | 973.8184 | 863.5306
2033 | 102.7672 | 0.98277 | 1.1801 | 0.90587 | 205.687 | 48.1724 | 56568.3635 | 56297.748 | 1121.955 | 1021.955 | 863.5733
2034 | 90.6867 | 0.98487 | 1.0106 | 0.93871 | 194.8707 | 49.5353 | 51612.8043 | 50740.438 | 1010.8088 | 910.8088 | 863.6203
2035 | 115.6248 | 1.0177 | 0.8019 | 0.88873 | 206.8673 | 47.8983 | 39978.6397 | 38833.5609 | 772.6712 | 672.6712 | 863.7213
2036 | 112.7681 | 0.95344 | 0.63509 | 1.0132 | 218.9657 | 51.7732 | 39099.2617 | 40168.463 | 799.3693 | 699.3693 | 863.8888
2037 | 102.8712 | 1.0743 | 0.84873 | 0.92007 | 221.9566 | 52.0309 | 46602.3411 | 46254.6014 | 921.092 | 821.092 | 863.9508
2038 | 106.2882 | 0.92007 | 0.9448 | 0.94036 | 189.8754 | 49.7708 | 49275.7495 | 49068.3666 | 977.3673 | 877.3673 | 864.0308
2039 | 108.9642 | 0.91094 | 0.97927 | 1.0461 | 189.9313 | 55.4425 | 62491.97 | 64002.7781 | 1276.0556 | 1176.0556 | 864.1424
2040 | 95.5362 | 1.0377 | 0.81019 | 0.93797 | 199.5611 | 51.3237 | 44289.9129 | 44209.1799 | 880.1836 | 780.1836 | 864.1685
2041 | 95.3168 | 1.0118 | 1.1516 | 0.93108 | 190.9603 | 47.6777 | 55910.5633 | 57080.6928 | 1137.6139 | 1037.6139 | 864.2066
2042 | 105.9935 | 1.0357 | 0.85183 | 0.96213 | 194.5091 | 44.0934 | 41172.7788 | 41679.8908 | 829.5978 | 729.5978 | 864.2585
2043 | 84.9538 | 0.98845 | 0.72279 | 1.0135 | 215.3051 | 51.8516 | 42553.5956 | 43564.7655 | 867.2953 | 767.2953 | 864.312
2044 | 91.5396 | 0.97336 | 0.58859 | 0.97614 | 213.0953 | 52.6355 | 35144.2429 | 34610.0083 | 688.2002 | 588.2002 | 864.3395
2045 | 89.1076 | 0.95201 | 1.1356 | 1.0503 | 196.2413 | 49.0975 | 62920.7665 | 60836.6182 | 1212.7324 | 1112.7324 | 864.3648
2046 | 99.5127 | 0.98427 | 1.1524 | 1.0043 | 216.425 | 46.829 | 58998.8016 | 57724.4761 | 1150.4895 | 1050.4895 | 864.4657
2047 | 110.8321 | 0.93954 | 1.0903 | 0.96846 | 199.0028 | 51.2317 | 59627.7492 | 60667.1398 | 1209.3428 | 1109.3428 | 864.4964
2048 | 109.5048 | 1.0525 | 1.1431 | 1.0873 | 198.9214 | 50.0761 | 68210.8413 | 67766.7685 | 1351.3354 | 1251.3354 | 864.5319
2049 | 93.0624 | 1.0427 | 1.0697 | 0.94922 | 207.3722 | 46.3998 | 51824.0799 | 50524.2988 | 1006.486 | 906.486 | 864.5699
2050 | 95.8403 | 0.84689 | 0.98945 | 0.96774 | 212.4031 | 50.1246 | 52276.4862 | 52207.607 | 1040.1521 | 940.1521 | 864.5927
2051 | 92.8642 | 0.98723 | 0.82671 | 1.0342 | 198.4664 | 50.8342 | 48321.2225 | 48815.8977 | 972.318 | 872.318 | 864.6364
2052 | 127.6579 | 0.93235 | 0.86851 | 0.97085 | 204.0937 | 49.4897 | 47823.9623 | 49021.7042 | 976.4341 | 876.4341 | 864.6486
2053 | 92.5822 | 1.0936 | 1.2264 | 0.99463 | 212.153 | 48.865 | 64763.1334 | 64915.1358 | 1294.3027 | 1194.3027 | 864.6916
2054 | 92.9034 | 0.9372 | 0.91759 | 0.91698 | 207.9244 | 50.0258 | 46655.8636 | 46056.3465 | 917.1269 | 817.1269 | 864.7372
2055 | 110.1797 | 0.96428 | 0.93698 | 0.99921 | 200.0106 | 51.1305 | 53502.7991 | 53419.1313 | 1064.3826 | 964.3826 | 864.8081
2056 | 80.5531 | 1.1191 | 0.68074 | 0.99705 | 208.4427 | 46.0442 | 35610.8006 | 36035.9677 | 716.7194 | 616.7194 | 864.8331
2057 | 111.7034 | 1.0054 | 1.1347 | 1.024 | 208.2472 | 52.5905 | 67220.5037 | 67257.6242 | 1341.1525 | 1241.1525 | 864.873
2058 | 103.4803 | 1.0206 | 0.95941 | 0.93425 | 208.5969 | 52.7979 | 53108.9349 | 53729.1271 | 1070.5825 | 970.5825 | 864.8991
2059 | 123.5606 | 0.98187 | 0.96543 | 1.0634 | 195.9549 | 52.5375 | 60508.2103 | 59446.2337 | 1184.9247 | 1084.9247 | 864.9005
2060 | 102.6226 | 1.0191 | 1.0969 | 1.0009 | 196.9879 | 55.0133 | 66349.4305 | 67143.1498 | 1338.863 | 1238.863 | 864.9465
2061 | 93.1058 | 0.85946 | 0.99929 | 0.95133 | 190.4616 | 49.8043 | 51522.4464 | 51785.3925 | 1031.7078 | 931.7078 | 864.9894
2062 | 103.2953 | 0.99521 | 1.0344 | 0.91193 | 205.3685 | 49.5958 | 52085.7807 | 52893.8221 | 1053.8764 | 953.8764 | 864.994
2063 | 96.7903 | 0.9883 | 0.85454 | 0.97627 | 192.8507 | 49.0607 | 45815.2134 | 45655.452 | 909.109 | 809.109 | 865.0408
2064 | 106.1693 | 1.0535 | 1.1226 | 1.0168 | 204.0796 | 50.3977 | 63367.6711 | 61747.5283 | 1230.9506 | 1130.9506 | 865.0533
2065 | 106.2774 | 1.0125 | 1.1606 | 1.0181 | 194.2504 | 49.8861 | 64512.0884 | 63990.6137 | 1275.8123 | 1175.8123 | 865.0935
2066 | 112.3293 | 0.9517 | 1.0745 | 0.99778 | 200.0705 | 52.4808 | 62078.4368 | 62999.0384 | 1255.9808 | 1155.9808 | 865.1144
2067 | 98.3281 | 0.98395 | 1.1022 | 1.0405 | 196.2378 | 47.4946 | 59259.7571 | 60428.9841 | 1204.5797 | 1104.5797 | 865.1233
2068 | 108.7351 | 1.0131 | 0.73265 | 1.0382 | 192.2533 | 48.3713 | 42314.8362 | 42698.6913 | 849.9738 | 749.9738 | 865.1304
2069 | 103.093 | 0.93726 | 1.0531 | 1.0012 | 210.4202 | 51.7687 | 59795.9104 | 60591.1826 | 1207.8237 | 1107.8237 | 865.1564
2070 | 96.827 | 1.0303 | 1.3789 | 1.0104 | 204.4267 | 48.1221 | 72049.4245 | 72346.8827 | 1442.9377 | 1342.9377 | 865.2213
2071 | 96.3257 | 1.0272 | 0.89901 | 1.0075 | 195.6072 | 55.9824 | 56442.2492 | 54717.5608 | 1090.3512 | 990.3512 | 865.2614
2072 | 106.4073 | 1.0261 | 1.1361 | 1.0596 | 196.4662 | 46.6966 | 61506.8451 | 62202.1348 | 1240.0427 | 1140.0427 | 865.3125
2073 | 93.4144 | 1.0831 | 1.0719 | 1.0207 | 200.4671 | 48.1517 | 57755.9477 | 57125.9377 | 1138.5188 | 1038.5188 | 865.4701
2074 | 96.0213 | 1.0048 | 0.83751 | 1.0628 | 191.5959 | 48.1258 | 47670.8389 | 48003.5754 | 956.0715 | 856.0715 | 865.5516
2075 | 90.518 | 0.9909 | 1.0514 | 1.0635 | 201.2687 | 49.6345 | 60154.9796 | 60666.4511 | 1209.329 | 1109.329 | 865.6307
2076 | 95.9471 | 1.0344 | 0.93969 | 1.0482 | 213.5723 | 47.1554 | 51341.8373 | 50558.6 | 1007.172 | 907.172 | 865.6787
2077 | 110.7002 | 1.0696 | 1.167 | 0.9527 | 193.3497 | 50.851 | 62749.1063 | 63512.6343 | 1266.2527 | 1166.2527 | 865.6906
2078 | 99.8921 | 1.042 | 1.0095 | 0.99602 | 193.8532 | 47.4313 | 52822.0091 | 53455.0391 | 1065.1008 | 965.1008 | 865.7455
2079 | 104.6634 | 0.98312 | 1.1624 | 0.96726 | 191.4212 | 50.0611 | 61628.1509 | 62244.6438 | 1240.8929 | 1140.8929 | 865.8646
2080 | 123.3674 | 1.0415 | 0.98949 | 1.064 | 217.6586 | 49.677 | 58901.8301 | 59075.0662 | 1177.5013 | 1077.5013 | 865.9004
2081 | 108.5797 | 0.95783 | 1.2296 | 0.977 | 203.2996 | 46.678 | 61132.2617 | 62801.7214 | 1252.0344 | 1152.0344 | 865.9047
2082 | 96.041 | 0.96324 | 1.0711 | 0.9506 | 205.0618 | 51.8122 | 57751.4292 | 58715.7456 | 1170.3149 | 1070.3149 | 865.937
2083 | 95.7943 | 0.96514 | 0.98204 | 1.0219 | 193.6896 | 49.9666 | 54959.3041 | 56495.7052 | 1125.9141 | 1025.9141 | 866.0902
2084 | 93.1694 | 0.9089 | 1.2921 | 1.0211 | 200.2749 | 46.6544 | 65700.6623 | 66562.6314 | 1327.2526 | 1227.2526 | 866.1733
2085 | 108.2747 | 0.94842 | 0.95285 | 1.0595 | 197.67 | 54.619 | 60946.8691 | 60487.6708 | 1205.7534 | 1105.7534 | 866.177
2086 | 88.9245 | 1.0338 | 1.2731 | 0.86568 | 219.6078 | 53.915 | 64598.0531 | 61870.3269 | 1233.4065 | 1133.4065 | 866.206
2087 | 109.1628 | 1.0117 | 0.89106 | 1.018 | 186.4278 | 49.706 | 50765.6858 | 49014.6452 | 976.2929 | 876.2929 | 866.2659
2088 | 102.4378 | 0.95799 | 1.0958 | 1.0384 | 201.2948 | 55.7564 | 69116.3752 | 68813.4384 | 1372.2688 | 1272.2688 | 866.3462
2089 | 115.1742 | 1.0482 | 1.1308 | 0.83241 | 194.9533 | 50.7311 | 54073.4223 | 55457.6859 | 1105.1537 | 1005.1537 | 866.4226
2090 | 91.029 | 1.1164 | 1.0318 | 1.0274 | 193.233 | 46.4235 | 54124.5167 | 53962.9612 | 1075.2592 | 975.2592 | 866.4565
2091 | 105.8133 | 0.95009 | 1.1197 | 1.0208 | 202.7681 | 49.3238 | 61540.927 | 59774.9604 | 1191.4992 | 1091.4992 | 866.4914
2092 | 90.4953 | 0.98729 | 1.0534 | 0.88005 | 210.2668 | 54.797 | 55904.1966 | 57765.6013 | 1151.312 | 1051.312 | 866.5284
2093 | 113.4962 | 0.95181 | 1.0854 | 1.012 | 192.0995 | 49.8357 | 60317.6648 | 60987.9519 | 1215.759 | 1115.759 | 866.5563
2094 | 99.5733 | 0.98872 | 0.82108 | 1.0858 | 203.8496 | 48.5488 | 48265.1113 | 46508.6593 | 926.1732 | 826.1732 | 866.6472
2095 | 109.2049 | 1.0363 | 1.1042 | 0.93648 | 202.4666 | 58.1027 | 66859.5737 | 64803.7481 | 1292.075 | 1192.075 | 866.6817
2096 | 95.6114 | 0.97772 | 1.0846 | 1.0515 | 205.2031 | 51.3996 | 63630.404 | 64207.5626 | 1280.1513 | 1180.1513 | 866.7515
2097 | 100.4556 | 0.99395 | 0.94169 | 1.0184 | 187.5205 | 52.1598 | 55417.901 | 55300.429 | 1102.0086 | 1002.0086 | 866.773
2098 | 85.669 | 1.0213 | 1.0258 | 0.94266 | 190.139 | 52.1476 | 55180.0108 | 53129.984 | 1058.5997 | 958.5997 | 866.8273
2099 | 91.0689 | 1.0811 | 0.98002 | 0.91571 | 199.7168 | 48.4313 | 48431.1187 | 49102.6854 | 978.0537 | 878.0537 | 866.8741
2100 | 107.5015 | 0.97856 | 0.98164 | 0.95284 | 194.025 | 51.2335 | 53504.9973 | 55098.8882 | 1097.9778 | 997.9778 | 866.8764
2101 | 107.4252 | 1.0891 | 1.0622 | 1.0756 | 203.2181 | 51.038 | 64485.2215 | 63234.5076 | 1260.6902 | 1160.6902 | 866.9449
2102 | 94.262 | 1.0134 | 0.85742 | 0.98012 | 196.7167 | 51.5809 | 48471.1721 | 49217.5562 | 980.3511 | 880.3511 | 867.0036
2103 | 106.3669 | 0.89841 | 1.0757 | 0.9951 | 200.5713 | 49.977 | 58471.5242 | 60741.1489 | 1210.823 | 1110.823 | 867.0684
2104 | 88.718 | 1.0209 | 1.3082 | 0.94563 | 210.1263 | 47.6277 | 63442.6747 | 65113.4724 | 1298.2694 | 1198.2694 | 867.0998
2105 | 80.3644 | 1.0435 | 1.1963 | 1.0186 | 203.1349 | 48.7183 | 63654.5549 | 62801.3358 | 1252.0267 | 1152.0267 | 867.1099
2106 | 93.5393 | 1.048 | 1.2981 | 1.0491 | 208.1337 | 49.4289 | 72368.0054 | 70034.9669 | 1396.6993 | 1296.6993 | 867.1224
2107 | 115.0463 | 0.98286 | 1.0928 | 1.0134 | 174.3934 | 49.0399 | 60032.9883 | 59507.8554 | 1186.1571 | 1086.1571 | 867.3591
2108 | 95.1249 | 1.0477 | 1.1872 | 0.96245 | 194.1593 | 45.9543 | 57281.3931 | 54931.9027 | 1094.6381 | 994.6381 | 867.4858
2109 | 99.1043 | 1.0554 | 1.0275 | 0.95671 | 209.0035 | 51.9585 | 56721.591 | 57198.2689 | 1139.9654 | 1039.9654 | 867.4934
2110 | 100.761 | 0.95612 | 1.1364 | 0.99963 | 214.9421 | 50.8208 | 62840.015 | 61808.3537 | 1232.1671 | 1132.1671 | 867.5105
2111 | 98.4045 | 1.0007 | 1.2643 | 1.0347 | 186.6746 | 46.4024 | 65456.9004 | 62134.2873 | 1238.6857 | 1138.6857 | 867.5392
2112 | 97.0574 | 1.0357 | 0.97205 | 1.07 | 188.4218 | 49.4468 | 56589.1376 | 56473.1167 | 1125.4623 | 1025.4623 | 867.5624
2113 | 102.2682 | 0.93566 | 1.0902 | 1.0725 | 175.0204 | 51.9202 | 65847.047 | 66023.7015 | 1316.474 | 1216.474 | 867.6194
2114 | 97.5109 | 1.0261 | 1.069 | 1.0812 | 216.7271 | 55.9202 | 70446.8554 | 68230.0905 | 1360.6018 | 1260.6018 | 867.6217
2115 | 104.6673 | 0.95132 | 1.042 | 1.0625 | 195.5461 | 51.0446 | 61789.6706 | 60097.6072 | 1197.9521 | 1097.9521 | 867.7203
2116 | 96.2086 | 0.95609 | 0.8232 | 1.0143 | 205.0537 | 48.6042 | 45257.846 | 44668.4626 | 889.3693 | 789.3693 | 867.9159
2117 | 91.7045 | 1.027 | 0.82515 | 0.9664 | 199.3387 | 48.8495 | 43753.8614 | 43652.4761 | 869.0495 | 769.0495 | 867.9687
2118 | 103.3669 | 0.99001 | 0.88497 | 1.0173 | 196.8609 | 47.4385 | 47759.5389 | 46835.0207 | 932.7004 | 832.7004 | 868.0003
2119 | 81.6009 | 0.99344 | 1.0957 | 0.99473 | 209.9919 | 47.677 | 56039.5639 | 55567.972 | 1107.3594 | 1007.3594 | 868.0831
2120 | 72.7932 | 1.0436 | 0.89513 | 1.0706 | 198.7098 | 46.0914 | 47872.3933 | 47337.7582 | 942.7552 | 842.7552 | 868.2766
2121 | 108.0911 | 0.98922 | 1.0583 | 0.95249 | 196.5791 | 51.0249 | 57085.8219 | 55736.9131 | 1110.7383 | 1010.7383 | 868.3661
2122 | 90.3948 | 0.93809 | 1.1842 | 1.0083 | 189.144 | 53.4796 | 68581.1411 | 67769.8241 | 1351.3965 | 1251.3965 | 868.4932
2123 | 102.3244 | 1.0542 | 1.0288 | 0.93334 | 187.3727 | 47.1771 | 50577.0874 | 51870.5076 | 1033.4102 | 933.4102 | 868.502
2124 | 96.0227 | 1.0606 | 1.0035 | 1.0329 | 182.8698 | 48.166 | 55014.8203 | 56013.2066 | 1116.2641 | 1016.2641 | 868.5609
2125 | 113.0058 | 1.0683 | 0.9343 | 0.8953 | 215.5587 | 49.5886 | 47682.2045 | 46768.5198 | 931.3704 | 831.3704 | 868.5815
2126 | 94.495 | 1.0018 | 0.88235 | 0.99041 | 205.9069 | 48.2842 | 46971.7828 | 46301.1987 | 922.024 | 822.024 | 868.6587
2127 | 93.6768 | 1.0279 | 0.88973 | 0.96232 | 176.5164 | 50.5423 | 48317.8172 | 47227.3031 | 940.5461 | 840.5461 | 868.8054
2128 | 99.1966 | 0.93113 | 0.83328 | 1.0558 | 200.6037 | 45.1494 | 44093.5907 | 44742.4285 | 890.8486 | 790.8486 | 868.8381
2129 | 96.6977 | 0.94776 | 1.1631 | 0.97875 | 201.7279 | 46.6311 | 57560.3387 | 58283.5598 | 1161.6712 | 1061.6712 | 868.8791
2130 | 89.8501 | 1.0292 | 0.97749 | 1.0383 | 192.4173 | 46.332 | 51500.1733 | 52041.4165 | 1036.8283 | 936.8283 | 868.8844
2131 | 87.4331 | 0.98644 | 1.025 | 0.9888 | 217.8539 | 49.5576 | 54720.5993 | 54952.1449 | 1095.0429 | 995.0429 | 868.9293
2132 | 96.7176 | 1.0027 | 1.0859 | 1.0153 | 181.2394 | 49.4819 | 59531.3332 | 59501.8155 | 1186.0363 | 1086.0363 | 869.0856
2133 | 95.6241 | 1.0175 | 1.1753 | 1.033 | 209.9418 | 45.6461 | 60068.65 | 59129.829 | 1178.5966 | 1078.5966 | 869.1114
2134 | 92.099 | 0.95241 | 1.0224 | 0.97103 | 192.6531 | 50.5191 | 54780.3206 | 55870.9219 | 1113.4184 | 1013.4184 | 869.1202
2135 | 99.286 | 0.99792 | 1.3151 | 1.0121 | 217.0865 | 46.6047 | 66862.7372 | 65530.8483 | 1306.617 | 1206.617 | 869.1314
2136 | 92.2679 | 0.97862 | 1.1231 | 1.0077 | 202.219 | 50.0775 | 61397.6413 | 61443.1447 | 1224.8629 | 1124.8629 | 869.1405
2137 | 110.6264 | 1.0057 | 0.97625 | 0.99352 | 202.4593 | 49.9458 | 54203.0351 | 55999.4704 | 1115.9894 | 1015.9894 | 869.2101
2138 | 92.7322 | 0.9783 | 1.1822 | 1.066 | 194.5192 | 49.7779 | 67438.7419 | 70952.2957 | 1415.0459 | 1315.0459 | 869.2376
2139 | 84.727 | 1.0984 | 0.94954 | 0.93252 | 195.5334 | 47.0473 | 46232.2946 | 47095.7863 | 937.9157 | 837.9157 | 869.2814
2140 | 105.1571 | 1.0893 | 1.0896 | 1.015 | 180.4098 | 53.1961 | 65108.4948 | 64505.1793 | 1286.1036 | 1186.1036 | 869.3204
2141 | 105.6953 | 0.96316 | 1.097 | 0.9782 | 212.397 | 49.5239 | 58398.2476 | 58521.0474 | 1166.4209 | 1066.4209 | 869.3715
2142 | 116.478 | 0.89362 | 1.008 | 0.95053 | 201.4495 | 48.1916 | 51393.8811 | 50117.259 | 998.3452 | 898.3452 | 869.4055
2143 | 95.1618 | 1.0143 | 1.209 | 1.0359 | 219.3528 | 47.7671 | 64649.0862 | 63677.71 | 1269.5542 | 1169.5542 | 869.4106
2144 | 104.5644 | 1.0026 | 1.07 | 0.96025 | 213.0929 | 47.7578 | 54288.0567 | 54108.9048 | 1078.1781 | 978.1781 | 869.4457
2145 | 95.9403 | 0.99348 | 0.69253 | 0.99916 | 213.8234 | 44.9099 | 35569.8654 | 36058.0272 | 717.1605 | 617.1605 | 869.6003
2146 | 103.9224 | 0.99001 | 1.0104 | 1.0036 | 211.2921 | 54.108 | 60641.5714 | 60429.5996 | 1204.592 | 1104.592 | 869.6399
2147 | 103.1391 | 1.0665 | 1.0421 | 0.96855 | 198.0195 | 50.7327 | 56984.8287 | 57123.738 | 1138.4748 | 1038.4748 | 869.6459
2148 | 108.4899 | 0.95871 | 0.98462 | 1.0064 | 199.4828 | 47.4756 | 52184.0741 | 53039.4466 | 1056.7889 | 956.7889 | 869.6876
2149 | 98.1292 | 0.96312 | 1.107 | 1.0246 | 197.836 | 57.1916 | 70474.3441 | 72218.9607 | 1440.3792 | 1340.3792 | 869.7081
2150 | 92.0286 | 0.99043 | 1.1947 | 0.91654 | 200.0279 | 52.2719 | 62202.6696 | 62813.1527 | 1252.2631 | 1152.2631 | 869.7656
2151 | 93.2007 | 0.94739 | 1.2582 | 1.0989 | 189.4635 | 51.654 | 76166.4004 | 72873.2529 | 1453.4651 | 1353.4651 | 869.7969
2152 | 97.5688 | 0.9832 | 1.0607 | 1.0289 | 215.4154 | 48.9838 | 58374.7024 | 57611.5474 | 1148.2309 | 1048.2309 | 869.8217
2153 | 94.8924 | 0.94966 | 0.94627 | 1.0025 | 208.4464 | 52.823 | 55078.4822 | 54570.2994 | 1087.406 | 987.406 | 869.8315
2154 | 111.2806 | 0.99262 | 0.99188 | 0.95195 | 191.3278 | 49.2998 | 52186.5001 | 53470.821 | 1065.4164 | 965.4164 | 869.834
2155 | 112.6124 | 0.97504 | 1.1424 | 1.0202 | 200.0842 | 44.8203 | 57355.9607 | 57950.3863 | 1155.0077 | 1055.0077 | 869.885
2156 | 110.4811 | 0.96773 | 0.90018 | 1.0261 | 211.7198 | 48.582 | 50279.7929 | 50304.9656 | 1002.0993 | 902.0993 | 869.9584
2157 | 90.7894 | 1.0121 | 0.90815 | 0.94621 | 195.5951 | 46.6459 | 44564.5191 | 45935.5445 | 914.7109 | 814.7109 | 869.9804
2158 | 82.4341 | 1.0466 | 0.83593 | 1.0066 | 219.0162 | 52.9916 | 49381.4906 | 49894.1434 | 993.8829 | 893.8829 | 870.0013
2159 | 91.4229 | 1.0408 | 1.1894 | 0.98966 | 212.2044 | 52.2163 | 66646.2786 | 69388.5141 | 1383.7703 | 1283.7703 | 870.0484
2160 | 72.9974 | 0.94809 | 0.68908 | 0.92247 | 185.349 | 51.7819 | 36684.1843 | 36094.5698 | 717.8914 | 617.8914 | 870.1425
2161 | 105.2947 | 0.95416 | 1.0505 | 1.0382 | 205.0925 | 45.5906 | 54510.5144 | 54634.766 | 1088.6953 | 988.6953 | 870.214
2162 | 93.6051 | 0.93298 | 0.93349 | 0.95486 | 194.8228 | 53.5438 | 52597.4959 | 52464.4315 | 1045.2886 | 945.2886 | 870.2704
2163 | 89.5556 | 0.97859 | 1.0254 | 1.0221 | 199.4847 | 52.7562 | 60114.565 | 59224.9204 | 1180.4984 | 1080.4984 | 870.3005
2164 | 90.434 | 1.0683 | 0.99358 | 1.0469 | 202.4126 | 44.2701 | 50526.1853 | 49546.4918 | 986.9298 | 886.9298 | 870.307
2165 | 101.4842 | 1.0246 | 0.97983 | 0.99261 | 198.8947 | 46.473 | 50230.4545 | 50309.9926 | 1002.1999 | 902.1999 | 870.3163
2166 | 105.1135 | 1.0191 | 0.98032 | 0.91633 | 205.7765 | 50.9238 | 51405.3097 | 52034.3088 | 1036.6862 | 936.6862 | 870.3821
2167 | 93.0522 | 1.0159 | 0.99858 | 1.0342 | 216.4393 | 48.4313 | 54810.9976 | 54711.6004 | 1090.232 | 990.232 | 870.6
2168 | 81.1425 | 1.0112 | 1.0023 | 0.95445 | 206.4605 | 49.9543 | 52096.0134 | 52750.3814 | 1051.0076 | 951.0076 | 870.6725
2169 | 100.3717 | 0.97799 | 1.1181 | 1.0117 | 202.9842 | 53.5077 | 65985.4848 | 65741.5197 | 1310.8304 | 1210.8304 | 870.8389
2170 | 86.2802 | 0.99545 | 1.2602 | 1.0865 | 201.0839 | 48.8112 | 71227.8438 | 69121.5871 | 1378.4317 | 1278.4317 | 870.9443
2171 | 113.0923 | 0.99586 | 1.0103 | 1.0773 | 198.3658 | 49.1912 | 59274.7605 | 57172.3664 | 1139.4473 | 1039.4473 | 870.9884
2172 | 98.4697 | 0.9876 | 0.99977 | 1.0089 | 219.2111 | 53.5606 | 59451.4419 | 61251.874 | 1221.0375 | 1121.0375 | 871.0332
2173 | 120.7513 | 1.0287 | 0.95421 | 0.96135 | 209.5288 | 50.929 | 53254.2938 | 53382.8759 | 1063.6575 | 963.6575 | 871.0572
2174 | 103.0395 | 0.97984 | 1.1016 | 0.92642 | 197.8784 | 50.0257 | 56302.0424 | 56145.3544 | 1118.9071 | 1018.9071 | 871.0589
2175 | 103.5898 | 0.99318 | 1.2035 | 1.0338 | 204.7233 | 47.5691 | 64284.0192 | 63681.4517 | 1269.629 | 1169.629 | 871.1103
2176 | 98.1991 | 0.96656 | 0.91048 | 0.98464 | 208.0252 | 46.7339 | 46540.6108 | 46945.8463 | 934.9169 | 834.9169 | 871.1447
2177 | 86.7358 | 0.96546 | 1.1976 | 0.9157 | 193.6384 | 48.3005 | 57205.9327 | 58084.5281 | 1157.6906 | 1057.6906 | 871.154
2178 | 90.9362 | 1.0287 | 0.71243 | 1.0053 | 213.3906 | 47.4861 | 38666.0646 | 38794.7985 | 771.896 | 671.896 | 871.346
2179 | 97.1418 | 1.0078 | 0.95304 | 1.0334 | 194.732 | 47.7999 | 51950.7779 | 51798.6827 | 1031.9737 | 931.9737 | 871.4601
2180 | 103.8353 | 1.003 | 0.88797 | 0.98104 | 197.3692 | 55.2292 | 54060.9535 | 52439.2038 | 1044.7841 | 944.7841 | 871.5036
2181 | 120.4026 | 0.97495 | 1.2483 | 0.97234 | 194.3338 | 52.466 | 70037.2711 | 69513.8408 | 1386.2768 | 1286.2768 | 871.5211
2182 | 99.5129 | 0.9933 | 0.86595 | 1.0202 | 203.8958 | 46.9681 | 46341.6275 | 47396.2745 | 943.9255 | 843.9255 | 871.5211
2183 | 107.7894 | 1.0107 | 1.1076 | 1.0406 | 198.1068 | 52.3893 | 66288.235 | 65695.1307 | 1309.9026 | 1209.9026 | 871.5948
2184 | 88.6718 | 0.98923 | 1.1273 | 0.97077 | 195.4444 | 44.3129 | 52578.1538 | 52313.2117 | 1042.2642 | 942.2642 | 871.597
2185 | 89.2498 | 0.97031 | 1.0542 | 0.96491 | 221.5274 | 50.8995 | 56402.5208 | 55249.2598 | 1100.9852 | 1000.9852 | 871.7091
2186 | 106.7914 | 1.004 | 0.85931 | 1.0043 | 218.0649 | 53.4255 | 52053.4974 | 51242.5927 | 1020.8519 | 920.8519 | 871.7791
2187 | 106.7066 | 1.0818 | 1.0974 | 0.88158 | 212.2355 | 49.7549 | 54092.4472 | 51981.7396 | 1035.6348 | 935.6348 | 871.7952
2188 | 93.5303 | 0.98584 | 0.99297 | 1.044 | 195.4871 | 51.836 | 58711.4309 | 57740.9424 | 1150.8188 | 1050.8188 | 871.9088
2189 | 103.9886 | 0.94037 | 1.0146 | 1.0555 | 199.6876 | 49.2752 | 57787.2541 | 57982.9377 | 1155.6588 | 1055.6588 | 871.954
2190 | 89.7397 | 1.0392 | 1.2078 | 0.97435 | 219.9622 | 50.0535 | 63790.6906 | 63902.6356 | 1274.0527 | 1174.0527 | 872.0616
2191 | 94.4697 | 0.97327 | 1.1623 | 1.0125 | 192.5785 | 48.3861 | 61584.0201 | 60840.8533 | 1212.8171 | 1112.8171 | 872.318
2192 | 103.783 | 0.96944 | 0.94326 | 0.96312 | 191.8862 | 51.117 | 51773.2909 | 51542.2399 | 1026.8448 | 926.8448 | 872.3755
2193 | 101.8887 | 1.0741 | 0.98389 | 0.97459 | 180.9337 | 48.1356 | 51605.0035 | 51192.156 | 1019.8431 | 919.8431 | 872.4069
2194 | 111.4898 | 0.96808 | 1.2001 | 0.98123 | 197.5731 | 51.9417 | 66970.6672 | 63879.1743 | 1273.5835 | 1173.5835 | 872.5221
2195 | 112.2872 | 0.97848 | 0.98534 | 0.99174 | 209.6332 | 50.9014 | 55543.3505 | 56028.0022 | 1116.56 | 1016.56 | 872.5324
2196 | 101.4792 | 1.0604 | 0.88512 | 0.96842 | 189.0077 | 50.8631 | 49260.0485 | 48818.7726 | 972.3755 | 872.3755 | 872.699
2197 | 85.3246 | 1.0434 | 1.2299 | 1.0644 | 202.7665 | 49.6683 | 69646.743 | 68464.2061 | 1365.2841 | 1265.2841 | 872.7419
2198 | 96.7784 | 0.98203 | 0.92975 | 0.9732 | 192.2367 | 50.791 | 50976.583 | 50854.9843 | 1013.0997 | 913.0997 | 872.8041
2199 | 89.4565 | 0.96044 | 0.99493 | 1.0275 | 210.851 | 50.4765 | 56149.353 | 57207.5134 | 1140.1503 | 1040.1503 | 872.8058
2200 | 103.8593 | 0.95865 | 0.83623 | 0.9264 | 199.4984 | 49.952 | 43869.9052 | 42913.4 | 854.268 | 754.268 | 872.9645
2201 | 97.7399 | 0.98488 | 0.99421 | 1.0209 | 186.7497 | 48.2869 | 53846.7576 | 53481.872 | 1065.6374 | 965.6374 | 873.013
2202 | 76.6649 | 0.94745 | 1.0558 | 0.95691 | 203.5715 | 52.0831 | 56606.9243 | 56258.3255 | 1121.1665 | 1021.1665 | 873.0885
2203 | 102.5218 | 0.93547 | 1.147 | 0.97816 | 203.4632 | 49.4859 | 60470.7939 | 61146.8524 | 1218.937 | 1118.937 | 873.1862
2204 | 102.3696 | 0.9687 | 0.84553 | 1.0779 | 204.9219 | 53.3817 | 54151.768 | 54247.2913 | 1080.9458 | 980.9458 | 873.2217
2205 | 101.7627 | 0.98639 | 1.0025 | 0.98337 | 203.8443 | 45.5789 | 49713.1316 | 47118.2538 | 938.3651 | 838.3651 | 873.2904
2206 | 95.678 | 1.0236 | 1.1258 | 0.94712 | 191.605 | 48.3022 | 56425.9848 | 57367.8043 | 1143.3561 | 1043.3561 | 873.3107
2207 | 80.3729 | 1.0688 | 0.6164 | 1.0144 | 176.0451 | 48.8251 | 34899.1459 | 36514.8266 | 726.2965 | 626.2965 | 873.3307
2208 | 95.4769 | 1.0496 | 1.0736 | 0.89153 | 180.7578 | 48.4481 | 51408.6131 | 51857.9051 | 1033.1581 | 933.1581 | 873.3459
2209 | 111.3042 | 0.91663 | 1.0721 | 1.073 | 185.8647 | 52.2771 | 65656.7392 | 64749.3585 | 1290.9872 | 1190.9872 | 873.4398
2210 | 91.5782 | 1.0552 | 0.92723 | 1.0991 | 219.2203 | 52.9991 | 59355.5507 | 59306.011 | 1182.1202 | 1082.1202 | 873.4658
2211 | 98.9693 | 1.0292 | 0.84977 | 0.91697 | 216.2648 | 51.4635 | 45558.9751 | 45501.2855 | 906.0257 | 806.0257 | 873.4753
2212 | 99.4827 | 1.0293 | 0.66429 | 0.93549 | 212.744 | 50.3002 | 36621.8148 | 35655.7185 | 709.1144 | 609.1144 | 873.4827
2213 | 98.8218 | 1.0326 | 0.95205 | 0.9627 | 203.4349 | 50.4196 | 51559.8448 | 50599.602 | 1007.992 | 907.992 | 873.5465
2214 | 98.9818 | 1.0705 | 0.96587 | 0.99539 | 173.3366 | 48.655 | 52106.8662 | 51669.9761 | 1029.3995 | 929.3995 | 873.6473
2215 | 98.6367 | 1.031 | 1.0338 | 1.0068 | 208.8034 | 49.0234 | 56216.2235 | 54966.6593 | 1095.3332 | 995.3332 | 873.6964
2216 | 98.9708 | 1.1002 | 1.1351 | 0.98894 | 191.8281 | 47.9379 | 59224.8124 | 58899.5687 | 1173.9914 | 1073.9914 | 873.7177
2217 | 99.8576 | 1.0273 | 0.84871 | 1.0097 | 202.1163 | 49.4702 | 47670.3655 | 48209.5583 | 960.1912 | 860.1912 | 873.7807
2218 | 101.2949 | 0.97614 | 0.96773 | 1.0192 | 175.5848 | 46.3355 | 50460.3744 | 51215.389 | 1020.3078 | 920.3078 | 873.8265
2219 | 106.2228 | 0.97015 | 0.77101 | 1.0016 | 181.5217 | 45.3944 | 39915.9067 | 41548.2828 | 826.9657 | 726.9657 | 873.8587
2220 | 106.2985 | 1.0584 | 1.0364 | 1.0438 | 193.7634 | 48.6364 | 58280.5597 | 59016.546 | 1176.3309 | 1076.3309 | 873.9146
2221 | 96.1543 | 0.97136 | 1.2645 | 1.0721 | 194.4443 | 52.2866 | 75959.7213 | 78783.7423 | 1571.6748 | 1471.6748 | 873.9503
2222 | 111.3883 | 1.011 | 1.1346 | 0.96517 | 207.5712 | 50.7941 | 61552.3789 | 61457.9868 | 1225.1597 | 1125.1597 | 873.9532
2223 | 97.5513 | 0.9511 | 0.8433 | 1.0487 | 194.0072 | 47.3802 | 46490.8632 | 46537.1249 | 926.7425 | 826.7425 | 873.9533
2224 | 114.0337 | 0.97641 | 0.82678 | 1.062 | 186.08 | 51.1149 | 50758.5711 | 52421.2348 | 1044.4247 | 944.4247 | 873.9722
2225 | 104.8441 | 1.0248 | 0.91215 | 0.95223 | 191.0763 | 47.2031 | 46262.6118 | 46343.5234 | 922.8705 | 822.8705 | 874.0617
2226 | 100.7397 | 0.99164 | 0.94061 | 1.0801 | 208.8419 | 49.265 | 55182.6432 | 53839.2858 | 1072.7857 | 972.7857 | 874.0982
2227 | 103.2034 | 1.0567 | 0.86478 | 0.9923 | 208.9516 | 46.6954 | 45371.4442 | 44661.1337 | 889.2227 | 789.2227 | 874.1176
2228 | 98.3217 | 1.0272 | 1.2253 | 1.0218 | 227.7919 | 52.772 | 71627.2799 | 71501.6416 | 1426.0328 | 1326.0328 | 874.1563
2229 | 87.4758 | 1.0485 | 1.0015 | 0.96292 | 202.1653 | 50.6822 | 53724.9058 | 54479.0984 | 1085.582 | 985.582 | 874.1586
2230 | 97.9417 | 1.0578 | 1.0426 | 0.9556 | 204.7308 | 53.7702 | 59348.3634 | 59047.1492 | 1176.943 | 1076.943 | 874.3997
2231 | 95.8215 | 0.98219 | 0.8195 | 0.98211 | 197.4309 | 51.7474 | 46716.0896 | 48444.955 | 964.8991 | 864.8991 | 874.4683
2232 | 91.085 | 1.0531 | 0.83255 | 0.96925 | 187.3946 | 49.351 | 44745.2006 | 45284.7462 | 901.6949 | 801.6949 | 874.4736
2233 | 79.35 | 1.0221 | 0.84814 | 1.0144 | 201.2885 | 52.432 | 49564.5271 | 49603.1944 | 988.0639 | 888.0639 | 874.5462
2234 | 96.5642 | 0.90654 | 1.0003 | 0.97027 | 197.3552 | 51.0339 | 54196.4676 | 54026.4103 | 1076.5282 | 976.5282 | 874.5947
2235 | 104.5228 | 1.0027 | 1.172 | 0.97124 | 185.7515 | 50.5104 | 62976.9396 | 64385.1427 | 1283.7029 | 1183.7029 | 874.6348
2236 | 82.3753 | 0.96641 | 1.1133 | 0.94388 | 191.9045 | 44.6011 | 50608.8285 | 50651.6172 | 1009.0323 | 909.0323 | 874.6414
2237 | 101.8633 | 1.0491 | 1.0367 | 0.95864 | 195.8637 | 49.6889 | 54886.3063 | 55867.127 | 1113.3425 | 1013.3425 | 874.6513
2238 | 83.3852 | 1.0757 | 0.81674 | 0.95895 | 197.2002 | 49.7378 | 43613.934 | 42765.703 | 851.3141 | 751.3141 | 874.6527
2239 | 87.1018 | 0.96198 | 1.1534 | 1.0724 | 206.121 | 53.8296 | 71296.1884 | 72325.3768 | 1442.5075 | 1342.5075 | 874.6549
2240 | 99.1485 | 1.062 | 1.2336 | 1.0283 | 201.298 | 49.3201 | 67959.2398 | 67614.5105 | 1348.2902 | 1248.2902 | 874.705
2241 | 105.2197 | 0.94074 | 0.93053 | 1.0949 | 205.2402 | 52.1848 | 58539.4718 | 57619.3334 | 1148.3867 | 1048.3867 | 874.7153
2242 | 92.9892 | 1.1354 | 1.1466 | 0.98918 | 198.0244 | 53.2805 | 66255.3891 | 67008.9596 | 1336.1792 | 1236.1792 | 874.7366
2243 | 107.583 | 1.0052 | 1.1078 | 0.98252 | 193.7104 | 48.6061 | 58355.5969 | 58612.7498 | 1168.255 | 1068.255 | 874.8313
2244 | 110.158 | 0.97978 | 0.73973 | 0.89265 | 191.7303 | 52.483 | 40511.584 | 41590.1314 | 827.8026 | 727.8026 | 874.8444
2245 | 120.2491 | 1.0205 | 1.0729 | 1.0491 | 213.9852 | 47.9082 | 60017.7307 | 57331.445 | 1142.6289 | 1042.6289 | 874.8468
2246 | 114.1354 | 0.94546 | 1.1317 | 1.029 | 195.726 | 46.5475 | 59427.1741 | 58502.4766 | 1166.0495 | 1066.0495 | 874.9424
2247 | 96.3322 | 0.99266 | 0.99993 | 0.98901 | 208.3268 | 49.2452 | 53618.0873 | 52910.0046 | 1054.2001 | 954.2001 | 874.9879
2248 | 98.6719 | 0.94681 | 0.65922 | 1.0524 | 199.8507 | 50.3301 | 39818.2396 | 39541.9044 | 786.8381 | 686.8381 | 875.0949
2249 | 87.1835 | 1.059 | 1.1077 | 0.98409 | 205.8059 | 49.6727 | 58937.1555 | 57032.5731 | 1136.6515 | 1036.6515 | 875.1566
2250 | 99.2003 | 0.96 | 1.2451 | 1.0394 | 207.0206 | 49.9394 | 69595.4204 | 69867.9121 | 1393.3582 | 1293.3582 | 875.2385
2251 | 91.7802 | 0.90234 | 1.1351 | 1.0016 | 200.8216 | 51.2285 | 62684.2949 | 61395.03 | 1223.9006 | 1123.9006 | 875.2442
2252 | 86.7973 | 0.97224 | 0.94159 | 1.0254 | 213.3278 | 50.0867 | 52799.89 | 53365.8694 | 1063.3174 | 963.3174 | 875.2705
2253 | 106.0544 | 0.97437 | 1.1155 | 1.0119 | 195.2373 | 52.1003 | 64387.4064 | 64063.0344 | 1277.2607 | 1177.2607 | 875.309
2254 | 102.5756 | 1.0278 | 1.0613 | 1.0145 | 215.2381 | 47.7398 | 56647.7629 | 57932.5541 | 1154.6511 | 1054.6511 | 875.4488
2255 | 83.1312 | 0.95732 | 0.93587 | 1.0532 | 195.0026 | 51.6231 | 55188.1748 | 54351.3739 | 1083.0275 | 983.0275 | 875.5446
2256 | 105.2022 | 0.96333 | 1.134 | 0.9272 | 197.8316 | 48.5192 | 56129.1159 | 57555.7487 | 1147.115 | 1047.115 | 875.5546
2257 | 102.7982 | 0.9272 | 0.98464 | 1.0404 | 216.8398 | 46.4098 | 52181.2837 | 54096.1628 | 1077.9233 | 977.9233 | 875.5738
2258 | 115.6842 | 1.0342 | 1.0619 | 1.0348 | 186.3697 | 47.4783 | 58038.7465 | 55023.492 | 1096.4698 | 996.4698 | 875.6254
2259 | 101.5183 | 0.99656 | 1.1002 | 0.99056 | 208.1565 | 50.4928 | 60345.9979 | 62252.6412 | 1241.0528 | 1141.0528 | 875.7264
2260 | 106.3125 | 1.0704 | 1.0034 | 1.0095 | 204.0451 | 51.0203 | 57690.043 | 57806.9628 | 1152.1393 | 1052.1393 | 875.7662
2261 | 83.9597 | 0.8505 | 0.89398 | 1.0493 | 191.1709 | 50.1117 | 50777.2289 | 51634.0081 | 1028.6802 | 928.6802 | 875.8614
2262 | 110.8737 | 1.1526 | 0.9395 | 0.96044 | 184.7697 | 48.2138 | 49850.7478 | 49485.0749 | 985.7015 | 885.7015 | 875.8949
2263 | 104.2516 | 1.0486 | 0.84889 | 0.98354 | 197.5117 | 48.6711 | 46154.6642 | 47132.3436 | 938.6469 | 838.6469 | 875.9453
2264 | 107.9836 | 0.8976 | 1.0259 | 0.99578 | 220.227 | 50.9801 | 57242.4592 | 55781.2322 | 1111.6246 | 1011.6246 | 876.0382
2265 | 100.8274 | 0.98975 | 0.73075 | 1.0842 | 207.8484 | 49.6479 | 44496.246 | 43915.0244 | 874.3005 | 774.3005 | 876.0455
2266 | 97.0377 | 0.96356 | 1.0234 | 0.97403 | 203.4151 | 49.4873 | 54159.6256 | 54253.3949 | 1081.0679 | 981.0679 | 876.1304
2267 | 104.6757 | 0.99206 | 0.97452 | 0.92582 | 185.7452 | 48.4823 | 48962.4679 | 50505.8489 | 1006.117 | 906.117 | 876.2139
2268 | 112.3406 | 1.0264 | 0.98723 | 1.0007 | 211.5981 | 47.1585 | 52237.6346 | 52216.6844 | 1040.3337 | 940.3337 | 876.2393
2269 | 100.0124 | 0.95131 | 1.2312 | 0.96185 | 220.5569 | 52.3673 | 67218.147 | 66631.5382 | 1328.6308 | 1228.6308 | 876.2586
2270 | 87.7059 | 0.98963 | 0.85073 | 1.0312 | 191.7436 | 54.2081 | 52452.7307 | 53358.8681 | 1063.1774 | 963.1774 | 876.2929
2271 | 108.8128 | 0.89692 | 1.0801 | 1.0006 | 229.7788 | 56.0814 | 66314.1463 | 64994.5195 | 1295.8904 | 1195.8904 | 876.3771
2272 | 106.1101 | 1.0171 | 0.7887 | 1.0229 | 203.0218 | 48.882 | 44915.9044 | 44582.1412 | 887.6428 | 787.6428 | 876.3773
2273 | 92.8861 | 1.0437 | 1.0286 | 1.1011 | 196.1063 | 49.6361 | 61225.8222 | 60400.4136 | 1204.0083 | 1104.0083 | 876.3785
2274 | 110.9405 | 0.95316 | 0.96946 | 1.0155 | 187.1097 | 54.2383 | 59319.7017 | 60133.5325 | 1198.6707 | 1098.6707 | 876.4222
2275 | 97.7401 | 0.86224 | 0.86453 | 1.0253 | 200.9202 | 46.8365 | 45663.4947 | 45476.329 | 905.5266 | 805.5266 | 876.4308
2276 | 106.7937 | 0.99364 | 0.96797 | 0.9282 | 205.4857 | 50.4445 | 50881.1386 | 51170.712 | 1019.4142 | 919.4142 | 876.4341
2277 | 99.7768 | 1.0885 | 0.96509 | 1.0095 | 187.235 | 48.1346 | 52309.2715 | 50038.3029 | 996.7661 | 896.7661 | 876.452
2278 | 84.9098 | 1.0177 | 0.97542 | 1.0333 | 201.0902 | 48.7348 | 53531.3736 | 52572.5555 | 1047.4511 | 947.4511 | 876.5248
2279 | 114.2767 | 0.99463 | 0.96467 | 1.0461 | 186.8324 | 55.4362 | 62431.1861 | 64534.3896 | 1286.6878 | 1186.6878 | 876.6933
2280 | 98.1574 | 1.026 | 1.2404 | 0.98103 | 203.7572 | 52.7321 | 69683.044 | 69928.4253 | 1394.5685 | 1294.5685 | 876.7555
2281 | 109.6951 | 1.016 | 0.95848 | 0.95324 | 193.8065 | 47.9982 | 49397.1706 | 49579.8339 | 987.5967 | 887.5967 | 876.7767
2282 | 96.1078 | 1.0933 | 0.988 | 1.0596 | 177.236 | 49.5297 | 57231.8828 | 58787.942 | 1171.7588 | 1071.7588 | 876.7799
2283 | 104.7391 | 0.90036 | 0.75854 | 1.0373 | 198.9286 | 48.8445 | 43237.2909 | 43394.0243 | 863.8805 | 763.8805 | 876.7863
2284 | 99.8983 | 1.0419 | 1.0014 | 0.91738 | 214.315 | 49.2209 | 50555.3299 | 49495.0226 | 985.9005 | 885.9005 | 876.8414
2285 | 100.2275 | 1.0749 | 0.9265 | 0.9778 | 203.5982 | 47.2912 | 48140.8835 | 48578.1183 | 967.5624 | 867.5624 | 876.8503
2286 | 92.7359 | 1.0803 | 1.1955 | 0.94167 | 195.7445 | 49.2442 | 60565.0263 | 58716.5713 | 1170.3314 | 1070.3314 | 876.9342
2287 | 92.7519 | 1.0546 | 0.97177 | 0.94189 | 199.6436 | 53.1688 | 54066.0266 | 54612.0564 | 1088.2411 | 988.2411 | 876.9718
2288 | 123.1007 | 1.0125 | 1.0413 | 1.0319 | 221.7703 | 49.5622 | 59653.3181 | 57927.6388 | 1154.5528 | 1054.5528 | 877.0216
2289 | 108.5694 | 1.0596 | 1.093 | 1.0046 | 199.4575 | 53.662 | 65295.7587 | 64644.996 | 1288.8999 | 1188.8999 | 877.0222
2290 | 105.9885 | 0.99607 | 0.89692 | 0.94961 | 194.3649 | 48.1509 | 46289.273 | 47577.5476 | 947.551 | 847.551 | 877.0651
2291 | 112.1344 | 1.0347 | 0.87029 | 1.0067 | 199.9675 | 44.5759 | 44426.6647 | 43489.1402 | 865.7828 | 765.7828 | 877.116
2292 | 112.3741 | 0.97281 | 1.0204 | 1.0309 | 199.8467 | 52.0333 | 60621.0923 | 60955.3229 | 1215.1065 | 1115.1065 | 877.1826
2293 | 109.0017 | 0.95148 | 0.88082 | 0.944 | 209.5077 | 53.332 | 50085.6638 | 50569.9389 | 1007.3988 | 907.3988 | 877.2188
2294 | 95.8834 | 0.98262 | 0.9598 | 1.0104 | 194.7455 | 51.7741 | 55283.8904 | 55579.469 | 1107.5894 | 1007.5894 | 877.2437
2295 | 107.1987 | 0.89456 | 1.0037 | 0.97916 | 209.612 | 51.42 | 55676.6963 | 55116.5097 | 1098.3302 | 998.3302 | 877.2614
2296 | 90.5215 | 1.0824 | 1.2169 | 1.0325 | 205.4548 | 48.6457 | 66092.1626 | 64694.7228 | 1289.8945 | 1189.8945 | 877.3673
2297 | 85.9001 | 0.95765 | 1.2071 | 1.0377 | 203.3874 | 50.5773 | 67718.447 | 66284.2505 | 1321.685 | 1221.685 | 877.4048
2298 | 88.2551 | 1.055 | 1.056 | 0.91973 | 204.0836 | 53.2042 | 56831.6931 | 58060.798 | 1157.216 | 1057.216 | 877.4679
2299 | 95.3472 | 1.0214 | 1.0516 | 1.0107 | 210.26 | 47.326 | 55120.5791 | 54912.2097 | 1094.2442 | 994.2442 | 877.5137
2300 | 107.0561 | 1.0028 | 1.0289 | 1.0047 | 184.0718 | 54.479 | 62345.1285 | 63345.554 | 1262.9111 | 1162.9111 | 877.5557
2301 | 110.1516 | 0.98163 | 1.2456 | 0.9809 | 203.4826 | 44.9532 | 59988.6493 | 59401.8995 | 1184.038 | 1084.038 | 877.5598
2302 | 100.333 | 0.99305 | 1.0337 | 1.0394 | 182.0637 | 45.09 | 53121.0181 | 52454.0634 | 1045.0813 | 945.0813 | 877.6386
2303 | 118.9172 | 0.92429 | 0.74557 | 0.95697 | 190.1726 | 51.6081 | 42684.1361 | 41992.6155 | 835.8523 | 735.8523 | 877.6448
2304 | 89.4229 | 0.94105 | 1.0569 | 0.94286 | 190.5423 | 46.7149 | 50674.8768 | 51739.0462 | 1030.7809 | 930.7809 | 877.6763
2305 | 90.8168 | 0.99898 | 0.8961 | 0.96129 | 199.1722 | 51.9563 | 49668.7275 | 48031.3316 | 956.6266 | 856.6266 | 877.7365
2306 | 96.4026 | 0.96347 | 0.88252 | 0.98997 | 188.7379 | 52.1647 | 50608.6501 | 51265.8829 | 1021.3177 | 921.3177 | 877.9129
2307 | 96.5536 | 1.0471 | 0.9751 | 1.0429 | 181.3909 | 53.1456 | 59602.6481 | 60201.9617 | 1200.0392 | 1100.0392 | 877.9176
2308 | 95.6356 | 0.95873 | 0.98179 | 1.0888 | 209.4829 | 49.5619 | 57732.2721 | 56698.1028 | 1129.9621 | 1029.9621 | 877.9463
2309 | 111.7527 | 1.0476 | 0.93845 | 1.0474 | 187.7713 | 47.5066 | 52443.0953 | 54123.181 | 1078.4636 | 978.4636 | 877.9489
2310 | 105.7968 | 1.0088 | 0.84722 | 1.0278 | 189.3277 | 50.9299 | 49973.5256 | 50457.6659 | 1005.1533 | 905.1533 | 877.9502
2311 | 92.1318 | 0.92668 | 1.0822 | 0.97964 | 186.1179 | 45.5939 | 52416.2776 | 52536.1374 | 1046.7227 | 946.7227 | 877.9886
2312 | 87.945 | 0.96097 | 0.97515 | 1.0887 | 205.7721 | 48.9965 | 56362.0329 | 56104.2796 | 1118.0856 | 1018.0856 | 878.0206
2313 | 90.1664 | 0.95413 | 0.78782 | 0.96068 | 189.2705 | 50.9787 | 43158.0289 | 43840.175 | 872.8035 | 772.8035 | 878.0317
2314 | 103.4402 | 1.0559 | 1.1899 | 1.0295 | 209.7432 | 51.0566 | 68333.0357 | 69074.2876 | 1377.4858 | 1277.4858 | 878.0537
2315 | 104.5821 | 1.0275 | 0.86252 | 1.0562 | 212.55 | 46.5199 | 47590.2391 | 47196.8166 | 939.9363 | 839.9363 | 878.1479
2316 | 106.8559 | 0.96783 | 0.83507 | 1.0947 | 196.0881 | 50.6172 | 51703.7807 | 50774.6109 | 1011.4922 | 911.4922 | 878.1668
2317 | 100.2557 | 0.93487 | 1.1564 | 1.0081 | 191.0239 | 48.8338 | 61697.6045 | 62473.9521 | 1245.479 | 1145.479 | 878.1989
2318 | 119.4288 | 0.94684 | 1.1109 | 1.0779 | 209.22 | 48.1767 | 63342.865 | 62459.9421 | 1245.1988 | 1145.1988 | 878.2057
2319 | 90.4264 | 1.0296 | 1.089 | 0.98446 | 194.9361 | 52.4564 | 61317.1951 | 60811.8863 | 1212.2377 | 1112.2377 | 878.2117
2320 | 100.2547 | 1.0082 | 0.86143 | 0.89439 | 175.5569 | 51.2374 | 44830.143 | 45997.7714 | 915.9554 | 815.9554 | 878.2487
2321 | 98.2654 | 0.8459 | 0.87024 | 0.94282 | 178.8373 | 49.9491 | 45312.949 | 46055.8379 | 917.1168 | 817.1168 | 878.2532
2322 | 115.3366 | 0.99243 | 0.78883 | 1.0115 | 195.6755 | 49.8253 | 45652.7782 | 45490.9492 | 905.819 | 805.819 | 878.2816
2323 | 115.5809 | 0.99785 | 1.0723 | 0.90381 | 198.7873 | 45.4955 | 49538.6458 | 49640.1108 | 988.8022 | 888.8022 | 878.3003
2324 | 102.0957 | 0.93994 | 1.0175 | 1.0931 | 205.2197 | 48.2118 | 58453.5242 | 56021.3987 | 1116.428 | 1016.428 | 878.3612
2325 | 103.2641 | 1.0516 | 1.0622 | 1.0044 | 213.3714 | 49.6063 | 58521.1757 | 57141.3557 | 1138.8271 | 1038.8271 | 878.3984
2326 | 81.3776 | 0.96645 | 1.2071 | 1.0337 | 195.4367 | 49.1513 | 65390.872 | 66176.3852 | 1319.5277 | 1219.5277 | 878.4174
2327 | 95.1167 | 0.99991 | 0.89789 | 0.94982 | 192.9455 | 46.1697 | 43959.0629 | 43166.9507 | 859.339 | 759.339 | 878.4454
2328 | 83.5876 | 1.0781 | 1.0737 | 0.9574 | 195.5824 | 54.0504 | 60627.8217 | 61165.3768 | 1219.3075 | 1119.3075 | 878.6016
2329 | 111.6906 | 0.98725 | 1.0443 | 0.98734 | 208.1113 | 49.2216 | 56385.1114 | 56544.7038 | 1126.8941 | 1026.8941 | 878.6306
2330 | 103.1192 | 1.0385 | 0.85484 | 1.024 | 199.7259 | 49.1305 | 48466.1219 | 48867.2971 | 973.3459 | 873.3459 | 878.6547
2331 | 92.0994 | 1.0668 | 0.87806 | 1.1467 | 195.7527 | 46.8618 | 51985.6984 | 52234.3203 | 1040.6864 | 940.6864 | 878.7254
2332 | 99.2631 | 0.97112 | 0.97333 | 0.99222 | 202.7254 | 50.9235 | 54291.5754 | 54190.842 | 1079.8168 | 979.8168 | 878.8435
2333 | 105.7835 | 1.0201 | 1.391 | 1.0038 | 189.8074 | 50.832 | 76652.7043 | 77887.0405 | 1553.7408 | 1453.7408 | 878.8468
2334 | 85.1429 | 0.94732 | 1.1346 | 1.0568 | 190.5858 | 46.6026 | 59828.8759 | 57757.2322 | 1151.1446 | 1051.1446 | 878.8632
2335 | 91.4602 | 0.98744 | 1.2913 | 1.0163 | 209.6177 | 47.2663 | 66509.726 | 65546.3363 | 1306.9267 | 1206.9267 | 878.986
2336 | 96.9033 | 1.0224 | 0.93469 | 1.0215 | 206.3253 | 50.9557 | 53908.4424 | 54647.3935 | 1088.9479 | 988.9479 | 879.0203
2337 | 96.6779 | 0.98614 | 1.0834 | 1.037 | 206.164 | 50.4644 | 61714.9331 | 62829.6863 | 1252.5937 | 1152.5937 | 879.0404
2338 | 83.1049 | 1.0161 | 0.96328 | 1.0264 | 186.9928 | 52.4542 | 56476.3083 | 56222.1042 | 1120.4421 | 1020.4421 | 879.0602
2339 | 99.5408 | 0.93198 | 0.79627 | 0.98876 | 183.528 | 50.5522 | 44673.7987 | 44667.2755 | 889.3455 | 789.3455 | 879.0815
2340 | 118.7048 | 0.91909 | 0.88066 | 1.0594 | 196.394 | 44.0829 | 46135.6746 | 47756.0426 | 951.1209 | 851.1209 | 879.1106
2341 | 105.9666 | 0.98483 | 1.007 | 1.0066 | 203.2773 | 50.5087 | 56671.5074 | 55806.6961 | 1112.1339 | 1012.1339 | 879.16
2342 | 92.28 | 0.99745 | 0.88943 | 0.95585 | 207.033 | 51.4141 | 48649.6297 | 49530.1543 | 986.6031 | 886.6031 | 879.184
2343 | 96.4399 | 0.93356 | 1.0244 | 1.0262 | 173.7611 | 47.4049 | 54273.3617 | 54121.967 | 1078.4393 | 978.4393 | 879.2021
2344 | 112.94 | 1.0326 | 1.219 | 0.97047 | 203.8442 | 50.7489 | 66156.5025 | 65608.0021 | 1308.16 | 1208.16 | 879.2268
2345 | 89.6285 | 1.01 | 0.93391 | 0.96149 | 183.5233 | 52.9683 | 52540.671 | 51921.4434 | 1034.4289 | 934.4289 | 879.2425
2346 | 123.9684 | 1.0251 | 1.2336 | 0.97022 | 211.5913 | 49.1788 | 65321.9759 | 64940.0195 | 1294.8004 | 1194.8004 | 879.3158
2347 | 95.1236 | 1.1236 | 1.028 | 1.0151 | 192.3213 | 49.7729 | 57447.397 | 57733.3163 | 1150.6663 | 1050.6663 | 879.4245
2348 | 78.3336 | 1.0168 | 0.91061 | 1.0545 | 202.7518 | 50.3917 | 52605.2489 | 52422.7853 | 1044.4557 | 944.4557 | 879.4429
2349 | 92.4861 | 0.99411 | 1.1687 | 0.99416 | 212.5056 | 54.4329 | 68460.9681 | 68861.6955 | 1373.2339 | 1273.2339 | 879.4507
2350 | 84.1972 | 1.031 | 0.89592 | 1.0422 | 210.73 | 46.3962 | 47559.1757 | 48534.0834 | 966.6817 | 866.6817 | 879.4683
2351 | 97.0565 | 1.024 | 1.2664 | 0.9365 | 203.9817 | 45.8602 | 59150.8347 | 57697.5113 | 1149.9502 | 1049.9502 | 879.4976
2352 | 91.8514 | 0.97828 | 0.83936 | 1.0793 | 187.5236 | 50.413 | 50387.4066 | 52292.2808 | 1041.8456 | 941.8456 | 879.548
2353 | 111.6873 | 0.98442 | 0.9345 | 1.0344 | 204.0004 | 48.1173 | 52006.2449 | 51034.9867 | 1016.6997 | 916.6997 | 879.575
2354 | 99.6687 | 0.93823 | 0.88212 | 0.98561 | 205.9368 | 50.5649 | 48896.7709 | 47899.93 | 953.9986 | 853.9986 | 879.5766
2355 | 105.5298 | 0.97911 | 1.1424 | 0.95587 | 199.604 | 56.1971 | 67374.7371 | 68045.5804 | 1356.9116 | 1256.9116 | 879.596
2356 | 80.5928 | 1.0408 | 0.9911 | 0.91932 | 201.7767 | 47.3302 | 47296.3349 | 48306.5772 | 962.1315 | 862.1315 | 879.6182
2357 | 95.3921 | 1.0422 | 0.87061 | 1.0118 | 216.4762 | 45.7752 | 45089.0995 | 44732.287 | 890.6457 | 790.6457 | 879.7379
2358 | 87.3084 | 1.042 | 1.02 | 0.97583 | 196.5408 | 47.6046 | 51911.8037 | 51160.6571 | 1019.2131 | 919.2131 | 879.7811
2359 | 105.1365 | 1.0131 | 0.84652 | 0.9637 | 183.9555 | 48.3003 | 44731.5691 | 42114.9588 | 838.2992 | 738.2992 | 879.8504
2360 | 103.32 | 0.95319 | 1.0617 | 0.99473 | 210.1977 | 47.6898 | 55272.6828 | 55100.9687 | 1098.0194 | 998.0194 | 879.9018
2361 | 113.3862 | 1.0214 | 1.2101 | 0.97413 | 199.2751 | 50.8553 | 66036.1519 | 64646.9166 | 1288.9383 | 1188.9383 | 880.0354
2362 | 86.547 | 1.0004 | 0.97023 | 0.90648 | 209.579 | 49.9876 | 48501.4568 | 49358.4418 | 983.1688 | 883.1688 | 880.0562
2363 | 76.5741 | 1.0547 | 0.88424 | 0.97464 | 185.7992 | 46.9314 | 44422.0497 | 45123.433 | 898.4687 | 798.4687 | 880.1891
2364 | 101.3799 | 0.98847 | 1.0842 | 0.9442 | 198.5658 | 50.3208 | 56752.2606 | 57817.4148 | 1152.3483 | 1052.3483 | 880.1898
2365 | 104.7309 | 0.92693 | 0.92815 | 0.91032 | 201.4229 | 48.1346 | 45543.8023 | 45285.4827 | 901.7097 | 801.7097 | 880.2075
2366 | 94.218 | 0.92751 | 0.96783 | 1.0443 | 186.157 | 51.027 | 56216.4137 | 58144.3389 | 1158.8868 | 1058.8868 | 880.3091
2367 | 104.8025 | 0.98581 | 0.81874 | 1.1023 | 180.5255 | 50.0823 | 50552.3817 | 50345.1366 | 1002.9027 | 902.9027 | 880.3167
2368 | 102.0931 | 0.93636 | 1.0482 | 1.0771 | 206.8775 | 53.0863 | 65216.8294 | 65144.492 | 1298.8898 | 1198.8898 | 880.3434
2369 | 95.5465 | 1.0003 | 0.94587 | 0.95429 | 201.885 | 50.0687 | 50181.0344 | 49573.8004 | 987.476 | 887.476 | 880.3511
2370 | 89.5658 | 1.0812 | 1.0194 | 0.97739 | 197.4435 | 49.5185 | 54328.0083 | 52729.0177 | 1050.5804 | 950.5804 | 880.3525
2371 | 100.8478 | 1.0546 | 0.93568 | 0.97085 | 201.6903 | 43.9041 | 44754.1287 | 45743.3089 | 910.8662 | 810.8662 | 880.38
2372 | 100.0175 | 0.89288 | 0.82288 | 1.0652 | 202.2705 | 50.4904 | 48968.7799 | 49225.1267 | 980.5025 | 880.5025 | 880.3804
2373 | 89.3242 | 1.0374 | 1.0643 | 0.94504 | 202.5203 | 47.8146 | 52726.9395 | 54126.0501 | 1078.521 | 978.521 | 880.4485
2374 | 88.3577 | 0.94941 | 0.89976 | 1.0179 | 195.1408 | 48.8469 | 49029.322 | 50849.0823 | 1012.9816 | 912.9816 | 880.4679
2375 | 103.0751 | 1.0347 | 0.88333 | 1.0294 | 197.4596 | 45.6648 | 46590.1625 | 48410.332 | 964.2066 | 864.2066 | 880.4684
2376 | 104.6669 | 1.0499 | 1.0723 | 1.0157 | 203.8306 | 52.6329 | 63315.8423 | 64286.2435 | 1281.7249 | 1181.7249 | 880.5025
2377 | 119.9726 | 1.015 | 0.84308 | 1.0225 | 203.5763 | 46.4912 | 45942.102 | 46434.2871 | 924.6857 | 824.6857 | 880.6053
2378 | 87.0878 | 0.9684 | 0.79213 | 0.94762 | 208.0459 | 48.0121 | 40296.9188 | 39325.5198 | 782.5104 | 682.5104 | 880.6667
2379 | 96.511 | 0.93892 | 0.74417 | 1.0597 | 190.8781 | 43.974 | 38853.7945 | 38561.4495 | 767.229 | 667.229 | 880.736
2380 | 102.784 | 1.1064 | 1.1096 | 1.0567 | 210.1424 | 51.5263 | 66488.5447 | 66660.7975 | 1329.216 | 1229.216 | 880.7885
2381 | 83.0291 | 1.0251 | 0.94607 | 0.98066 | 192.5666 | 48.7893 | 49610.686 | 49962.4735 | 995.2495 | 895.2495 | 880.8374
2382 | 89.0833 | 0.90976 | 0.90325 | 1.0975 | 191.8812 | 48.1377 | 51811.7437 | 52914.2732 | 1054.2855 | 954.2855 | 881.0705
2383 | 102.2823 | 0.98548 | 0.90068 | 1.0607 | 196.0606 | 52.4074 | 55543.9344 | 56499.3243 | 1125.9865 | 1025.9865 | 881.0829
2384 | 102.9214 | 0.98891 | 0.98429 | 0.94088 | 192.0974 | 48.781 | 50332.8252 | 49733.6084 | 990.6722 | 890.6722 | 881.1189
2385 | 103.7504 | 0.93924 | 0.72863 | 0.97401 | 188.9299 | 53.3602 | 43258.2558 | 42186.9272 | 839.7385 | 739.7385 | 881.1573
2386 | 109.7115 | 0.99626 | 0.90124 | 0.96334 | 192.4682 | 54.7268 | 53687.6182 | 54340.2159 | 1082.8043 | 982.8043 | 881.177
2387 | 105.3582 | 0.9899 | 0.93246 | 1.0246 | 205.6657 | 49.391 | 52547.2488 | 53189.5689 | 1059.7914 | 959.7914 | 881.2046
2388 | 102.977 | 1.0078 | 1.0372 | 0.97263 | 207.6899 | 49.7669 | 55580.1278 | 54097.279 | 1077.9456 | 977.9456 | 881.2892
2389 | 84.1714 | 0.98937 | 0.87672 | 0.9269 | 200.6519 | 51.0131 | 45903.5659 | 45411.331 | 904.2266 | 804.2266 | 881.3114
2390 | 113.624 | 0.95006 | 1.051 | 0.96428 | 199.9072 | 54.5239 | 61344.9751 | 60077.611 | 1197.5522 | 1097.5522 | 881.3649
2391 | 88.3429 | 0.91932 | 0.80004 | 1.0025 | 209.4822 | 49.7428 | 44144.5885 | 44002.7329 | 876.0547 | 776.0547 | 881.5131
2392 | 84.6812 | 0.95768 | 1.1221 | 0.91506 | 191.29 | 51.6357 | 57400.2387 | 57117.4475 | 1138.3489 | 1038.3489 | 881.5786
2393 | 94.6213 | 0.92821 | 1.1136 | 0.93871 | 199.9703 | 46.6716 | 53089.4994 | 53370.8879 | 1063.4178 | 963.4178 | 881.657
2394 | 106.11 | 1.1049 | 0.9232 | 1.1065 | 185.3841 | 53.1979 | 60765.9063 | 58816.9099 | 1172.3382 | 1072.3382 | 881.7948
2395 | 107.6036 | 0.95036 | 1.2277 | 0.9813 | 198.7544 | 48.5208 | 63613.0494 | 63806.0941 | 1272.1219 | 1172.1219 | 881.8134
2396 | 98.3052 | 0.98808 | 0.57567 | 0.95978 | 193.2385 | 47.6943 | 31177.8546 | 31079.4339 | 617.5887 | 517.5887 | 881.9031
2397 | 85.829 | 0.98389 | 1.1227 | 0.99012 | 187.7674 | 49.6678 | 59592.7259 | 60827.8432 | 1212.5569 | 1112.5569 | 881.9103
2398 | 92.4995 | 0.93052 | 0.96582 | 0.95396 | 212.1775 | 51.8292 | 52426.2458 | 54122.7063 | 1078.4541 | 978.4541 | 881.9148
2399 | 107.9648 | 0.96838 | 1.1619 | 1.0613 | 201.2329 | 50.254 | 67425.5198 | 67461.1114 | 1345.2222 | 1245.2222 | 881.9255
2400 | 99.2335 | 0.92512 | 0.87695 | 0.89898 | 214.0476 | 46.3737 | 41030.2858 | 40743.3131 | 810.8663 | 710.8663 | 881.9349
2401 | 105.0428 | 1.0376 | 1.0782 | 0.98328 | 203.5443 | 49.5524 | 58138.645 | 57225.534 | 1140.5107 | 1040.5107 | 881.9938
2402 | 98.2091 | 0.99323 | 1.1609 | 1.0644 | 194.3726 | 51.9112 | 69403.2731 | 70299.478 | 1401.9896 | 1301.9896 | 882.0473
2403 | 94.2515 | 0.94736 | 0.86433 | 1.0069 | 217.0124 | 49.1644 | 47395.076 | 46641.6187 | 928.8324 | 828.8324 | 882.1425
2404 | 110.865 | 1.0386 | 1.097 | 1.0554 | 184.5376 | 53.2824 | 68006.0558 | 66374.5444 | 1323.4909 | 1223.4909 | 882.2816
2405 | 96.6216 | 0.98203 | 0.95527 | 1.0642 | 217.6783 | 45.3383 | 50608.8256 | 47320.2874 | 942.4057 | 842.4057 | 882.2857
2406 | 99.349 | 0.97346 | 1.217 | 0.94982 | 205.1962 | 49.9832 | 62813.9141 | 64568.09 | 1287.3618 | 1187.3618 | 882.3665
2407 | 104.4969 | 1.0169 | 1.0667 | 0.99625 | 200.762 | 50.3636 | 59071.5966 | 60954.7258 | 1215.0945 | 1115.0945 | 882.4535
2408 | 104.8527 | 0.94216 | 1.2019 | 0.99407 | 180.9999 | 52.2988 | 67833.0632 | 68283.6356 | 1361.6727 | 1261.6727 | 882.4565
2409 | 94.5123 | 0.97749 | 0.74697 | 1.013 | 202.0318 | 47.7704 | 40761.578 | 43205.0218 | 860.1004 | 760.1004 | 882.4993
2410 | 111.492 | 0.99553 | 1.0428 | 0.98254 | 203.2217 | 53.0429 | 60438.8934 | 60168.0822 | 1199.3616 | 1099.3616 | 882.6478
2411 | 105.911 | 0.97074 | 0.99149 | 0.94325 | 186.2584 | 53.3889 | 55605.9718 | 55449.17 | 1104.9834 | 1004.9834 | 882.6721
2412 | 93.9592 | 1.0005 | 1.0773 | 0.93075 | 208.0122 | 52.366 | 57638.8228 | 56492.3594 | 1125.8472 | 1025.8472 | 882.7288
2413 | 92.7235 | 0.99387 | 0.87408 | 0.99483 | 208.0002 | 51.7681 | 49994.0375 | 50372.3326 | 1003.4467 | 903.4467 | 882.7337
2414 | 98.4195 | 1.0362 | 0.9918 | 1.0406 | 203.4292 | 51.2455 | 58320.0404 | 57999.9406 | 1155.9988 | 1055.9988 | 882.7682
2415 | 107.9605 | 0.98449 | 0.90582 | 1.0112 | 190.8656 | 49.2667 | 50555.8756 | 49741.8804 | 990.8376 | 890.8376 | 882.785
2416 | 95.4905 | 1.01 | 0.96648 | 1.0264 | 203.8533 | 51.6754 | 56449.3027 | 56313.6451 | 1122.2729 | 1022.2729 | 882.7986
2417 | 85.8833 | 1.0224 | 1.1261 | 1.0032 | 185.301 | 46.3467 | 56611.6943 | 57357.6461 | 1143.1529 | 1043.1529 | 882.8159
2418 | 92.0421 | 0.99645 | 0.91536 | 1.0032 | 184.664 | 50.7419 | 51432.0268 | 50465.7181 | 1005.3144 | 905.3144 | 882.8998
2419 | 94.129 | 1.022 | 1.1564 | 1.0369 | 186.9606 | 48.2202 | 62645.0605 | 62690.7928 | 1249.8159 | 1149.8159 | 882.9742
2420 | 87.8337 | 0.94946 | 0.91146 | 0.91328 | 216.9428 | 48.0332 | 44206.7043 | 45612.7022 | 908.254 | 808.254 | 882.9919
2421 | 89.017 | 1.0277 | 1.0293 | 1.0151 | 199.8763 | 55.9138 | 63736.4104 | 63647.3674 | 1268.9473 | 1168.9473 | 883.0017
2422 | 79.1295 | 0.95547 | 1.0778 | 0.97419 | 201.2742 | 53.1609 | 60036.9634 | 59472.9664 | 1185.4593 | 1085.4593 | 883.11
2423 | 96.1811 | 0.99148 | 0.99759 | 1.0275 | 210.4515 | 50.0783 | 56316.2526 | 58484.6114 | 1165.6922 | 1065.6922 | 883.1112
2424 | 107.5741 | 0.97171 | 1.0596 | 0.91332 | 202.257 | 51.0905 | 54984.5682 | 54732.1304 | 1090.6426 | 990.6426 | 883.1688
2425 | 84.3525 | 0.96182 | 1.26 | 1.1063 | 195.3611 | 51.2386 | 75772.6824 | 74610.4617 | 1488.2092 | 1388.2092 | 883.2383
2426 | 98.128 | 1.0976 | 1.1756 | 0.99804 | 204.1479 | 52.5838 | 67563.2366 | 66887.3133 | 1333.7463 | 1233.7463 | 883.2874
2427 | 112.4884 | 1.0566 | 0.90795 | 0.9377 | 179.3219 | 51.6076 | 50250.6507 | 50666.7599 | 1009.3352 | 909.3352 | 883.335
2428 | 86.23 | 1.0431 | 1.0494 | 0.95816 | 184.0694 | 46.0606 | 50640.6308 | 50658.6016 | 1009.172 | 909.172 | 883.3906
2429 | 108.1636 | 1.0241 | 1.0388 | 0.92741 | 215.6564 | 50.0643 | 53992.2962 | 55072.6189 | 1097.4524 | 997.4524 | 883.4372
2430 | 89.3532 | 1.0489 | 1.0448 | 1.0338 | 203.6912 | 50.3439 | 59298.4985 | 57751.9837 | 1151.0397 | 1051.0397 | 883.5394
2431 | 105.6384 | 1.0255 | 0.90158 | 0.93651 | 204.8242 | 50.8514 | 48649.2513 | 47876.5517 | 953.531 | 853.531 | 883.6456
2432 | 83.6599 | 1.0548 | 0.89922 | 0.97807 | 180.4366 | 46.2899 | 44977.1308 | 44725.7903 | 890.5158 | 790.5158 | 883.6467
2433 | 84.1786 | 1.0075 | 0.95831 | 1.0448 | 203.6934 | 51.4277 | 56057.5565 | 56827.9133 | 1132.5583 | 1032.5583 | 883.6618
2434 | 105.3239 | 1.0485 | 1.235 | 0.96361 | 193.7827 | 49.9138 | 65107.8611 | 63697.5663 | 1269.9513 | 1169.9513 | 883.7195
2435 | 102.5041 | 0.9953 | 0.88348 | 1.0321 | 195.1752 | 49.3876 | 50265.97 | 50113.0299 | 998.2606 | 898.2606 | 883.7209
2436 | 94.5162 | 0.92794 | 0.84399 | 0.99533 | 212.3841 | 52.067 | 48517.7268 | 49996.9614 | 995.9392 | 895.9392 | 883.826
2437 | 86.3845 | 1.0079 | 1.1203 | 1.0763 | 201.9886 | 48.6965 | 63157.6833 | 64654.8061 | 1289.0961 | 1189.0961 | 883.829
2438 | 103.7218 | 1.0559 | 0.91483 | 1.0396 | 210.6829 | 55.4971 | 59067.5104 | 58859.2982 | 1173.186 | 1073.186 | 883.8986
2439 | 104.8264 | 1.0426 | 1.064 | 1.0008 | 189.7169 | 50.8425 | 59881.135 | 58734.9833 | 1170.6997 | 1070.6997 | 884.0135
2440 | 92.3061 | 1.0041 | 0.9207 | 1.1029 | 200.0385 | 55.1248 | 61283.1807 | 61329.6045 | 1222.5921 | 1122.5921 | 884.0228
2441 | 103.3988 | 1.0172 | 1.3262 | 1.0472 | 179.005 | 51.4022 | 76974.7176 | 77918.0817 | 1554.3616 | 1454.3616 | 884.0393
2442 | 97.2132 | 0.98409 | 0.98548 | 0.98815 | 216.5745 | 51.53 | 55326.4098 | 56590.0345 | 1127.8007 | 1027.8007 | 884.1746
2443 | 111.8682 | 1.0503 | 0.88825 | 0.95347 | 199.7374 | 50.3522 | 48759.9851 | 46893.6476 | 933.873 | 833.873 | 884.2999
2444 | 103.4122 | 0.95228 | 0.86197 | 0.99476 | 200.3902 | 50.937 | 48892.2139 | 49641.1565 | 988.8231 | 888.8231 | 884.3897
2445 | 99.9317 | 0.98841 | 1.3239 | 0.9937 | 194.2277 | 51.0144 | 72345.0633 | 72282.7575 | 1441.6551 | 1341.6551 | 884.437
2446 | 94.4045 | 0.9959 | 1.0929 | 0.97872 | 216.0505 | 50.6006 | 59098.2223 | 58969.9406 | 1175.3988 | 1075.3988 | 884.5616
2447 | 96.4357 | 1.0091 | 0.85779 | 0.96377 | 204.5742 | 49.0143 | 45495.2323 | 45749.7843 | 910.9957 | 810.9957 | 884.6092
2448 | 115.6919 | 0.96182 | 1.1309 | 1.034 | 208.5443 | 45.6382 | 58654.4073 | 60077.5394 | 1197.5508 | 1097.5508 | 884.7352
2449 | 103.9784 | 1.0868 | 1.0941 | 0.93997 | 210.2984 | 52.2264 | 59824.614 | 60788.6836 | 1211.7737 | 1111.7737 | 884.8716
2450 | 95.2952 | 0.9106 | 1.017 | 0.90671 | 192.8084 | 55.6609 | 56347.204 | 57590.0311 | 1147.8006 | 1047.8006 | 884.8836
2451 | 110.7829 | 1.0877 | 1.2057 | 1.0336 | 220.1787 | 49.6721 | 68109.2848 | 69059.715 | 1377.1943 | 1277.1943 | 884.9059
2452 | 94.5323 | 0.94502 | 1.0831 | 1.015 | 194.2689 | 50.4331 | 60141.2524 | 59786.4467 | 1191.7289 | 1091.7289 | 884.9626
2453 | 97.1687 | 0.94784 | 1.1763 | 1.0278 | 187.4124 | 47.8742 | 62475.6977 | 63087.8571 | 1257.7571 | 1157.7571 | 884.9841
2454 | 95.0015 | 0.98003 | 1.0672 | 0.96966 | 196.188 | 52.242 | 59119.751 | 59192.4925 | 1179.8498 | 1079.8498 | 885.0746
2455 | 115.1876 | 1.01 | 0.86652 | 1.0068 | 214.6253 | 49.9725 | 49624.0419 | 50416.6535 | 1004.3331 | 904.3331 | 885.1073
2456 | 105.9929 | 0.91913 | 0.87784 | 0.95031 | 200.3241 | 48.5456 | 45427.6745 | 47016.721 | 936.3344 | 836.3344 | 885.1467
2457 | 85.9907 | 0.92682 | 0.82544 | 1.0416 | 212.0595 | 45.4617 | 42922.9902 | 42229.9319 | 840.5986 | 740.5986 | 885.1748
2458 | 121.9899 | 1.1029 | 1.0424 | 0.99855 | 198.6361 | 47.8767 | 56474.3591 | 55729.0138 | 1110.5803 | 1010.5803 | 885.234
2459 | 102.6161 | 1.008 | 1.0734 | 1.0132 | 197.0177 | 52.6907 | 62949.0254 | 63061.9293 | 1257.2386 | 1157.2386 | 885.2837
2460 | 67.8832 | 0.98021 | 0.70388 | 1.096 | 193.249 | 53.1498 | 44731.8309 | 43131.2479 | 858.625 | 758.625 | 885.2919
2461 | 97.5504 | 0.98542 | 1.246 | 0.99135 | 199.0764 | 46.7963 | 62501.9692 | 62271.9614 | 1241.4392 | 1141.4392 | 885.2933
2462 | 110.8652 | 0.96896 | 0.98318 | 1.0102 | 204.9393 | 52.7974 | 58315.1027 | 59193.4218 | 1179.8684 | 1079.8684 | 885.3235
2463 | 117.7658 | 1.033 | 0.93106 | 0.98481 | 196.0229 | 51.9223 | 54120.9926 | 54836.148 | 1092.723 | 992.723 | 885.3649
2464 | 117.0674 | 1.016 | 0.81511 | 1.018 | 205.3695 | 48.1229 | 45861.1621 | 44679.1298 | 889.5826 | 789.5826 | 885.4241
2465 | 104.2289 | 0.96833 | 1.0481 | 1.0961 | 196.5101 | 47.952 | 60126.0984 | 60205.7559 | 1200.1151 | 1100.1151 | 885.425
2466 | 76.3484 | 0.95783 | 1.0834 | 1.0123 | 202.8228 | 49.3864 | 57981.6649 | 59163.5054 | 1179.2701 | 1079.2701 | 885.4712
2467 | 90.8685 | 1.0448 | 1.1834 | 1.0852 | 186.6942 | 51.0806 | 70637.6063 | 72469.7525 | 1445.3951 | 1345.3951 | 885.501
2468 | 100.7748 | 0.99772 | 0.97676 | 0.93316 | 216.0106 | 50.208 | 51027.3453 | 51253.4081 | 1021.0682 | 921.0682 | 885.5439
2469 | 109.5688 | 0.92497 | 0.89866 | 1.063 | 206.172 | 50.3219 | 53378.3901 | 55457.3109 | 1105.1462 | 1005.1462 | 885.6015
2470 | 94.957 | 1.053 | 1.0631 | 0.98705 | 198.9245 | 50.0785 | 57755.9887 | 58091.4189 | 1157.8284 | 1057.8284 | 885.6162
2471 | 91.1704 | 1.0427 | 1.1415 | 1.1074 | 201.4838 | 49.5555 | 67559.2679 | 68402.3963 | 1364.0479 | 1264.0479 | 885.6552
2472 | 99.9688 | 1.0915 | 1.0707 | 0.90395 | 213.6935 | 50.0528 | 54117.0837 | 53066.0349 | 1057.3207 | 957.3207 | 885.6729
2473 | 107.4774 | 0.96498 | 1.0089 | 0.98994 | 186.143 | 54.0419 | 59766.3987 | 59063.8822 | 1177.2776 | 1077.2776 | 885.7015
2474 | 111.725 | 1.0097 | 0.9899 | 1.0903 | 185.9674 | 47.6276 | 56962.2779 | 55916.4597 | 1114.3292 | 1014.3292 | 885.7196
2475 | 98.3893 | 0.95518 | 0.82114 | 1.0399 | 215.9144 | 45.2162 | 43074.2069 | 42383.9404 | 843.6788 | 743.6788 | 885.7303
2476 | 99.9255 | 0.98904 | 0.65329 | 0.95894 | 199.5909 | 55.5994 | 40525.2732 | 40858.1695 | 813.1634 | 713.1634 | 885.7477
2477 | 101.6002 | 0.96596 | 0.86433 | 0.96486 | 192.2553 | 51.1657 | 47883.5842 | 47846.0629 | 952.9213 | 852.9213 | 885.7727
2478 | 99.8859 | 1.0702 | 0.928 | 1.1282 | 194.4209 | 45.4469 | 52636.1718 | 51828.9157 | 1032.5783 | 932.5783 | 885.9005
2479 | 110.6683 | 1.053 | 0.84869 | 1.0343 | 198.7601 | 48.2433 | 48167.0162 | 48358.5684 | 963.1714 | 863.1714 | 886.049
2480 | 84.4921 | 0.97982 | 1.1252 | 1.0849 | 214.344 | 47.4789 | 62104.0615 | 60311.5948 | 1202.2319 | 1102.2319 | 886.0682
2481 | 78.3082 | 0.97044 | 0.87777 | 1.0586 | 203.5215 | 47.5024 | 47951.8431 | 49051.1119 | 977.0222 | 877.0222 | 886.1158
2482 | 115.676 | 0.96751 | 1.2927 | 0.89803 | 200.3645 | 48.4427 | 61859.2794 | 61450.8355 | 1225.0167 | 1125.0167 | 886.1297
2483 | 110.9275 | 0.97486 | 1.1103 | 1.0293 | 212.4205 | 49.1136 | 61651.5919 | 62059.3241 | 1237.1865 | 1137.1865 | 886.15
2484 | 107.466 | 0.97645 | 0.82876 | 0.94456 | 191.4108 | 50.2142 | 44768.8592 | 44242.9309 | 880.8586 | 780.8586 | 886.1541
2485 | 103.4746 | 0.97234 | 0.93062 | 0.99366 | 198.7477 | 50.6218 | 52102.8817 | 50571.5137 | 1007.4303 | 907.4303 | 886.5582
2486 | 96.4164 | 1.0237 | 1.2128 | 0.94191 | 188.7483 | 52.4729 | 65311.0272 | 67073.2843 | 1337.4657 | 1237.4657 | 886.5636
2487 | 98.3787 | 1.0445 | 1.0679 | 1.0217 | 202.6163 | 46.7881 | 56061.4733 | 53846.6052 | 1072.9321 | 972.9321 | 886.5974
2488 | 72.0537 | 0.95524 | 1.2114 | 0.95927 | 201.8946 | 47.4865 | 58651.3479 | 58558.8435 | 1167.1769 | 1067.1769 | 886.6031
2489 | 101.1973 | 1.0707 | 1.3011 | 0.96367 | 197.5555 | 51.8655 | 70849.7909 | 71811.1592 | 1432.2232 | 1332.2232 | 886.7061
2490 | 114.829 | 0.98151 | 0.89118 | 1.0008 | 181.3841 | 48.0979 | 48498.6766 | 47905.7728 | 954.1155 | 854.1155 | 886.7076
2491 | 104.0971 | 0.95568 | 1.1872 | 0.90853 | 220.0306 | 50.116 | 59262.7378 | 58429.2514 | 1164.585 | 1064.585 | 886.7789
2492 | 100.3746 | 1.0349 | 0.82322 | 1.0368 | 183.8562 | 50.0608 | 48111.5349 | 48362.0828 | 963.2417 | 863.2417 | 886.825
2493 | 102.5329 | 1.0647 | 1.0067 | 0.9469 | 202.6448 | 50.5711 | 53930.6195 | 53412.3773 | 1064.2475 | 964.2475 | 886.8306
2494 | 98.5504 | 0.99278 | 0.71958 | 0.89707 | 188.5537 | 46.7837 | 34965.5236 | 35941.1556 | 714.8231 | 614.8231 | 886.8358
2495 | 103.8987 | 1.0256 | 1.0567 | 1.0225 | 198.4339 | 44.238 | 52710.258 | 51859.9826 | 1033.1997 | 933.1997 | 886.919
2496 | 107.3591 | 0.95336 | 1.1861 | 0.94171 | 193.2339 | 53.3943 | 65297.0383 | 65021.3045 | 1296.4261 | 1196.4261 | 886.9298
2497 | 120.0628 | 1.0025 | 1.0702 | 1.0127 | 219.4812 | 49.3531 | 59648.3617 | 60279.4196 | 1201.5884 | 1101.5884 | 886.965
2498 | 94.4291 | 1.0106 | 1.1883 | 0.89998 | 195.9171 | 47.9951 | 56104.3993 | 58708.9917 | 1170.1798 | 1070.1798 | 886.9987
2499 | 98.4402 | 1.0062 | 1.1857 | 1.0267 | 187.4611 | 53.5072 | 70627.0446 | 69088.0637 | 1377.7613 | 1277.7613 | 887.0707
2500 | 87.399 | 0.97107 | 0.77024 | 1.0456 | 196.283 | 46.3595 | 41466.8505 | 41077.3442 | 817.5469 | 717.5469 | 887.099
2501 | 109.4497 | 1.0209 | 0.83854 | 0.9874 | 218.3666 | 50.1391 | 47334.4733 | 46787.9008 | 931.758 | 831.758 | 887.1126
2502 | 101.2692 | 1.0634 | 1.103 | 0.97484 | 204.3602 | 49.5838 | 58860.8835 | 58640.6388 | 1168.8128 | 1068.8128 | 887.2557
2503 | 108.7651 | 1.0116 | 1.1052 | 1.0447 | 191.9452 | 48.1049 | 61028.9138 | 61051.3877 | 1217.0278 | 1117.0278 | 887.3422
2504 | 92.3349 | 1.1298 | 1.1424 | 1.071 | 207.1648 | 51.2501 | 68259.3484 | 66555.5605 | 1327.1112 | 1227.1112 | 887.476
2505 | 98.9748 | 1.0311 | 0.8751 | 0.97997 | 204.3189 | 52.8512 | 50921.7427 | 50836.6097 | 1012.7322 | 912.7322 | 887.4983
2506 | 110.3489 | 1.055 | 0.9444 | 0.98102 | 207.2874 | 53.0029 | 55484.0344 | 54113.9589 | 1078.2792 | 978.2792 | 887.5153
2507 | 103.5701 | 0.98565 | 1.0885 | 0.99599 | 199.0073 | 52.2792 | 62214.2991 | 61050.7443 | 1217.0149 | 1117.0149 | 887.5507
2508 | 104.5993 | 1.0199 | 0.82328 | 0.97433 | 188.124 | 50.6955 | 46261.6997 | 45386.5437 | 903.7309 | 803.7309 | 887.5967
2509 | 104.3547 | 0.98176 | 1.1386 | 0.99712 | 218.5569 | 54.4233 | 67582.9288 | 65978.8386 | 1315.5768 | 1215.5768 | 887.6028
2510 | 85.9088 | 0.94912 | 1.0855 | 1.0203 | 207.2718 | 50.2963 | 60012.3487 | 60405.1843 | 1204.1037 | 1104.1037 | 887.6393
2511 | 102.0114 | 1.0005 | 0.93149 | 1.0575 | 199.8168 | 51.4989 | 56183.2998 | 56375.9088 | 1123.5182 | 1023.5182 | 887.6397
2512 | 118.1517 | 1.0471 | 0.96633 | 1.0073 | 205.3386 | 49.9791 | 55038.9197 | 54846.537 | 1092.9307 | 992.9307 | 887.6449
2513 | 101.9556 | 0.98411 | 0.96507 | 0.91685 | 208.8289 | 46.7022 | 46217.8164 | 46988.1748 | 935.7635 | 835.7635 | 887.7455
2514 | 109.8364 | 0.9716 | 0.88123 | 0.94521 | 189.1745 | 52.86 | 49860.1206 | 48697.9219 | 969.9584 | 869.9584 | 887.7525
2515 | 103.6102 | 1.054 | 0.96519 | 0.94923 | 205.6828 | 53.0264 | 54578.1116 | 56903.0047 | 1134.0601 | 1034.0601 | 887.8554
2516 | 98.7613 | 0.93172 | 0.94817 | 0.99095 | 217.4249 | 49.6585 | 51445.5586 | 50817.202 | 1012.344 | 912.344 | 887.908
2517 | 105.7057 | 0.98183 | 1.2408 | 1.0785 | 204.1451 | 47.2399 | 68325.3656 | 69910.5354 | 1394.2107 | 1294.2107 | 887.9704
2518 | 99.8228 | 0.99695 | 1.0773 | 0.88855 | 199.8758 | 48.68 | 51641.1033 | 51678.2447 | 1029.5649 | 929.5649 | 888.0639
2519 | 79.497 | 1.092 | 1.0966 | 1.0092 | 184.7287 | 54.0577 | 64704.1997 | 66666.8588 | 1329.3372 | 1229.3372 | 888.2928
2520 | 105.4253 | 0.98473 | 1.0099 | 0.96893 | 192.6509 | 52.9378 | 57487.1463 | 57668.8695 | 1149.3774 | 1049.3774 | 888.3516
2521 | 102.2048 | 1.0356 | 0.85054 | 0.97397 | 207.9675 | 50.0408 | 46957.9588 | 46793.5088 | 931.8702 | 831.8702 | 888.4376
2522 | 100.5217 | 1.0449 | 0.87729 | 1.0583 | 209.0625 | 49.2159 | 51074.1371 | 49230.2632 | 980.6053 | 880.6053 | 888.5151
2523 | 116.9575 | 1.0451 | 0.82197 | 0.97146 | 207.8016 | 45.5096 | 42110.3015 | 41213.6638 | 820.2733 | 720.2733 | 888.5201
2524 | 101.5382 | 1.0433 | 1.0071 | 0.91546 | 194.7811 | 49.9123 | 51500.1945 | 51369.8818 | 1023.3976 | 923.3976 | 888.6381
2525 | 98.9358 | 1.011 | 0.92557 | 1.0409 | 193.799 | 47.6476 | 50863.9712 | 51537.8121 | 1026.7562 | 926.7562 | 888.6792
2526 | 102.7805 | 0.93885 | 0.91452 | 0.93516 | 203.5389 | 47.5736 | 45480.5782 | 45038.3852 | 896.7677 | 796.7677 | 888.7944
2527 | 98.3297 | 1.0583 | 0.76346 | 1.0491 | 198.876 | 50.6288 | 46018.9648 | 48196.5823 | 959.9316 | 859.9316 | 888.8022
2528 | 95.5969 | 1.0103 | 0.81324 | 0.96909 | 187.9194 | 50.8321 | 45158.594 | 46101.1754 | 918.0235 | 818.0235 | 888.8231
2529 | 88.2328 | 0.95069 | 1.0335 | 1.0047 | 193.41 | 45.6488 | 51419.1244 | 50173.7686 | 999.4754 | 899.4754 | 888.8237
2530 | 101.3256 | 1.0736 | 0.88822 | 1.0399 | 197.2388 | 47.7813 | 49528.7607 | 49688.8636 | 989.7773 | 889.7773 | 888.8289
2531 | 94.4681 | 1.0617 | 1.2241 | 1.0241 | 198.4461 | 44.9891 | 61108.2818 | 61460.9263 | 1225.2185 | 1125.2185 | 888.8603
2532 | 118.9629 | 1.0201 | 1.1843 | 1.0619 | 179.351 | 46.6377 | 64492.5115 | 64222.4156 | 1280.4483 | 1180.4483 | 888.9241
2533 | 97.7926 | 1.0371 | 0.94761 | 1.064 | 206.8341 | 50.6209 | 56380.7387 | 57066.8821 | 1137.3376 | 1037.3376 | 888.9469
2534 | 97.9047 | 1.0949 | 1.028 | 1.0394 | 183.9493 | 47.0782 | 55530.0234 | 56490.5588 | 1125.8112 | 1025.8112 | 888.948
2535 | 108.4073 | 0.99498 | 1.0141 | 0.9966 | 207.2736 | 52.5073 | 58935.2669 | 60282.6559 | 1201.6531 | 1101.6531 | 888.9621
2536 | 97.5547 | 0.95807 | 1.0656 | 0.97312 | 201.6138 | 55.6223 | 63079.3344 | 65395.5807 | 1303.9116 | 1203.9116 | 889.1334
2537 | 107.7551 | 0.97699 | 1.1411 | 0.92658 | 199.0227 | 49.3374 | 57556.836 | 57140.6252 | 1138.8125 | 1038.8125 | 889.1585
2538 | 101.1302 | 0.86895 | 1.1583 | 1.0188 | 200.7681 | 50.7801 | 64586.996 | 65054.6596 | 1297.0932 | 1197.0932 | 889.2084
2539 | 94.2859 | 0.94492 | 0.95846 | 0.99833 | 204.1738 | 48.6854 | 51126.5759 | 51390.0862 | 1023.8017 | 923.8017 | 889.2317
2540 | 104.3093 | 0.99552 | 1.0648 | 0.97601 | 197.4168 | 46.3843 | 53219.3154 | 51686.3663 | 1029.7273 | 929.7273 | 889.2464
2541 | 119.0775 | 1.0031 | 0.94728 | 0.96955 | 199.0311 | 53.9868 | 56230.531 | 56033.8256 | 1116.6765 | 1016.6765 | 889.2984
2542 | 99.2302 | 1.0122 | 1.2607 | 0.91732 | 210.8864 | 49.2315 | 62089.4866 | 62520.7894 | 1246.4158 | 1146.4158 | 889.4912
2543 | 108.3491 | 0.93714 | 0.96799 | 0.9993 | 199.9672 | 51.9885 | 55767.6339 | 55670.5546 | 1109.4111 | 1009.4111 | 889.4952
2544 | 95.4641 | 1.0317 | 0.71059 | 0.98629 | 206.7245 | 47.7352 | 38363.3318 | 37384.9888 | 743.6998 | 643.6998 | 889.5238
2545 | 86.9545 | 1.009 | 0.98598 | 0.97108 | 205.3514 | 46.6239 | 48936.4203 | 49643.0154 | 988.8603 | 888.8603 | 889.56
2546 | 80.5898 | 1.0396 | 1.0373 | 1.0184 | 205.4421 | 46.1391 | 52811.1261 | 52221.1926 | 1040.4239 | 940.4239 | 889.7219
2547 | 104.1511 | 0.9578 | 1.2275 | 0.9877 | 195.199 | 49.3408 | 64940.503 | 63008.6779 | 1256.1736 | 1156.1736 | 889.7303
2548 | 86.0099 | 1.0264 | 0.71997 | 1.0147 | 219.2145 | 51.5527 | 42431.1728 | 42702.8274 | 850.0565 | 750.0565 | 889.7728
2549 | 104.3774 | 0.91139 | 0.90348 | 0.95123 | 200.5219 | 53.2961 | 51074.2789 | 50052.4123 | 997.0482 | 897.0482 | 889.7773
2550 | 103.3712 | 1.13 | 1.1015 | 0.97428 | 206.3601 | 47.2948 | 56486.7992 | 56159.0751 | 1119.1815 | 1019.1815 | 889.8077
2551 | 101.3476 | 0.9713 | 1.2416 | 0.95741 | 208.8484 | 48.3745 | 62475.725 | 64165.6011 | 1279.312 | 1179.312 | 889.8604
2552 | 105.4771 | 1.0487 | 1.2424 | 1.0038 | 193.8533 | 46.2849 | 63034.6764 | 63512.1754 | 1266.2435 | 1166.2435 | 889.9304
2553 | 105.5139 | 0.98007 | 0.93724 | 0.95248 | 211.8053 | 53.7536 | 53756.2153 | 53226.4173 | 1060.5283 | 960.5283 | 889.939
2554 | 126.238 | 0.96204 | 1.2888 | 1.0335 | 205.6444 | 48.5373 | 70745.2979 | 69661.1576 | 1389.2232 | 1289.2232 | 889.9714
2555 | 89.122 | 1.0239 | 1.0534 | 0.9633 | 190.5811 | 53.417 | 59267.0427 | 58831.9256 | 1172.6385 | 1072.6385 | 889.9849
2556 | 96.9011 | 1.0578 | 1.1931 | 1.0617 | 205.111 | 50.8569 | 69837.7006 | 69278.8582 | 1381.5772 | 1281.5772 | 890.0377
2557 | 88.7688 | 0.94356 | 0.84683 | 0.88195 | 180.7763 | 49.5962 | 41376.1951 | 41562.2155 | 827.2443 | 727.2443 | 890.0907
2558 | 93.3161 | 0.96217 | 0.92304 | 1.0216 | 211.9131 | 49.6002 | 51435.9013 | 52169.3984 | 1039.388 | 939.388 | 890.1564
2559 | 106.0683 | 0.97385 | 1.0088 | 0.98845 | 210.7798 | 53.7444 | 59355.746 | 60026.4683 | 1196.5294 | 1096.5294 | 890.1786
2560 | 91.6256 | 0.98694 | 1.0661 | 0.99991 | 182.5962 | 51.2985 | 59506.3563 | 58003.0737 | 1156.0615 | 1056.0615 | 890.2515
2561 | 96.7621 | 0.97462 | 1.3118 | 1.0794 | 207.4549 | 53.0458 | 80318.8497 | 80223.1712 | 1600.4634 | 1500.4634 | 890.3727
2562 | 98.9562 | 1.0272 | 1.3127 | 1.0121 | 201.8862 | 48.9747 | 70248.4787 | 71607.5623 | 1428.1512 | 1328.1512 | 890.374
2563 | 108.8694 | 0.91438 | 0.80324 | 0.97173 | 202.1095 | 49.6536 | 43901.4177 | 42400.9326 | 844.0187 | 744.0187 | 890.5244
2564 | 91.974 | 0.92398 | 0.90529 | 0.98157 | 196.904 | 50.18 | 49051.354 | 48684.3784 | 969.6876 | 869.6876 | 890.6088
2565 | 108.3745 | 0.94328 | 0.98757 | 0.96477 | 188.7529 | 53.0657 | 56173.3849 | 57225.7524 | 1140.515 | 1040.515 | 890.615
2566 | 95.7555 | 1.0424 | 0.95956 | 1.0423 | 201.8477 | 48.4175 | 53458.9031 | 53638.4514 | 1068.769 | 968.769 | 890.6705
2567 | 106.5437 | 1.0883 | 1.0033 | 1.0383 | 203.1391 | 49.3445 | 57328.4735 | 57934.2231 | 1154.6845 | 1054.6845 | 890.6722
2568 | 82.1018 | 1.056 | 1.0453 | 0.93278 | 177.9258 | 51.2774 | 54620.5299 | 55605.2771 | 1108.1055 | 1008.1055 | 890.682
2569 | 82.2959 | 1.0759 | 0.88722 | 1.0754 | 195.2691 | 47.4732 | 49692.9153 | 51500.935 | 1026.0187 | 926.0187 | 890.7663
2570 | 103.5797 | 0.92888 | 0.9593 | 1.0654 | 193.6434 | 53.0555 | 59521.8705 | 56715.4374 | 1130.3087 | 1030.3087 | 890.8376
2571 | 100.0609 | 1.1099 | 0.99178 | 0.89669 | 195.7493 | 47.169 | 47382.5982 | 47476.6766 | 945.5335 | 845.5335 | 890.9482
2572 | 109.0926 | 0.95735 | 1.0637 | 0.92774 | 203.4874 | 50.3181 | 55113.8294 | 57197.2247 | 1139.9445 | 1039.9445 | 890.9606
2573 | 102.327 | 0.93954 | 0.82621 | 0.9948 | 204.2996 | 48.4718 | 44704.155 | 45697.8357 | 909.9567 | 809.9567 | 891.0076
2574 | 109.8157 | 0.97788 | 0.89876 | 0.99473 | 208.4234 | 49.8065 | 50085.115 | 49989.441 | 995.7888 | 895.7888 | 891.0291
2575 | 98.6103 | 0.96471 | 1.0781 | 0.97265 | 198.198 | 51.4836 | 59079.9539 | 59813.2539 | 1192.2651 | 1092.2651 | 891.0413
2576 | 93.0998 | 0.99809 | 0.90661 | 1.0199 | 202.8235 | 49.9154 | 50995.9802 | 50508.959 | 1006.1792 | 906.1792 | 891.0538
2577 | 104.1538 | 1.0691 | 0.97896 | 0.98832 | 193.935 | 49.9984 | 54136.3533 | 53767.1542 | 1071.3431 | 971.3431 | 891.1324
2578 | 95.5404 | 0.98939 | 0.97901 | 1.038 | 194.1659 | 52.0326 | 57990.2316 | 59357.2678 | 1183.1454 | 1083.1454 | 891.2354
2579 | 97.9849 | 0.97624 | 1.0315 | 0.97469 | 201.2728 | 51.4841 | 56889.2677 | 57452.8899 | 1145.0578 | 1045.0578 | 891.2444
2580 | 116.7972 | 1.0125 | 0.9915 | 1.0007 | 206.9576 | 47.0132 | 52411.8229 | 49860.9742 | 993.2195 | 893.2195 | 891.2923
2581 | 98.5287 | 1.0054 | 1.2126 | 0.92393 | 197.9732 | 49.6206 | 60706.0581 | 61562.1293 | 1227.2426 | 1127.2426 | 891.2982
2582 | 109.8388 | 0.93517 | 0.97829 | 1.0659 | 213.9025 | 52.0034 | 59784.3713 | 58680.6577 | 1169.6132 | 1069.6132 | 891.31
2583 | 97.8171 | 0.96981 | 0.83356 | 1.119 | 209.2303 | 50.8513 | 52465.6623 | 54370.2836 | 1083.4057 | 983.4057 | 891.3236
2584 | 112.3188 | 1.0384 | 0.86115 | 0.99644 | 187.4128 | 52.194 | 51061.5998 | 51548.1809 | 1026.9636 | 926.9636 | 891.3982
2585 | 102.7518 | 1.0808 | 1.0161 | 1.0528 | 198.0435 | 56.0734 | 66407.3874 | 64369.3894 | 1283.3878 | 1183.3878 | 891.4068
2586 | 86.5872 | 1.0333 | 1.1752 | 0.90822 | 202.6389 | 50.9343 | 59122.8149 | 58914.334 | 1174.2867 | 1074.2867 | 891.4127
2587 | 100.2108 | 0.96054 | 1.2761 | 0.90821 | 204.0242 | 55.3243 | 69649.4883 | 70603.689 | 1408.0738 | 1308.0738 | 891.4144
2588 | 91.2103 | 0.98763 | 1.0994 | 0.9594 | 197.8825 | 46.4577 | 53385.4299 | 51501.9133 | 1026.0383 | 926.0383 | 891.4565
2589 | 99.5894 | 0.95507 | 1.293 | 0.99127 | 205.4322 | 52.6103 | 72638.5767 | 71727.7651 | 1430.5553 | 1330.5553 | 891.5355
2590 | 102.9021 | 1.0266 | 1.1776 | 1.0394 | 209.0022 | 50.0328 | 66730.9218 | 66706.0059 | 1330.1201 | 1230.1201 | 891.5499
2591 | 96.5974 | 0.92127 | 1.1283 | 1.0345 | 208.911 | 51.1782 | 64499.7176 | 61713.972 | 1230.2794 | 1130.2794 | 891.5994
2592 | 88.3866 | 1.0273 | 1.0399 | 0.96892 | 184.2552 | 48.337 | 53279.1078 | 53058.7238 | 1057.1745 | 957.1745 | 891.6082
2593 | 94.1941 | 1.005 | 0.90784 | 1.0119 | 169.8389 | 47.3417 | 48141.7848 | 48201.2786 | 960.0256 | 860.0256 | 891.6654
2594 | 108.931 | 1.0277 | 0.81051 | 1.0528 | 201.1645 | 50.3912 | 48842.7389 | 50486.0328 | 1005.7207 | 905.7207 | 891.8684
2595 | 75.7241 | 0.95002 | 1.0733 | 1.0059 | 210.1445 | 46.141 | 53344.4384 | 52978.6898 | 1055.5738 | 955.5738 | 891.9945
2596 | 105.3329 | 1.0374 | 1.049 | 1.0896 | 200.449 | 50.911 | 63955.2475 | 64673.2282 | 1289.4646 | 1189.4646 | 892.0043
2597 | 93.0333 | 1.0532 | 0.772 | 1.0253 | 204.2784 | 49.5112 | 44243.1508 | 45793.1836 | 911.8637 | 811.8637 | 892.1131
2598 | 98.7836 | 0.99513 | 0.90815 | 1.001 | 188.4078 | 50.5337 | 51093.2201 | 51057.8649 | 1017.1573 | 917.1573 | 892.1331
2599 | 107.6988 | 1.0364 | 0.7858 | 1.0483 | 213.0772 | 50.5882 | 47531.9274 | 46995.9849 | 935.9197 | 835.9197 | 892.1343
2600 | 100.9521 | 1.0448 | 1.1166 | 0.96058 | 196.2175 | 48.8614 | 57756.3235 | 58464.9483 | 1165.299 | 1065.299 | 892.1434
2601 | 107.6536 | 1.0229 | 0.96926 | 1.036 | 211.0292 | 46.1683 | 51653.6805 | 51760.7627 | 1031.2153 | 931.2153 | 892.1697
2602 | 84.1657 | 0.99963 | 1.1547 | 1.0847 | 198.5747 | 49.4249 | 66259.5526 | 67398.0025 | 1343.9601 | 1243.9601 | 892.1996
2603 | 107.9228 | 0.93741 | 0.73402 | 1.1283 | 205.6748 | 52.2127 | 48730.4963 | 48457.818 | 965.1564 | 865.1564 | 892.2431
2604 | 109.7724 | 1.0394 | 1.1389 | 0.98567 | 202.6879 | 56.4315 | 69989.9481 | 69750.734 | 1391.0147 | 1291.0147 | 892.3291
2605 | 97.7728 | 1.0786 | 1.0141 | 0.95715 | 182.7401 | 48.57 | 52449.7471 | 52623.0205 | 1048.4604 | 948.4604 | 892.3891
2606 | 103.2898 | 0.99005 | 1.0406 | 1.0032 | 200.2015 | 52.7051 | 60609.0159 | 62878.8772 | 1253.5775 | 1153.5775 | 892.4468
2607 | 96.1476 | 1.0361 | 1.0309 | 1.0326 | 208.1781 | 49.6675 | 58031.336 | 58172.2602 | 1159.4452 | 1059.4452 | 892.4982
2608 | 110.1316 | 1.0047 | 0.93729 | 1.0069 | 195.1281 | 46.9213 | 49670.3211 | 50300.2754 | 1002.0055 | 902.0055 | 892.588
2609 | 95.5823 | 0.96624 | 1.1606 | 0.97094 | 205.1917 | 51.9699 | 63566.3078 | 63310.962 | 1262.2192 | 1162.2192 | 892.6904
2610 | 103.8577 | 0.9816 | 1.1564 | 1.0347 | 200.3694 | 48.9557 | 63768.376 | 63826.2336 | 1272.5247 | 1172.5247 | 892.6994
2611 | 94.2928 | 1.0677 | 0.93473 | 0.98583 | 206.6802 | 52.0074 | 53366.283 | 55222.1644 | 1100.4433 | 1000.4433 | 892.7347
2612 | 90.6238 | 1.0281 | 0.88173 | 0.95147 | 195.8517 | 47.7919 | 44743.5891 | 44481.2004 | 885.624 | 785.624 | 892.7995
2613 | 105.1945 | 0.99896 | 0.93691 | 0.9934 | 197.2656 | 48.8158 | 50761.2589 | 51577.9901 | 1027.5598 | 927.5598 | 892.8314
2614 | 99.5283 | 0.98972 | 1.0473 | 0.89165 | 194.7276 | 46.3842 | 48077.1262 | 46873.5077 | 933.4702 | 833.4702 | 892.8521
2615 | 93.2987 | 1.034 | 0.86282 | 0.98386 | 205.3998 | 50.1237 | 47590.4355 | 47742.3082 | 950.8462 | 850.8462 | 892.8522
2616 | 108.8294 | 1.0284 | 1.0781 | 0.98171 | 170.7153 | 51.1474 | 60028.5924 | 60440.4528 | 1204.8091 | 1104.8091 | 892.8771
2617 | 103.708 | 0.96774 | 0.76735 | 0.99428 | 203.4967 | 50.4277 | 43738.5225 | 43669.54 | 869.3908 | 769.3908 | 892.9606
2618 | 103.0821 | 0.97027 | 1.0881 | 1.0454 | 189.4662 | 50.4802 | 62658.9204 | 61353.6787 | 1223.0736 | 1123.0736 | 893.0207
2619 | 108.2576 | 0.9556 | 0.96725 | 0.96139 | 196.9879 | 51.3359 | 53244.8124 | 54059.3337 | 1077.1867 | 977.1867 | 893.0427
2620 | 101.7146 | 0.98126 | 0.92961 | 1.0326 | 206.1239 | 51.647 | 54937.0064 | 53664.4212 | 1069.2884 | 969.2884 | 893.066
2621 | 116.0273 | 0.93923 | 0.95394 | 1.0329 | 193.1382 | 53.0111 | 58203.3845 | 56179.5641 | 1119.5913 | 1019.5913 | 893.0721
2622 | 92.5646 | 1.0122 | 1.157 | 1.0067 | 204.3471 | 50.4105 | 63640.7316 | 64344.9029 | 1282.8981 | 1182.8981 | 893.1464
2623 | 106.5851 | 1.0608 | 0.9459 | 0.92842 | 205.1325 | 50.2407 | 50006.6821 | 50567.4643 | 1007.3493 | 907.3493 | 893.1795
2624 | 81.5178 | 0.96976 | 0.998 | 0.9903 | 202.8111 | 50.6406 | 54255.331 | 54192.6384 | 1079.8528 | 979.8528 | 893.2195
2625 | 101.9894 | 1.0732 | 1.2069 | 1.0687 | 192.0048 | 49.7623 | 69820.0623 | 68225.7742 | 1360.5155 | 1260.5155 | 893.2892
2626 | 92.1865 | 1.031 | 1.0703 | 0.95313 | 189.0845 | 50.1609 | 56127.7098 | 56378.3833 | 1123.5677 | 1023.5677 | 893.3032
2627 | 98.1229 | 1.0403 | 0.87098 | 0.96357 | 191.9852 | 52.2886 | 49412.5836 | 49466.1731 | 985.3235 | 885.3235 | 893.3201
2628 | 91.901 | 0.99124 | 1.1264 | 0.95967 | 198.8289 | 53.7022 | 63142.7651 | 63301.8083 | 1262.0362 | 1162.0362 | 893.4024
2629 | 97.2959 | 1.0748 | 1.2424 | 1.0299 | 194.4642 | 54.9368 | 76237.7721 | 76447.6213 | 1524.9524 | 1424.9524 | 893.6876
2630 | 104.3618 | 0.94097 | 0.83594 | 0.97115 | 184.7217 | 48.3575 | 44191.6266 | 43375.1156 | 863.5023 | 763.5023 | 893.7443
2631 | 97.3198 | 1.001 | 0.95905 | 0.98472 | 197.1564 | 47.3626 | 49540.321 | 49824.9082 | 992.4982 | 892.4982 | 893.8829
2632 | 109.4418 | 1.0163 | 1.0738 | 0.90963 | 183.0239 | 48.0659 | 52480.4119 | 53496.8419 | 1065.9368 | 965.9368 | 893.8993
2633 | 105.1195 | 0.95988 | 1.2295 | 0.98195 | 184.7316 | 50.9298 | 66812.2236 | 66796.2929 | 1331.9259 | 1231.9259 | 893.9929
2634 | 90.1721 | 1.0128 | 0.94878 | 0.97502 | 206.3742 | 52.5945 | 53663.5883 | 53860.2458 | 1073.2049 | 973.2049 | 893.9968
2635 | 98.4945 | 1.0539 | 0.99433 | 0.95029 | 194.7103 | 46.2906 | 48739.8323 | 48664.0696 | 969.2814 | 869.2814 | 894.1019
2636 | 94.3125 | 0.90427 | 0.88133 | 0.99027 | 213.8115 | 53.8386 | 51793.232 | 50822.678 | 1012.4536 | 912.4536 | 894.1899
2637 | 82.3753 | 1.057 | 0.9786 | 0.99673 | 208.6592 | 45.1312 | 48158.9564 | 48757.7022 | 971.154 | 871.154 | 894.2816
2638 | 87.8888 | 1.0229 | 1.2031 | 1.0147 | 202.9265 | 51.3732 | 67542.3031 | 67067.6312 | 1337.3526 | 1237.3526 | 894.2904
2639 | 107.8173 | 1.0628 | 1.1798 | 1.0797 | 206.5066 | 53.6742 | 74728.4865 | 73938.7934 | 1474.7759 | 1374.7759 | 894.3044
2640 | 98.6063 | 0.96293 | 1.136 | 1.0446 | 198.1475 | 45.4368 | 58432.2836 | 60640.0973 | 1208.8019 | 1108.8019 | 894.3335
2641 | 87.8363 | 0.98682 | 0.97832 | 1.0511 | 202.958 | 49.5707 | 55473.5243 | 56218.1591 | 1120.3632 | 1020.3632 | 894.4006
2642 | 100.1761 | 1.0282 | 0.99961 | 0.99205 | 213.3419 | 50.9825 | 56021.5154 | 54728.4507 | 1090.569 | 990.569 | 894.437
2643 | 80.3522 | 0.90215 | 0.90077 | 0.80931 | 203.4017 | 48.1773 | 38817.1122 | 39452.0003 | 785.04 | 685.04 | 894.5688
2644 | 109.8461 | 1.0476 | 1.1053 | 0.97304 | 208.0853 | 54.35 | 64915.5171 | 66177.4997 | 1319.55 | 1219.55 | 894.6122
2645 | 121.3427 | 0.9318 | 1.1919 | 0.99824 | 196.3078 | 57.8252 | 75532.1029 | 75571.7201 | 1507.4344 | 1407.4344 | 894.6961
2646 | 84.1623 | 1.0461 | 1.0196 | 1.033 | 201.709 | 51.0095 | 58420.0387 | 57289.1475 | 1141.783 | 1041.783 | 894.7077
2647 | 102.5975 | 0.96774 | 1.2682 | 1.0229 | 217.5246 | 45.4906 | 63745.4282 | 64748.991 | 1290.9798 | 1190.9798 | 894.771
2648 | 104.1181 | 0.99658 | 0.99099 | 0.93889 | 191.4438 | 47.1404 | 48943.5148 | 49295.7388 | 981.9148 | 881.9148 | 894.7926
2649 | 100.7833 | 0.97295 | 0.86347 | 1.0512 | 207.2233 | 47.578 | 48057.1602 | 49253.5226 | 981.0705 | 881.0705 | 894.9559
2650 | 98.843 | 1.0236 | 0.8619 | 1.0808 | 212.5714 | 52.783 | 54723.0985 | 55923.3626 | 1114.4673 | 1014.4673 | 894.979
2651 | 99.4568 | 0.97866 | 0.95648 | 1.037 | 190.702 | 52.0556 | 56891.6771 | 56869.3909 | 1133.3878 | 1033.3878 | 895.2239
2652 | 89.6459 | 0.99531 | 0.87822 | 1.0347 | 190.2758 | 45.3017 | 45397.4436 | 43983.9535 | 875.6791 | 775.6791 | 895.2495
2653 | 105.9417 | 0.97373 | 0.93248 | 0.93885 | 203.4462 | 52.9178 | 51989.6118 | 52480.9092 | 1045.6182 | 945.6182 | 895.3498
2654 | 101.5287 | 1.044 | 0.91073 | 0.97488 | 187.5748 | 53.2958 | 53155.4028 | 53215.3746 | 1060.3075 | 960.3075 | 895.4013
2655 | 108.2784 | 1.0469 | 1.0303 | 1.0581 | 211.5893 | 52.1188 | 62940.6721 | 60645.4644 | 1208.9093 | 1108.9093 | 895.4068
2656 | 111.0754 | 0.94837 | 1.2348 | 0.96111 | 196.4698 | 51.5115 | 66754.9453 | 63599.1833 | 1267.9837 | 1167.9837 | 895.4106
2657 | 101.0577 | 0.97195 | 0.78922 | 0.90553 | 204.0973 | 49.8041 | 40689.3659 | 40931.6705 | 814.6334 | 714.6334 | 895.4161
2658 | 103.5867 | 0.94728 | 1.0332 | 1.0292 | 207.0693 | 51.3138 | 59806.8297 | 58742.684 | 1170.8537 | 1070.8537 | 895.418
2659 | 87.5649 | 1.0177 | 1.0052 | 0.93592 | 219.2067 | 53.4519 | 55270.0884 | 56761.1598 | 1131.2232 | 1031.2232 | 895.4471
2660 | 90.7476 | 0.97425 | 0.86123 | 0.92602 | 204.5658 | 49.1278 | 43727.907 | 44194.9696 | 879.8994 | 779.8994 | 895.4494
2661 | 89.9913 | 1.0171 | 1.0593 | 1.043 | 194.2545 | 52.0217 | 62433.8359 | 63347.0629 | 1262.9413 | 1162.9413 | 895.4501
2662 | 104.1315 | 1.0066 | 1.1679 | 1.0038 | 194.303 | 55.7466 | 71390.9144 | 73206.2959 | 1460.1259 | 1360.1259 | 895.5459
2663 | 99.9149 | 0.97802 | 0.76957 | 1.0191 | 210.6126 | 46.9589 | 41628.552 | 41547.4176 | 826.9484 | 726.9484 | 895.6203
2664 | 116.9349 | 1.0212 | 0.89431 | 1.0439 | 201.3086 | 50.1315 | 52990.1651 | 52848.5105 | 1052.9702 | 952.9702 | 895.7312
2665 | 101.365 | 0.96347 | 0.90568 | 0.9817 | 200.9516 | 52.0289 | 51541.2964 | 51552.8429 | 1027.0569 | 927.0569 | 895.7888
2666 | 76.1456 | 0.93071 | 1.1186 | 0.93585 | 201.1318 | 49.2051 | 55197.3889 | 54389.0844 | 1083.7817 | 983.7817 | 895.8252
2667 | 112.7346 | 1.0419 | 1.0162 | 0.96703 | 211.6379 | 48.4327 | 53495.3536 | 53700.1347 | 1070.0027 | 970.0027 | 895.9095
2668 | 94.3036 | 0.99903 | 1.0677 | 1.0115 | 200.7243 | 51.3173 | 60453.4316 | 58582.8521 | 1167.657 | 1067.657 | 895.9392
2669 | 108.9371 | 0.95992 | 0.91632 | 0.92416 | 202.1529 | 48.0681 | 45934.1439 | 44876.0515 | 893.521 | 793.521 | 895.9392
2670 | 89.4168 | 1.0079 | 0.84478 | 1.0158 | 196.4932 | 50.0305 | 47638.4861 | 45815.1854 | 912.3037 | 812.3037 | 895.9754
2671 | 96.283 | 1.0429 | 0.83846 | 0.95935 | 198.2241 | 45.382 | 41259.4228 | 40009.8768 | 796.1975 | 696.1975 | 895.993
2672 | 93.0079 | 0.97138 | 1.0915 | 1.0448 | 206.4219 | 51.498 | 63585.8301 | 62630.2814 | 1248.6056 | 1148.6056 | 896.0417
2673 | 99.2134 | 1.0123 | 1.1435 | 0.93097 | 193.5971 | 48.5909 | 56802.4347 | 56669.3297 | 1129.3866 | 1029.3866 | 896.0633
2674 | 88.5433 | 1.0509 | 1.1162 | 0.95333 | 199.9481 | 49.9149 | 57961.1868 | 56988.2198 | 1135.7644 | 1035.7644 | 896.177
2675 | 106.2823 | 1.0208 | 1.1149 | 0.99686 | 204.0743 | 49.1119 | 60114.9387 | 58565.3349 | 1167.3067 | 1067.3067 | 896.2973
2676 | 107.1274 | 1.0256 | 1.0668 | 0.92779 | 191.7789 | 51.9095 | 57273.547 | 56300.9485 | 1122.019 | 1022.019 | 896.3707
2677 | 106.7012 | 0.96352 | 1.2094 | 0.9329 | 228.4477 | 52.3537 | 64680.8197 | 64126.644 | 1278.5329 | 1178.5329 | 896.4106
2678 | 106.4586 | 0.98341 | 0.6612 | 1.0064 | 200.7557 | 50.4865 | 39082.1487 | 39296.1707 | 781.9234 | 681.9234 | 896.4218
2679 | 94.2561 | 1.0562 | 0.97412 | 1.0732 | 191.228 | 51.8011 | 59501.6375 | 58063.6088 | 1157.2722 | 1057.2722 | 896.4467
2680 | 106.9314 | 0.91884 | 0.98272 | 1.0158 | 196.4574 | 47.2515 | 52005.8448 | 52595.7434 | 1047.9149 | 947.9149 | 896.4744
2681 | 108.0542 | 0.98498 | 1.0383 | 0.99962 | 194.266 | 50.1518 | 57584.9148 | 58137.6823 | 1158.7536 | 1058.7536 | 896.4996
2682 | 89.31 | 0.94256 | 1.0971 | 0.905 | 209.5412 | 48.8269 | 52797.054 | 53423.4206 | 1064.4684 | 964.4684 | 896.507
2683 | 103.1006 | 1.004 | 0.93246 | 0.94025 | 196.8614 | 47.0995 | 46366.4326 | 45940.017 | 914.8003 | 814.8003 | 896.528
2684 | 108.236 | 1.0468 | 1.2217 | 0.93386 | 208.4259 | 50.9115 | 64060.8725 | 63288.6481 | 1261.773 | 1161.773 | 896.5732
2685 | 104.805 | 0.99188 | 1.1443 | 0.99434 | 198.0907 | 47.8103 | 59567.3651 | 60555.8342 | 1207.1167 | 1107.1167 | 896.7661
2686 | 96.3864 | 1.0224 | 1.0922 | 1.0087 | 217.9231 | 48.7421 | 58718.5724 | 59089.2123 | 1177.7842 | 1077.7842 | 896.7901
2687 | 100.2384 | 0.87058 | 1.0785 | 0.99984 | 203.2697 | 46.4988 | 54402.5287 | 54284.508 | 1081.6902 | 981.6902 | 896.8835
2688 | 84.7687 | 1.0484 | 1.1244 | 1.0041 | 196.5279 | 55.9582 | 68349.4355 | 67553.8639 | 1347.0773 | 1247.0773 | 896.9215
2689 | 109.9183 | 0.95771 | 1.0902 | 0.9269 | 202.3668 | 47.7836 | 53519.4986 | 54693.6613 | 1089.8732 | 989.8732 | 896.9281
2690 | 91.8226 | 1.0292 | 0.79492 | 0.99431 | 188.1259 | 51.8257 | 46048.1609 | 45082.0085 | 897.6402 | 797.6402 | 896.9371
2691 | 112.7088 | 1.016 | 1.039 | 1.0408 | 215.696 | 53.7197 | 64457.1678 | 64930.8053 | 1294.6161 | 1194.6161 | 897.0482
2692 | 74.0871 | 1.0706 | 0.8228 | 1.0606 | 209.721 | 53.0482 | 50708.5493 | 51476.9212 | 1025.5384 | 925.5384 | 897.1291
2693 | 87.4463 | 1.0723 | 0.79366 | 1.003 | 216.6838 | 54.0382 | 48302.5239 | 47401.0978 | 944.022 | 844.022 | 897.1356
2694 | 111.7054 | 1.0026 | 1.0981 | 1.0019 | 205.5807 | 47.3615 | 57615.219 | 58108.0692 | 1158.1614 | 1058.1614 | 897.2336
2695 | 95.7382 | 1.0931 | 1.014 | 0.90952 | 183.3152 | 52.8445 | 54448.5127 | 55575.5992 | 1107.512 | 1007.512 | 897.2362
2696 | 116.9876 | 1.0379 | 1.0726 | 1.0029 | 203.9713 | 48.8094 | 58635.1592 | 59484.4545 | 1185.6891 | 1085.6891 | 897.2898
2697 | 88.896 | 0.98067 | 0.94714 | 1.0261 | 195.2349 | 51.284 | 54507.324 | 54603.6634 | 1088.0733 | 988.0733 | 897.385
2698 | 99.5234 | 1.0798 | 0.99876 | 0.9997 | 197.2947 | 48.9262 | 54306.2621 | 54957.3738 | 1095.1475 | 995.1475 | 897.5186
2699 | 111.4717 | 1.0293 | 1.0495 | 1.0018 | 200.749 | 52.5551 | 61488.5009 | 62241.9964 | 1240.8399 | 1140.8399 | 897.6471
2700 | 99.2305 | 0.9576 | 1.1689 | 0.98402 | 199.4139 | 50.2461 | 62769.6903 | 62278.7144 | 1241.5743 | 1141.5743 | 897.789
2701 | 115.1241 | 0.9881 | 0.73887 | 0.97163 | 204.1536 | 50.1399 | 41903.5292 | 42830.6237 | 852.6125 | 752.6125 | 897.8139
2702 | 101.2841 | 0.94512 | 1.1782 | 0.99834 | 222.9995 | 50.2243 | 64104.6349 | 64011.7816 | 1276.2356 | 1176.2356 | 897.8859
2703 | 108.6352 | 1.0597 | 0.99777 | 1.0629 | 183.2877 | 53.9922 | 63659.2095 | 64692.4467 | 1289.8489 | 1189.8489 | 897.8952
2704 | 104.7324 | 1.0599 | 0.97703 | 1.0542 | 178.4629 | 49.3753 | 56514.1933 | 57005.2845 | 1136.1057 | 1036.1057 | 897.9331
2705 | 102.1756 | 0.93801 | 0.9448 | 1.0852 | 201.0979 | 49.4243 | 55613.4008 | 55423.4744 | 1104.4695 | 1004.4695 | 897.9916
2706 | 103.9071 | 0.93312 | 0.94616 | 1.0159 | 187.0062 | 49.6809 | 52758.8192 | 53916.3228 | 1074.3265 | 974.3265 | 898.032
2707 | 96.1046 | 0.97318 | 0.88597 | 0.97913 | 207.9161 | 46.6539 | 45042.6455 | 44373.4179 | 883.4684 | 783.4684 | 898.0644
2708 | 97.1926 | 0.94493 | 1.0218 | 0.94311 | 208.6922 | 52.0108 | 55108.544 | 54788.2379 | 1091.7648 | 991.7648 | 898.1525
2709 | 105.3643 | 0.95915 | 1.0352 | 1.026 | 203.344 | 50.6277 | 59093.0359 | 58190.2393 | 1159.8048 | 1059.8048 | 898.1899
2710 | 114.6525 | 0.98686 | 0.8769 | 0.97861 | 201.3807 | 49.4388 | 48220.7472 | 48175.1057 | 959.5021 | 859.5021 | 898.2469
2711 | 101.4998 | 0.96697 | 1.0707 | 1.0788 | 208.1872 | 49.4664 | 62200.0927 | 61551.3251 | 1227.0265 | 1127.0265 | 898.2606
2712 | 100.7619 | 1.0069 | 0.69674 | 0.98195 | 201.0471 | 50.4801 | 39859.1428 | 40898.5376 | 813.9708 | 713.9708 | 898.3452
2713 | 96.3197 | 0.99382 | 1.0497 | 1.0336 | 183.7848 | 53.0734 | 62849.4799 | 61850.0828 | 1233.0017 | 1133.0017 | 898.3668
2714 | 94.5715 | 0.94965 | 0.97699 | 0.97309 | 203.5336 | 52.9275 | 55275.1784 | 56871.8776 | 1133.4376 | 1033.4376 | 898.3735
2715 | 87.6017 | 1.0248 | 0.97243 | 1.0457 | 200.8119 | 48.7746 | 54177.2727 | 56257.1044 | 1121.1421 | 1021.1421 | 898.4778
2716 | 104.8401 | 0.93472 | 0.97923 | 0.9683 | 187.7254 | 50.5108 | 53031.2273 | 54322.9574 | 1082.4591 | 982.4591 | 898.5132
2717 | 89.1531 | 0.99167 | 1.0955 | 0.97198 | 189.6189 | 47.2675 | 54699.1327 | 54945.5594 | 1094.9112 | 994.9112 | 898.5512
2718 | 111.421 | 1.0212 | 0.95995 | 1.0232 | 203.8712 | 50.2956 | 55328.3974 | 55842.8492 | 1112.857 | 1012.857 | 898.5841
2719 | 107.7032 | 1.0549 | 0.61073 | 0.97492 | 193.7834 | 48.4369 | 34536.6222 | 34264.9134 | 681.2983 | 581.2983 | 898.65
2720 | 106.437 | 1.0008 | 1.1611 | 1.0336 | 207.5441 | 46.1327 | 60488.2938 | 59989.1368 | 1195.7827 | 1095.7827 | 898.8546
2721 | 104.0383 | 0.93027 | 1.0593 | 0.99538 | 203.4516 | 49.7763 | 57506.045 | 56953.033 | 1135.0607 | 1035.0607 | 898.8657
2722 | 98.1268 | 0.98937 | 1.0373 | 0.91233 | 198.945 | 51.8823 | 54335.1081 | 55454.2634 | 1105.0853 | 1005.0853 | 898.8677
2723 | 87.9853 | 1.0378 | 1.0563 | 0.9383 | 197.1548 | 48.5152 | 52713.9598 | 53494.0168 | 1065.8803 | 965.8803 | 898.8679
2724 | 95.5718 | 0.88197 | 0.96681 | 0.95944 | 203.6462 | 50.4568 | 51259.9381 | 49538.9437 | 986.7789 | 886.7789 | 898.877
2725 | 94.9904 | 0.93984 | 0.84268 | 0.96921 | 205.6267 | 50.23 | 45714.6466 | 44617.7912 | 888.3558 | 788.3558 | 899.0056
2726 | 86.1221 | 1.0252 | 1.095 | 0.97695 | 194.4964 | 48.8069 | 56715.0279 | 55885.8523 | 1113.717 | 1013.717 | 899.1155
2727 | 103.3845 | 1.004 | 0.86209 | 0.88477 | 201.5203 | 50.3561 | 43837.354 | 43066.3309 | 857.3266 | 757.3266 | 899.1547
2728 | 117.4679 | 1.0166 | 0.89653 | 0.95001 | 195.8483 | 51.2506 | 49967.0608 | 49848.0301 | 992.9606 | 892.9606 | 899.1913
2729 | 107.8916 | 1.0363 | 0.78413 | 0.96244 | 199.1497 | 48.1188 | 41893.2524 | 41514.0937 | 826.2819 | 726.2819 | 899.2031
2730 | 116.4547 | 1.0119 | 1.3449 | 1.0067 | 200.2681 | 49.7617 | 73438.735 | 72126.5237 | 1438.5305 | 1338.5305 | 899.2256
2731 | 111.7427 | 1.0102 | 1.1905 | 1.0237 | 181.5123 | 50.9552 | 68031.4678 | 69496.4716 | 1385.9294 | 1285.9294 | 899.2892
2732 | 88.2087 | 0.99279 | 1.0123 | 1.0145 | 214.8787 | 48.0735 | 53796.5357 | 55217.0411 | 1100.3408 | 1000.3408 | 899.3718
2733 | 110.3747 | 1.0572 | 1.0237 | 1.0316 | 187.7529 | 49.13 | 57805.9629 | 59572.429 | 1187.4486 | 1087.4486 | 899.4754
2734 | 122.797 | 1.0857 | 1.0883 | 1.0104 | 205.4215 | 48.4171 | 59904.1979 | 60263.8722 | 1201.2774 | 1101.2774 | 899.501
2735 | 105.2948 | 1.0495 | 1.243 | 0.95929 | 194.4565 | 54.0802 | 70656.7313 | 71176.1872 | 1419.5237 | 1319.5237 | 899.5046
2736 | 125.4195 | 0.8899 | 0.93652 | 0.94427 | 203.9767 | 48.7012 | 48707.448 | 49131.5307 | 978.6306 | 878.6306 | 899.5619
2737 | 90.7644 | 1.0013 | 0.64916 | 1.0226 | 214.9158 | 52.0629 | 39508.3902 | 38652.9946 | 769.0599 | 669.0599 | 899.7046
2738 | 120.3404 | 1.0014 | 0.5179 | 1.033 | 182.5244 | 51.6593 | 34045.393 | 32523.721 | 646.4744 | 546.4744 | 899.7324
2739 | 101.8084 | 1.0385 | 0.94937 | 0.84145 | 204.5366 | 51.3582 | 46661.6958 | 47020.2971 | 936.4059 | 836.4059 | 899.7745
2740 | 92.3883 | 1.0535 | 1.1615 | 0.96819 | 194.3852 | 47.0239 | 57651.3691 | 57379.5671 | 1143.5913 | 1043.5913 | 899.8226
2741 | 108.6294 | 0.98298 | 0.77724 | 0.95807 | 189.0573 | 52.4574 | 44852.6523 | 43967.0752 | 875.3415 | 775.3415 | 899.8902
2742 | 91.5392 | 1.0426 | 1.0546 | 1.0494 | 206.3982 | 51.1104 | 61651.7217 | 61835.9551 | 1232.7191 | 1132.7191 | 899.96
2743 | 96.4498 | 1.0486 | 1.0577 | 1.0393 | 196.7543 | 45.9285 | 55329.2647 | 53596.7227 | 1067.9345 | 967.9345 | 899.9931
2744 | 87.1497 | 0.98092 | 1.2071 | 1.0362 | 211.6025 | 47.8366 | 64138.9067 | 65631.1598 | 1308.6232 | 1208.6232 | 900.0575
2745 | 96.628 | 0.98073 | 0.81325 | 0.99784 | 206.7019 | 48.7349 | 44373.1613 | 45460.529 | 905.2106 | 805.2106 | 900.0594
2746 | 86.0615 | 0.97099 | 1.1892 | 1.015 | 198.2488 | 52.3719 | 67786.5767 | 67717.6532 | 1350.3531 | 1250.3531 | 900.1511
2747 | 104.1293 | 1.0346 | 1.316 | 1.0215 | 192.9311 | 49.1128 | 71502.7685 | 70136.7582 | 1398.7352 | 1298.7352 | 900.1641
2748 | 91.4168 | 0.95147 | 0.9107 | 1.024 | 187.7859 | 52.0216 | 53226.424 | 54149.4768 | 1078.9895 | 978.9895 | 900.1713
2749 | 94.9545 | 1.0016 | 1.1093 | 1.0469 | 218.033 | 50.6608 | 63867.0942 | 63920.5277 | 1274.4106 | 1174.4106 | 900.2087
2750 | 99.8769 | 0.9331 | 1.2214 | 0.93084 | 193.4201 | 51.5363 | 63588.9793 | 62267.4297 | 1241.3486 | 1141.3486 | 900.2551
2751 | 105.853 | 0.99462 | 1.1491 | 0.84162 | 198.2688 | 53.469 | 57536.1277 | 59081.5047 | 1177.6301 | 1077.6301 | 900.2814
2752 | 93.9634 | 0.94757 | 0.93419 | 0.9696 | 187.1776 | 50.7593 | 50683.7983 | 49339.9289 | 982.7986 | 882.7986 | 900.3255
2753 | 107.145 | 0.99248 | 1.1099 | 0.91958 | 193.8332 | 49.0155 | 55434.167 | 54649.3747 | 1088.9875 | 988.9875 | 900.3411
2754 | 105.5587 | 0.91779 | 0.8756 | 1.1214 | 193.0336 | 52.3211 | 56635.8462 | 55599.5747 | 1107.9915 | 1007.9915 | 900.4265
2755 | 92.3306 | 0.94173 | 1.2021 | 1.0498 | 207.8051 | 51.2648 | 69359.4669 | 69125.8658 | 1378.5173 | 1278.5173 | 900.557
2756 | 88.2217 | 1.0616 | 1.0458 | 0.96275 | 210.0389 | 48.0919 | 53135.6768 | 52163.653 | 1039.2731 | 939.2731 | 900.5816
2757 | 89.0509 | 0.92353 | 0.69226 | 0.98581 | 203.7714 | 47.5521 | 36565.7158 | 36152.4051 | 719.0481 | 619.0481 | 900.5875
2758 | 107.1807 | 1.0141 | 0.90481 | 0.95769 | 215.2428 | 50.3015 | 49270.6233 | 48180.9216 | 959.6184 | 859.6184 | 900.59
2759 | 93.4075 | 0.94144 | 1.0217 | 0.97906 | 199.4073 | 50.3837 | 55027.1894 | 56139.1652 | 1118.7833 | 1018.7833 | 900.6613
2760 | 90.7867 | 1.0049 | 0.68623 | 1.0103 | 185.2276 | 48.4791 | 38219.304 | 38253.2967 | 761.0659 | 661.0659 | 900.6765
2761 | 94.1005 | 1.0954 | 1.0122 | 1.0394 | 209.683 | 49.0737 | 56896.138 | 56751.3966 | 1131.0279 | 1031.0279 | 900.6903
2762 | 103.4323 | 1.0957 | 1.1543 | 1.0254 | 203.3523 | 53.262 | 69283.753 | 70426.172 | 1404.5234 | 1304.5234 | 900.6912
2763 | 89.3268 | 0.95601 | 1.0892 | 0.91878 | 193.9257 | 49.5154 | 53975.0783 | 53895.4671 | 1073.9093 | 973.9093 | 900.7083
2764 | 110.949 | 1.061 | 0.83605 | 0.96475 | 201.9203 | 45.0015 | 41796.4691 | 41790.3258 | 831.8065 | 731.8065 | 900.7643
2765 | 98.2631 | 0.99579 | 0.92834 | 1.0898 | 208.253 | 48.5217 | 54045.4641 | 53363.905 | 1063.2781 | 963.2781 | 900.7961
2766 | 90.8838 | 1.0332 | 0.72457 | 1.0028 | 172.607 | 54.0549 | 44524.7514 | 45728.1385 | 910.5628 | 810.5628 | 900.8921
2767 | 107.9105 | 0.91665 | 0.97328 | 1.01 | 193.4965 | 54.9815 | 59678.1953 | 60301.4989 | 1202.03 | 1102.03 | 900.8986
2768 | 90.1696 | 1.0176 | 1.0676 | 0.99026 | 203.0733 | 50.1006 | 57768.8473 | 57636.8262 | 1148.7365 | 1048.7365 | 901.0548
2769 | 91.0475 | 0.87031 | 0.95033 | 1.0404 | 202.9903 | 53.5334 | 57375.6326 | 57045.2278 | 1136.9046 | 1036.9046 | 901.1371
2770 | 97.3036 | 1.0022 | 0.73014 | 0.98438 | 204.3369 | 51.5576 | 42288.195 | 42807.822 | 852.1564 | 752.1564 | 901.1445
2771 | 114.4752 | 1.0491 | 1.0621 | 0.99207 | 207.8075 | 50.1435 | 59066.6961 | 57909.6585 | 1154.1932 | 1054.1932 | 901.1642
2772 | 78.6579 | 1.0317 | 1.2388 | 1.0119 | 203.9927 | 49.4037 | 66144.8073 | 66935.1558 | 1334.7031 | 1234.7031 | 901.4276
2773 | 83.6304 | 1.0058 | 1.0759 | 1.0485 | 182.2995 | 53.815 | 65414.8635 | 65703.6118 | 1310.0722 | 1210.0722 | 901.4987
2774 | 109.806 | 0.98231 | 0.9539 | 0.93762 | 205.8206 | 43.6362 | 43940.3289 | 43809.8178 | 872.1964 | 772.1964 | 901.5247
2775 | 98.8995 | 1.0148 | 0.87979 | 1.0984 | 178.5807 | 52.1914 | 55853.8636 | 53099.1733 | 1057.9835 | 957.9835 | 901.5825
2776 | 97.1287 | 1.005 | 1.0395 | 0.95581 | 187.2471 | 49.3323 | 54016.2019 | 53896.8362 | 1073.9367 | 973.9367 | 901.6234
2777 | 94.6485 | 1.0569 | 0.99086 | 0.96555 | 190.025 | 50.3531 | 53401.3375 | 52696.0455 | 1049.9209 | 949.9209 | 901.6259
2778 | 101.633 | 1.041 | 0.94272 | 0.9664 | 201.9955 | 47.833 | 48840.6266 | 49318.3256 | 982.3665 | 882.3665 | 901.6668
2779 | 122.0774 | 1.0515 | 0.9141 | 0.97117 | 209.3575 | 53.9207 | 54998.9753 | 57258.5684 | 1141.1714 | 1041.1714 | 901.725
2780 | 80.534 | 1.0407 | 1.1747 | 1.0222 | 198.8124 | 50.9878 | 65696.3944 | 66003.1679 | 1316.0634 | 1216.0634 | 901.7445
2781 | 107.5982 | 1.0235 | 0.99994 | 1.0068 | 206.5102 | 52.4345 | 58767.7942 | 59355.8776 | 1183.1176 | 1083.1176 | 901.776
2782 | 91.5041 | 0.98531 | 1.1838 | 1.0042 | 195.9007 | 48.4248 | 62125.9966 | 63043.5228 | 1256.8705 | 1156.8705 | 901.9102
2783 | 100.2423 | 0.8878 | 0.94596 | 0.95149 | 201.8048 | 51.0983 | 50741.4143 | 51677.4117 | 1029.5482 | 929.5482 | 901.9318
2784 | 102.9761 | 1.05 | 0.7639 | 0.9438 | 186.9811 | 52.5841 | 43784.1634 | 43914.9201 | 874.2984 | 774.2984 | 901.957
2785 | 107.9236 | 1.0149 | 1.1108 | 1.0278 | 193.1839 | 52.3967 | 65756.7588 | 65903.1199 | 1314.0624 | 1214.0624 | 902.0055
2786 | 92.044 | 0.94991 | 0.9278 | 1.0453 | 211.4549 | 53.254 | 56513.151 | 54959.5984 | 1095.192 | 995.192 | 902.0116
2787 | 103.0715 | 1.1008 | 0.96627 | 0.99986 | 190.9958 | 49.4685 | 53596.8011 | 54906.5743 | 1094.1315 | 994.1315 | 902.0712
2788 | 104.31 | 1.0046 | 1.0837 | 1.0312 | 201.5831 | 50.9076 | 62424.8004 | 61435.0856 | 1224.7017 | 1124.7017 | 902.0993
2789 | 105.0124 | 0.9627 | 0.99973 | 0.94211 | 190.6913 | 48.6097 | 50887.9828 | 51006.2247 | 1016.1245 | 916.1245 | 902.1217
2790 | 110.2617 | 1.0012 | 0.88317 | 1.0394 | 206.3114 | 49.3246 | 50928.2243 | 51266.9235 | 1021.3385 | 921.3385 | 902.1999
2791 | 111.3101 | 1.053 | 1.0546 | 1.0258 | 203.5902 | 50.9316 | 61274.1578 | 62604.3146 | 1248.0863 | 1148.0863 | 902.236
2792 | 97.3009 | 0.97664 | 1.0928 | 0.98331 | 199.543 | 48.4487 | 56863.5112 | 57163.3113 | 1139.2662 | 1039.2662 | 902.2972
2793 | 94.208 | 1.0313 | 1.1495 | 1.102 | 211.8634 | 48.3077 | 66102.6965 | 65430.7251 | 1304.6145 | 1204.6145 | 902.3145
2794 | 77.6786 | 1.0447 | 1.0063 | 0.9939 | 210.6689 | 49.321 | 53542.1937 | 52133.2275 | 1038.6645 | 938.6645 | 902.3797
2795 | 98.0812 | 1.1439 | 1.1153 | 0.94991 | 198.8708 | 52.9149 | 62194.8384 | 61718.0955 | 1230.3619 | 1130.3619 | 902.4357
2796 | 97.3305 | 0.99791 | 1.1005 | 1.0783 | 192.3042 | 47.4409 | 61094.1907 | 62129.8518 | 1238.597 | 1138.597 | 902.4718
2797 | 106.3148 | 0.95971 | 1.0512 | 0.89587 | 194.647 | 52.9714 | 55486.1979 | 55966.7368 | 1115.3347 | 1015.3347 | 902.5847
2798 | 83.7798 | 0.90647 | 1.1109 | 0.94609 | 196.2563 | 47.9711 | 54255.3978 | 54411.5981 | 1084.232 | 984.232 | 902.6571
2799 | 111.7637 | 1.0006 | 1.1517 | 1.0692 | 198.0258 | 48.5918 | 65466.0993 | 66028.7231 | 1316.5745 | 1216.5745 | 902.681
2800 | 125.8942 | 0.99574 | 1.0573 | 0.96854 | 196.8962 | 51.8867 | 59834.2849 | 59323.2546 | 1182.4651 | 1082.4651 | 902.7192
2801 | 112.2945 | 0.99691 | 1.0465 | 1.0826 | 191.2643 | 48.4669 | 60530.1875 | 62586.1733 | 1247.7235 | 1147.7235 | 902.7509
2802 | 104.1969 | 0.93657 | 0.99216 | 0.90702 | 206.8457 | 51.3271 | 51405.259 | 49614.6397 | 988.2928 | 888.2928 | 902.8422
2803 | 96.7503 | 1.0216 | 1.1609 | 1.0142 | 206.6618 | 50.1851 | 64254.3718 | 64115.6216 | 1278.3124 | 1178.3124 | 902.8764
2804 | 85.5011 | 0.9798 | 0.90141 | 0.99639 | 200.0843 | 49.4098 | 48717.0259 | 47495.4704 | 945.9094 | 845.9094 | 902.9027
2805 | 107.8505 | 1.0728 | 1.2078 | 1.0421 | 202.376 | 51.9957 | 71662.5539 | 69104.9671 | 1378.0993 | 1278.0993 | 903.0355
2806 | 115.9011 | 0.98884 | 0.87228 | 1.0324 | 186.3388 | 50.0733 | 51017.7901 | 52163.3413 | 1039.2668 | 939.2668 | 903.0771
2807 | 108.5074 | 1.0714 | 1.0909 | 0.94884 | 185.9019 | 50.1631 | 57942.5243 | 58942.4864 | 1174.8497 | 1074.8497 | 903.0881
2808 | 88.7126 | 0.9762 | 1.0483 | 1.1185 | 191.7201 | 52.2329 | 65955.8334 | 66473.5585 | 1325.4712 | 1225.4712 | 903.1008
2809 | 103.395 | 1.0537 | 1.188 | 0.98848 | 215.3246 | 49.0401 | 63148.5001 | 65298.6391 | 1301.9728 | 1201.9728 | 903.1739
2810 | 75.7194 | 1.0387 | 1.2247 | 0.93823 | 192.7399 | 54.9139 | 67610.0404 | 69067.2881 | 1377.3458 | 1277.3458 | 903.2252
2811 | 105.728 | 0.97214 | 0.84757 | 1.0747 | 192.7552 | 49.2982 | 50163.1527 | 50492.2453 | 1005.8449 | 905.8449 | 903.2764
2812 | 100.2461 | 0.99749 | 0.82299 | 1.0122 | 218.0378 | 51.0103 | 47811.6637 | 47144.4771 | 938.8895 | 838.8895 | 903.3981
2813 | 85.1464 | 0.95131 | 0.58619 | 0.91857 | 216.1176 | 46.8906 | 29262.6454 | 28570.4454 | 567.4089 | 467.4089 | 903.4195
2814 | 101.9207 | 0.98257 | 1.11 | 1.0167 | 204.2268 | 52.7026 | 64957.8031 | 65162.6908 | 1299.2538 | 1199.2538 | 903.4467
2815 | 101.4379 | 0.99671 | 0.89901 | 1.0128 | 198.6926 | 52.4504 | 53259.6966 | 54994.3435 | 1095.8869 | 995.8869 | 903.4487
2816 | 94.9267 | 1.0663 | 0.77995 | 0.94843 | 206.3151 | 45.0458 | 38087.518 | 37671.5821 | 749.4316 | 649.4316 | 903.4536
2817 | 99.8977 | 0.92313 | 1.1127 | 1.0136 | 185.7454 | 46.3207 | 56700.4383 | 57476.3633 | 1145.5273 | 1045.5273 | 903.4805
2818 | 100.5471 | 1.0446 | 0.99897 | 1.037 | 208.4266 | 53.5591 | 61317.4885 | 62783.3928 | 1251.6679 | 1151.6679 | 903.5039
2819 | 89.2689 | 1.0228 | 1.02 | 1.0444 | 205.1913 | 49.8413 | 57851.8221 | 55975.4212 | 1115.5084 | 1015.5084 | 903.5315
2820 | 91.7751 | 0.91964 | 1.3009 | 0.87702 | 194.5329 | 49.8359 | 61259.587 | 62189.7509 | 1239.795 | 1139.795 | 903.5587
2821 | 109.5696 | 1.02 | 1.138 | 1.0813 | 186.2204 | 58.7126 | 78992.105 | 80122.142 | 1598.4428 | 1498.4428 | 903.5876
2822 | 91.2635 | 0.91088 | 1.181 | 1.0151 | 202.0417 | 45.007 | 57901.1867 | 56767.8409 | 1131.3568 | 1031.3568 | 903.743
2823 | 103.5677 | 1.0721 | 0.90518 | 0.98836 | 192.8155 | 50.377 | 50855.9802 | 51991.23 | 1035.8246 | 935.8246 | 903.813
2824 | 88.9971 | 0.98844 | 1.1498 | 1.0163 | 220.8557 | 48.7205 | 61442.9187 | 62981.3105 | 1255.6262 | 1155.6262 | 903.8447
2825 | 94.2276 | 0.9842 | 0.93229 | 1.0177 | 210.5392 | 50.184 | 52478.2608 | 53667.1817 | 1069.3436 | 969.3436 | 903.933
2826 | 80.3843 | 1.0117 | 0.92339 | 0.96445 | 200.2523 | 45.6007 | 44518.668 | 45608.2534 | 908.1651 | 808.1651 | 903.9603
2827 | 107.4066 | 0.99794 | 1.0693 | 1.0241 | 204.5806 | 49.1201 | 59256.4352 | 58058.4989 | 1157.17 | 1057.17 | 903.988
2828 | 90.0939 | 1.0264 | 0.94068 | 0.99083 | 205.6531 | 50.7493 | 52199.8807 | 51977.3752 | 1035.5475 | 935.5475 | 904.1996
2829 | 85.5488 | 1.0197 | 0.96259 | 1.0902 | 209.8577 | 51.8334 | 59128.049 | 60099.705 | 1197.9941 | 1097.9941 | 904.2028
2830 | 109.8788 | 0.97143 | 0.68142 | 0.9281 | 187.224 | 48.6156 | 36122.147 | 36184.2848 | 719.6857 | 619.6857 | 904.266
2831 | 112.1202 | 0.94553 | 1.1349 | 1.0223 | 210.3763 | 48.2172 | 61262.4449 | 60740.8838 | 1210.8177 | 1110.8177 | 904.2664
2832 | 87.2299 | 0.9777 | 1.196 | 0.9858 | 204.0121 | 47.9787 | 60863.4903 | 59217.6164 | 1180.3523 | 1080.3523 | 904.2999
2833 | 103.4513 | 0.98913 | 0.95392 | 0.9516 | 179.1783 | 50.3673 | 51053.9356 | 48873.2895 | 973.4658 | 873.4658 | 904.3329
2834 | 103.3533 | 1.0081 | 1.097 | 1.0538 | 186.6169 | 51.3185 | 64860.11 | 65189.8082 | 1299.7962 | 1199.7962 | 904.3331
2835 | 89.87 | 1.0772 | 1.0678 | 0.94839 | 199.3202 | 54.7874 | 60984.9789 | 62053.0285 | 1237.0606 | 1137.0606 | 904.34
2836 | 79.6015 | 1.0564 | 1.2736 | 1.0142 | 193.3098 | 50.3974 | 69531.6621 | 69617.6269 | 1388.3525 | 1288.3525 | 904.3564
2837 | 97.225 | 0.9503 | 0.91504 | 0.94868 | 205.4703 | 53.6438 | 51728.8148 | 52461.609 | 1045.2322 | 945.2322 | 904.3887
2838 | 119.1697 | 1.0045 | 1.008 | 1.0301 | 199.1179 | 48.4916 | 56356.3387 | 58135.6463 | 1158.7129 | 1058.7129 | 904.5416
2839 | 88.995 | 0.97449 | 0.90338 | 0.97252 | 207.3305 | 48.8958 | 47405.8409 | 48452.0408 | 965.0408 | 865.0408 | 904.5524
2840 | 105.6631 | 1.1433 | 1.228 | 0.99387 | 198.7462 | 49.7261 | 66896.5004 | 68046.0152 | 1356.9203 | 1256.9203 | 904.5647
2841 | 98.0547 | 0.99707 | 0.76201 | 0.96805 | 205.724 | 49.8813 | 41878.2405 | 42674.6961 | 849.4939 | 749.4939 | 904.6098
2842 | 90.7048 | 0.90224 | 1.2467 | 0.94992 | 199.0921 | 49.2099 | 62501.5238 | 63417.7732 | 1264.3555 | 1164.3555 | 904.6761
2843 | 91.8467 | 0.9405 | 1.239 | 1.0349 | 211.9677 | 50.9264 | 69914.1729 | 70896.3222 | 1413.9264 | 1313.9264 | 904.6904
2844 | 104.1428 | 1.0189 | 0.94179 | 1.0289 | 220.4143 | 48.0949 | 51928.8784 | 52347.8311 | 1042.9566 | 942.9566 | 904.7058
2845 | 111.7043 | 0.96987 | 1.0823 | 1.0467 | 206.0346 | 48.352 | 60223.9897 | 58917.593 | 1174.3519 | 1074.3519 | 904.7091
2846 | 92.2932 | 1.0224 | 0.77607 | 0.92413 | 220.737 | 46.8632 | 38252.8236 | 37768.0104 | 751.3602 | 651.3602 | 904.7658
2847 | 112.5784 | 0.96967 | 1.1131 | 1.0017 | 197.1361 | 51.1164 | 62772.4426 | 64080.4715 | 1277.6094 | 1177.6094 | 904.9404
2848 | 92.163 | 0.87671 | 0.91992 | 1.0639 | 190.4726 | 48.4047 | 51477.2693 | 52876.504 | 1053.5301 | 953.5301 | 904.947
2849 | 90.8402 | 1.01 | 0.97121 | 0.96815 | 205.8272 | 49.3059 | 51090.3589 | 52210.4492 | 1040.209 | 940.209 | 904.9605
2850 | 99.5026 | 0.96967 | 1.0744 | 1.0007 | 182.5634 | 51.1058 | 60059.0348 | 59679.0693 | 1189.5814 | 1089.5814 | 904.9898
2851 | 109.596 | 0.97103 | 1.2185 | 0.93553 | 196.7179 | 47.9357 | 59942.0411 | 59486.307 | 1185.7261 | 1085.7261 | 905.0056
2852 | 124.1933 | 0.97148 | 1.0474 | 1.0429 | 203.3549 | 48.7013 | 59273.4843 | 58951.3985 | 1175.028 | 1075.028 | 905.0319
2853 | 87.1719 | 0.95697 | 0.95726 | 1.0816 | 201.1829 | 52.1797 | 58578.4631 | 58778.1838 | 1171.5637 | 1071.5637 | 905.0487
2854 | 117.1885 | 0.95234 | 1.042 | 1.0244 | 218.7586 | 52.3667 | 61961.3202 | 62994.4156 | 1255.8883 | 1155.8883 | 905.0644
2855 | 105.145 | 1.0022 | 1.0993 | 1.0704 | 206.1735 | 45.9126 | 59067.8199 | 59834.9997 | 1192.7 | 1092.7 | 905.1215
2856 | 104.7438 | 1.1226 | 1.1329 | 0.94364 | 210.5048 | 50.0739 | 59629.8693 | 58521.7219 | 1166.4344 | 1066.4344 | 905.1224
2857 | 91.8223 | 0.95125 | 0.93006 | 0.99828 | 203.0388 | 47.9465 | 48907.4546 | 49268.2461 | 981.3649 | 881.3649 | 905.1533
2858 | 105.5222 | 0.9963 | 1.1162 | 0.94773 | 206.7771 | 49.8886 | 58224.6638 | 57299.5478 | 1141.991 | 1041.991 | 905.2484
2859 | 79.0036 | 1.0438 | 0.94569 | 1.0611 | 187.069 | 47.525 | 51794.3215 | 53711.0749 | 1070.2215 | 970.2215 | 905.2761
2860 | 116.2287 | 1.0502 | 1.0408 | 0.97477 | 210.9333 | 48.414 | 55236.3483 | 56824.2244 | 1132.4845 | 1032.4845 | 905.3144
2861 | 91.0094 | 1.0689 | 0.77331 | 0.90418 | 199.0299 | 51.5329 | 41244.9533 | 41114.8434 | 818.2969 | 718.2969 | 905.4092
2862 | 106.4642 | 1.0541 | 1.4198 | 1.0158 | 194.4779 | 47.4655 | 73976.3172 | 73225.6835 | 1460.5137 | 1360.5137 | 905.4146
2863 | 107.0586 | 0.9396 | 1.0356 | 1.0149 | 186.6875 | 50.7256 | 58600.6418 | 58517.074 | 1166.3415 | 1066.3415 | 905.419
2864 | 93.4209 | 1.0282 | 0.78234 | 1.0196 | 200.317 | 47.8745 | 42988.1722 | 43370.1795 | 863.4036 | 763.4036 | 905.5301
2865 | 104.1902 | 0.98224 | 1.0483 | 1.0236 | 204.6415 | 47.5599 | 56106.6966 | 54951.8762 | 1095.0375 | 995.0375 | 905.5863
2866 | 101.1853 | 0.96258 | 1.1612 | 0.94891 | 222.2049 | 47.9813 | 57766.8282 | 57247.6785 | 1140.9536 | 1040.9536 | 905.626
2867 | 93.0372 | 1.0423 | 1.0699 | 1.1216 | 201.4202 | 52.1379 | 67819.3082 | 67807.2713 | 1352.1454 | 1252.1454 | 905.7207
2868 | 104.113 | 1.028 | 0.75324 | 0.96605 | 200.2093 | 50.1765 | 42082.0385 | 41136.875 | 818.7375 | 718.7375 | 905.7696
2869 | 98.7326 | 0.9115 | 1.058 | 0.98608 | 188.6429 | 47.4562 | 53969.0593 | 53838.8235 | 1072.7765 | 972.7765 | 905.8449
2870 | 117.3888 | 0.94539 | 0.99643 | 0.96429 | 220.2628 | 48.4813 | 52184.1137 | 51828.2307 | 1032.5646 | 932.5646 | 905.9798
2871 | 104.0502 | 0.93148 | 0.90075 | 0.96229 | 190.3498 | 50.7853 | 49132.2269 | 48513.2955 | 966.2659 | 866.2659 | 905.9907
2872 | 99.6997 | 0.96605 | 0.84963 | 0.94963 | 193.2691 | 47.4622 | 43058.678 | 44336.3108 | 882.7262 | 782.7262 | 906.0121
2873 | 91.524 | 1.0894 | 0.93731 | 0.99022 | 172.9847 | 47.6631 | 49163.6465 | 49920.0317 | 994.4006 | 894.4006 | 906.0193
2874 | 100.8025 | 0.97717 | 1.078 | 1.0452 | 216.2569 | 48.1841 | 59250.5672 | 59217.8409 | 1180.3568 | 1080.3568 | 906.117
2875 | 91.3464 | 1.0774 | 1.0817 | 1.0424 | 206.689 | 51.8657 | 63796.5353 | 64743.8218 | 1290.8764 | 1190.8764 | 906.15
2876 | 87.6453 | 1.001 | 0.88681 | 0.97905 | 194.7721 | 48.5775 | 46632.9393 | 46693.6863 | 929.8737 | 829.8737 | 906.1792
2877 | 110.2788 | 0.98262 | 1.3082 | 1.0565 | 193.163 | 50.8819 | 76033.2241 | 72276.6363 | 1441.5327 | 1341.5327 | 906.2113
2878 | 101.0771 | 1.0432 | 1.0347 | 0.98998 | 199.2646 | 52.8304 | 59886.0613 | 59963.6679 | 1195.2734 | 1095.2734 | 906.2699
2879 | 109.0906 | 0.94277 | 0.93614 | 0.97312 | 177.8424 | 51.1176 | 52002.0785 | 52584.1281 | 1047.6826 | 947.6826 | 906.3682
2880 | 101.3333 | 0.96318 | 0.8667 | 0.99292 | 206.868 | 51.7109 | 49754.8824 | 50258.2092 | 1001.1642 | 901.1642 | 906.4694
2881 | 90.7929 | 0.86264 | 0.77127 | 1.0187 | 195.9035 | 46.8147 | 40643.3913 | 38996.7687 | 775.9354 | 675.9354 | 906.486
2882 | 87.0671 | 0.97532 | 1.0397 | 1.0069 | 215.6692 | 50.9695 | 57902.7406 | 57467.5696 | 1145.3514 | 1045.3514 | 906.4928
2883 | 95.539 | 1.0258 | 1.0151 | 0.97181 | 208.2442 | 50.3837 | 54846.6894 | 55706.4928 | 1110.1299 | 1010.1299 | 906.5897
2884 | 105.1853 | 1.048 | 0.76427 | 1.0691 | 198.9754 | 50.4291 | 46963.2258 | 46477.4043 | 925.5481 | 825.5481 | 906.6239
2885 | 98.3285 | 1.0266 | 0.88361 | 0.95351 | 207.4824 | 52.1755 | 49433.7002 | 49401.9636 | 984.0393 | 884.0393 | 906.6357
2886 | 81.9476 | 0.93881 | 0.99865 | 0.96647 | 195.283 | 53.4122 | 55855.8426 | 55776.9806 | 1111.5396 | 1011.5396 | 906.8012
2887 | 108.7634 | 1.0046 | 1.0155 | 0.96267 | 211.8446 | 48.6712 | 53112.0459 | 51325.9018 | 1022.518 | 922.518 | 906.8397
2888 | 103.3546 | 0.97006 | 0.99776 | 0.94888 | 193.4208 | 51.4806 | 54094.2936 | 54751.0694 | 1091.0214 | 991.0214 | 906.8652
2889 | 99.9642 | 1.0076 | 0.68939 | 0.95049 | 203.3969 | 51.3749 | 39042.0651 | 39492.2154 | 785.8443 | 685.8443 | 906.922
2890 | 107.2873 | 0.96643 | 0.98881 | 1.0336 | 205.6962 | 45.1462 | 51029.6471 | 49772.8263 | 991.4565 | 891.4565 | 906.9277
2891 | 98.513 | 0.99091 | 0.65051 | 0.98272 | 198.6888 | 52.9937 | 39248.9479 | 39637.7686 | 788.7554 | 688.7554 | 907.1109
2892 | 103.3293 | 1.0506 | 0.92147 | 1.0675 | 184.3985 | 48.7653 | 53447.6109 | 53163.7304 | 1059.2746 | 959.2746 | 907.172
2893 | 95.3417 | 1.0917 | 0.97946 | 0.92884 | 217.3944 | 48.4257 | 49313.1468 | 48797.7013 | 971.954 | 871.954 | 907.2619
2894 | 96.3661 | 0.9811 | 0.92898 | 0.94321 | 196.5973 | 53.4717 | 52105.2174 | 52415.4045 | 1044.3081 | 944.3081 | 907.2749
2895 | 110.2307 | 1.0473 | 1.0677 | 0.98797 | 220.2249 | 47.058 | 55291.7352 | 55546.5174 | 1106.9303 | 1006.9303 | 907.2977
2896 | 88.6555 | 0.98486 | 1.0333 | 0.98373 | 208.64 | 53.3069 | 59048.5951 | 59442.5201 | 1184.8504 | 1084.8504 | 907.3493
2897 | 109.2476 | 1.0697 | 1.0966 | 0.91273 | 205.8166 | 52.1621 | 58510.3971 | 59299.4846 | 1181.9897 | 1081.9897 | 907.3497
2898 | 114.1203 | 1.0871 | 0.74027 | 1.0597 | 185.2808 | 50.1152 | 45714.3285 | 46942.7912 | 934.8558 | 834.8558 | 907.3682
2899 | 77.8416 | 0.96132 | 1.0355 | 1.0234 | 183.7936 | 50.9769 | 58021.9825 | 60042.9418 | 1196.8588 | 1096.8588 | 907.3988
2900 | 91.5051 | 1.0409 | 0.84033 | 1.034 | 208.8708 | 49.8748 | 48296.0256 | 47715.876 | 950.3175 | 850.3175 | 907.4303
2901 | 96.9787 | 1.0072 | 0.82372 | 1.0341 | 200.6724 | 50.0267 | 47699.5049 | 46481.2495 | 925.625 | 825.625 | 907.4625
2902 | 96.7084 | 0.99657 | 0.97396 | 0.92925 | 194.3074 | 49.4295 | 49694.9106 | 48656.5682 | 969.1314 | 869.1314 | 907.464
2903 | 106.0101 | 0.98011 | 1.0047 | 0.87762 | 189.0912 | 48.9533 | 48439.3618 | 48580.9723 | 967.6194 | 867.6194 | 907.475
2904 | 104.9625 | 1.0233 | 1.0809 | 0.96828 | 198.6417 | 52.8621 | 61203.1704 | 61039.684 | 1216.7937 | 1116.7937 | 907.502
2905 | 91.6825 | 1.0414 | 1.1984 | 1.036 | 211.7561 | 51.8344 | 69513.7584 | 65888.0753 | 1313.7615 | 1213.7615 | 907.5248
2906 | 95.3695 | 0.95071 | 1.036 | 0.98421 | 203.657 | 45.9042 | 51171.6889 | 49507.7072 | 986.1541 | 886.1541 | 907.6647
2907 | 103.688 | 0.98333 | 0.91687 | 0.96942 | 200.017 | 48.2367 | 47992.2418 | 48826.1028 | 972.5221 | 872.5221 | 907.6663
2908 | 100.1971 | 1.0466 | 0.90321 | 0.96988 | 207.6089 | 43.8713 | 43239.4378 | 44184.8526 | 879.6971 | 779.6971 | 907.7068
2909 | 110.8846 | 0.96779 | 1.0297 | 1.1043 | 225.4426 | 51.0414 | 63742.1 | 62067.6382 | 1237.3528 | 1137.3528 | 907.7403
2910 | 93.7219 | 1.0651 | 1.1341 | 1.0039 | 186.7376 | 48.1216 | 59777.0492 | 60217.6943 | 1200.3539 | 1100.3539 | 907.7411
2911 | 90.927 | 0.98153 | 0.93612 | 1.0318 | 181.8489 | 46.5278 | 49275.4469 | 48477.5815 | 965.5516 | 865.5516 | 907.7514
2912 | 101.5862 | 0.99321 | 1.0366 | 0.96302 | 216.7918 | 50.5047 | 55730.9184 | 56567.8774 | 1127.3575 | 1027.3575 | 907.9375
2913 | 108.5897 | 1.0768 | 1.1046 | 1.0121 | 204.6357 | 50.2152 | 62218.0203 | 61951.3345 | 1235.0267 | 1135.0267 | 907.992
2914 | 103.6986 | 0.99771 | 1.0951 | 0.95744 | 195.8867 | 48.8414 | 56460.1714 | 55354.756 | 1103.0951 | 1003.0951 | 907.999
2915 | 98.2246 | 0.93188 | 1.0787 | 1.0394 | 207.8705 | 51.8622 | 63103.5801 | 61624.7668 | 1228.4953 | 1128.4953 | 908.1502
2916 | 102.2189 | 1.0953 | 1.0284 | 0.95584 | 216.0444 | 46.3433 | 50957.5002 | 49223.4222 | 980.4684 | 880.4684 | 908.1606
2917 | 101.1112 | 0.96275 | 1.0388 | 1.1291 | 209.5018 | 48.6522 | 62009.9701 | 62537.758 | 1246.7552 | 1146.7552 | 908.1607
2918 | 94.379 | 0.99863 | 0.88057 | 0.98957 | 186.476 | 50.303 | 48761.1565 | 49018.8561 | 976.3771 | 876.3771 | 908.1645
2919 | 105.1709 | 0.97966 | 1.1047 | 1.0068 | 204.304 | 50.4049 | 61459.622 | 62775.3799 | 1251.5076 | 1151.5076 | 908.1745
2920 | 93.1317 | 0.97938 | 1.1031 | 0.97542 | 213.5284 | 51.4276 | 60239.6409 | 61049.1116 | 1216.9822 | 1116.9822 | 908.2383
2921 | 97.8366 | 1.091 | 0.86479 | 0.94619 | 197.2281 | 52.4522 | 48715.4346 | 47097.4511 | 937.949 | 837.949 | 908.2772
2922 | 87.1821 | 1.0388 | 0.8948 | 1.0218 | 205.1051 | 49.787 | 50236.2198 | 50813.0304 | 1012.2606 | 912.2606 | 908.3049
2923 | 89.8167 | 0.9433 | 1.1613 | 1.0259 | 190.8003 | 47.1328 | 60335.4078 | 60507.9901 | 1206.1598 | 1106.1598 | 908.3094
2924 | 96.6454 | 1.0872 | 0.89839 | 1.0228 | 193.2597 | 51.3902 | 52814.8559 | 54118.712 | 1078.3742 | 978.3742 | 908.3779
2925 | 103.9249 | 1.0299 | 0.88629 | 0.98662 | 202.279 | 49.2408 | 48530.2691 | 46808.6286 | 932.1726 | 832.1726 | 908.4131
2926 | 109.4217 | 0.97497 | 0.89566 | 1.0443 | 215.188 | 50.8108 | 53161.5801 | 52255.3214 | 1041.1064 | 941.1064 | 908.4811
2927 | 91.2508 | 0.96061 | 1.0956 | 0.97432 | 202.0748 | 49.42 | 57290.2314 | 56830.2875 | 1132.6057 | 1032.6057 | 908.5638
2928 | 104.9891 | 0.93353 | 1.2495 | 0.93085 | 194.9036 | 48.8004 | 61738.5889 | 61395.5116 | 1223.9102 | 1123.9102 | 908.5839
2929 | 121.7529 | 1.036 | 1.0744 | 1.1148 | 210.4951 | 49.494 | 65737.8011 | 63992.172 | 1275.8434 | 1175.8434 | 908.6212
2930 | 90.4634 | 1.0563 | 1.1873 | 1.0036 | 200.9975 | 50.9659 | 65802.6867 | 65071.5606 | 1297.4312 | 1197.4312 | 908.639
2931 | 88.6139 | 1.0111 | 0.88888 | 1.1127 | 196.3857 | 50.0814 | 54217.5655 | 54296.0784 | 1081.9216 | 981.9216 | 908.676
2932 | 101.9154 | 0.83371 | 1.0659 | 0.9429 | 206.1165 | 45.8277 | 50158.1135 | 51419.5331 | 1024.3907 | 924.3907 | 908.6811
2933 | 77.3235 | 1.1227 | 1.02 | 0.9942 | 213.4636 | 51.5054 | 56913.6488 | 55661.1032 | 1109.2221 | 1009.2221 | 908.7006
2934 | 101.8482 | 1.0901 | 1.1215 | 1.0269 | 198.605 | 53.8019 | 68134.5706 | 67920.8541 | 1354.4171 | 1254.4171 | 908.7044
2935 | 92.0822 | 1.0813 | 0.85401 | 0.98792 | 188.13 | 51.9245 | 49166.2107 | 48103.7281 | 958.0746 | 858.0746 | 908.7058
2936 | 97.8502 | 0.8945 | 0.96939 | 0.98934 | 198.7442 | 50.2947 | 52836.5835 | 53012.6432 | 1056.2529 | 956.2529 | 908.8342
2937 | 100.9207 | 1.0769 | 0.93797 | 1.0575 | 204.9288 | 45.9383 | 50764.7313 | 49241.868 | 980.8374 | 880.8374 | 908.8529
2938 | 93.1429 | 0.96203 | 1.0036 | 1.0178 | 203.6804 | 47.2211 | 52667.444 | 53327.3015 | 1062.546 | 962.546 | 908.8809
2939 | 115.1119 | 1.004 | 1.1136 | 0.99257 | 199.0362 | 44.9349 | 55061.181 | 54844.5633 | 1092.8913 | 992.8913 | 908.8946
2940 | 81.1635 | 1.0066 | 0.90296 | 1.0251 | 199.0911 | 46.7034 | 47243.3237 | 46482.6149 | 925.6523 | 825.6523 | 908.9704
2941 | 96.7541 | 0.95663 | 0.87779 | 0.96739 | 195.8791 | 52.1783 | 49333.4715 | 52134.9948 | 1038.6999 | 938.6999 | 909.0149
2942 | 94.5326 | 1.0239 | 1.2761 | 0.89859 | 189.2493 | 52.4583 | 65422.0797 | 66712.1297 | 1330.2426 | 1230.2426 | 909.0258
2943 | 117.4714 | 1.0156 | 1.0486 | 1.0437 | 187.8241 | 48.4401 | 58984.932 | 57978.7015 | 1155.574 | 1055.574 | 909.0323
2944 | 97.9789 | 1.0459 | 1.2308 | 0.99845 | 197.6581 | 48.9975 | 65431.0127 | 64166.8901 | 1279.3378 | 1179.3378 | 909.0391
2945 | 94.7881 | 1.0372 | 0.85829 | 0.9857 | 205.4714 | 53.6758 | 50893.3434 | 51711.3232 | 1030.2265 | 930.2265 | 909.0505
2946 | 109.9814 | 0.97634 | 0.79698 | 0.89492 | 222.2144 | 47.2103 | 38963.4565 | 38280.2456 | 761.6049 | 661.6049 | 909.1104
2947 | 87.7213 | 1.0002 | 1.0777 | 1.0776 | 195.8966 | 51.1447 | 64078.9537 | 62024.4671 | 1236.4893 | 1136.4893 | 909.1113
2948 | 96.4194 | 0.9801 | 0.80833 | 0.94476 | 195.621 | 48.4586 | 41781.879 | 41447.6324 | 824.9526 | 724.9526 | 909.172
2949 | 83.932 | 0.98404 | 0.86913 | 1.0225 | 206.0933 | 49.0421 | 47840.0955 | 46155.0131 | 919.1003 | 819.1003 | 909.3257
2950 | 86.0165 | 1.0181 | 0.90794 | 1.0006 | 190.1419 | 52.7004 | 52681.0783 | 52160.5871 | 1039.2117 | 939.2117 | 909.3352
2951 | 116.2978 | 1.0866 | 0.92667 | 0.9321 | 213.7104 | 50.877 | 50587.4167 | 48947.1181 | 974.9424 | 874.9424 | 909.3575
2952 | 102.3146 | 0.93061 | 0.64503 | 0.90892 | 182.4515 | 51.504 | 35282.0479 | 36601.3763 | 728.0275 | 628.0275 | 909.4094
2953 | 99.1671 | 1.0278 | 1.0882 | 1.0231 | 196.0629 | 48.5383 | 59183.6239 | 58654.0175 | 1169.0804 | 1069.0804 | 909.4526
2954 | 111.9721 | 0.99482 | 1.2611 | 0.99921 | 191.8954 | 52.7601 | 72553.3911 | 69104.9679 | 1378.0994 | 1278.0994 | 909.46
2955 | 118.7758 | 1.0357 | 1.0555 | 0.98355 | 204.0513 | 47.6977 | 55588.4974 | 55090.6505 | 1097.813 | 997.813 | 909.6764
2956 | 88.2993 | 0.95902 | 0.93325 | 1.0798 | 212.3042 | 49.7581 | 54568.3084 | 54169.4836 | 1079.3897 | 979.3897 | 909.7088
2957 | 108.7297 | 1.0089 | 0.8759 | 1.0455 | 208.2784 | 51.9849 | 53515.0164 | 51296.9534 | 1021.9391 | 921.9391 | 909.7898
2958 | 90.7817 | 1.0419 | 1.0549 | 1.0529 | 199.295 | 48.6742 | 58862.8034 | 58115.7562 | 1158.3151 | 1058.3151 | 910.0856
2959 | 72.2696 | 1.0314 | 0.77925 | 1.1369 | 190.071 | 48.1837 | 46470.4519 | 45266.1265 | 901.3225 | 801.3225 | 910.1388
2960 | 107.6041 | 1.0012 | 0.81635 | 0.97279 | 196.6439 | 49.3007 | 44659.6615 | 44206.8812 | 880.1376 | 780.1376 | 910.2017
2961 | 102.2483 | 1.0002 | 1.1717 | 1.0527 | 193.1577 | 51.1131 | 68461.7665 | 66964.663 | 1335.2933 | 1235.2933 | 910.2287
2962 | 99.1376 | 0.98282 | 0.87106 | 0.96362 | 210.9584 | 50.8169 | 47816.5266 | 48625.1018 | 968.502 | 868.502 | 910.2344
2963 | 115.6561 | 0.99484 | 1.1114 | 1.0368 | 197.6688 | 52.0491 | 66158.3408 | 68943.5819 | 1374.8716 | 1274.8716 | 910.2704
2964 | 94.2797 | 1.0085 | 0.71598 | 0.8978 | 190.4838 | 52.6835 | 39065.0374 | 38945.1796 | 774.9036 | 674.9036 | 910.2883
2965 | 98.7123 | 1.0476 | 1.0812 | 0.9683 | 199.4873 | 45.9798 | 53089.667 | 53629.4875 | 1068.5898 | 968.5898 | 910.3412
2966 | 112.651 | 1.1093 | 0.87939 | 1.0398 | 227.6971 | 51.4656 | 53719.203 | 52917.5377 | 1054.3508 | 954.3508 | 910.352
2967 | 89.0122 | 0.97549 | 0.87685 | 0.96243 | 194.2278 | 45.7398 | 42766.0085 | 43933.9902 | 874.6798 | 774.6798 | 910.3556
2968 | 117.1016 | 0.98643 | 1.0083 | 0.8973 | 187.2831 | 49.389 | 50578.7887 | 52683.8345 | 1049.6767 | 949.6767 | 910.4103
2969 | 108.1881 | 1.0324 | 0.95854 | 0.97931 | 203.2851 | 46.3068 | 48844.4029 | 48426.5938 | 964.5319 | 864.5319 | 910.4117
2970 | 83.8619 | 0.99515 | 1.1215 | 0.99785 | 199.8223 | 50.523 | 60955.8738 | 63559.7364 | 1267.1947 | 1167.1947 | 910.5588
2971 | 97.1184 | 1.0134 | 1.1725 | 1.0358 | 185.282 | 49.4428 | 65093.4933 | 64759.7395 | 1291.1948 | 1191.1948 | 910.5617
2972 | 85.1853 | 1.0262 | 1.1088 | 1.0157 | 198.6892 | 46.4484 | 56566.7353 | 55420.9989 | 1104.42 | 1004.42 | 910.5784
2973 | 93.627 | 0.99594 | 0.9896 | 1.0747 | 197.7071 | 52.6074 | 61051.5602 | 60060.3127 | 1197.2063 | 1097.2063 | 910.6666
2974 | 90.6744 | 0.97732 | 0.84292 | 0.99251 | 210.4413 | 45.2355 | 42063.7118 | 40976.5648 | 815.5313 | 715.5313 | 910.7476
2975 | 96.7982 | 0.92929 | 0.88561 | 1.0314 | 221.3632 | 49.1813 | 49568.948 | 48755.5133 | 971.1103 | 871.1103 | 910.8088
2976 | 103.0091 | 0.93867 | 1.017 | 1.0205 | 202.3367 | 52.2607 | 59494.8186 | 60469.0622 | 1205.3812 | 1105.3812 | 910.8569
2977 | 102.1239 | 1.0062 | 0.93089 | 0.88813 | 190.9072 | 52.3502 | 48850.9487 | 48641.9056 | 968.8381 | 868.8381 | 910.8794
2978 | 103.8551 | 1.0188 | 0.78861 | 0.98664 | 195.5294 | 50.8118 | 45107.1921 | 44574.4051 | 887.4881 | 787.4881 | 910.9842
2979 | 106.8523 | 1.0106 | 0.84904 | 0.95168 | 194.7269 | 50.8207 | 46746.376 | 46144.3337 | 918.8867 | 818.8867 | 911.0677
2980 | 84.7692 | 0.93886 | 0.99455 | 0.95469 | 198.4088 | 49.6354 | 51276.9127 | 50761.2646 | 1011.2253 | 911.2253 | 911.1881
2981 | 105.0772 | 0.99905 | 1.0418 | 1.0306 | 193.5721 | 54.9699 | 64983.1082 | 64025.5941 | 1276.5119 | 1176.5119 | 911.2253
2982 | 115.3805 | 1.0295 | 1.1908 | 0.93748 | 199.297 | 53.4025 | 66157.266 | 63290.7611 | 1261.8152 | 1161.8152 | 911.3558
2983 | 90.033 | 1.0175 | 1.0748 | 0.96141 | 208.6682 | 49.8913 | 56332.0789 | 56192.3447 | 1119.8469 | 1019.8469 | 911.447
2984 | 118.7652 | 0.91845 | 1.023 | 1.07 | 204.9469 | 46.3616 | 56012.0872 | 56677.4856 | 1129.5497 | 1029.5497 | 911.4922
2985 | 104.4287 | 0.99244 | 1.0987 | 1.0176 | 209.6012 | 51.8264 | 63527.13 | 64329.2101 | 1282.5842 | 1182.5842 | 911.5153
2986 | 90.1855 | 1.0498 | 0.86049 | 1.0738 | 213.4902 | 47.8876 | 48996.7495 | 49217.626 | 980.3525 | 880.3525 | 911.7208
2987 | 105.6687 | 1.0329 | 0.96731 | 1.0729 | 181.1176 | 49.1344 | 56536.9939 | 57533.4956 | 1146.6699 | 1046.6699 | 911.7338
2988 | 95.3204 | 0.99805 | 0.97729 | 1.0329 | 192.4916 | 50.8946 | 56410.4756 | 58357.8989 | 1163.158 | 1063.158 | 911.7534
2989 | 94.9587 | 0.96492 | 0.9046 | 1.0174 | 211.5807 | 51.9503 | 52781.6758 | 52072.6261 | 1037.4525 | 937.4525 | 911.7649
2990 | 101.1998 | 0.97749 | 1.0182 | 0.89574 | 194.3678 | 49.4156 | 50150.9431 | 49414.9933 | 984.2999 | 884.2999 | 911.769
2991 | 110.0892 | 1.0189 | 1.1105 | 1.0194 | 203.9382 | 53.9853 | 67375.748 | 65106.649 | 1298.133 | 1198.133 | 911.7696
2992 | 99.9242 | 0.99425 | 0.97651 | 0.94557 | 203.903 | 52.0587 | 53444.8031 | 53213.792 | 1060.2758 | 960.2758 | 911.79
2993 | 97.5609 | 1.0324 | 0.92909 | 0.97887 | 188.2554 | 46.1161 | 46774.0315 | 46840.8094 | 932.8162 | 832.8162 | 911.8015
2994 | 116.2248 | 0.91824 | 0.71905 | 1.0131 | 196.986 | 49.8726 | 41849.0379 | 42192.0622 | 839.8412 | 739.8412 | 911.8232
2995 | 90.5603 | 0.98644 | 1.0861 | 0.93429 | 208.8828 | 54.6451 | 60542.2046 | 60682.0089 | 1209.6402 | 1109.6402 | 911.8249
2996 | 111.0127 | 1.0408 | 0.84352 | 0.96871 | 215.3738 | 49.7396 | 46605.934 | 48429.6347 | 964.5927 | 864.5927 | 911.8308
2997 | 124.4221 | 0.99223 | 1.0122 | 1.0212 | 205.3413 | 47.5316 | 55202.1362 | 56204.901 | 1120.098 | 1020.098 | 911.8318
2998 | 109.132 | 1.0259 | 0.84218 | 1.011 | 195.1086 | 49.9019 | 48269.0682 | 45637.3359 | 908.7467 | 808.7467 | 911.8362
2999 | 95.9942 | 0.89111 | 1.0781 | 0.94345 | 186.687 | 52.2217 | 57768.0069 | 59616.9452 | 1188.3389 | 1088.3389 | 911.8481
3000 | 93.1659 | 0.98906 | 0.81331 | 1.0154 | 180.8529 | 51.0506 | 47044.0541 | 45766.1346 | 911.3227 | 811.3227 | 911.8907
3001 | 99.0439 | 1.0278 | 1.3049 | 0.99417 | 196.3222 | 54.6162 | 76609.41 | 75707.2435 | 1510.1449 | 1410.1449 | 911.9714
3002 | 94.1078 | 0.98029 | 1.039 | 0.95717 | 182.5752 | 47.0607 | 51324.6654 | 51328.0944 | 1022.5619 | 922.5619 | 911.9769
3003 | 116.4127 | 1.0456 | 1.0579 | 0.97152 | 178.2734 | 52.8363 | 60912.5855 | 61417.4994 | 1224.35 | 1124.35 | 912.0248
3004 | 100.1111 | 1.0079 | 1.0676 | 0.96841 | 182.2456 | 45.2254 | 51503.3372 | 50479.3158 | 1005.5863 | 905.5863 | 912.1414
3005 | 136.7752 | 0.92843 | 0.71153 | 0.91378 | 193.294 | 51.5951 | 40291.2908 | 40433.0439 | 804.6609 | 704.6609 | 912.1581
3006 | 88.9256 | 0.94613 | 0.94843 | 1.0577 | 209.3997 | 49.1318 | 53629.4678 | 53459.1263 | 1065.1825 | 965.1825 | 912.1605
3007 | 112.5582 | 1.0103 | 0.96624 | 1.0236 | 214.2129 | 50.6814 | 56105.6823 | 59547.7443 | 1186.9549 | 1086.9549 | 912.1653
3008 | 88.8424 | 0.86828 | 0.89009 | 0.98966 | 217.4095 | 52.647 | 50654.546 | 52420.1393 | 1044.4028 | 944.4028 | 912.2606
3009 | 101.4504 | 0.99675 | 0.88022 | 1.0276 | 216.1899 | 47.6721 | 48155.9129 | 47500.2384 | 946.0048 | 846.0048 | 912.2697
3010 | 94.5528 | 0.92762 | 1.3431 | 1.0507 | 199.3553 | 53.6541 | 80621.8743 | 79696.1151 | 1589.9223 | 1489.9223 | 912.32
3011 | 102.738 | 0.93082 | 0.83435 | 0.99269 | 202.9171 | 52.1103 | 48346.8143 | 47883.0505 | 953.661 | 853.661 | 912.344
3012 | 102.7529 | 0.95631 | 1.1632 | 0.93132 | 203.7602 | 52.8457 | 62643.3634 | 62877.877 | 1253.5575 | 1153.5575 | 912.4394
3013 | 114.5009 | 1.0356 | 1.054 | 0.98971 | 196.7216 | 46.2253 | 53896.4848 | 55415.3343 | 1104.3067 | 1004.3067 | 912.4536
3014 | 99.7189 | 1.0267 | 1.0759 | 1.1078 | 195.7679 | 49.6043 | 64398.694 | 64051.9041 | 1277.0381 | 1177.0381 | 912.4612
3015 | 85.6312 | 0.953 | 1.0843 | 1.0405 | 179.8952 | 50.5975 | 61392.6877 | 62007.8759 | 1236.1575 | 1136.1575 | 912.5178
3016 | 107.625 | 1.0458 | 0.96413 | 0.99267 | 212.8065 | 45.7945 | 49195.3881 | 50252.7387 | 1001.0548 | 901.0548 | 912.5298
3017 | 93.0833 | 1.0343 | 1.0695 | 1.0511 | 193.9963 | 50.7693 | 62155.8544 | 64115.9628 | 1278.3193 | 1178.3193 | 912.5682
3018 | 103.9642 | 1.0431 | 0.99229 | 0.96005 | 200.0921 | 54.1649 | 57673.977 | 58988.2872 | 1175.7657 | 1075.7657 | 912.6125
3019 | 111.1016 | 1.0144 | 0.87555 | 0.97615 | 216.9745 | 47.825 | 46481.5127 | 47155.8314 | 939.1166 | 839.1166 | 912.7322
3020 | 116.3393 | 0.98228 | 1.0031 | 0.96732 | 204.1451 | 48.5185 | 52829.411 | 53205.5168 | 1060.1103 | 960.1103 | 912.7379
3021 | 89.4507 | 1.0077 | 0.81171 | 1.0266 | 203.3173 | 51.8944 | 48126.6558 | 47407.2109 | 944.1442 | 844.1442 | 912.7928
3022 | 99.0994 | 1.0047 | 0.98838 | 0.98104 | 181.2679 | 49.1558 | 52738.8217 | 50642.6456 | 1008.8529 | 908.8529 | 912.7959
3023 | 98.3904 | 0.97101 | 0.8532 | 1.039 | 187.7501 | 52.0535 | 51304.9883 | 52163.624 | 1039.2725 | 939.2725 | 912.8448
3024 | 92.3319 | 0.94651 | 1.05 | 1.0057 | 208.0562 | 54.6679 | 62713.358 | 62485.7975 | 1245.716 | 1145.716 | 912.9406
3025 | 97.2441 | 1.0178 | 0.80598 | 0.97447 | 199.6052 | 51.8633 | 46066.176 | 47347.7767 | 942.9555 | 842.9555 | 912.9814
3026 | 91.185 | 0.9995 | 0.95753 | 1.0512 | 190.9533 | 45.3804 | 50003.3256 | 48689.8425 | 969.7969 | 869.7969 | 912.9816
3027 | 95.9809 | 1.0887 | 1.0097 | 0.99372 | 203.104 | 50.6098 | 56269.556 | 57252.416 | 1141.0483 | 1041.0483 | 913.0745
3028 | 84.3249 | 1.0256 | 1.1349 | 0.98262 | 206.6634 | 51.2012 | 61734.6839 | 61891.7666 | 1233.8353 | 1133.8353 | 913.088
3029 | 116.2424 | 0.88079 | 1.2348 | 1.0315 | 205.2064 | 49.9609 | 68955.5787 | 68145.8586 | 1358.9172 | 1258.9172 | 913.0997
3030 | 101.8955 | 1.0049 | 1.1127 | 1.0064 | 185.5857 | 47.0605 | 57699.6944 | 57720.6347 | 1150.4127 | 1050.4127 | 913.2146
3031 | 93.6509 | 0.983 | 1.0665 | 0.98684 | 184.9111 | 53.6642 | 61606.9557 | 60954.5177 | 1215.0904 | 1115.0904 | 913.2688
3032 | 87.1806 | 1.0261 | 0.81345 | 1.0973 | 208.9041 | 46.743 | 46114.5015 | 46536.3428 | 926.7269 | 826.7269 | 913.3092
3033 | 104.3338 | 1.0255 | 1.2033 | 1.0406 | 203.3379 | 46.365 | 63223.9501 | 64876.681 | 1293.5336 | 1193.5336 | 913.3276
3034 | 102.4189 | 0.98304 | 0.77884 | 0.95395 | 191.8679 | 46.8054 | 39679.956 | 37877.0267 | 753.5405 | 653.5405 | 913.3518
3035 | 98.7439 | 1.0413 | 1.1212 | 1.0088 | 193.7988 | 47.5648 | 58882.6129 | 57569.2707 | 1147.3854 | 1047.3854 | 913.3883
3036 | 82.6734 | 1.0231 | 0.96106 | 0.9271 | 207.2977 | 46.5093 | 45580.9499 | 47912.0866 | 954.2417 | 854.2417 | 913.4186
3037 | 89.5763 | 1.0523 | 0.7831 | 0.94812 | 201.3529 | 49.7735 | 41848.4891 | 40878.8221 | 813.5764 | 713.5764 | 913.5062
3038 | 97.7168 | 0.93214 | 0.89062 | 1.0389 | 212.6105 | 48.0106 | 49006.6641 | 48574.672 | 967.4934 | 867.4934 | 913.9219
3039 | 94.0023 | 1.0203 | 1.0905 | 1.0947 | 206.33 | 53.6401 | 69386.4956 | 68049.7699 | 1356.9954 | 1256.9954 | 913.9409
3040 | 97.8434 | 1.0822 | 1.1937 | 1.0133 | 197.3705 | 49.8739 | 65805.2548 | 64631.5235 | 1288.6305 | 1188.6305 | 914.0769
3041 | 85.5648 | 1.0213 | 0.87345 | 0.92621 | 190.3253 | 50.6306 | 45574.6225 | 45169.4629 | 899.3893 | 799.3893 | 914.1802
3042 | 106.2289 | 0.98101 | 0.98363 | 0.91843 | 196.4897 | 51.8458 | 52436.9539 | 52028.4103 | 1036.5682 | 936.5682 | 914.1941
3043 | 92.7 | 1.0361 | 1.0761 | 0.97593 | 206.9684 | 46.8542 | 53912.6053 | 53231.4949 | 1060.6299 | 960.6299 | 914.2034
3044 | 101.4238 | 0.94677 | 1.2349 | 0.92872 | 207.2114 | 47.3281 | 59029.7848 | 57704.4474 | 1150.0889 | 1050.0889 | 914.2075
3045 | 110.1932 | 0.98106 | 0.87577 | 1.0084 | 200.024 | 46.4402 | 46232.8025 | 47804.2879 | 952.0858 | 852.0858 | 914.2196
3046 | 103.061 | 0.98862 | 0.91198 | 0.96428 | 199.351 | 49.2865 | 48563.584 | 48344.1696 | 962.8834 | 862.8834 | 914.2484
3047 | 96.1422 | 1.0177 | 0.87855 | 1.0386 | 201.9804 | 52.4032 | 53146.4129 | 53372.029 | 1063.4406 | 963.4406 | 914.3448
3048 | 105.2942 | 0.98794 | 0.88025 | 0.97409 | 206.051 | 47.4103 | 45789.7248 | 46299.7109 | 921.9942 | 821.9942 | 914.6623
3049 | 98.5566 | 1.0807 | 1.2347 | 1.044 | 193.6906 | 51.9086 | 72634.3144 | 72767.6234 | 1451.3525 | 1351.3525 | 914.7355
3050 | 102.2888 | 1.006 | 0.94747 | 1.0122 | 209.7275 | 54.0856 | 57643.0807 | 57848.2968 | 1152.9659 | 1052.9659 | 914.8287
3051 | 88.2663 | 1.0942 | 1.0258 | 0.93639 | 207.7622 | 54.0115 | 57304.375 | 57682.9596 | 1149.6592 | 1049.6592 | 914.829
3052 | 104.8471 | 0.96017 | 1.2058 | 1.0402 | 195.4616 | 48.9145 | 66473.7096 | 65698.3083 | 1309.9662 | 1209.9662 | 914.9522
3053 | 108.7293 | 0.96013 | 1.179 | 0.95946 | 209.4925 | 49.0059 | 60761.6436 | 58264.5431 | 1161.2909 | 1061.2909 | 915.1292
3054 | 108.3307 | 1.0325 | 0.9719 | 0.99704 | 213.1371 | 50.8783 | 55206.0906 | 54905.2075 | 1094.1041 | 994.1041 | 915.131
3055 | 98.1015 | 0.96288 | 0.84838 | 0.95559 | 206.9996 | 45.0105 | 40948.9729 | 40473.2472 | 805.4649 | 705.4649 | 915.1926
3056 | 107.8317 | 1.0099 | 0.90754 | 1.0374 | 184.2099 | 52.4419 | 55269.9846 | 54119.7894 | 1078.3958 | 978.3958 | 915.2311
3057 | 98.7379 | 1.0663 | 1.1914 | 1.1459 | 213.7107 | 54.6952 | 80645.383 | 80860.8999 | 1613.218 | 1513.218 | 915.2705
3058 | 104.1053 | 0.94618 | 1.2551 | 1.0237 | 185.4842 | 47.9698 | 66543.543 | 66988.3117 | 1335.7662 | 1235.7662 | 915.2841
3059 | 95.3486 | 0.9448 | 0.78213 | 1.0191 | 200.5097 | 50.7986 | 45266.0156 | 45125.3705 | 898.5074 | 798.5074 | 915.3136
3060 | 103.4895 | 0.94953 | 0.98192 | 0.92942 | 187.5487 | 49.2539 | 49977.4901 | 49324.9628 | 982.4993 | 882.4993 | 915.4248
3061 | 92.4151 | 1.0013 | 1.211 | 1.0547 | 207.1571 | 47.4667 | 65227.276 | 66790.6548 | 1331.8131 | 1231.8131 | 915.4731
3062 | 97.7364 | 1.0459 | 1.1083 | 1.0347 | 205.7008 | 50.7165 | 63550.2458 | 63692.3106 | 1269.8462 | 1169.8462 | 915.5154
3063 | 101.8938 | 1.0646 | 0.98628 | 0.99004 | 188.1149 | 46.2016 | 50314.1056 | 49972.5064 | 995.4501 | 895.4501 | 915.5645
3064 | 99.9055 | 0.94816 | 0.90819 | 0.98976 | 212.1897 | 46.6381 | 46552.4198 | 46938.7689 | 934.7754 | 834.7754 | 915.5791
3065 | 104.5604 | 1.0212 | 0.97434 | 0.98657 | 207.9523 | 47.6738 | 51124.9969 | 51203.2412 | 1020.0648 | 920.0648 | 915.5889
3066 | 99.1986 | 1.0335 | 1.1358 | 0.92932 | 207.375 | 51.2083 | 59510.1765 | 59158.7186 | 1179.1744 | 1079.1744 | 915.7697
3067 | 113.0474 | 1.0106 | 0.94085 | 1.0005 | 223.6073 | 49.305 | 52270.0589 | 52569.088 | 1047.3818 | 947.3818 | 915.8594
3068 | 88.3526 | 0.97345 | 0.99564 | 1.0392 | 203.1094 | 47.4591 | 53390.5136 | 53380.911 | 1063.6182 | 963.6182 | 915.9895
3069 | 101.2665 | 1.011 | 0.99955 | 1.007 | 194.0484 | 48.905 | 54424.4583 | 52127.2853 | 1038.5457 | 938.5457 | 916.0337
3070 | 101.82 | 1.0299 | 1.0514 | 0.99423 | 198.4117 | 48.225 | 55664.756 | 53761.7578 | 1071.2352 | 971.2352 | 916.1245
3071 | 87.3961 | 0.97093 | 1.0663 | 1.0523 | 204.1701 | 50.1417 | 60722.482 | 60646.9462 | 1208.9389 | 1108.9389 | 916.1921
3072 | 97.4546 | 0.98782 | 1.0531 | 0.97929 | 202.4739 | 49.3579 | 55855.5445 | 55787.1988 | 1111.744 | 1011.744 | 916.2322
3073 | 116.6925 | 0.97602 | 1.3269 | 0.99595 | 196.6141 | 46.821 | 67404.6545 | 67920.0148 | 1354.4003 | 1254.4003 | 916.2641
3074 | 98.187 | 1.0444 | 1.0791 | 1.0359 | 207.8394 | 49.6499 | 60797.4699 | 60378.0234 | 1203.5605 | 1103.5605 | 916.3132
3075 | 97.6674 | 1.0127 | 1.031 | 0.95469 | 201.6829 | 51.1794 | 55637.6138 | 56881.956 | 1133.6391 | 1033.6391 | 916.3589
3076 | 101.846 | 1.0808 | 1.0307 | 1.0014 | 179.4777 | 51.7456 | 59288.0248 | 59084.9907 | 1177.6998 | 1077.6998 | 916.5207
3077 | 108.1739 | 0.97101 | 0.97088 | 1.01 | 193.7709 | 49.2126 | 53620.9294 | 55694.0157 | 1109.8803 | 1009.8803 | 916.6448
3078 | 87.7548 | 0.96303 | 0.70597 | 1.0669 | 202.1591 | 48.8578 | 41132.2815 | 40000.2344 | 796.0047 | 696.0047 | 916.6675
3079 | 106.9506 | 0.96161 | 0.86607 | 1.0762 | 180.2416 | 45.569 | 47341.0203 | 46279.6782 | 921.5936 | 821.5936 | 916.6997
3080 | 111.4192 | 0.95117 | 0.96181 | 0.94234 | 209.6025 | 52.8466 | 53707.9661 | 55459.9284 | 1105.1986 | 1005.1986 | 916.7852
3081 | 96.7795 | 1.139 | 0.913 | 0.99979 | 210.8149 | 45.978 | 47247.8842 | 47858.6258 | 953.1725 | 853.1725 | 916.8056
3082 | 106.8834 | 0.9882 | 0.87847 | 0.99771 | 188.0425 | 48.3003 | 47622.6291 | 48029.8033 | 956.5961 | 856.5961 | 916.8417
3083 | 112.7041 | 1.0209 | 1.1489 | 1.0037 | 198.8034 | 51.5221 | 65537.7641 | 62870.6616 | 1253.4132 | 1153.4132 | 916.8929
3084 | 107.0842 | 1.0646 | 0.94148 | 0.99353 | 224.1085 | 51.5303 | 54299.7327 | 53573.755 | 1067.4751 | 967.4751 | 917.0226
3085 | 78.4843 | 1.0464 | 0.94556 | 0.97942 | 178.2397 | 46.6996 | 47262.435 | 46783.1623 | 931.6632 | 831.6632 | 917.0604
3086 | 101.5662 | 1.0851 | 1.017 | 0.97716 | 217.202 | 51.5125 | 57086.9601 | 56029.1311 | 1116.5826 | 1016.5826 | 917.0959
3087 | 111.2689 | 0.94333 | 1.2199 | 1.0259 | 200.3008 | 49.5242 | 67374.8335 | 68991.8585 | 1375.8372 | 1275.8372 | 917.137
3088 | 99.6143 | 1.0545 | 0.74038 | 0.98241 | 210.3595 | 50.8979 | 42577.4951 | 42507.4489 | 846.149 | 746.149 | 917.1573
3089 | 103.101 | 1.0283 | 0.7665 | 0.91263 | 211.2613 | 49.4688 | 40061.0075 | 40156.614 | 799.1323 | 699.1323 | 917.2813
3090 | 96.3586 | 1.0751 | 1.0647 | 1.0656 | 194.9897 | 46.7271 | 58048.3102 | 59769.3972 | 1191.3879 | 1091.3879 | 917.2873
3091 | 80.4295 | 0.98497 | 1.0978 | 1.096 | 201.2177 | 50.544 | 65019.93 | 63800.2511 | 1272.005 | 1172.005 | 917.3912
3092 | 103.4364 | 1.0299 | 1.0132 | 1.0148 | 185.6545 | 53.9033 | 61351.7556 | 61949.8795 | 1234.9976 | 1134.9976 | 917.4129
3093 | 104.7864 | 0.87766 | 0.84866 | 0.9726 | 187.9507 | 48.5823 | 44756.312 | 44701.3566 | 890.0271 | 790.0271 | 917.5026
3094 | 99.9197 | 1.088 | 0.939 | 1.0078 | 190.4546 | 50.8758 | 53865.0961 | 53754.5843 | 1071.0917 | 971.0917 | 917.5362
3095 | 114.0568 | 0.89259 | 0.9095 | 1.0279 | 191.5182 | 50.7559 | 52809.3406 | 53738.883 | 1070.7777 | 970.7777 | 917.5515
3096 | 96.861 | 0.90357 | 1.0017 | 1.0422 | 205.5666 | 54.2019 | 61533.4484 | 61990.7067 | 1235.8141 | 1135.8141 | 917.7208
3097 | 97.7032 | 0.99065 | 0.81375 | 1.0957 | 192.9838 | 51.0987 | 50698.8991 | 50655.5208 | 1009.1104 | 909.1104 | 917.7229
3098 | 85.7528 | 0.97127 | 0.66098 | 0.90444 | 192.8807 | 51.5489 | 35303.3161 | 34734.625 | 690.6925 | 590.6925 | 917.7433
3099 | 111.8301 | 1.0787 | 1.0376 | 0.98552 | 217.324 | 42.7925 | 49136.7956 | 49215.8362 | 980.3167 | 880.3167 | 917.7544
3100 | 96.6174 | 0.8949 | 0.93121 | 0.96974 | 207.5143 | 46.8498 | 46565.0332 | 46783.8029 | 931.6761 | 831.6761 | 917.8143
3101 | 112.6629 | 0.92752 | 0.93508 | 0.99327 | 199.5139 | 49.006 | 50836.3326 | 51802.7629 | 1032.0553 | 932.0553 | 917.8442
3102 | 88.5991 | 1.0094 | 1.0331 | 0.98938 | 201.7993 | 53.2168 | 59356.9423 | 58644.2392 | 1168.8848 | 1068.8848 | 917.9386
3103 | 78.9419 | 0.91888 | 1.1395 | 0.98576 | 196.7692 | 49.4369 | 59312.3254 | 58089.1301 | 1157.7826 | 1057.7826 | 917.9618
3104 | 93.3726 | 0.88658 | 0.89042 | 1.05 | 196.9204 | 50.6097 | 51704.7162 | 52907.0769 | 1054.1415 | 954.1415 | 917.9626
3105 | 90.1487 | 1.0902 | 0.72965 | 0.96258 | 197.9689 | 48.1104 | 38716.4263 | 38622.8347 | 768.4567 | 668.4567 | 917.9981
3106 | 109.3557 | 0.97774 | 1.1258 | 0.99313 | 193.5091 | 49.1327 | 60378.5093 | 61017.6975 | 1216.354 | 1116.354 | 918.002
3107 | 86.7567 | 0.93105 | 1.2424 | 1.0939 | 211.0968 | 49.7338 | 71819.4425 | 73139.7186 | 1458.7944 | 1358.7944 | 918.0382
3108 | 108.2313 | 0.92913 | 0.90956 | 1.033 | 185.5713 | 51.2303 | 53470.4817 | 52505.3147 | 1046.1063 | 946.1063 | 918.165
3109 | 88.5655 | 1.0421 | 1.1044 | 0.98301 | 194.9455 | 53.6058 | 63339.0304 | 63753.77 | 1271.0754 | 1171.0754 | 918.1915
3110 | 90.4795 | 1.0702 | 1.0961 | 0.95761 | 188.3406 | 51.9924 | 59796.7926 | 62089.0512 | 1237.781 | 1137.781 | 918.2203
3111 | 94.2374 | 1.0178 | 0.80455 | 1.0672 | 208.9267 | 49.3866 | 47350.6719 | 45204.4755 | 900.0895 | 800.0895 | 918.2899
3112 | 87.0413 | 0.9795 | 1.0017 | 0.94226 | 188.6314 | 49.0019 | 50616.9793 | 48656.0094 | 969.1202 | 869.1202 | 918.2922
3113 | 105.0782 | 1.0308 | 1.0087 | 1.0222 | 205.679 | 49.9047 | 57066.9315 | 56549.453 | 1126.9891 | 1026.9891 | 918.4009
3114 | 95.1525 | 0.95576 | 0.88567 | 0.86319 | 217.6614 | 47.8313 | 41134.6925 | 40111.4342 | 798.2287 | 698.2287 | 918.5134
3115 | 98.6714 | 1.0953 | 1.0975 | 0.95358 | 188.7957 | 49.2128 | 57011.7736 | 57365.6814 | 1143.3136 | 1043.3136 | 918.5369
3116 | 122.5029 | 0.92051 | 1.0227 | 1.0078 | 198.4335 | 44.9614 | 51606.1719 | 51240.3596 | 1020.8072 | 920.8072 | 918.5754
3117 | 107.082 | 0.94464 | 1.0612 | 1.0066 | 205.3401 | 48.2004 | 56568.8276 | 55130.1332 | 1098.6027 | 998.6027 | 918.6724
3118 | 103.6489 | 0.98329 | 1.0216 | 0.88567 | 200.1043 | 57.717 | 58302.876 | 59483.535 | 1185.6707 | 1085.6707 | 918.6855
3119 | 100.8289 | 1.012 | 0.69965 | 0.94123 | 206.1583 | 49.6487 | 37967.5222 | 37782.9163 | 751.6583 | 651.6583 | 918.7604
3120 | 100.2753 | 0.88631 | 1.1943 | 0.98396 | 202.8508 | 51.5989 | 65424.4863 | 65054.1666 | 1297.0833 | 1197.0833 | 918.7802
3121 | 114.5212 | 0.98823 | 0.65444 | 1.1165 | 199.9966 | 42.8376 | 36348.3142 | 37395.6803 | 743.9136 | 643.9136 | 918.7937
3122 | 86.7793 | 1.042 | 1.1101 | 0.99461 | 195.1718 | 47.6266 | 57085.4711 | 58571.355 | 1167.4271 | 1067.4271 | 918.8096
3123 | 89.6122 | 0.94523 | 0.72076 | 1.0028 | 199.3885 | 47.337 | 38421.9506 | 38052.8972 | 757.0579 | 657.0579 | 918.829
3124 | 100.3858 | 0.93498 | 0.98097 | 0.95478 | 201.1533 | 46.6137 | 48234.8053 | 47025.2082 | 936.5042 | 836.5042 | 918.855
3125 | 108.1787 | 0.89606 | 0.92395 | 0.97473 | 202.8522 | 49.815 | 49895.2295 | 50667.8732 | 1009.3575 | 909.3575 | 918.8564
3126 | 111.1904 | 1.0801 | 0.88628 | 1.0087 | 212.9521 | 50.3531 | 51274.967 | 51844.0233 | 1032.8805 | 932.8805 | 919.0033
3127 | 101.5924 | 1.0673 | 0.94486 | 1.0143 | 211.9262 | 52.0463 | 55735.8957 | 55230.3944 | 1100.6079 | 1000.6079 | 919.0777
3128 | 108.008 | 0.98626 | 1.0247 | 0.94938 | 207.8783 | 45.9383 | 49792.1873 | 48526.4222 | 966.5284 | 866.5284 | 919.0797
3129 | 108.0549 | 0.90976 | 1.2492 | 1.0294 | 198.7691 | 49.516 | 68740.8392 | 68799.5714 | 1371.9914 | 1271.9914 | 919.1018
3130 | 107.2901 | 1.0028 | 0.92882 | 0.99063 | 192.7107 | 48.315 | 49846.031 | 48206.8193 | 960.1364 | 860.1364 | 919.1691
3131 | 98.4574 | 1.0005 | 1.3752 | 0.96086 | 186.9111 | 52.1501 | 74233.6586 | 74474.9646 | 1485.4993 | 1385.4993 | 919.1793
3132 | 110.7894 | 0.95064 | 0.89466 | 1.0255 | 203.1037 | 49.478 | 50810.2788 | 51872.1568 | 1033.4431 | 933.4431 | 919.1811
3133 | 123.6287 | 1.0431 | 0.89068 | 0.95339 | 202.9435 | 49.7402 | 48855.1741 | 48688.2819 | 969.7656 | 869.7656 | 919.2131
3134 | 108.1287 | 1.0681 | 0.839 | 1.0689 | 197.5599 | 48.0914 | 48881.5551 | 50427.619 | 1004.5524 | 904.5524 | 919.2407
3135 | 84.0648 | 1.0074 | 1.2692 | 0.98646 | 199.3305 | 53.1514 | 71249.5407 | 73366.6013 | 1463.332 | 1363.332 | 919.2605
3136 | 96.8685 | 0.95843 | 0.99706 | 1.0706 | 195.0905 | 50.3587 | 58627.1047 | 57645.8451 | 1148.9169 | 1048.9169 | 919.2733
3137 | 95.062 | 0.9007 | 1.0366 | 0.97031 | 202.6853 | 47.5523 | 52104.3596 | 51179.5087 | 1019.5902 | 919.5902 | 919.3057
3138 | 98.1365 | 0.98954 | 1.1656 | 0.87294 | 207.2408 | 49.8672 | 55788.7597 | 56483.5308 | 1125.6706 | 1025.6706 | 919.3141
3139 | 115.1931 | 1.0063 | 0.87798 | 0.93385 | 209.959 | 53.5109 | 50286.4263 | 49914.0794 | 994.2816 | 894.2816 | 919.3775
3140 | 99.9162 | 1.0359 | 1.0295 | 0.96059 | 183.9964 | 52.6376 | 57685.0311 | 56981.199 | 1135.624 | 1035.624 | 919.4142
3141 | 104.004 | 0.97922 | 1.2079 | 1.0272 | 206.8917 | 51.296 | 69071.9961 | 66722.1561 | 1330.4431 | 1230.4431 | 919.4695
3142 | 107.4718 | 1.0947 | 0.88712 | 1.067 | 189.6247 | 51.9857 | 55512.2555 | 53600.3772 | 1068.0075 | 968.0075 | 919.5548
3143 | 93.3605 | 0.9692 | 0.98496 | 1.0872 | 204.6638 | 48.7064 | 56767.0529 | 58039.9762 | 1156.7995 | 1056.7995 | 919.5761
3144 | 107.4328 | 1.0979 | 0.93407 | 0.96641 | 207.6529 | 53.3216 | 54629.9839 | 54612.9302 | 1088.2586 | 988.2586 | 919.5902
3145 | 101.7166 | 1.0004 | 1.0953 | 0.98346 | 191.8649 | 50.3227 | 59521.7968 | 58691.8054 | 1169.8361 | 1069.8361 | 919.5969
3146 | 94.2211 | 0.97979 | 1.0772 | 0.97785 | 200.8999 | 49.2291 | 56600.844 | 57212.4733 | 1140.2495 | 1040.2495 | 919.6415
3147 | 105.6576 | 0.93703 | 1.0678 | 0.94402 | 183.6366 | 49.1134 | 54555.1851 | 52226.2429 | 1040.5249 | 940.5249 | 919.7409
3148 | 110.028 | 1.0311 | 1.1799 | 1.0677 | 191.965 | 46.8698 | 64556.1388 | 63380.312 | 1263.6062 | 1163.6062 | 919.7427
3149 | 92.1614 | 1.0733 | 1.0672 | 0.97772 | 206.7509 | 50.5483 | 57950.3228 | 56477.5415 | 1125.5508 | 1025.5508 | 919.8431
3150 | 111.8679 | 1.02 | 0.92294 | 0.96111 | 189.9507 | 49.9546 | 50201.7114 | 48859.3086 | 973.1862 | 873.1862 | 919.8469
3151 | 109.5853 | 0.98381 | 1.1236 | 1.0089 | 199.0202 | 52.4766 | 65346.7688 | 65775.9179 | 1311.5184 | 1211.5184 | 919.8488
3152 | 93.6836 | 1.0044 | 0.73492 | 0.97607 | 200.4943 | 51.186 | 41734.5476 | 42040.9292 | 836.8186 | 736.8186 | 919.9579
3153 | 106.8103 | 0.9778 | 1.0309 | 0.90232 | 202.6868 | 49.6758 | 51598.5978 | 52193.2958 | 1039.8659 | 939.8659 | 919.9718
3154 | 110.1044 | 0.94666 | 0.87771 | 0.95849 | 202.3479 | 52.2334 | 49589.6595 | 50564.8863 | 1007.2977 | 907.2977 | 919.9916
3155 | 95.1832 | 1.064 | 0.88109 | 1.0688 | 194.304 | 50.3317 | 52690.3759 | 52703.0766 | 1050.0615 | 950.0615 | 920.0163
3156 | 94.433 | 1.008 | 0.94806 | 0.96271 | 183.525 | 45.9544 | 46500.769 | 48021.5904 | 956.4318 | 856.4318 | 920.0648
3157 | 87.4938 | 1.0168 | 0.93101 | 1.0221 | 205.1452 | 52.6857 | 55026.9719 | 55101.9572 | 1098.0391 | 998.0391 | 920.0662
3158 | 86.0824 | 1.0374 | 1.1118 | 0.9812 | 196.1368 | 48.8745 | 57879.7695 | 58620.4624 | 1168.4092 | 1068.4092 | 920.0917
3159 | 115.3796 | 0.95281 | 1.1385 | 0.9672 | 204.5342 | 49.0076 | 59558.8333 | 59024.715 | 1176.4943 | 1076.4943 | 920.0934
3160 | 111.6588 | 0.97792 | 0.90688 | 1.0029 | 189.826 | 46.4889 | 47549.3553 | 47051.7615 | 937.0352 | 837.0352 | 920.1077
3161 | 95.9238 | 0.85105 | 0.93772 | 1.0275 | 205.7687 | 48.0346 | 50407.4205 | 49295.5163 | 981.9103 | 881.9103 | 920.1569
3162 | 103.6319 | 0.99128 | 1.1702 | 0.96534 | 201.6139 | 50.4177 | 62333.8145 | 62543.8316 | 1246.8766 | 1146.8766 | 920.1687
3163 | 117.1829 | 0.99168 | 1.245 | 1.0085 | 203.0294 | 47.4714 | 65327.0494 | 64775.5864 | 1291.5117 | 1191.5117 | 920.1764
3164 | 110.4565 | 0.9113 | 1.1306 | 0.98607 | 206.615 | 47.0568 | 57405.949 | 57186.2422 | 1139.7248 | 1039.7248 | 920.2007
3165 | 91.0532 | 0.95212 | 0.87755 | 1.0297 | 197.8941 | 49.6257 | 49342.0255 | 49322.6743 | 982.4535 | 882.4535 | 920.2032
3166 | 112.0981 | 1.0938 | 0.90662 | 0.94635 | 187.8814 | 46.9882 | 46263.8953 | 46430.2773 | 924.6055 | 824.6055 | 920.2147
3167 | 76.8702 | 0.95625 | 0.9928 | 1.0528 | 223.4528 | 53.0889 | 59614.182 | 59828.4564 | 1192.5691 | 1092.5691 | 920.2474
3168 | 112.7305 | 1.0199 | 0.84867 | 0.99554 | 218.6252 | 49.8832 | 48099.2544 | 49042.0678 | 976.8414 | 876.8414 | 920.3078
3169 | 102.6251 | 0.98387 | 0.90929 | 1.0512 | 188.8866 | 50.2858 | 53332.1803 | 53416.9034 | 1064.3381 | 964.3381 | 920.3291
3170 | 105.2511 | 1.0145 | 0.83536 | 0.90399 | 214.1419 | 48.9359 | 42393.9407 | 41916.8597 | 834.3372 | 734.3372 | 920.3803
3171 | 107.8255 | 0.9631 | 0.96319 | 1.0389 | 211.0282 | 48.5056 | 53785.6202 | 53560.6551 | 1067.2131 | 967.2131 | 920.4747
3172 | 108.4609 | 1.0244 | 0.99502 | 1.0308 | 190.5617 | 46.6638 | 53235.3557 | 53970.4516 | 1075.409 | 975.409 | 920.5658
3173 | 83.6246 | 0.98992 | 1.124 | 0.97701 | 198.3845 | 51.521 | 61040.5649 | 61638.2731 | 1228.7655 | 1128.7655 | 920.6345
3174 | 120.6552 | 0.99541 | 0.89416 | 0.91273 | 208.2033 | 53.1273 | 49947.6125 | 50287.2241 | 1001.7445 | 901.7445 | 920.6489
3175 | 100.2057 | 0.95849 | 1.1506 | 1.0228 | 195.7219 | 50.0506 | 63904.9888 | 65903.0209 | 1314.0604 | 1214.0604 | 920.6523
3176 | 110.943 | 1.068 | 0.9337 | 0.97821 | 197.026 | 50.808 | 52622.8345 | 53441.8095 | 1064.8362 | 964.8362 | 920.6837
3177 | 107.0334 | 0.94988 | 1.0731 | 0.92727 | 201.3716 | 52.6559 | 57948.7334 | 57787.2964 | 1151.7459 | 1051.7459 | 920.6851
3178 | 117.7854 | 0.95983 | 1.2315 | 0.9858 | 192.8574 | 49.7373 | 66196.8183 | 67568.4326 | 1347.3687 | 1247.3687 | 920.689
3179 | 95.6554 | 0.8969 | 0.80395 | 1.0513 | 200.201 | 47.2054 | 44147.6744 | 43707.0655 | 870.1413 | 770.1413 | 920.7166
3180 | 95.7455 | 0.95746 | 0.8433 | 1.0551 | 203.6706 | 47.0315 | 46363.4953 | 45574.5396 | 907.4908 | 807.4908 | 920.746
3181 | 99.4686 | 0.96144 | 1.0739 | 1.0339 | 218.9358 | 49.5666 | 59992.6196 | 59728.5314 | 1190.5706 | 1090.5706 | 920.7909
3182 | 98.5961 | 0.93814 | 1.1674 | 1.0037 | 210.2122 | 50.1855 | 63657.2577 | 63269.4132 | 1261.3883 | 1161.3883 | 920.8072
3183 | 99.9218 | 1.003 | 1.1083 | 1.0286 | 206.5775 | 48.2351 | 60026.524 | 60231.7472 | 1200.6349 | 1100.6349 | 920.8519
3184 | 90.1973 | 1.0756 | 0.92997 | 0.97891 | 210.4015 | 49.938 | 50516.4558 | 50575.0985 | 1007.502 | 907.502 | 920.9094
3185 | 106.5283 | 0.94441 | 1.1979 | 1.001 | 194.2752 | 50.0958 | 65301.989 | 64818.9207 | 1292.3784 | 1192.3784 | 920.9281
3186 | 91.8744 | 0.9517 | 0.96824 | 0.94373 | 204.3716 | 48.5498 | 48812.2729 | 47205.9749 | 940.1195 | 840.1195 | 920.9422
3187 | 84.8897 | 0.95124 | 0.91528 | 1.0351 | 196.99 | 50.4535 | 52069.0847 | 52973.2245 | 1055.4645 | 955.4645 | 921.0682
3188 | 110.396 | 1.0026 | 0.97346 | 1.0743 | 209.4607 | 51.568 | 59844.6465 | 60373.5174 | 1203.4703 | 1103.4703 | 921.0841
3189 | 113.216 | 0.91833 | 1.0558 | 1.0521 | 182.0817 | 51.2165 | 62399.1731 | 62614.0107 | 1248.2802 | 1148.2802 | 921.3177
3190 | 96.833 | 0.95013 | 1.2048 | 0.96754 | 184.6901 | 49.2932 | 62182.2553 | 62144.5733 | 1238.8915 | 1138.8915 | 921.3385
3191 | 107.4999 | 0.99244 | 1.1833 | 1.0318 | 192.7485 | 50.1504 | 66776.4362 | 66678.7205 | 1329.5744 | 1229.5744 | 921.3764
3192 | 103.4607 | 0.96145 | 1.2214 | 1.0634 | 179.0594 | 46.6456 | 65406.8631 | 63096.2998 | 1257.926 | 1157.926 | 921.4143
3193 | 103.587 | 1.0124 | 0.82053 | 1.0009 | 208.6589 | 53.6887 | 49932.6509 | 50717.0596 | 1010.3412 | 910.3412 | 921.5267
3194 | 88.9678 | 1.0457 | 0.85541 | 0.98562 | 194.4371 | 49.0333 | 46096.1616 | 46162.6416 | 919.2528 | 819.2528 | 921.6076
3195 | 94.8363 | 0.91497 | 1.0613 | 0.96265 | 197.9824 | 47.8184 | 53202.374 | 53247.1476 | 1060.943 | 960.943 | 921.6276
3196 | 116.382 | 1.0333 | 0.92626 | 0.90646 | 196.4038 | 49.9109 | 48104.8506 | 48495.2371 | 965.9047 | 865.9047 | 921.7081
3197 | 94.3174 | 0.98166 | 1.0468 | 1.0341 | 215.0674 | 49.2031 | 58033.2331 | 57803.574 | 1152.0715 | 1052.0715 | 921.7143
3198 | 96.2217 | 1.0454 | 0.97435 | 1.0358 | 184.7514 | 47.3835 | 52771.8292 | 54146.5183 | 1078.9304 | 978.9304 | 921.7445
3199 | 101.4553 | 1.0146 | 0.85928 | 1.0406 | 200.8973 | 48.1433 | 48205.9094 | 47388.978 | 943.7796 | 843.7796 | 921.7807
3200 | 90.8037 | 1.0555 | 0.71933 | 0.97852 | 191.2364 | 47.0932 | 37852.3305 | 37168.8678 | 739.3774 | 639.3774 | 921.8445
3201 | 90.1209 | 0.96917 | 1.0825 | 0.92113 | 201.2744 | 52.869 | 57535.9537 | 57858.7496 | 1153.175 | 1053.175 | 921.9351
3202 | 108.9385 | 0.99601 | 0.99192 | 1.0135 | 201.684 | 52.0308 | 58154.0019 | 60732.0765 | 1210.6415 | 1110.6415 | 921.9391
3203 | 87.0873 | 1.0219 | 1.1215 | 1.1378 | 196.1949 | 51.0829 | 69926.4636 | 70443.6882 | 1404.8738 | 1304.8738 | 921.9994
3204 | 87.3283 | 0.99082 | 0.93583 | 0.99254 | 211.29 | 50.4378 | 51424.3311 | 49761.7718 | 991.2354 | 891.2354 | 922.0021
3205 | 123.6794 | 1.0908 | 0.85755 | 0.99597 | 211.1951 | 48.6976 | 48373.3801 | 49108.3397 | 978.1668 | 878.1668 | 922.0262
3206 | 98.5315 | 0.95357 | 0.84303 | 0.9622 | 229.058 | 49.1529 | 44718.398 | 45373.0964 | 903.4619 | 803.4619 | 922.0301
3207 | 77.7098 | 0.95017 | 0.99699 | 0.91725 | 202.2317 | 48.1008 | 47741.7008 | 46651.402 | 929.028 | 829.028 | 922.0789
3208 | 104.1119 | 1.0611 | 1.0262 | 0.9519 | 196.449 | 47.5785 | 51930.7644 | 50419.4366 | 1004.3887 | 904.3887 | 922.1009
3209 | 106.2322 | 1.0056 | 1.0906 | 1.0708 | 185.8136 | 46.4885 | 59439.445 | 58769.7866 | 1171.3957 | 1071.3957 | 922.1745
3210 | 106.812 | 0.95374 | 0.88005 | 0.9477 | 193.0415 | 47.2496 | 44413.5819 | 44370.7211 | 883.4144 | 783.4144 | 922.2667
3211 | 87.337 | 1.0101 | 0.98721 | 0.98536 | 197.8782 | 48.3506 | 51496.8421 | 51844.5485 | 1032.891 | 932.891 | 922.2947
3212 | 104.6873 | 0.98528 | 1.004 | 0.9517 | 185.0271 | 49.0652 | 52126.1115 | 52460.7374 | 1045.2147 | 945.2147 | 922.3071
3213 | 92.3193 | 0.93913 | 1.0189 | 0.96454 | 203.3878 | 45.2377 | 48585.9897 | 48238.2558 | 960.7651 | 860.7651 | 922.329
3214 | 71.1054 | 1.0326 | 1.0078 | 1.0464 | 161.6173 | 51.1775 | 57890.3379 | 56857.8616 | 1133.1572 | 1033.1572 | 922.3309
3215 | 95.1708 | 0.99571 | 1.0921 | 0.98632 | 190.5116 | 45.6015 | 53631.053 | 52982.6747 | 1055.6535 | 955.6535 | 922.3796
3216 | 124.3042 | 1.0734 | 0.73587 | 1.0418 | 204.5751 | 49.8849 | 45102.2776 | 44736.5447 | 890.7309 | 790.7309 | 922.4528
3217 | 102.2248 | 1.0219 | 0.66338 | 0.96304 | 200.577 | 48.3414 | 36134.4003 | 36353.6025 | 723.0721 | 623.0721 | 922.5049
3218 | 108.4902 | 0.88284 | 0.91034 | 1.0329 | 207.992 | 51.6551 | 53725.3939 | 53202.8495 | 1060.057 | 960.057 | 922.5146
3219 | 105.2557 | 1.0918 | 1.0143 | 1.1016 | 196.079 | 46.3897 | 57361.4338 | 59070.6128 | 1177.4123 | 1077.4123 | 922.518
3220 | 81.7618 | 1.0293 | 0.88373 | 0.99476 | 211.1855 | 50.141 | 48509.7935 | 49712.5762 | 990.2515 | 890.2515 | 922.5619
3221 | 107.4096 | 0.98087 | 1.0339 | 1.0111 | 203.0102 | 47.9116 | 55335.0734 | 53315.1773 | 1062.3035 | 962.3035 | 922.6084
3222 | 107.7189 | 0.90749 | 1.1528 | 1.0138 | 205.1183 | 47.2533 | 60048.7969 | 60130.026 | 1198.6005 | 1098.6005 | 922.6862
3223 | 88.6419 | 0.98156 | 0.78531 | 0.94954 | 201.1253 | 49.81 | 41677.3289 | 41562.7448 | 827.2549 | 727.2549 | 922.7298
3224 | 98.8091 | 0.95727 | 1.2821 | 0.99656 | 202.0345 | 55.0004 | 75679.5007 | 76186.9644 | 1519.7393 | 1419.7393 | 922.7343
3225 | 101.9304 | 0.99612 | 1.016 | 1.0251 | 188.8652 | 47.8531 | 54887.2108 | 55519.466 | 1106.3893 | 1006.3893 | 922.8986
3226 | 104.2497 | 0.92155 | 0.84795 | 1.008 | 194.0063 | 50.1331 | 47859.7187 | 46945.5183 | 934.9104 | 834.9104 | 922.9079
3227 | 105.3351 | 1.0679 | 1.0996 | 1.0241 | 209.4933 | 45.3349 | 56360.3852 | 56713.6285 | 1130.2726 | 1030.2726 | 923.0034
3228 | 123.6902 | 0.9842 | 0.9499 | 0.95804 | 218.4488 | 51.9515 | 53821.079 | 52890.0856 | 1053.8017 | 953.8017 | 923.108
3229 | 91.1811 | 0.99231 | 1.23 | 0.96796 | 194.1018 | 56.5635 | 72656.1087 | 74905.1247 | 1494.1025 | 1394.1025 | 923.129
3230 | 91.9833 | 1.0184 | 0.80118 | 0.95906 | 207.1032 | 54.7736 | 47425.1644 | 46594.7116 | 927.8942 | 827.8942 | 923.2254
3231 | 94.0813 | 1.0645 | 0.99017 | 0.91518 | 188.3873 | 50.6129 | 51121.9917 | 49816.4557 | 992.3291 | 892.3291 | 923.2311
3232 | 87.0917 | 1.0683 | 1.1731 | 0.97148 | 196.4532 | 47.0385 | 58181.2989 | 58176.423 | 1159.5285 | 1059.5285 | 923.2895
3233 | 102.6793 | 1.001 | 1.2985 | 0.92059 | 211.4363 | 49.9413 | 65046.0145 | 65179.3198 | 1299.5864 | 1199.5864 | 923.3885
3234 | 106.3197 | 1.0838 | 1.1452 | 1.0098 | 201.922 | 47.5453 | 60661.8923 | 61908.2462 | 1234.1649 | 1134.1649 | 923.3976
3235 | 107.7982 | 1.0099 | 1.016 | 1.0146 | 210.9524 | 51.3784 | 58770.0886 | 59813.8903 | 1192.2778 | 1092.2778 | 923.4013
3236 | 88.2569 | 1.0236 | 1.0758 | 1.0493 | 208.2075 | 49.3469 | 60366.1389 | 58376.7429 | 1163.5349 | 1063.5349 | 923.4782
3237 | 97.1083 | 1.0031 | 0.72065 | 0.98643 | 191.5105 | 53.0356 | 43058.9079 | 43019.8956 | 856.3979 | 756.3979 | 923.5181
3238 | 110.8299 | 0.98593 | 0.97257 | 0.9112 | 218.9286 | 55.0753 | 55045.2345 | 54690.9456 | 1089.8189 | 989.8189 | 923.6036
3239 | 113.1317 | 0.94234 | 1.096 | 0.90703 | 193.4519 | 50.0016 | 55230.0543 | 55913.6314 | 1114.2726 | 1014.2726 | 923.6383
3240 | 109.8283 | 1.0603 | 1.0896 | 0.99653 | 211.0786 | 50.7893 | 61272.5045 | 62016.0419 | 1236.3208 | 1136.3208 | 923.6816
3241 | 101.1556 | 0.96452 | 0.97777 | 1.045 | 205.7539 | 42.6896 | 47991.7464 | 47135.3535 | 938.7071 | 838.7071 | 923.7695
3242 | 104.5983 | 0.91315 | 1.1274 | 1.0314 | 208.1361 | 49.9707 | 63087.6524 | 63671.5593 | 1269.4312 | 1169.4312 | 923.8017
3243 | 112.2564 | 1.0591 | 1.1077 | 1.0368 | 192.0332 | 47.1012 | 59885.5353 | 59791.2685 | 1191.8254 | 1091.8254 | 923.828
3244 | 101.4072 | 0.97653 | 0.84573 | 1.0772 | 189.088 | 49.5748 | 50262.3696 | 49464.6642 | 985.2933 | 885.2933 | 923.8328
3245 | 106.6701 | 0.96415 | 0.81989 | 0.9127 | 192.2473 | 54.1611 | 46292.1553 | 46182.8922 | 919.6578 | 819.6578 | 923.8778
3246 | 88.387 | 1.0023 | 1.0369 | 0.99881 | 207.0291 | 50.5571 | 57043.7513 | 56695.2792 | 1129.9056 | 1029.9056 | 923.8988
3247 | 82.7218 | 1.0551 | 1.1093 | 1.0151 | 188.8383 | 48.0728 | 58518.2591 | 57779.8883 | 1151.5978 | 1051.5978 | 923.9459
3248 | 111.6871 | 1.0672 | 1.1068 | 0.99675 | 189.3182 | 48.2692 | 59195.2874 | 59724.4752 | 1190.4895 | 1090.4895 | 923.994
3249 | 109.045 | 1.0146 | 0.84679 | 0.90096 | 206.9484 | 52.7602 | 46296.1119 | 45050.2522 | 897.005 | 797.005 | 924.0203
3250 | 88.9231 | 0.93185 | 0.88485 | 0.99765 | 203.0296 | 48.1869 | 46733.9168 | 47166.6004 | 939.332 | 839.332 | 924.0312
3251 | 87.2944 | 1.0885 | 0.96551 | 0.99371 | 203.4149 | 50.515 | 53469.481 | 53629.7747 | 1068.5955 | 968.5955 | 924.0623
3252 | 109.1012 | 0.93495 | 0.95466 | 1.0283 | 200.3887 | 56.8632 | 61821.8836 | 61822.5471 | 1232.4509 | 1132.4509 | 924.2398
3253 | 104.1167 | 1.0121 | 1.1346 | 0.98969 | 191.2632 | 46.7114 | 57564.7521 | 56871.4358 | 1133.4287 | 1033.4287 | 924.3079
3254 | 100.2012 | 0.93897 | 1.1221 | 1.059 | 199.588 | 51.9743 | 66850.6516 | 65750.577 | 1311.0115 | 1211.0115 | 924.3907
3255 | 95.0767 | 0.9313 | 1.0748 | 1.0174 | 215.3824 | 52.6949 | 62502.1195 | 62721.7307 | 1250.4346 | 1150.4346 | 924.3929
3256 | 105.6841 | 0.92509 | 0.86128 | 0.97082 | 193.0904 | 48.574 | 45557.1141 | 45046.2483 | 896.925 | 796.925 | 924.402
3257 | 88.9843 | 1.0954 | 1.0751 | 0.94543 | 204.1974 | 46.3008 | 51778.8416 | 49857.3177 | 993.1464 | 893.1464 | 924.4305
3258 | 93.1124 | 0.9634 | 0.87957 | 0.99623 | 212.5058 | 52.4652 | 50891.6658 | 52227.8223 | 1040.5564 | 940.5564 | 924.4521
3259 | 95.8148 | 1.0503 | 0.89885 | 0.99159 | 211.4081 | 51.6707 | 51464.7027 | 50704.2816 | 1010.0856 | 910.0856 | 924.4548
3260 | 94.8794 | 0.86412 | 1.0921 | 0.9942 | 187.13 | 51.5707 | 60406.5525 | 61038.3786 | 1216.7676 | 1116.7676 | 924.5802
3261 | 94.3194 | 0.99686 | 1.0124 | 0.98175 | 208.8496 | 47.426 | 51804.0435 | 50281.1724 | 1001.6234 | 901.6234 | 924.6696
3262 | 90.446 | 0.96219 | 1.0818 | 1.0737 | 215.9971 | 46.7513 | 58587.1357 | 59510.9267 | 1186.2185 | 1086.2185 | 924.8046
3263 | 99.0654 | 1.0693 | 1.0644 | 0.94027 | 202.2339 | 49.9719 | 55507.7388 | 53869.2888 | 1073.3858 | 973.3858 | 924.9459
3264 | 96.4991 | 1.0335 | 0.98635 | 0.94034 | 199.8939 | 48.7995 | 50329.1365 | 48449.701 | 964.994 | 864.994 | 924.9491
3265 | 100.9072 | 1.0071 | 1.1871 | 1.0463 | 188.1747 | 51.6611 | 69604.7097 | 71087.0759 | 1417.7415 | 1317.7415 | 924.979
3266 | 79.4389 | 1.019 | 1.0558 | 0.94729 | 183.2542 | 50.1261 | 54373.3164 | 52617.1193 | 1048.3424 | 948.3424 | 925.0252
3267 | 98.9063 | 0.94339 | 0.96642 | 1.1086 | 191.3171 | 48.6114 | 56810.5475 | 57985.3304 | 1155.7066 | 1055.7066 | 925.0426
3268 | 104.5914 | 0.9468 | 1.0641 | 1.0355 | 216.2577 | 51.3392 | 61871.6886 | 61546.507 | 1226.9301 | 1126.9301 | 925.0588
3269 | 99.5641 | 0.91754 | 0.77818 | 0.92106 | 205.1075 | 50.1417 | 40724.7201 | 41156.5555 | 819.1311 | 719.1311 | 925.1488
3270 | 94.3429 | 0.96784 | 1.1314 | 0.92292 | 200.1971 | 48.1616 | 54886.7115 | 55895.9731 | 1113.9195 | 1013.9195 | 925.1758
3271 | 81.8874 | 0.98192 | 1.187 | 0.95548 | 203.2413 | 53.3734 | 65030.0463 | 65912.3442 | 1314.2469 | 1214.2469 | 925.2318
3272 | 99.8818 | 0.96015 | 1.0922 | 0.98783 | 202.2967 | 53.5838 | 63153.8612 | 60005.5537 | 1196.1111 | 1096.1111 | 925.2352
3273 | 102.8138 | 1.0742 | 1.0557 | 0.97015 | 193.1373 | 47.4558 | 54038.8378 | 51847.1645 | 1032.9433 | 932.9433 | 925.2596
3274 | 96.5552 | 0.93497 | 0.87429 | 0.9577 | 195.3131 | 49.4562 | 46070.172 | 45222.6256 | 900.4525 | 800.4525 | 925.3393
3275 | 107.0197 | 0.94573 | 0.96455 | 0.97988 | 192.765 | 50.2461 | 52767.9455 | 53172.4248 | 1059.4485 | 959.4485 | 925.3869
3276 | 104.6624 | 0.98367 | 1.0293 | 1.0098 | 204.2125 | 50.5721 | 57976.9439 | 57340.3536 | 1142.8071 | 1042.8071 | 925.3899
3277 | 96.8505 | 1.0164 | 0.92352 | 1.0287 | 194.3315 | 49.8216 | 52432.8188 | 54311.4896 | 1082.2298 | 982.2298 | 925.3925
3278 | 102.4446 | 1.0511 | 1.1676 | 1.0213 | 190.0875 | 56.3996 | 73514.3086 | 72601.0128 | 1448.0203 | 1348.0203 | 925.4809
3279 | 107.8622 | 0.97418 | 0.98764 | 1.0315 | 197.3784 | 48.7871 | 55023.5847 | 54644.5662 | 1088.8913 | 988.8913 | 925.5128
3280 | 107.1093 | 0.96853 | 0.92437 | 0.92011 | 207.2466 | 52.4205 | 50229.6569 | 51869.5591 | 1033.3912 | 933.3912 | 925.5384
3281 | 106.7404 | 0.89997 | 0.90501 | 1.0579 | 198.8386 | 48.5498 | 51342.3656 | 49718.6993 | 990.374 | 890.374 | 925.6216
3282 | 117.9426 | 1.1028 | 0.98918 | 1.133 | 218.0272 | 49.0335 | 61551.8325 | 63101.8336 | 1258.0367 | 1158.0367 | 925.743
3283 | 91.8033 | 1.0259 | 1.0618 | 0.91817 | 205.8085 | 48.59 | 52154.1049 | 52274.1677 | 1041.4834 | 941.4834 | 925.8396
3284 | 89.3651 | 0.99524 | 0.98887 | 0.98132 | 203.8208 | 53.783 | 57177.9228 | 57145.9007 | 1138.918 | 1038.918 | 925.8552
3285 | 118.6234 | 0.98142 | 0.79506 | 1.0604 | 195.7041 | 50.7131 | 48853.2297 | 49042.5165 | 976.8503 | 876.8503 | 925.8776
3286 | 107.1586 | 0.95448 | 1.0317 | 0.96837 | 216.3664 | 51.2409 | 56650.3624 | 57306.0892 | 1142.1218 | 1042.1218 | 925.9088
3287 | 103.0801 | 1.0001 | 0.95996 | 1.0078 | 207.254 | 52.1062 | 55986.6512 | 56341.0428 | 1122.8209 | 1022.8209 | 925.9102
3288 | 105.3249 | 0.94176 | 1.064 | 1.0466 | 206.0073 | 51.9344 | 63190.3478 | 63685.5554 | 1269.7111 | 1169.7111 | 926.0187
3289 | 113.4749 | 1.0434 | 0.92166 | 0.96979 | 210.6938 | 54.5816 | 55458.8991 | 53419.9899 | 1064.3998 | 964.3998 | 926.0271
3290 | 102.0067 | 1.0786 | 1.1237 | 0.96878 | 206.3139 | 51.546 | 61991.3121 | 62625.9613 | 1248.5192 | 1148.5192 | 926.0383
3291 | 117.776 | 0.93508 | 1.1825 | 1.0868 | 218.1399 | 49.2311 | 68908.1446 | 67067.5086 | 1337.3502 | 1237.3502 | 926.137
3292 | 98.2361 | 1.0006 | 0.98651 | 0.94146 | 194.1893 | 50.9654 | 52538.3442 | 52567.6112 | 1047.3522 | 947.3522 | 926.2355
3293 | 92.4502 | 1.0016 | 1.0716 | 1.1527 | 215.8128 | 48.7894 | 64999.4688 | 64672.7628 | 1289.4553 | 1189.4553 | 926.3276
3294 | 100.2899 | 1.0237 | 1.037 | 1.1014 | 200.7734 | 51.3394 | 64109.8028 | 64207.9276 | 1280.1586 | 1180.1586 | 926.3348
3295 | 109.8577 | 1.0397 | 1.1774 | 1.0672 | 202.5142 | 47.9928 | 65987.0631 | 65636.9027 | 1308.7381 | 1208.7381 | 926.3638
3296 | 106.8372 | 0.96407 | 1.0273 | 1.0309 | 215.7977 | 46.8014 | 54599.9813 | 56481.1289 | 1125.6226 | 1025.6226 | 926.4653
3297 | 112.9813 | 0.99243 | 1.0345 | 0.99333 | 204.5205 | 52.5667 | 60115.5143 | 58573.7427 | 1167.4749 | 1067.4749 | 926.5145
3298 | 97.7132 | 1.0827 | 1.0757 | 1.0671 | 190.0333 | 50.71 | 63767.2275 | 63517.7844 | 1266.3557 | 1166.3557 | 926.5376
3299 | 94.1045 | 0.96584 | 0.97613 | 0.99729 | 201.0388 | 46.903 | 50123.5936 | 49664.9194 | 989.2984 | 889.2984 | 926.5723
3300 | 99.0837 | 0.95039 | 0.90679 | 1.0777 | 188.1885 | 48.9524 | 52636.0213 | 51280.3783 | 1021.6076 | 921.6076 | 926.635
3301 | 105.2645 | 1.0933 | 0.93968 | 1.0576 | 222.1957 | 49.696 | 55329.3651 | 53196.4335 | 1059.9287 | 959.9287 | 926.6455
3302 | 115.4317 | 1.0421 | 0.82619 | 0.97373 | 201.506 | 53.5488 | 49722.2546 | 50396.6477 | 1003.933 | 903.933 | 926.673
3303 | 102.6126 | 1.0872 | 0.99502 | 1.1139 | 214.7919 | 49.7785 | 60941.1453 | 59930.3916 | 1194.6078 | 1094.6078 | 926.7562
3304 | 94.4086 | 0.92018 | 1.0291 | 1.0146 | 193.4476 | 49.5683 | 56257.3191 | 55068.5162 | 1097.3703 | 997.3703 | 926.7715
3305 | 110.9612 | 1.0441 | 0.90699 | 0.98439 | 202.7686 | 52.4864 | 53144.6855 | 52472.9776 | 1045.4596 | 945.4596 | 926.7973
3306 | 100.2097 | 0.93807 | 1.0333 | 1.035 | 210.7364 | 50.2504 | 58674.9169 | 59147.9038 | 1178.9581 | 1078.9581 | 926.8448
3307 | 109.7976 | 1.071 | 1.0366 | 0.92899 | 217.4256 | 52.3669 | 56805.4772 | 58718.1422 | 1170.3628 | 1070.3628 | 926.9038
3308 | 95.244 | 1.0485 | 0.72597 | 1.0628 | 213.4132 | 49.8336 | 43637.8709 | 43581.5424 | 867.6308 | 767.6308 | 926.9606
3309 | 115.2088 | 1.0425 | 1.1245 | 1.0906 | 213.7603 | 48.779 | 65893.0914 | 66795.5922 | 1331.9118 | 1231.9118 | 926.9636
3310 | 109.5599 | 0.9793 | 0.89237 | 1.0321 | 178.3143 | 51.4947 | 53128.4193 | 54962.2609 | 1095.2452 | 995.2452 | 927.0275
3311 | 88.2519 | 1.005 | 1.0983 | 1.0439 | 198.5965 | 48.0828 | 59593.7265 | 58300.1782 | 1162.0036 | 1062.0036 | 927.0492
3312 | 99.6167 | 0.98701 | 1.3282 | 0.92393 | 198.9625 | 46.3278 | 61605.7034 | 61609.5124 | 1228.1902 | 1128.1902 | 927.0569
3313 | 110.3799 | 1.0202 | 1.2802 | 0.92995 | 218.8717 | 52.7997 | 69023.3463 | 66811.789 | 1332.2358 | 1232.2358 | 927.0923
3314 | 113.4834 | 1.0369 | 1.1627 | 1.0189 | 209.3558 | 55.6453 | 72676.3332 | 73886.5074 | 1473.7301 | 1373.7301 | 927.1049
3315 | 93.7952 | 0.99163 | 0.8287 | 1.0152 | 184.4949 | 44.245 | 41523.0711 | 41495.1111 | 825.9022 | 725.9022 | 927.1623
3316 | 101.2408 | 1.0363 | 1.0027 | 0.96108 | 212.7058 | 48.6653 | 52215.6381 | 51762.5261 | 1031.2505 | 931.2505 | 927.2229
3317 | 99.5955 | 1.0298 | 0.94867 | 0.9546 | 189.6425 | 54.6062 | 55241.5453 | 53786.5906 | 1071.7318 | 971.7318 | 927.2245
3318 | 99.3533 | 1.0292 | 0.93322 | 0.99282 | 195.8258 | 49.4391 | 51057.3909 | 50374.0256 | 1003.4805 | 903.4805 | 927.2707
3319 | 112.1083 | 1.0524 | 0.98919 | 1.0071 | 201.0462 | 49.8725 | 55768.0712 | 53448.9188 | 1064.9784 | 964.9784 | 927.2707
3320 | 86.1843 | 0.97765 | 0.92153 | 1.0868 | 200.9633 | 49.3815 | 53818.0183 | 54564.8279 | 1087.2966 | 987.2966 | 927.4034
3321 | 113.1344 | 0.93888 | 1.1746 | 0.97361 | 210.1214 | 48.6708 | 61038.9881 | 60225.9558 | 1200.5191 | 1100.5191 | 927.4268
3322 | 102.1215 | 1.025 | 1.2197 | 0.99333 | 190.3554 | 52.0061 | 68641.7852 | 67479.8816 | 1345.5976 | 1245.5976 | 927.428
3323 | 102.9551 | 1.0516 | 1.2282 | 1.0423 | 201.0535 | 47.8059 | 66577.9599 | 66396.3173 | 1323.9263 | 1223.9263 | 927.5239
3324 | 90.9543 | 0.96237 | 1.2641 | 0.9898 | 206.4323 | 46.6403 | 62645.874 | 61655.9534 | 1229.1191 | 1129.1191 | 927.5385
3325 | 99.7324 | 0.99069 | 1.2512 | 0.95144 | 215.3686 | 51.9059 | 67136.2601 | 66456.0513 | 1325.121 | 1225.121 | 927.5437
3326 | 96.0351 | 0.99315 | 0.97071 | 1.0365 | 222.9423 | 53.4722 | 59124.7945 | 60041.9651 | 1196.8393 | 1096.8393 | 927.5598
3327 | 107.6842 | 0.96921 | 0.80495 | 1.0144 | 195.3063 | 45.5925 | 42182.1185 | 41736.3857 | 830.7277 | 730.7277 | 927.5601
3328 | 104.3432 | 1.0356 | 0.98084 | 1.0024 | 206.656 | 50.6265 | 55453.9318 | 56612.6159 | 1128.2523 | 1028.2523 | 927.5765
3329 | 102.7193 | 1.0124 | 0.96103 | 0.90489 | 191.9811 | 50.2105 | 49077.8862 | 50311.8015 | 1002.236 | 902.236 | 927.5891
3330 | 101.9945 | 1.0305 | 0.87114 | 1.0837 | 205.0747 | 52.8371 | 55640.2164 | 54944.6042 | 1094.8921 | 994.8921 | 927.6421
3331 | 102.9203 | 1.049 | 1.021 | 0.98311 | 196.1262 | 53.3266 | 59482.1644 | 58228.1162 | 1160.5623 | 1060.5623 | 927.6782
3332 | 108.353 | 0.99644 | 0.80479 | 1.0453 | 200.3308 | 51.1342 | 48736.5351 | 49219.0194 | 980.3804 | 880.3804 | 927.7227
3333 | 102.5804 | 0.9827 | 1.2033 | 0.96519 | 206.3855 | 44.9536 | 56947.6029 | 56653.9772 | 1129.0795 | 1029.0795 | 927.7413
3334 | 109.4891 | 1.0269 | 0.97593 | 1.0014 | 185.5324 | 49.9323 | 54597.9662 | 54690.5525 | 1089.8111 | 989.8111 | 927.8044
3335 | 106.9302 | 0.99335 | 0.91453 | 0.97024 | 183.1872 | 52.8179 | 52659.3395 | 53121.7723 | 1058.4354 | 958.4354 | 927.9408
3336 | 89.4972 | 1.0523 | 1.0475 | 1.0376 | 197.5276 | 43.1939 | 51210.501 | 50303.5593 | 1002.0712 | 902.0712 | 927.9438
3337 | 103.31 | 1.0166 | 1.1248 | 0.9668 | 216.9364 | 49.5214 | 59270.5956 | 59812.5969 | 1192.2519 | 1092.2519 | 927.9445
3338 | 109.5257 | 0.96989 | 1.0489 | 0.92353 | 204.6261 | 53.9175 | 58160.5695 | 57536.1107 | 1146.7222 | 1046.7222 | 928.0367
3339 | 115.063 | 1.0391 | 1.1575 | 0.96635 | 215.6908 | 47.0208 | 58430.8866 | 55707.5098 | 1110.1502 | 1010.1502 | 928.0979
3340 | 101.4486 | 0.96365 | 0.71709 | 1.0565 | 198.9473 | 51.3537 | 44124.0163 | 44851.8534 | 893.0371 | 793.0371 | 928.1601
3341 | 92.1522 | 1.0141 | 1.2928 | 0.97782 | 198.1797 | 50.7361 | 69077.1487 | 70099.2805 | 1397.9856 | 1297.9856 | 928.1684
3342 | 101.614 | 0.93417 | 1.5079 | 0.93725 | 211.7926 | 47.8036 | 72308.1233 | 71720.8224 | 1430.4164 | 1330.4164 | 928.2101
3343 | 112.1783 | 0.97181 | 1.1545 | 0.95199 | 212.6479 | 52.3799 | 63491.7416 | 64512.6751 | 1286.2535 | 1186.2535 | 928.3118
3344 | 110.8618 | 1.0216 | 1.0497 | 0.9634 | 217.7439 | 49.0624 | 55390.3466 | 55474.0275 | 1105.4805 | 1005.4805 | 928.3139
3345 | 102.2887 | 1.0592 | 1.1179 | 0.94513 | 188.6467 | 50.1324 | 58586.0163 | 59951.9168 | 1195.0383 | 1095.0383 | 928.3381
3346 | 110.4665 | 0.92323 | 1.0096 | 1.0065 | 189.1625 | 50.0887 | 56195.0706 | 55493.4909 | 1105.8698 | 1005.8698 | 928.3393
3347 | 104.3703 | 0.98714 | 0.98072 | 1.0155 | 193.914 | 47.4336 | 52319.9748 | 52110.2791 | 1038.2056 | 938.2056 | 928.3625
3348 | 103.8821 | 0.95836 | 0.84366 | 0.96367 | 204.5674 | 44.1313 | 40476.9417 | 40720.7448 | 810.4149 | 710.4149 | 928.43
3349 | 110.5071 | 0.90243 | 0.99847 | 0.94427 | 188.2154 | 47.7253 | 49944.2343 | 49086.8266 | 977.7365 | 877.7365 | 928.4893
3350 | 100.2197 | 1.0507 | 1.1012 | 0.98093 | 207.7443 | 52.3343 | 62251.1135 | 60516.099 | 1206.322 | 1106.322 | 928.5222
3351 | 99.6283 | 0.95918 | 0.87374 | 0.9979 | 213.8514 | 50.0652 | 48650.2196 | 49730.7503 | 990.615 | 890.615 | 928.5346
3352 | 98.4377 | 0.93895 | 1.0069 | 1.0092 | 211.5894 | 45.7585 | 50937.3456 | 51674.3914 | 1029.4878 | 929.4878 | 928.5435
3353 | 97.3512 | 1.0275 | 0.69181 | 0.94673 | 200.7246 | 50.4324 | 38276.6836 | 37761.3528 | 751.2271 | 651.2271 | 928.5464
3354 | 93.3067 | 1.0043 | 1.1916 | 0.99697 | 208.8259 | 52.5857 | 67607.0395 | 68892.4934 | 1373.8499 | 1273.8499 | 928.5554
3355 | 88.7532 | 1.0155 | 0.98862 | 1.0332 | 212.8305 | 50.8353 | 56718.7077 | 59406.431 | 1184.1286 | 1084.1286 | 928.5674
3356 | 105.917 | 0.92105 | 0.9083 | 1.0782 | 194.0803 | 48.2103 | 52110.5013 | 52258.6996 | 1041.174 | 941.174 | 928.6522
3357 | 105.1975 | 1.0171 | 1.0402 | 1.0123 | 201.5338 | 47.3935 | 55178.0852 | 55038.7473 | 1096.7749 | 996.7749 | 928.6802
3358 | 92.0362 | 1.0586 | 1.0598 | 0.96466 | 198.7371 | 53.0836 | 59638.1591 | 57994.5678 | 1155.8914 | 1055.8914 | 928.7408
3359 | 103.193 | 1.017 | 1.1336 | 0.94116 | 192.6051 | 52.6801 | 61928.2498 | 61629.5086 | 1228.5902 | 1128.5902 | 928.859
3360 | 85.0286 | 1.0409 | 0.61788 | 1.0136 | 205.5498 | 48.1473 | 34620.4113 | 35732.2077 | 710.6442 | 610.6442 | 929.0348
3361 | 117.5174 | 0.91122 | 0.85376 | 0.96371 | 214.1211 | 46.7598 | 43694.3068 | 44246.1813 | 880.9236 | 780.9236 | 929.0636
3362 | 98.2413 | 1.0011 | 1.0777 | 1.0207 | 189.3982 | 49.0639 | 58983.2898 | 58873.8118 | 1173.4762 | 1073.4762 | 929.0918
3363 | 105.8169 | 1.0505 | 1.0818 | 0.95685 | 197.8088 | 49.9789 | 57486.1159 | 57440.7256 | 1144.8145 | 1044.8145 | 929.1223
3364 | 105.3431 | 0.9004 | 1.0289 | 0.98829 | 190.1515 | 51.8137 | 57793.329 | 56225.1319 | 1120.5026 | 1020.5026 | 929.1615
3365 | 86.7175 | 1.037 | 1.0877 | 0.91942 | 190.7148 | 47.7696 | 52260.0253 | 52412.3647 | 1044.2473 | 944.2473 | 929.1782
3366 | 96.3892 | 0.98489 | 1.0251 | 1.0142 | 202.7138 | 46.1263 | 52534.1669 | 52234.6692 | 1040.6934 | 940.6934 | 929.3313
3367 | 86.9171 | 0.95448 | 1.0125 | 1.0938 | 203.794 | 50.8278 | 60714.3701 | 60564.6034 | 1207.2921 | 1107.2921 | 929.3415
3368 | 91.4581 | 1.1054 | 1.1626 | 1.0594 | 193.033 | 50.0837 | 66940.5954 | 66328.4173 | 1322.5683 | 1222.5683 | 929.3445
3369 | 91.4678 | 0.94622 | 0.95564 | 1.0056 | 202.8635 | 50.4798 | 53084.6225 | 53331.0666 | 1062.6213 | 962.6213 | 929.355
3370 | 94.3042 | 0.91127 | 0.91904 | 0.92976 | 198.9801 | 49.9232 | 47147.7813 | 46439.5019 | 924.79 | 824.79 | 929.3995
3371 | 109.3646 | 0.9582 | 0.85787 | 0.91903 | 195.0442 | 46.8002 | 41997.0215 | 42110.1471 | 838.2029 | 738.2029 | 929.4245
3372 | 93.359 | 0.97641 | 1.1864 | 0.96795 | 197.2918 | 47.9491 | 59630.6858 | 59858.9066 | 1193.1781 | 1093.1781 | 929.4878
3373 | 119.1656 | 0.95625 | 1.3589 | 1.0133 | 189.2156 | 52.6511 | 78688.3383 | 80377.0873 | 1603.5417 | 1503.5417 | 929.5482
3374 | 107.5799 | 0.97553 | 1.0568 | 0.96466 | 187.9205 | 51.9757 | 58629.4978 | 58243.7597 | 1160.8752 | 1060.8752 | 929.5649
3375 | 104.7687 | 1.0633 | 1.0903 | 1.0199 | 180.3927 | 49.6787 | 60954.9113 | 60098.7012 | 1197.974 | 1097.974 | 929.6607
3376 | 94.4571 | 1.0864 | 0.84478 | 1.0056 | 196.3466 | 47.4815 | 45405.1454 | 47157.1989 | 939.144 | 839.144 | 929.7273
3377 | 107.2991 | 0.93761 | 1.0691 | 0.99249 | 211.7827 | 45.8561 | 53481.6513 | 52308.1585 | 1042.1632 | 942.1632 | 929.7514
3378 | 93.1461 | 0.98066 | 1.1732 | 1.0693 | 193.6296 | 55.2408 | 74539.932 | 74054.211 | 1477.0842 | 1377.0842 | 929.9233
3379 | 110.4351 | 1.0547 | 0.73322 | 0.9246 | 206.3081 | 51.0896 | 40792.622 | 42208.6593 | 840.1732 | 740.1732 | 929.9454
3380 | 92.2101 | 0.98383 | 1.0321 | 1.0269 | 211.6187 | 51.248 | 59175.4672 | 61632.2622 | 1228.6452 | 1128.6452 | 929.9495
3381 | 95.5638 | 1.0062 | 1.2377 | 1.0281 | 208.4623 | 48.9473 | 67202.3543 | 65423.1082 | 1304.4622 | 1204.4622 | 930.0188
3382 | 92.811 | 0.9522 | 0.99992 | 1.0291 | 201.0758 | 50.0137 | 56084.6287 | 55084.4755 | 1097.6895 | 997.6895 | 930.0868
3383 | 103.7927 | 1.0107 | 0.79129 | 1.0049 | 206.1264 | 49.2884 | 44569.4767 | 44266.3651 | 881.3273 | 781.3273 | 930.1141
3384 | 92.2801 | 1.0365 | 1.1312 | 0.94041 | 185.4144 | 50.5253 | 58767.3022 | 57333.7143 | 1142.6743 | 1042.6743 | 930.1213
3385 | 110.3485 | 0.99133 | 0.92682 | 0.94058 | 219.7275 | 49.8649 | 49143.9421 | 49853.6071 | 993.0721 | 893.0721 | 930.1304
3386 | 107.9135 | 1.031 | 1.0716 | 1.0327 | 203.3065 | 47.4829 | 58031.9069 | 59260.8629 | 1181.2173 | 1081.2173 | 930.1346
3387 | 115.8455 | 1.0395 | 0.93104 | 0.9164 | 205.2675 | 46.9128 | 45880.4827 | 45093.4581 | 897.8692 | 797.8692 | 930.1351
3388 | 81.6457 | 1.0015 | 1.0701 | 0.97915 | 204.8835 | 49.4469 | 56057.524 | 55624.7055 | 1108.4941 | 1008.4941 | 930.1614
3389 | 100.3826 | 1.0175 | 0.86651 | 0.9455 | 196.8086 | 50.6127 | 46832.8918 | 45852.4452 | 913.0489 | 813.0489 | 930.1779
3390 | 92.9565 | 1.0447 | 0.8713 | 0.99033 | 208.7437 | 48.5543 | 46820.3134 | 46011.9733 | 916.2395 | 816.2395 | 930.1782
3391 | 89.4549 | 1.0493 | 0.8207 | 1.0107 | 206.3526 | 52.2027 | 48407.3017 | 48364.7744 | 963.2955 | 863.2955 | 930.2253
3392 | 90.0775 | 0.89387 | 1.2538 | 1.0503 | 195.5905 | 51.2288 | 71781.85 | 70591.5605 | 1407.8312 | 1307.8312 | 930.2265
3393 | 117.0443 | 0.96201 | 1.1395 | 0.94002 | 195.6872 | 44.7987 | 53226.0056 | 52915.6489 | 1054.313 | 954.313 | 930.3532
3394 | 83.2169 | 1.0359 | 1.1635 | 1.0236 | 203.9063 | 46.697 | 59845.907 | 62140.1193 | 1238.8024 | 1138.8024 | 930.3879
3395 | 112.9057 | 1.0179 | 1.1452 | 0.97998 | 188.2052 | 48.9874 | 60796.4596 | 61497.1568 | 1225.9431 | 1125.9431 | 930.3956
3396 | 107.0212 | 1.0354 | 1.1262 | 1.0361 | 184.837 | 49.2847 | 63152.2709 | 63916.4625 | 1274.3293 | 1174.3293 | 930.3998
3397 | 80.548 | 0.95487 | 0.855 | 1.0059 | 193.3408 | 48.3381 | 45486.15 | 45213.2301 | 900.2646 | 800.2646 | 930.4351
3398 | 102.3168 | 1.0552 | 0.93186 | 1.0404 | 197.293 | 51.7731 | 55979.6697 | 55010.168 | 1096.2034 | 996.2034 | 930.5105
3399 | 111.3471 | 0.98765 | 0.83474 | 0.97155 | 196.7591 | 51.2088 | 47358.5222 | 44343.9787 | 882.8796 | 782.8796 | 930.5331
3400 | 111.9191 | 1.0157 | 1.1647 | 1.092 | 185.7174 | 53.5349 | 74359.2212 | 73475.7745 | 1465.5155 | 1365.5155 | 930.6208
3401 | 98.5212 | 0.99313 | 0.87012 | 0.87506 | 203.5397 | 52.5795 | 45382.7165 | 44636.7128 | 888.7343 | 788.7343 | 930.6519
3402 | 95.7495 | 0.92816 | 0.71568 | 1.0461 | 193.2058 | 52.8198 | 44432.4797 | 45424.717 | 904.4943 | 804.4943 | 930.7673
3403 | 105.6556 | 1.0497 | 1.1067 | 1.0515 | 187.2667 | 48.8916 | 62505.2141 | 62163.6884 | 1239.2738 | 1139.2738 | 930.7809
3404 | 110.7018 | 1.0585 | 0.9601 | 0.98952 | 203.5259 | 49.2885 | 52805.1259 | 53499.322 | 1065.9864 | 965.9864 | 930.8673
3405 | 71.6931 | 1.0189 | 0.81176 | 0.99279 | 197.0276 | 49.8646 | 44025.845 | 45100.7781 | 898.0156 | 798.0156 | 930.9109
3406 | 94.4121 | 1.0605 | 0.88939 | 1.0176 | 188.2655 | 48.6167 | 49056.9055 | 47566.9804 | 947.3396 | 847.3396 | 930.9546
3407 | 82.9656 | 0.99707 | 1.1195 | 1.0271 | 202.349 | 49.375 | 61057.2413 | 62406.7677 | 1244.1354 | 1144.1354 | 930.9876
3408 | 119.4409 | 1.0044 | 0.92561 | 0.94332 | 198.5244 | 49.3753 | 49233.4445 | 49625.7544 | 988.5151 | 888.5151 | 931.0011
3409 | 90.7072 | 0.95072 | 1.1656 | 0.98316 | 200.7055 | 53.3693 | 65965.3278 | 64579.2245 | 1287.5845 | 1187.5845 | 931.0078
3410 | 94.3924 | 1.0391 | 0.86275 | 0.97044 | 217.0127 | 47.82 | 44944.3087 | 45707.4705 | 910.1494 | 810.1494 | 931.1468
3411 | 99.9202 | 1.0031 | 1.02 | 0.99584 | 200.0486 | 47.4214 | 53122.9611 | 53920.6766 | 1074.4135 | 974.4135 | 931.1564
3412 | 102.3086 | 1.0173 | 0.99044 | 0.94818 | 191.9166 | 49.7585 | 52099.9162 | 51380.1783 | 1023.6036 | 923.6036 | 931.1685
3413 | 96.3796 | 0.98015 | 1.3549 | 1.0761 | 203.5215 | 48.8273 | 76007.7336 | 78396.6354 | 1563.9327 | 1463.9327 | 931.1915
3414 | 93.5134 | 0.98351 | 1.1047 | 0.94583 | 220.6216 | 50.197 | 57284.2798 | 56469.5037 | 1125.3901 | 1025.3901 | 931.2104
3415 | 94.6909 | 1.0418 | 0.9844 | 0.94694 | 192.9158 | 43.9571 | 45504.8544 | 44291.7919 | 881.8358 | 781.8358 | 931.2153
3416 | 84.2674 | 1.0175 | 0.90492 | 0.99042 | 197.8729 | 52.1521 | 51410.7379 | 51350.1716 | 1023.0034 | 923.0034 | 931.2505
3417 | 86.6051 | 1.0156 | 1.0058 | 0.98746 | 221.2433 | 58.9769 | 63984.3227 | 64247.0317 | 1280.9406 | 1180.9406 | 931.2516
3418 | 86.9873 | 1.0683 | 0.84227 | 1.023 | 192.3586 | 44.6666 | 42829.9781 | 43594.8718 | 867.8974 | 767.8974 | 931.3354
3419 | 111.9759 | 0.97679 | 0.96775 | 0.99039 | 201.0361 | 45.2069 | 48473.9797 | 50428.2346 | 1004.5647 | 904.5647 | 931.3752
3420 | 104.412 | 1.0053 | 1.0231 | 0.92172 | 208.4322 | 51.8765 | 54573.4282 | 54207.2491 | 1080.145 | 980.145 | 931.3988
3421 | 107.5487 | 1.004 | 0.96826 | 0.99323 | 189.0337 | 47.1854 | 50662.9901 | 51092.2102 | 1017.8442 | 917.8442 | 931.4097
3422 | 90.8218 | 0.99713 | 0.83717 | 1.0348 | 184.1032 | 50.5846 | 48585.0474 | 48751.6612 | 971.0332 | 871.0332 | 931.4886
3423 | 91.1015 | 1.0341 | 1.0723 | 0.97264 | 190.8689 | 49.2784 | 56227.116 | 55094.0891 | 1097.8818 | 997.8818 | 931.5296
3424 | 110.2885 | 1.0266 | 1.2101 | 0.91863 | 185.3563 | 48.8405 | 60007.4468 | 61119.693 | 1218.3939 | 1118.3939 | 931.5484
3425 | 105.5998 | 1.0431 | 1.1031 | 1.0683 | 205.7855 | 53.3776 | 68991.0193 | 66350.0289 | 1323.0006 | 1223.0006 | 931.5724
3426 | 97.4555 | 1.0481 | 0.94931 | 0.97596 | 189.7987 | 51.0045 | 52654.5162 | 51601.8346 | 1028.0367 | 928.0367 | 931.7078
3427 | 92.2216 | 0.88129 | 0.95509 | 1.0119 | 215.4557 | 47.2071 | 49676.1494 | 49939.6304 | 994.7926 | 894.7926 | 931.7221
3428 | 99.6209 | 0.99299 | 0.81065 | 0.99718 | 196.5186 | 47.6898 | 43464.6054 | 44663.3997 | 889.268 | 789.268 | 931.7255
3429 | 101.4309 | 0.98584 | 1.1517 | 0.97439 | 191.072 | 49.1712 | 60287.9593 | 60680.5205 | 1209.6104 | 1109.6104 | 931.7696
3430 | 98.6227 | 1.0122 | 0.97498 | 0.97094 | 191.2873 | 50.9554 | 53514.6713 | 52192.87 | 1039.8574 | 939.8574 | 931.8143
3431 | 96.5232 | 0.96201 | 1.1581 | 0.97552 | 211.443 | 53.0449 | 65065.8647 | 66521.7044 | 1326.4341 | 1226.4341 | 931.8535
3432 | 107.7535 | 1.0841 | 0.95205 | 1.0274 | 199.7736 | 48.8779 | 53718.4313 | 54738.0581 | 1090.7612 | 990.7612 | 931.8847
3433 | 81.4058 | 0.98757 | 1.2176 | 0.95501 | 212.8128 | 50.6166 | 63139.4248 | 62937.5498 | 1254.751 | 1154.751 | 931.8931
3434 | 97.5348 | 0.94078 | 0.97941 | 1.0316 | 210.4986 | 48.6378 | 53812.9199 | 53385.0272 | 1063.7005 | 963.7005 | 931.8991
3435 | 90.2476 | 0.98581 | 1.0275 | 0.9823 | 188.8152 | 51.1764 | 56393.8726 | 54655.4738 | 1089.1095 | 989.1095 | 931.9013
3436 | 101.5646 | 0.93619 | 1.0684 | 1.0005 | 188.4269 | 48.638 | 56801.3072 | 58529.0862 | 1166.5817 | 1066.5817 | 931.9737
3437 | 97.377 | 0.99799 | 1.1151 | 0.943 | 215.8459 | 51.014 | 58817.2929 | 58785.0599 | 1171.7012 | 1071.7012 | 931.9996
3438 | 120.4723 | 1.0045 | 1.1297 | 0.99841 | 196.8186 | 48.8108 | 61158.553 | 58570.0091 | 1167.4002 | 1067.4002 | 932.0553
3439 | 96.0556 | 1.0041 | 0.77725 | 0.96605 | 198.7749 | 47.2911 | 40268.7025 | 40563.6098 | 807.2722 | 707.2722 | 932.0699
3440 | 117.2269 | 0.98021 | 0.9537 | 1.0612 | 192.4794 | 56.4448 | 63803.4987 | 64441.1626 | 1284.8233 | 1184.8233 | 932.0702
3441 | 108.9415 | 0.97434 | 1.0499 | 1.0425 | 181.9669 | 45.6214 | 54958.1105 | 54243.7483 | 1080.875 | 980.875 | 932.1151
3442 | 116.7356 | 0.91915 | 1.0123 | 0.98557 | 210.2727 | 44.1717 | 49018.4336 | 47062.1623 | 937.2432 | 837.2432 | 932.1524
3443 | 91.5111 | 0.99242 | 0.86408 | 0.91956 | 210.7199 | 52.8973 | 47045.7322 | 48323.5142 | 962.4703 | 862.4703 | 932.1646
3444 | 88.4651 | 1.0671 | 1.1305 | 0.94198 | 216.8089 | 49.9766 | 58153.2042 | 57881.9463 | 1153.6389 | 1053.6389 | 932.1878
3445 | 109.7416 | 0.8966 | 1.0272 | 1.0196 | 205.5351 | 53.8908 | 61949.9685 | 63075.0558 | 1257.5011 | 1157.5011 | 932.1889
3446 | 104.6231 | 0.94926 | 1.1138 | 0.93318 | 188.5773 | 45.2643 | 51732.0633 | 51381.9173 | 1023.6383 | 923.6383 | 932.1965
3447 | 119.588 | 1.0686 | 0.96492 | 0.97171 | 212.3969 | 52.648 | 56304.7355 | 55914.2305 | 1114.2846 | 1014.2846 | 932.2145
3448 | 113.17 | 1.0938 | 1.3614 | 0.96126 | 188.5119 | 51.096 | 73379.6296 | 73931.7798 | 1474.6356 | 1374.6356 | 932.223
3449 | 92.3625 | 0.99247 | 0.85735 | 1.0016 | 217.1825 | 47.5575 | 45416.4198 | 43099.7796 | 857.9956 | 757.9956 | 932.2296
3450 | 115.268 | 1.0421 | 1.0082 | 0.91864 | 202.7206 | 47.3648 | 49759.0876 | 48487.2735 | 965.7455 | 865.7455 | 932.255
3451 | 109.4073 | 0.98676 | 0.98498 | 0.9761 | 207.8621 | 50.7145 | 54441.9144 | 54031.4301 | 1076.6286 | 976.6286 | 932.4335
3452 | 98.0087 | 0.95844 | 0.77453 | 1.0399 | 215.0684 | 48.1534 | 43521.3534 | 43686.5293 | 869.7306 | 769.7306 | 932.5478
3453 | 85.9888 | 1.0039 | 1.0964 | 0.96605 | 186.6631 | 50.6921 | 58256.3274 | 60055.2662 | 1197.1053 | 1097.1053 | 932.5487
3454 | 98.2111 | 1.0606 | 0.79524 | 0.98688 | 187.7836 | 53.2315 | 47508.9561 | 47884.6394 | 953.6928 | 853.6928 | 932.5646
3455 | 117.7194 | 0.94565 | 0.74015 | 0.96115 | 181.0993 | 53.4048 | 44117.9319 | 44216.2143 | 880.3243 | 780.3243 | 932.5783
3456 | 108.0425 | 1.0377 | 1.0197 | 1.0286 | 214.0143 | 46.7895 | 54539.9645 | 54712.8243 | 1090.2565 | 990.2565 | 932.5889
3457 | 82.6888 | 0.98121 | 0.79466 | 0.8919 | 186.0592 | 47.2666 | 37521.548 | 36785.5273 | 731.7105 | 631.7105 | 932.5891
3458 | 100.1054 | 1.0855 | 1.1874 | 1.013 | 213.1979 | 51.0338 | 67141.2552 | 68643.004 | 1368.8601 | 1268.8601 | 932.6609
3459 | 104.2412 | 0.96746 | 1.1281 | 0.93391 | 187.5953 | 49.2311 | 57018.7376 | 55669.0378 | 1109.3808 | 1009.3808 | 932.7791
3460 | 101.3504 | 0.96822 | 1.1457 | 1.0241 | 192.2095 | 47.425 | 60489.8083 | 60341.4277 | 1202.8286 | 1102.8286 | 932.8372
3461 | 93.6803 | 1.0387 | 0.56464 | 1.012 | 188.4482 | 54.5504 | 36667.1364 | 36902.3746 | 734.0475 | 634.0475 | 932.8388
3462 | 102.7771 | 1.0186 | 0.99261 | 0.98657 | 197.5012 | 54.79 | 59587.744 | 60704.107 | 1210.0821 | 1110.0821 | 932.8805
3463 | 105.635 | 0.96664 | 0.96954 | 0.96637 | 203.0267 | 46.7432 | 48771.5593 | 48374.66 | 963.4932 | 863.4932 | 932.891
3464 | 107.8968 | 0.97593 | 0.89389 | 1.0279 | 191.2706 | 49.9153 | 51313.0775 | 51098.0901 | 1017.9618 | 917.9618 | 932.925
3465 | 99.972 | 1.0603 | 0.72512 | 0.96851 | 195.295 | 50.9415 | 41370.4813 | 41403.7382 | 824.0748 | 724.0748 | 932.9433
3466 | 106.2804 | 1.0095 | 0.81007 | 1.0506 | 200.2608 | 53.2326 | 51214.9947 | 50608.2229 | 1008.1645 | 908.1645 | 932.9746
3467 | 102.2105 | 1.0545 | 0.98152 | 1.0575 | 209.3604 | 53.0268 | 60965.3265 | 61868.6928 | 1233.3739 | 1133.3739 | 933.0294
3468 | 96.1005 | 1.0718 | 1.0939 | 0.94485 | 187.8325 | 51.0219 | 58177.992 | 58779.6445 | 1171.5929 | 1071.5929 | 933.0827
3469 | 113.1147 | 0.96641 | 1.2055 | 0.92197 | 182.8544 | 46.2391 | 56629.1336 | 56556.7541 | 1127.1351 | 1027.1351 | 933.1531
3470 | 106.9582 | 1.0687 | 1.0629 | 0.94581 | 185.3871 | 50.9344 | 57213.9299 | 59255.2071 | 1181.1041 | 1081.1041 | 933.1581
3471 | 81.0015 | 0.9953 | 1.0504 | 0.97852 | 204.3393 | 49.8912 | 55507.5911 | 54158.4466 | 1079.1689 | 979.1689 | 933.1997
3472 | 99.7034 | 1.0115 | 1.1605 | 1.0088 | 189.6346 | 48.3035 | 61613.413 | 64209.5641 | 1280.1913 | 1180.1913 | 933.2127
3473 | 77.0836 | 1.0283 | 1.3139 | 0.91499 | 208.6615 | 52.0043 | 66852.8217 | 68876.465 | 1373.5293 | 1273.5293 | 933.2731
3474 | 85.6739 | 0.97379 | 0.79513 | 1.1108 | 195.3076 | 54.1115 | 52503.049 | 51158.456 | 1019.1691 | 919.1691 | 933.2845
3475 | 96.9622 | 0.98271 | 0.80848 | 0.97419 | 211.8937 | 55.5387 | 49247.2633 | 48889.0352 | 973.7807 | 873.7807 | 933.3655
3476 | 92.2553 | 0.96499 | 1.0906 | 1.0755 | 203.7179 | 50.3506 | 63746.9607 | 64216.3878 | 1280.3278 | 1180.3278 | 933.3891
3477 | 99.7107 | 1.0024 | 0.86066 | 0.95053 | 167.8668 | 50.6755 | 46689.5194 | 46707.1968 | 930.1439 | 830.1439 | 933.3912
3478 | 114.0213 | 0.94751 | 1.0017 | 1.0613 | 210.6459 | 50.1202 | 58910.9514 | 58266.2462 | 1161.3249 | 1061.3249 | 933.4102
3479 | 91.3686 | 0.92736 | 1.0825 | 1.0147 | 189.1603 | 49.1529 | 58347.2597 | 61359.2584 | 1223.1852 | 1123.1852 | 933.4431
3480 | 118.0989 | 1.0612 | 0.72724 | 0.90495 | 190.9799 | 47.9711 | 37773.3589 | 37612.2734 | 748.2455 | 648.2455 | 933.4539
3481 | 127.7549 | 0.95503 | 0.86923 | 1.0172 | 192.0934 | 49.6208 | 50120.0671 | 49408.7278 | 984.1746 | 884.1746 | 933.4786
3482 | 93.8681 | 0.95595 | 0.95325 | 1.049 | 199.7536 | 51.647 | 56480.1947 | 54322.0088 | 1082.4402 | 982.4402 | 933.5569
3483 | 83.1087 | 0.95768 | 0.94517 | 1.0367 | 207.976 | 48.7239 | 51830.025 | 51697.4749 | 1029.9495 | 929.9495 | 933.5625
3484 | 103.9528 | 1.0041 | 0.97836 | 1.0754 | 193.1274 | 51.2341 | 59444.3137 | 58312.6172 | 1162.2523 | 1062.2523 | 933.5849
3485 | 121.4317 | 1.0489 | 1.1066 | 1.0778 | 199.5803 | 48.0254 | 63592.0917 | 60893.0453 | 1213.8609 | 1113.8609 | 933.6587
3486 | 100.3314 | 0.9733 | 0.83745 | 0.97216 | 188.2615 | 47.5876 | 43577.7914 | 44579.2995 | 887.586 | 787.586 | 933.7029
3487 | 117.6141 | 1.0385 | 1.0043 | 1.0402 | 217.0651 | 50.324 | 58934.1189 | 59956.0812 | 1195.1216 | 1095.1216 | 933.7141
3488 | 110.0193 | 1.0488 | 0.84103 | 0.95432 | 205.0111 | 53.3755 | 49203.6524 | 49209.4903 | 980.1898 | 880.1898 | 933.8502
3489 | 101.8881 | 1.0319 | 0.72515 | 1.0309 | 202.0876 | 46.3742 | 39743.8105 | 40572.2237 | 807.4445 | 707.4445 | 933.8569
3490 | 116.3467 | 0.99904 | 1.1049 | 0.93669 | 210.0386 | 50.2575 | 58065.3304 | 59600.7302 | 1188.0146 | 1088.0146 | 933.8624
3491 | 92.4602 | 1.0515 | 1.1718 | 1.0184 | 199.2506 | 48.724 | 63080.834 | 63795.2893 | 1271.9058 | 1171.9058 | 933.8778
3492 | 108.2842 | 0.96195 | 0.96881 | 0.9116 | 184.9136 | 49.6304 | 49186.3068 | 47791.8034 | 951.8361 | 851.8361 | 934.0597
3493 | 127.8117 | 0.94994 | 1.037 | 0.97611 | 187.461 | 48.9457 | 55673.5874 | 55532.7626 | 1106.6553 | 1006.6553 | 934.0919
3494 | 99.7184 | 1.0563 | 0.9036 | 1.0095 | 202.785 | 47.6165 | 48651.9643 | 48927.312 | 974.5462 | 874.5462 | 934.1943
3495 | 96.2389 | 0.94366 | 1.1886 | 1.0595 | 197.9239 | 50.595 | 68508.8256 | 68266.8065 | 1361.3361 | 1261.3361 | 934.2025
3496 | 101.6649 | 1.0431 | 0.97576 | 1.0478 | 199.5689 | 48.9044 | 55383.7472 | 55240.5056 | 1100.8101 | 1000.8101 | 934.2444
3497 | 108.3492 | 0.96897 | 1.2103 | 1.1172 | 203.7651 | 53.1688 | 77676.2887 | 75893.4857 | 1513.8697 | 1413.8697 | 934.2941
3498 | 101.1802 | 0.88146 | 1.0746 | 1.0246 | 203.6332 | 50.0555 | 59782.1483 | 59633.1032 | 1188.6621 | 1088.6621 | 934.4001
3499 | 114.9447 | 1.0013 | 1.0789 | 1.0026 | 205.069 | 52.2725 | 62764.0482 | 63377.1546 | 1263.5431 | 1163.5431 | 934.4289
3500 | 101.0368 | 1.0198 | 0.92895 | 1.03 | 184.9935 | 45.1052 | 47988.8477 | 48575.5251 | 967.5105 | 867.5105 | 934.5088
3501 | 93.9321 | 1.0057 | 1.0355 | 1.008 | 206.3625 | 48.5421 | 55455.8994 | 57470.3469 | 1145.4069 | 1045.4069 | 934.5262
3502 | 96.7212 | 0.97556 | 1.0659 | 0.96935 | 204.1344 | 52.6234 | 59540.6653 | 60626.3138 | 1208.5263 | 1108.5263 | 934.5606
3503 | 107.0948 | 1.0446 | 0.811 | 1.0274 | 204.7707 | 49.1634 | 46666.9823 | 48747.2169 | 970.9443 | 870.9443 | 934.617
3504 | 108.2587 | 1.0142 | 1.077 | 1.0204 | 203.4816 | 49.0124 | 59448.6554 | 59858.4535 | 1193.1691 | 1093.1691 | 934.7703
3505 | 120.6811 | 1.0071 | 1.1767 | 0.95877 | 196.9648 | 50.2846 | 63039.6775 | 62549.1673 | 1246.9833 | 1146.9833 | 934.7907
3506 | 96.4083 | 0.99049 | 0.99518 | 1.0126 | 204.4174 | 49.9717 | 55331.7669 | 54744.4256 | 1090.8885 | 990.8885 | 934.8104
3507 | 119.593 | 0.9622 | 1.0423 | 1.0461 | 193.0827 | 43.0529 | 52088.9417 | 52902.1432 | 1054.0429 | 954.0429 | 934.9011
3508 | 103.5688 | 1.05 | 1.0688 | 1.1016 | 194.7221 | 56.4883 | 72847.0276 | 71695.5186 | 1429.9104 | 1329.9104 | 934.9553
3509 | 89.7647 | 1.015 | 0.98457 | 0.98654 | 211.9962 | 50.4304 | 53790.8847 | 52878.6699 | 1053.5734 | 953.5734 | 934.9778
3510 | 80.9004 | 0.93268 | 1.0774 | 0.93457 | 191.457 | 47.7743 | 51901.3041 | 51420.1012 | 1024.402 | 924.402 | 934.9863
3511 | 117.0638 | 0.9679 | 1.1265 | 0.90928 | 190.5766 | 49.0714 | 56012.4268 | 56725.2003 | 1130.504 | 1030.504 | 935.0168
3512 | 89.9509 | 1.0642 | 1.0109 | 1.0336 | 181.2885 | 52.3164 | 59854.4386 | 60028.0958 | 1196.5619 | 1096.5619 | 935.0667
3513 | 105.6343 | 1.0604 | 0.76025 | 0.98139 | 214.4326 | 50.0529 | 43165.7763 | 41950.7966 | 835.0159 | 735.0159 | 935.0674
3514 | 92.3427 | 0.98672 | 1.2445 | 1.0574 | 208.0942 | 47.8654 | 67555.7195 | 67559.9218 | 1347.1984 | 1247.1984 | 935.0701
3515 | 97.8422 | 0.94849 | 1.1324 | 1.0564 | 193.2864 | 49.446 | 63929.4575 | 62129.5784 | 1238.5916 | 1138.5916 | 935.1135
3516 | 91.2733 | 1.0181 | 0.97502 | 0.983 | 214.3884 | 49.5859 | 52347.6818 | 51721.7542 | 1030.4351 | 930.4351 | 935.1852
3517 | 92.9906 | 0.99498 | 1.0051 | 0.9342 | 206.4596 | 48.9335 | 50679.8016 | 49355.5625 | 983.1112 | 883.1112 | 935.3139
3518 | 100.1699 | 0.99174 | 0.84196 | 1.0193 | 196.0936 | 48.5788 | 46711.5875 | 46556.3798 | 927.1276 | 827.1276 | 935.3884
3519 | 84.8184 | 1.0037 | 1.0484 | 0.97383 | 189.3251 | 49.9341 | 55421.4997 | 57292.3088 | 1141.8462 | 1041.8462 | 935.394
3520 | 108.0924 | 1.0794 | 1.1299 | 1.0228 | 218.6986 | 50.6712 | 64690.2345 | 60930.3981 | 1214.608 | 1114.608 | 935.3983
3521 | 103.3419 | 1.002 | 1.1535 | 1.0352 | 203.4704 | 50.8249 | 66159.1635 | 65712.3923 | 1310.2478 | 1210.2478 | 935.4633
3522 | 109.3887 | 0.94637 | 0.9724 | 0.9148 | 195.3607 | 49.2473 | 49101.561 | 49445.2974 | 984.9059 | 884.9059 | 935.4815
3523 | 112.8129 | 1.0176 | 1.1313 | 1.0514 | 210.7087 | 53.3494 | 69790.4004 | 71161.3421 | 1419.2268 | 1319.2268 | 935.5475
3524 | 104.6889 | 1.0208 | 1.0952 | 1.0669 | 193.0487 | 47.9267 | 61313.1723 | 61540.6104 | 1226.8122 | 1126.8122 | 935.5674
3525 | 86.3922 | 0.90425 | 1.2055 | 0.89765 | 194.5094 | 52.9466 | 61626.8703 | 60077.5878 | 1197.5518 | 1097.5518 | 935.5897
3526 | 72.2921 | 0.94279 | 1.0065 | 1.0326 | 198.6138 | 51.9772 | 57762.0471 | 57938.5749 | 1154.7715 | 1054.7715 | 935.6348
3527 | 99.1148 | 1.0645 | 0.94739 | 1.021 | 203.6081 | 51.5432 | 55499.8099 | 55113.8992 | 1098.278 | 998.278 | 935.6439
3528 | 101.1475 | 0.92888 | 0.86323 | 0.99303 | 210.099 | 54.5508 | 52096.8822 | 51597.1893 | 1027.9438 | 927.9438 | 935.6447
3529 | 101.7851 | 0.97028 | 0.68372 | 1.0288 | 200.8004 | 49.3163 | 39761.9069 | 39940.7164 | 794.8143 | 694.8143 | 935.6637
3530 | 99.6098 | 1.0634 | 0.79764 | 1.0119 | 187.5739 | 47.799 | 43830.5006 | 43315.7315 | 862.3146 | 762.3146 | 935.6934
3531 | 106.373 | 0.95398 | 0.98875 | 1.1069 | 213.1658 | 48.4687 | 58176.2768 | 59621.6839 | 1188.4337 | 1088.4337 | 935.7025
3532 | 110.5896 | 0.96095 | 0.85004 | 0.97013 | 185.4698 | 48.2744 | 45125.1791 | 43829.7905 | 872.5958 | 772.5958 | 935.7255
3533 | 95.1595 | 1.0382 | 1.1932 | 0.99194 | 199.0067 | 49.3507 | 63486.4684 | 62591.0848 | 1247.8217 | 1147.8217 | 935.8246
3534 | 107.8125 | 0.97526 | 0.95912 | 0.94802 | 207.0923 | 50.8743 | 51814.6035 | 51141.4523 | 1018.829 | 918.829 | 935.8246
3535 | 91.2975 | 1.028 | 1.0581 | 1.0235 | 203.4939 | 49.2208 | 58129.4062 | 58087.9341 | 1157.7587 | 1057.7587 | 935.8545
3536 | 93.4995 | 0.96947 | 1.0539 | 0.98791 | 203.0687 | 50.4072 | 57254.3131 | 58943.3032 | 1174.8661 | 1074.8661 | 935.9143
3537 | 100.8581 | 1.0709 | 1.021 | 0.9647 | 182.4732 | 51.6684 | 56655.2147 | 56573.5611 | 1127.4712 | 1027.4712 | 936.0456
3538 | 105.5405 | 0.90465 | 0.95171 | 1.0733 | 206.6029 | 49.2835 | 55251.5483 | 56116.009 | 1118.3202 | 1018.3202 | 936.0809
3539 | 102.6608 | 1.0219 | 0.70808 | 0.98245 | 211.913 | 50.3955 | 40556.7301 | 41253.7475 | 821.0749 | 721.0749 | 936.1565
3540 | 94.8702 | 0.964 | 0.83807 | 0.96024 | 188.6007 | 49.2631 | 44338.312 | 43056.5962 | 857.1319 | 757.1319 | 936.1694
3541 | 108.8609 | 0.97924 | 0.88151 | 0.9556 | 217.182 | 46.8391 | 44666.2505 | 44980.7997 | 895.616 | 795.616 | 936.1858
3542 | 98.3705 | 1.0367 | 0.94768 | 0.87439 | 208.0883 | 46.9493 | 43899.9727 | 43539.6504 | 866.793 | 766.793 | 936.2078
3543 | 107.7665 | 1.0058 | 1.0886 | 0.97673 | 206.1747 | 46.6812 | 54902.4621 | 56018.9888 | 1116.3798 | 1016.3798 | 936.2207
3544 | 99.9179 | 0.95622 | 0.9235 | 1.0554 | 181.1114 | 48.5024 | 52090.2586 | 51554.6157 | 1027.0923 | 927.0923 | 936.2752
3545 | 93.3927 | 0.95779 | 0.95758 | 1.0381 | 195.6672 | 52.2039 | 56758.9027 | 55956.5615 | 1115.1312 | 1015.1312 | 936.3088
3546 | 101.705 | 0.99823 | 0.96087 | 1.0008 | 201.0988 | 50.0655 | 53430.5515 | 53042.5278 | 1056.8506 | 956.8506 | 936.3438
3547 | 97.1984 | 0.99129 | 0.91266 | 0.99587 | 187.3311 | 50.8652 | 51319.4675 | 51210.7374 | 1020.2147 | 920.2147 | 936.3867
3548 | 93.8448 | 0.98348 | 1.1747 | 0.97697 | 203.2386 | 51.3538 | 63878.7332 | 63382.9613 | 1263.6592 | 1163.6592 | 936.3898
3549 | 113.2026 | 1.0717 | 1.135 | 1.0567 | 218.1032 | 50.8223 | 67335.9922 | 64649.9175 | 1288.9984 | 1188.9984 | 936.4276
3550 | 111.0951 | 1.0118 | 0.98082 | 1.0635 | 200.2937 | 51.4427 | 59641.6419 | 58116.9763 | 1158.3395 | 1058.3395 | 936.4441
3551 | 94.8672 | 1.0197 | 0.87598 | 0.99898 | 188.853 | 50.5258 | 49290.9061 | 52582.1669 | 1047.6433 | 947.6433 | 936.567
3552 | 114.1479 | 0.94484 | 0.94345 | 1.1006 | 206.7664 | 51.5658 | 59310.9507 | 59702.6447 | 1190.0529 | 1090.0529 | 936.5682
3553 | 99.3212 | 0.97502 | 0.99115 | 0.963 | 183.7864 | 48.2523 | 50912.1522 | 51457.4382 | 1025.1488 | 925.1488 | 936.6256
3554 | 109.9792 | 1.0471 | 1.0911 | 1.0415 | 204.451 | 51.1115 | 64174.8524 | 63825.6453 | 1272.5129 | 1172.5129 | 936.6862
3555 | 113.2478 | 1.0202 | 1.2532 | 1.02 | 187.1716 | 50.3069 | 70303.9316 | 71444.6473 | 1424.8929 | 1324.8929 | 936.729
3556 | 97.536 | 0.9143 | 1.083 | 0.90365 | 205.643 | 48.1341 | 51605.3036 | 50711.4328 | 1010.2287 | 910.2287 | 936.733
3557 | 85.3407 | 1.0506 | 1.0794 | 1.0452 | 212.1603 | 47.4883 | 58046.9296 | 56968.4042 | 1135.3681 | 1035.3681 | 936.7351
3558 | 109.4339 | 1.0278 | 0.97499 | 1.0183 | 209.6314 | 50.6085 | 56149.6802 | 54955.1663 | 1095.1033 | 995.1033 | 936.758
3559 | 111.1738 | 1.0091 | 1.2027 | 1.026 | 207.6394 | 48.9094 | 66044.0988 | 66829.0775 | 1332.5815 | 1232.5815 | 936.7634
3560 | 102.9086 | 0.94102 | 1.1438 | 1.0108 | 198.0125 | 47.7349 | 60013.911 | 59216.3187 | 1180.3264 | 1080.3264 | 936.8283
3561 | 109.4349 | 1.0258 | 1.0698 | 1.0598 | 205.2423 | 47.4047 | 59269.9559 | 59896.1755 | 1193.9235 | 1093.9235 | 936.8285
3562 | 110.4312 | 1.0189 | 0.51457 | 0.97515 | 211.0202 | 53.0152 | 32778.205 | 32961.5011 | 655.23 | 555.23 | 936.8737
3563 | 96.0155 | 1.0438 | 0.96913 | 1.0622 | 201.7475 | 49.7677 | 56420.9691 | 55764.0905 | 1111.2818 | 1011.2818 | 937.0109
3564 | 91.9284 | 0.92437 | 1.1939 | 1.0561 | 200.6885 | 46.2902 | 62500.4093 | 61111.0713 | 1218.2214 | 1118.2214 | 937.1482
3565 | 100.1425 | 1.1134 | 0.94626 | 1.0347 | 197.2713 | 54.0546 | 59147.2766 | 59417.5052 | 1184.3501 | 1084.3501 | 937.2654
3566 | 96.7628 | 0.98552 | 1.1077 | 1.0131 | 189.8237 | 48.5132 | 59259.8378 | 58896.0402 | 1173.9208 | 1073.9208 | 937.3042
3567 | 95.9968 | 0.98309 | 0.99478 | 0.9175 | 193.3014 | 53.2512 | 53821.8662 | 53541.4937 | 1066.8299 | 966.8299 | 937.4173
3568 | 85.479 | 1.0008 | 1.1729 | 0.96362 | 195.8937 | 55.6302 | 67828.0709 | 68248.3213 | 1360.9664 | 1260.9664 | 937.4389
3569 | 99.6074 | 1.002 | 0.91469 | 1.0163 | 197.749 | 52.0912 | 53823.0488 | 52893.6767 | 1053.8735 | 953.8735 | 937.4525
3570 | 94.1669 | 0.95662 | 0.98548 | 1.0362 | 201.8073 | 53.5779 | 59737.8859 | 60634.4591 | 1208.6892 | 1108.6892 | 937.4634
3571 | 109.359 | 0.97492 | 1.1577 | 1.0213 | 214.3025 | 50.8423 | 65749.7613 | 64113.0448 | 1278.2609 | 1178.2609 | 937.4852
3572 | 108.7433 | 0.94904 | 1.1245 | 0.93505 | 198.123 | 50.0174 | 57952.0491 | 58605.3975 | 1168.108 | 1068.108 | 937.5209
3573 | 101.5855 | 1.0088 | 1.0385 | 0.92451 | 204.6098 | 46.5674 | 49687.387 | 50125.6603 | 998.5132 | 898.5132 | 937.5348
3574 | 119.9403 | 0.99158 | 1.0157 | 0.98595 | 211.7141 | 46.9465 | 52809.4487 | 52352.8238 | 1043.0565 | 943.0565 | 937.5392
3575 | 110.0918 | 0.99509 | 0.93383 | 1.0234 | 193.2135 | 50.2727 | 53745.2947 | 54246.4407 | 1080.9288 | 980.9288 | 937.6176
3576 | 113.4039 | 0.93291 | 1.0331 | 1.0617 | 204.9514 | 49.6242 | 59882.4765 | 58907.1804 | 1174.1436 | 1074.1436 | 937.6188
3577 | 105.2453 | 1.0464 | 0.93367 | 1.0027 | 204.1246 | 47.7497 | 50166.4819 | 51790.7169 | 1031.8143 | 931.8143 | 937.7415
3578 | 103.1786 | 0.89998 | 1.0939 | 0.92602 | 200.1819 | 46.897 | 52059.0393 | 51759.5752 | 1031.1915 | 931.1915 | 937.8525
3579 | 87.3996 | 0.91534 | 1.1456 | 0.96132 | 200.3145 | 54.6074 | 64705.2559 | 66378.2053 | 1323.5641 | 1223.5641 | 937.9115
3580 | 103.6316 | 1.0101 | 0.7555 | 1.0826 | 190.9874 | 49.5522 | 45906.9967 | 45667.9306 | 909.3586 | 809.3586 | 938.0592
3581 | 103.7306 | 0.98782 | 0.79142 | 0.94386 | 210.419 | 52.4418 | 44757.4661 | 45428.9281 | 904.5786 | 804.5786 | 938.1234
3582 | 101.021 | 1.0904 | 1.0725 | 0.9996 | 208.9926 | 46.9796 | 55749.3025 | 53212.4007 | 1060.248 | 960.248 | 938.2004
3583 | 109.3429 | 1.0041 | 1.3724 | 1.0454 | 201.2342 | 53.5815 | 82954.814 | 82247.7006 | 1640.954 | 1540.954 | 938.2056
3584 | 100.1627 | 0.99192 | 0.97436 | 0.97568 | 199.847 | 51.1861 | 53946.2848 | 54718.5586 | 1090.3712 | 990.3712 | 938.2446
3585 | 106.7328 | 0.98455 | 0.98285 | 1.0077 | 189.167 | 52.4215 | 57619.156 | 58264.3883 | 1161.2878 | 1061.2878 | 938.3241
3586 | 103.7913 | 1.0171 | 0.86766 | 1.0758 | 203.0635 | 53.8305 | 56132.2637 | 56114.7805 | 1118.2956 | 1018.2956 | 938.3481
3587 | 107.1434 | 1.0128 | 1.0048 | 1.065 | 187.6209 | 47.8284 | 56558.11 | 55295.8873 | 1101.9177 | 1001.9177 | 938.4063
3588 | 106.1184 | 1.0196 | 0.93811 | 1.027 | 199.5673 | 50.0029 | 53785.056 | 53745.6879 | 1070.9138 | 970.9138 | 938.524
3589 | 83.6261 | 1.0114 | 0.90174 | 0.9684 | 181.8131 | 49.3562 | 47456.804 | 47995.7961 | 955.9159 | 855.9159 | 938.5457
3590 | 95.0858 | 0.98508 | 0.79496 | 1.044 | 202.523 | 52.6948 | 48872.4039 | 48267.1503 | 961.343 | 861.343 | 938.5545
3591 | 98.1511 | 0.91801 | 1.1053 | 1.0107 | 203.6948 | 51.9187 | 62884.4371 | 63764.7461 | 1271.2949 | 1171.2949 | 938.6645
3592 | 83.8998 | 0.8991 | 0.98094 | 1.0012 | 217.2584 | 52.6958 | 55946.3313 | 56348.6484 | 1122.973 | 1022.973 | 938.6785
3593 | 105.3449 | 0.98246 | 0.97208 | 0.96444 | 205.2862 | 49.3606 | 51590.2449 | 51706.7566 | 1030.1351 | 930.1351 | 938.6999
3594 | 83.0351 | 0.90045 | 0.93457 | 1.1074 | 205.4522 | 50.2894 | 56010.7705 | 56798.4226 | 1131.9685 | 1031.9685 | 938.8023
3595 | 123.3545 | 1.0202 | 0.67135 | 0.95016 | 197.1019 | 51.4134 | 39463.3301 | 40786.0472 | 811.7209 | 711.7209 | 938.8196
3596 | 96.2691 | 0.99265 | 1.0454 | 1.0311 | 195.6899 | 53.198 | 62619.2367 | 63082.97 | 1257.6594 | 1157.6594 | 938.9242
3597 | 119.0635 | 1.0559 | 1.0124 | 1.0361 | 209.1747 | 48.4423 | 57111.882 | 54706.0435 | 1090.1209 | 990.1209 | 938.9569
3598 | 74.8684 | 0.98169 | 1.0422 | 0.98083 | 195.571 | 49.5709 | 54512.9078 | 53055.5413 | 1057.1108 | 957.1108 | 938.998
3599 | 99.6432 | 0.94448 | 0.85373 | 1.0155 | 182.8958 | 49.4627 | 47721.0842 | 47270.5501 | 941.411 | 841.411 | 939.0829
3600 | 94.5697 | 1.0123 | 0.89995 | 1.0223 | 189.38 | 50.6366 | 51621.8354 | 51001.683 | 1016.0337 | 916.0337 | 939.1794
3601 | 123.4514 | 0.89316 | 0.97699 | 0.94928 | 205.6327 | 51.5825 | 53732.8048 | 53669.0947 | 1069.3819 | 969.3819 | 939.2076
3602 | 95.9053 | 0.90039 | 0.80974 | 0.99053 | 200.8398 | 46.9621 | 41923.0648 | 41828.4227 | 832.5685 | 732.5685 | 939.2117
3603 | 97.8193 | 0.91054 | 0.92583 | 1.0312 | 191.7452 | 53.0034 | 55515.0478 | 56978.6646 | 1135.5733 | 1035.5733 | 939.2668
3604 | 103.3926 | 1.1031 | 0.81762 | 1.01 | 217.9335 | 45.9181 | 43373.3221 | 43311.3144 | 862.2263 | 762.2263 | 939.2721
3605 | 96.9834 | 1.0524 | 1.0723 | 1.0282 | 196.0819 | 43.5976 | 52710.5456 | 51325.7291 | 1022.5146 | 922.5146 | 939.2725
3606 | 99.582 | 1.0071 | 0.84955 | 0.99121 | 216.9039 | 53.3379 | 50481.1331 | 49582.2441 | 987.6449 | 887.6449 | 939.2731
3607 | 111.5186 | 1.0685 | 1.0116 | 0.99508 | 199.7084 | 53.4021 | 60317.207 | 58533.5005 | 1166.67 | 1066.67 | 939.2879
3608 | 113.7481 | 1.0194 | 1.0842 | 1.1629 | 210.4539 | 50.5845 | 69851.2834 | 71850.453 | 1433.0091 | 1333.0091 | 939.3231
3609 | 98.0213 | 1.0387 | 0.97964 | 1.0716 | 200.9195 | 45.2327 | 52288.6403 | 51316.5429 | 1022.3309 | 922.3309 | 939.3341
3610 | 90.1009 | 1.0283 | 1.0956 | 0.96335 | 194.8721 | 49.9051 | 57489.4398 | 56934.9144 | 1134.6983 | 1034.6983 | 939.3506
3611 | 109.9795 | 1.0256 | 1.0156 | 1.0178 | 191.6339 | 47.8284 | 55024.0141 | 56022.8966 | 1116.4579 | 1016.4579 | 939.3684
3612 | 99.5718 | 0.99378 | 1.0872 | 0.89069 | 202.6655 | 47.8836 | 51311.1632 | 51120.0439 | 1018.4009 | 918.4009 | 939.382
3613 | 102.3468 | 1.0712 | 0.87802 | 1.0556 | 200.9368 | 50.8363 | 52891.3865 | 54097.7123 | 1077.9542 | 977.9542 | 939.3841
3614 | 94.0668 | 0.99914 | 1.2237 | 1.0048 | 194.7084 | 50.0753 | 66475.8019 | 66994.6996 | 1335.894 | 1235.894 | 939.3851
3615 | 96.2085 | 0.90825 | 1.0816 | 0.96949 | 199.9983 | 47.36 | 53998.1728 | 53477.1504 | 1065.543 | 965.543 | 939.388
3616 | 102.1175 | 0.98336 | 0.8918 | 1.0324 | 201.421 | 51.3964 | 52684.1678 | 51909.714 | 1034.1943 | 934.1943 | 939.4314
3617 | 92.5305 | 0.84751 | 1.0223 | 1.1226 | 184.5056 | 49.5668 | 60957.659 | 62810.1666 | 1252.2033 | 1152.2033 | 939.5093
3618 | 97.8414 | 0.99662 | 1.013 | 0.98039 | 201.0843 | 51.2119 | 56053.7147 | 55584.7464 | 1107.6949 | 1007.6949 | 939.5262
3619 | 91.2893 | 1.0791 | 1.0274 | 1.0344 | 197.211 | 49.9268 | 58173.5767 | 60681.9204 | 1209.6384 | 1109.6384 | 939.5443
3620 | 100.9072 | 0.91216 | 1.0152 | 0.96168 | 197.4087 | 54.2116 | 58114.3785 | 57616.3574 | 1148.3271 | 1048.3271 | 939.8062
3621 | 116.5033 | 0.96926 | 0.94173 | 1.0375 | 190.2417 | 48.5623 | 53119.5485 | 52722.9109 | 1050.4582 | 950.4582 | 939.832
3622 | 93.5302 | 0.96125 | 1.0281 | 1.1084 | 203.0725 | 48.6571 | 60020.102 | 59154.4174 | 1179.0883 | 1079.0883 | 939.8462
3623 | 102.8824 | 0.98189 | 0.97088 | 1.0038 | 210.2295 | 52.9231 | 57133.5957 | 58303.6462 | 1162.0729 | 1062.0729 | 939.8574
3624 | 95.6169 | 0.99381 | 1.1579 | 1.0275 | 209.0601 | 53.2213 | 68584.3138 | 66447.7864 | 1324.9557 | 1224.9557 | 939.8659
3625 | 93.1585 | 1.0245 | 1.0061 | 0.96244 | 177.9857 | 53.7793 | 57384.9841 | 57199.6578 | 1139.9932 | 1039.9932 | 939.8891
3626 | 103.3593 | 1.046 | 0.84865 | 0.98652 | 196.7752 | 51.6873 | 49058.3955 | 48224.707 | 960.4941 | 860.4941 | 939.8951
3627 | 106.4299 | 0.99067 | 1.2734 | 0.98446 | 204.9154 | 47.7627 | 65116.8799 | 65099.1707 | 1297.9834 | 1197.9834 | 939.9657
3628 | 94.3063 | 0.99205 | 1.1071 | 1.0077 | 197.9082 | 47.6322 | 57789.5397 | 58247.2448 | 1160.9449 | 1060.9449 | 940.0219
3629 | 106.1064 | 0.94694 | 0.82534 | 1.0095 | 187.4509 | 49.9312 | 46806.537 | 47010.2427 | 936.2049 | 836.2049 | 940.0482
3630 | 95.1095 | 0.99534 | 1.0327 | 1.0758 | 211.7233 | 46.0951 | 55785.5471 | 58027.4153 | 1156.5483 | 1056.5483 | 940.0555
3631 | 108.1425 | 1.0203 | 0.8424 | 1.0676 | 199.2249 | 47.8804 | 48543.6698 | 48707.1251 | 970.1425 | 870.1425 | 940.1521
3632 | 98.8766 | 0.9795 | 1.0088 | 0.95855 | 186.9141 | 49.1688 | 52496.5277 | 53033.1448 | 1056.6629 | 956.6629 | 940.1778
3633 | 105.1172 | 1.037 | 1.0662 | 0.99595 | 208.4371 | 46.184 | 54283.3766 | 54272.4883 | 1081.4498 | 981.4498 | 940.1793
3634 | 90.905 | 0.97404 | 0.85011 | 1.0171 | 210.4351 | 44.4175 | 42548.9131 | 43051.8725 | 857.0375 | 757.0375 | 940.209
3635 | 103.0465 | 0.96312 | 0.82594 | 1.0676 | 206.475 | 52.1912 | 51409.0858 | 51959.259 | 1035.1852 | 935.1852 | 940.2113
3636 | 103.9358 | 0.99979 | 0.95541 | 1.0275 | 192.5143 | 50.1809 | 54668.1976 | 55916.0793 | 1114.3216 | 1014.3216 | 940.2759
3637 | 116.9195 | 1.0287 | 0.86669 | 1.0262 | 181.5111 | 50.2714 | 50939.0004 | 50599.951 | 1007.999 | 907.999 | 940.305
3638 | 116.3372 | 0.99357 | 1.0365 | 0.94939 | 213.1085 | 50.5748 | 55828.984 | 56454.1612 | 1125.0832 | 1025.0832 | 940.3062
3639 | 99.5369 | 1.0624 | 0.93347 | 1.1093 | 201.8377 | 46.8811 | 53703.8576 | 53731.5448 | 1070.6309 | 970.6309 | 940.3192
3640 | 111.4942 | 0.99514 | 0.75093 | 1.0079 | 199.9596 | 49.8838 | 43490.9971 | 44358.6992 | 883.174 | 783.174 | 940.3337
3641 | 106.6411 | 0.99221 | 0.80476 | 0.98056 | 203.9494 | 58.0742 | 52176.1737 | 52505.8951 | 1046.1179 | 946.1179 | 940.3696
3642 | 96.2031 | 1.058 | 0.7579 | 0.95632 | 192.5421 | 47.4849 | 39442.1506 | 38012.5362 | 756.2507 | 656.2507 | 940.3754
3643 | 94.2476 | 1.0042 | 1.0318 | 1.0541 | 198.0002 | 51.3796 | 60941.9895 | 59478.1713 | 1185.5634 | 1085.5634 | 940.4239
3644 | 103.872 | 0.94289 | 0.80922 | 1.0594 | 191.8343 | 49.6693 | 47637.221 | 46576.1808 | 927.5236 | 827.5236 | 940.4866
3645 | 106.1788 | 1.0151 | 1.0653 | 0.89678 | 212.0327 | 46.6391 | 49795.8747 | 51159.0536 | 1019.1811 | 919.1811 | 940.508
3646 | 95.421 | 1.0078 | 1.1835 | 1.106 | 187.73 | 48.2839 | 68031.689 | 67794.9789 | 1351.8996 | 1251.8996 | 940.5249
3647 | 105.2119 | 0.99926 | 0.87016 | 0.94925 | 206.7691 | 48.651 | 45507.3911 | 46700.3811 | 930.0076 | 830.0076 | 940.5564
3648 | 118.6476 | 0.99809 | 1.058 | 1.0539 | 203.5254 | 49.2189 | 60911.062 | 60240.3284 | 1200.8066 | 1100.8066 | 940.6647
3649 | 91.6673 | 1.0762 | 0.67545 | 1.0023 | 195.7935 | 48.8189 | 38060.8949 | 37153.8655 | 739.0773 | 639.0773 | 940.6864
3650 | 91.8121 | 1.0248 | 0.87162 | 0.94564 | 199.2166 | 49.5152 | 45670.2977 | 44420.5578 | 884.4112 | 784.4112 | 940.6934
3651 | 105.6007 | 0.96836 | 0.86002 | 0.99167 | 202.5987 | 49.6221 | 47597.5465 | 48047.4826 | 956.9497 | 856.9497 | 940.7831
3652 | 99.6192 | 1.043 | 0.91878 | 1.0091 | 198.261 | 46.9585 | 48615.7043 | 47750.1583 | 951.0032 | 851.0032 | 940.8094
3653 | 115.4278 | 1.1005 | 0.92402 | 1.0171 | 180.9321 | 47.8272 | 51204.485 | 50633.7976 | 1008.676 | 908.676 | 940.8114
3654 | 120.8912 | 0.96896 | 1.2504 | 0.90097 | 195.549 | 43.7462 | 54602.7153 | 55564.643 | 1107.2929 | 1007.2929 | 940.8189
3655 | 110.8501 | 1.0422 | 1.1318 | 0.96018 | 200.1885 | 48.4609 | 58463.5423 | 59742.3272 | 1190.8465 | 1090.8465 | 940.8778
3656 | 104.4643 | 1.0787 | 0.87177 | 0.96004 | 201.4225 | 48.78 | 46524.0214 | 45928.6112 | 914.5722 | 814.5722 | 940.9153
3657 | 90.6513 | 0.98343 | 0.83822 | 0.91775 | 212.5931 | 50.1034 | 43222.7116 | 43882.4227 | 873.6485 | 773.6485 | 940.952
3658 | 83.0355 | 1.0796 | 1.0846 | 0.9889 | 201.9624 | 52.0914 | 60740.516 | 59494.2428 | 1185.8849 | 1085.8849 | 940.9704
3659 | 85.8534 | 0.9924 | 0.97279 | 1.0333 | 178.9764 | 44.3035 | 48486.7575 | 49001.9116 | 976.0382 | 876.0382 | 940.9715
3660 | 96.8961 | 1.0097 | 1.0822 | 1.1051 | 206.8625 | 46.8492 | 60820.3253 | 60715.2183 | 1210.3044 | 1110.3044 | 941.0995
3661 | 120.3563 | 1.0111 | 1.0246 | 0.99511 | 204.8078 | 53.9428 | 61770.309 | 61631.8256 | 1228.6365 | 1128.6365 | 941.1064
3662 | 104.6723 | 1.041 | 1.0024 | 1.0755 | 206.323 | 49.9595 | 59512.488 | 60026.216 | 1196.5243 | 1096.5243 | 941.1118
3663 | 108.5783 | 1.0443 | 0.87925 | 1.0219 | 192.21 | 47.5347 | 48292.9068 | 47676.1251 | 949.5225 | 849.5225 | 941.174
3664 | 95.557 | 1.0318 | 1.1727 | 1.0619 | 195.9003 | 50.0225 | 67422.97 | 65853.5615 | 1313.0712 | 1213.0712 | 941.2754
3665 | 84.0742 | 0.96766 | 0.91149 | 0.98916 | 214.9784 | 48.3477 | 47738.8617 | 47338.4912 | 942.7698 | 842.7698 | 941.2868
3666 | 104.948 | 0.99714 | 0.96722 | 0.95926 | 190.4948 | 49.3236 | 51115.2417 | 51090.7132 | 1017.8143 | 917.8143 | 941.2903
3667 | 104.7551 | 0.99234 | 0.98949 | 1.0023 | 203.4113 | 46.2285 | 50857.0372 | 53323.5415 | 1062.4708 | 962.4708 | 941.3272
3668 | 94.1155 | 1.0041 | 0.76918 | 0.98682 | 196.8534 | 49.8098 | 42711.7917 | 42228.5121 | 840.5702 | 740.5702 | 941.4091
3669 | 92.3761 | 0.97086 | 1.2688 | 0.99166 | 205.9792 | 51.6697 | 69853.1446 | 69555.4167 | 1387.1083 | 1287.1083 | 941.4387
3670 | 100.7219 | 0.9243 | 1.1539 | 0.89948 | 206.034 | 47.2702 | 53668.4883 | 54789.9785 | 1091.7996 | 991.7996 | 941.4623
3671 | 95.5012 | 0.93669 | 1.0092 | 0.98813 | 210.5515 | 48.0637 | 52439.0919 | 51627.3186 | 1028.5464 | 928.5464 | 941.4834
3672 | 111.9654 | 0.96359 | 1.0271 | 0.98656 | 210.8554 | 50.1861 | 56477.1369 | 57044.0402 | 1136.8808 | 1036.8808 | 941.5226
3673 | 108.1406 | 0.98466 | 0.964 | 0.93063 | 200.3261 | 47.5929 | 47964.6871 | 48157.6478 | 959.153 | 859.153 | 941.5427
3674 | 90.2367 | 1.0279 | 0.98729 | 0.96295 | 195.7433 | 46.3566 | 48567.268 | 49021.1084 | 976.4222 | 876.4222 | 941.561
3675 | 89.4948 | 1.0205 | 1.2111 | 1.0165 | 193.1818 | 50.7994 | 67370.6784 | 68853.8596 | 1373.0772 | 1273.0772 | 941.7165
3676 | 96.4945 | 1.0338 | 1.0284 | 0.93294 | 203.6998 | 52.8683 | 56199.2386 | 56626.3727 | 1128.5275 | 1028.5275 | 941.7484
3677 | 88.4987 | 1.0475 | 1.0515 | 0.99187 | 208.1245 | 48.2576 | 55009.6291 | 53139.4298 | 1058.7886 | 958.7886 | 941.772
3678 | 99.9731 | 0.98078 | 0.83652 | 1.0507 | 187.9385 | 48.7174 | 47783.6274 | 48942.2182 | 974.8444 | 874.8444 | 941.8189
3679 | 120.0986 | 1.0222 | 0.90322 | 0.98011 | 193.3045 | 48.9101 | 49495.6163 | 49765.5012 | 991.31 | 891.31 | 941.8456
3680 | 101.5476 | 1.0308 | 1.0832 | 1.0191 | 194.4985 | 53.3236 | 64635.7816 | 65022.6619 | 1296.4532 | 1196.4532 | 941.853
3681 | 104.3427 | 1.0702 | 0.9946 | 1.0324 | 204.7388 | 54.8841 | 62691.0879 | 61802.1885 | 1232.0438 | 1132.0438 | 941.9386
3682 | 104.7067 | 0.90998 | 0.80888 | 0.99504 | 213.4326 | 47.0247 | 42542.6441 | 42510.633 | 846.2127 | 746.2127 | 942.0626
3683 | 104.7156 | 0.95418 | 0.84973 | 1.0408 | 197.8638 | 49.5887 | 49010.6336 | 48866.5343 | 973.3307 | 873.3307 | 942.0772
3684 | 105.3839 | 1.0058 | 0.93476 | 1.0669 | 209.7853 | 49.7178 | 55062.5631 | 55733.6697 | 1110.6734 | 1010.6734 | 942.1632
3685 | 85.61 | 1.0479 | 0.98766 | 0.97024 | 204.7132 | 47.0559 | 49518.4472 | 51075.1286 | 1017.5026 | 917.5026 | 942.2521
3686 | 106.0421 | 0.91779 | 0.87497 | 1.0198 | 207.5363 | 50.1442 | 49829.4098 | 49633.9595 | 988.6792 | 888.6792 | 942.2642
3687 | 106.4334 | 0.99081 | 0.86826 | 1.0587 | 190.3705 | 47.4509 | 48810.9936 | 48134.8088 | 958.6962 | 858.6962 | 942.2835
3688 | 99.9994 | 1.0761 | 1.0001 | 0.99244 | 192.2244 | 44.9124 | 49602.8155 | 48628.0471 | 968.5609 | 868.5609 | 942.5455
3689 | 100.4937 | 1.0383 | 0.98727 | 0.99392 | 192.8229 | 54.0735 | 58895.562 | 57748.6425 | 1150.9728 | 1050.9728 | 942.556
3690 | 82.2965 | 0.9917 | 0.98977 | 1.0888 | 199.4884 | 48.4576 | 56376.112 | 56447.1498 | 1124.943 | 1024.943 | 942.8545
3691 | 101.4909 | 1.0487 | 0.97029 | 1.016 | 194.2752 | 46.4918 | 50974.3685 | 51506.8506 | 1026.137 | 926.137 | 942.8988
3692 | 103.4938 | 1.037 | 0.86347 | 1.0946 | 192.2687 | 47.7649 | 50465.8693 | 50869.4136 | 1013.3883 | 913.3883 | 942.9477
3693 | 87.4565 | 0.9819 | 1.1128 | 0.95143 | 213.1088 | 48.5058 | 55733.7381 | 56775.355 | 1131.5071 | 1031.5071 | 942.9566
3694 | 86.893 | 1.048 | 1.0128 | 0.96517 | 195.7 | 53.8414 | 57728.912 | 56925.6802 | 1134.5136 | 1034.5136 | 942.9911
3695 | 93.2917 | 0.98419 | 1.2742 | 1.0753 | 190.5506 | 48.3175 | 70826.7451 | 72269.3549 | 1441.3871 | 1341.3871 | 943.0032
3696 | 108.2825 | 1.0775 | 1.1003 | 1.0495 | 195.2986 | 50.8331 | 64827.2396 | 62888.2796 | 1253.7656 | 1153.7656 | 943.008
3697 | 92.176 | 1.1466 | 0.98193 | 1.007 | 185.6236 | 47.5689 | 52249.5351 | 51974.0732 | 1035.4815 | 935.4815 | 943.0438
3698 | 110.0138 | 0.98743 | 0.89507 | 0.98415 | 206.9829 | 49.8932 | 49577.1707 | 48672.287 | 969.4457 | 869.4457 | 943.0565
3699 | 98.7801 | 1.0611 | 1.0443 | 0.97497 | 208.7661 | 47.8614 | 53957.3076 | 53972.5168 | 1075.4503 | 975.4503 | 943.0799
3700 | 97.249 | 1.0295 | 0.97189 | 1.0127 | 200.7661 | 49.7217 | 54118.2998 | 52002.2789 | 1036.0456 | 936.0456 | 943.0856
3701 | 93.1511 | 0.92883 | 1.0224 | 0.9831 | 212.0329 | 49.0791 | 53788.1804 | 55728.7703 | 1110.5754 | 1010.5754 | 943.1084
3702 | 82.5309 | 0.9702 | 0.88643 | 1.052 | 195.358 | 53.8728 | 54745.1666 | 56521.7436 | 1126.4349 | 1026.4349 | 943.1134
3703 | 92.2683 | 1.0607 | 1.0185 | 0.97657 | 191.4398 | 52.6996 | 57767.9055 | 58945.062 | 1174.9012 | 1074.9012 | 943.1164
3704 | 118.4259 | 0.97719 | 1.205 | 1.0325 | 204.0895 | 50.3566 | 68681.8646 | 69977.4526 | 1395.5491 | 1295.5491 | 943.1374
3705 | 110.4205 | 1.0497 | 1.0342 | 0.94932 | 194.943 | 49.0233 | 54008.8335 | 54762.5633 | 1091.2513 | 991.2513 | 943.1437
3706 | 92.6747 | 1.0359 | 0.83519 | 1.0481 | 198.0458 | 43.9151 | 42856.432 | 44031.0849 | 876.6217 | 776.6217 | 943.1765
3707 | 85.5237 | 0.98665 | 0.92755 | 0.93559 | 189.2434 | 51.3466 | 49080.6935 | 48988.3093 | 975.7662 | 875.7662 | 943.1779
3708 | 108.199 | 0.94357 | 0.87793 | 0.96985 | 195.6906 | 47.3356 | 45332.9217 | 44256.5506 | 881.131 | 781.131 | 943.207
3709 | 90.2869 | 1.0117 | 0.92903 | 1.0304 | 173.7737 | 47.8983 | 50401.0684 | 50039.5045 | 996.7901 | 896.7901 | 943.3412
3710 | 101.9676 | 0.97514 | 1.0282 | 0.9476 | 220.0166 | 49.2052 | 53054.4824 | 53734.112 | 1070.6822 | 970.6822 | 943.481
3711 | 90.4388 | 1.0377 | 0.98423 | 1.0766 | 197.9089 | 48.5229 | 56170.0425 | 56820.9156 | 1132.4183 | 1032.4183 | 943.5003
3712 | 110.5127 | 0.90876 | 1.2677 | 1.0422 | 197.4876 | 52.8424 | 75320.2375 | 77005.8413 | 1536.1168 | 1436.1168 | 943.5368
3713 | 88.468 | 1.04 | 0.97974 | 0.96192 | 198.2569 | 47.6348 | 49473.2598 | 50541.9842 | 1006.8397 | 906.8397 | 943.6462
3714 | 89.2456 | 1.0409 | 0.98608 | 0.92938 | 208.4035 | 49.8676 | 50541.81 | 50361.2577 | 1003.2252 | 903.2252 | 943.761
3715 | 89.7163 | 0.99518 | 1.2636 | 0.92313 | 191.7853 | 50.2021 | 63231.0366 | 64854.6253 | 1293.0925 | 1193.0925 | 943.8111
3716 | 80.3488 | 0.99837 | 1.0603 | 0.96689 | 203.1564 | 53.4374 | 59274.4051 | 61270.1791 | 1221.4036 | 1121.4036 | 943.8287
3717 | 91.136 | 0.93855 | 1.1728 | 0.96379 | 206.3082 | 53.7513 | 65559.2673 | 67150.5615 | 1339.0112 | 1239.0112 | 944.0107
3718 | 102.6446 | 1.0196 | 0.74132 | 1.0423 | 187.707 | 50.0041 | 44057.416 | 44948.7103 | 894.9742 | 794.9742 | 944.1024
3719 | 99.5624 | 1.0425 | 0.9847 | 0.93339 | 216.0261 | 49.2696 | 50614.3818 | 50021.091 | 996.4218 | 896.4218 | 944.1333
3720 | 109.493 | 0.98555 | 1.0568 | 1.0352 | 208.9027 | 55.1045 | 66439.8516 | 65249.4287 | 1300.9886 | 1200.9886 | 944.1791
3721 | 110.8866 | 0.93487 | 1.1895 | 1.0769 | 206.7258 | 49.4499 | 68681.1988 | 69634.2762 | 1388.6855 | 1288.6855 | 944.2142
3722 | 99.2298 | 0.96988 | 0.89223 | 1.0163 | 186.9145 | 50.5388 | 50879.4008 | 50451.5949 | 1005.0319 | 905.0319 | 944.2473
3723 | 103.9188 | 0.92389 | 1.0127 | 0.94201 | 198.0779 | 53.2922 | 56152.3846 | 55040.178 | 1096.8036 | 996.8036 | 944.3081
3724 | 97.8009 | 0.94623 | 1.2338 | 1.0158 | 206.3116 | 48.4262 | 65379.5031 | 65475.5266 | 1305.5105 | 1205.5105 | 944.4028
3725 | 117.9598 | 0.9806 | 1.1889 | 0.92664 | 204.7207 | 53.1786 | 64941.8284 | 63922.9654 | 1274.4593 | 1174.4593 | 944.4186
3726 | 92.2326 | 1.0296 | 0.86204 | 1.0207 | 192.4026 | 48.0545 | 47037.5975 | 45660.4902 | 909.2098 | 809.2098 | 944.4247
3727 | 112.7503 | 1.0287 | 0.99169 | 1.038 | 209.183 | 47.0805 | 54135.1596 | 54377.0674 | 1083.5413 | 983.5413 | 944.4557
3728 | 112.4868 | 1.0348 | 0.95586 | 0.98912 | 192.464 | 51.7001 | 55090.6535 | 54332.1795 | 1082.6436 | 982.6436 | 944.5214
3729 | 101.115 | 0.97541 | 0.9808 | 0.89052 | 180.5064 | 47.1931 | 46054.6847 | 45542.7229 | 906.8545 | 806.8545 | 944.5329
3730 | 94.3268 | 1.0257 | 0.84958 | 1.0158 | 207.4337 | 51.0624 | 49212.6274 | 50369.9042 | 1003.3981 | 903.3981 | 944.5363
3731 | 104.3141 | 1.0756 | 0.57237 | 1.0021 | 202.1598 | 49.7592 | 34323.9963 | 34181.9461 | 679.6389 | 579.6389 | 944.6345
3732 | 104.9729 | 0.98555 | 1.3616 | 1.0102 | 193.9142 | 48.8827 | 72492.1622 | 72375.7441 | 1443.5149 | 1343.5149 | 944.6964
3733 | 108.9889 | 0.99255 | 0.92798 | 1.0047 | 202.4211 | 52.2063 | 54523.0748 | 55664.6763 | 1109.2935 | 1009.2935 | 944.6969
3734 | 104.5281 | 1.0317 | 0.78581 | 1.0318 | 202.4887 | 48.1304 | 44415.4009 | 44472.5037 | 885.4501 | 785.4501 | 944.735
3735 | 104.7761 | 1.043 | 1.0684 | 1.0741 | 186.4569 | 50.1744 | 63251.5985 | 65308.734 | 1302.1747 | 1202.1747 | 944.7589
3736 | 101.1238 | 0.99217 | 0.91034 | 0.89066 | 197.5678 | 50.7122 | 46403.1749 | 46836.5764 | 932.7315 | 832.7315 | 944.7841
3737 | 91.0477 | 0.96848 | 0.87183 | 0.91107 | 188.8366 | 48.1954 | 42719.8783 | 42929.8288 | 854.5966 | 754.5966 | 944.8073
3738 | 97.9983 | 0.97968 | 1.1662 | 0.95615 | 200.5082 | 52.1004 | 63298.6589 | 62643.441 | 1248.8688 | 1148.8688 | 944.808
3739 | 98.6901 | 1.019 | 0.5948 | 1.0561 | 207.5501 | 48.5834 | 35612.347 | 36102.4973 | 718.0499 | 618.0499 | 944.8175
3740 | 127.629 | 1.0392 | 1.1085 | 0.95232 | 198.5309 | 49.9112 | 59509.3427 | 63815.0119 | 1272.3002 | 1172.3002 | 944.8246
3741 | 116.2385 | 1.0085 | 1.086 | 1.0013 | 210.8192 | 48.4947 | 58627.2608 | 58728.4842 | 1170.5697 | 1070.5697 | 944.8538
3742 | 107.847 | 1.0206 | 1.0207 | 1.0766 | 201.3707 | 49.5206 | 60068.6151 | 60396.3483 | 1203.927 | 1103.927 | 944.8819
3743 | 99.2852 | 0.98153 | 1.0144 | 1.0206 | 204.4471 | 54.4353 | 61864.9249 | 60971.8526 | 1215.4371 | 1115.4371 | 944.8988
3744 | 95.8163 | 0.96438 | 0.97536 | 1.0463 | 206.9203 | 50.6789 | 56610.4426 | 58766.1441 | 1171.3229 | 1071.3229 | 944.8998
3745 | 93.4854 | 1.0084 | 0.85729 | 0.98034 | 203.4445 | 53.2121 | 49940.842 | 49928.4387 | 994.5688 | 894.5688 | 944.9147
3746 | 99.745 | 1.0157 | 1.0359 | 0.96188 | 190.1797 | 50.149 | 55239.0293 | 55114.8018 | 1098.296 | 998.296 | 944.9269
3747 | 108.076 | 0.94744 | 1.1009 | 0.9385 | 185.467 | 48.68 | 55466.5147 | 56341.508 | 1122.8302 | 1022.8302 | 945.066
3748 | 97.1028 | 1.1429 | 1.0885 | 0.95283 | 196.2304 | 47.9469 | 55245.5535 | 56637.5975 | 1128.7519 | 1028.7519 | 945.0813
3749 | 111.5477 | 0.94414 | 0.98539 | 1.0005 | 187.5706 | 47.0562 | 51534.9839 | 51671.2269 | 1029.4245 | 929.4245 | 945.1693
3750 | 107.1453 | 1.0547 | 1.1777 | 1.0171 | 198.8595 | 50.4082 | 66278.7173 | 65584.3547 | 1307.6871 | 1207.6871 | 945.1774
3751 | 94.4457 | 0.91753 | 1.0655 | 1.0907 | 200.7343 | 44.1708 | 55360.2207 | 54765.9809 | 1091.3196 | 991.3196 | 945.2147
3752 | 106.6234 | 0.98266 | 0.80903 | 0.97752 | 199.7698 | 52.1969 | 46948.224 | 47006.5772 | 936.1315 | 836.1315 | 945.2322
3753 | 105.0218 | 0.893 | 1.0722 | 1.0235 | 208.8105 | 49.2845 | 58914.3729 | 59192.2499 | 1179.845 | 1079.845 | 945.236
3754 | 93.412 | 0.99358 | 0.9805 | 1.0896 | 215.4411 | 48.4664 | 56492.7107 | 57811.6311 | 1152.2326 | 1052.2326 | 945.2445
3755 | 106.3967 | 1.0223 | 0.99939 | 1.0205 | 203.4961 | 46.8964 | 53134.5366 | 53981.3995 | 1075.628 | 975.628 | 945.2654
3756 | 102.7018 | 0.9482 | 1.0014 | 1.0204 | 208.5233 | 48.8324 | 54861.2874 | 55377.075 | 1103.5415 | 1003.5415 | 945.2886
3757 | 102.2652 | 1.0608 | 1.1458 | 1.0234 | 196.4556 | 51.4608 | 66125.5134 | 67296.3662 | 1341.9273 | 1241.9273 | 945.3793
3758 | 113.2512 | 0.99599 | 0.92687 | 1.0313 | 206.6842 | 48.9315 | 52497.3463 | 52154.1426 | 1039.0829 | 939.0829 | 945.4037
3759 | 97.7658 | 1.0768 | 1.1297 | 1.038 | 206.2621 | 50.7403 | 65047.2392 | 65103.9643 | 1298.0793 | 1198.0793 | 945.4214
3760 | 95.5924 | 0.98034 | 1.1812 | 0.92785 | 201.88 | 48.8188 | 58279.6797 | 58495.3074 | 1165.9061 | 1065.9061 | 945.4596
3761 | 104.8438 | 1.0018 | 1.2901 | 1.0189 | 200.9798 | 50.371 | 71706.2554 | 71225.0655 | 1420.5013 | 1320.5013 | 945.5165
3762 | 83.4479 | 0.94907 | 1.0332 | 0.97363 | 207.3889 | 51.3088 | 55883.4062 | 57001.6955 | 1136.0339 | 1036.0339 | 945.5274
3763 | 100.0684 | 1.0053 | 1.2727 | 1.0129 | 224.6462 | 44.3892 | 61913.0678 | 61985.0947 | 1235.7019 | 1135.7019 | 945.5286
3764 | 101.8961 | 0.99103 | 1.0691 | 0.83774 | 203.3688 | 50.1205 | 50154.8238 | 49661.5843 | 989.2317 | 889.2317 | 945.5866
3765 | 108.5423 | 1.0436 | 0.83726 | 1.0418 | 180.8105 | 48.0307 | 47516.3385 | 48278.3737 | 961.5675 | 861.5675 | 945.6047
3766 | 94.5701 | 0.95544 | 0.76569 | 1.0323 | 186.6755 | 53.6831 | 47470.0417 | 47880.9811 | 953.6196 | 853.6196 | 945.6181
3767 | 103.285 | 1.0223 | 1.0716 | 1.0317 | 195.423 | 51.5126 | 62582.9931 | 62811.0926 | 1252.2219 | 1152.2219 | 945.6182
3768 | 122.2881 | 0.87248 | 0.78604 | 0.9877 | 197.9987 | 46.9192 | 41630.4542 | 41445.7487 | 824.915 | 724.915 | 945.6296
3769 | 95.349 | 1.0012 | 1.0666 | 0.91847 | 186.2657 | 47.9574 | 51746.5595 | 52614.9132 | 1048.2983 | 948.2983 | 945.7383
3770 | 123.614 | 0.94187 | 1.05 | 0.98064 | 194.4187 | 46.2907 | 53248.4525 | 53074.0728 | 1057.4815 | 957.4815 | 945.8452
3771 | 95.9944 | 0.99178 | 0.96929 | 1.0246 | 214.4174 | 52.3468 | 57183.7745 | 56015.0033 | 1116.3001 | 1016.3001 | 945.9319
3772 | 101.9036 | 1.016 | 1.043 | 0.97566 | 202.0887 | 52.3125 | 58854.2028 | 58882.2819 | 1173.6456 | 1073.6456 | 946.0037
3773 | 110.4679 | 0.92509 | 0.89107 | 1.0814 | 217.5978 | 52.4447 | 56112.4366 | 54595.1387 | 1087.9028 | 987.9028 | 946.0348
3774 | 107.2038 | 1.097 | 0.97987 | 0.96301 | 211.7544 | 47.5645 | 50688.3927 | 50791.2473 | 1011.8249 | 911.8249 | 946.0864
3775 | 82.1716 | 1.0484 | 1.1372 | 1.0346 | 203.2667 | 47.9913 | 60798.2861 | 62692.2823 | 1249.8456 | 1149.8456 | 946.1063
3776 | 87.245 | 1.0637 | 1.0594 | 0.91693 | 194.0774 | 51.5476 | 55049.4269 | 54767.0306 | 1091.3406 | 991.3406 | 946.1179
3777 | 109.2377 | 1.0121 | 1.1256 | 0.93037 | 197.9487 | 52.287 | 60735.5839 | 58288.4263 | 1161.7685 | 1061.7685 | 946.165
3778 | 111.5086 | 0.98657 | 0.70893 | 1.0945 | 184.0394 | 50.7697 | 45162.9079 | 45521.011 | 906.4202 | 806.4202 | 946.2657
3779 | 110.207 | 0.95993 | 1.1933 | 0.95852 | 205.3266 | 47.5776 | 59659.5916 | 57703.1198 | 1150.0624 | 1050.0624 | 946.3424
3780 | 106.2033 | 0.98931 | 1.1957 | 0.96343 | 209.2574 | 51.4552 | 64888.3345 | 64603.2808 | 1288.0656 | 1188.0656 | 946.3529
3781 | 97.437 | 0.91519 | 1.0403 | 1.0649 | 185.9921 | 49.5094 | 59448.7858 | 60741.3768 | 1210.8275 | 1110.8275 | 946.4261
3782 | 117.4175 | 0.94775 | 0.93051 | 1.0698 | 201.0247 | 50.3731 | 55951.7331 | 54682.3279 | 1089.6466 | 989.6466 | 946.4334
3783 | 105.0562 | 0.9978 | 0.87104 | 1.0115 | 190.5178 | 50.937 | 50407.964 | 50090.6948 | 997.8139 | 897.8139 | 946.4479
3784 | 84.8454 | 0.97928 | 1.1126 | 1.0026 | 199.1203 | 51.2068 | 61576.3585 | 61404.3954 | 1224.0879 | 1124.0879 | 946.4705
3785 | 81.567 | 1.0417 | 0.82331 | 0.93276 | 198.1836 | 50.8076 | 43532.7173 | 42970.6132 | 855.4123 | 755.4123 | 946.5037
3786 | 104.9163 | 1.0549 | 1.248 | 0.93345 | 197.1449 | 47.91 | 61311.7978 | 62294.8705 | 1241.8974 | 1141.8974 | 946.5516
3787 | 95.18 | 0.98467 | 0.72088 | 0.95594 | 202.0815 | 52.0072 | 40915.3866 | 41017.4533 | 816.3491 | 716.3491 | 946.6468
3788 | 98.5253 | 1.0199 | 0.90226 | 0.94151 | 196.7832 | 52.2083 | 49793.6542 | 48521.1285 | 966.4226 | 866.4226 | 946.6482
3789 | 104.0982 | 0.99004 | 0.92557 | 0.98892 | 187.4738 | 54.4866 | 55675.3267 | 55784.8896 | 1111.6978 | 1011.6978 | 946.711
3790 | 99.0487 | 1.0109 | 1.1513 | 1.0896 | 213.3179 | 49.8734 | 67773.6293 | 66229.7886 | 1320.5958 | 1220.5958 | 946.7227
3791 | 97.0541 | 1.01 | 1.0122 | 0.96573 | 200.8713 | 48.2374 | 52081.8402 | 53427.4305 | 1064.5486 | 964.5486 | 946.8304
3792 | 107.296 | 1.0503 | 0.86047 | 0.93576 | 213.1058 | 53.5915 | 49404.0273 | 52073.1707 | 1037.4634 | 937.4634 | 946.8347
3793 | 96.4555 | 0.97425 | 0.96736 | 1.1367 | 197.8538 | 50.5592 | 60544.5639 | 61606.8381 | 1228.1368 | 1128.1368 | 946.8618
3794 | 106.3198 | 1.036 | 1.1424 | 1.0118 | 183.133 | 50.3731 | 63951.9705 | 62625.8134 | 1248.5163 | 1148.5163 | 947.0746
3795 | 103.1997 | 1.0485 | 0.87765 | 0.96178 | 201.701 | 47.5363 | 45470.8667 | 45163.6012 | 899.272 | 799.272 | 947.2039
3796 | 95.4467 | 0.99208 | 1.1258 | 0.91094 | 216.7957 | 51.911 | 58369.9626 | 57891.3244 | 1153.8265 | 1053.8265 | 947.2523
3797 | 98.2747 | 1.0463 | 0.98173 | 0.94885 | 190.2353 | 50.9823 | 52922.8555 | 51401.0155 | 1024.0203 | 924.0203 | 947.2607
3798 | 87.4102 | 0.98565 | 1.0005 | 0.93152 | 203.887 | 45.2348 | 46259.2483 | 45764.7447 | 911.2949 | 811.2949 | 947.3522
3799 | 100.8068 | 0.90612 | 0.97273 | 0.98752 | 181.5222 | 49.4177 | 52165.6447 | 51803.5098 | 1032.0702 | 932.0702 | 947.3818
3800 | 107.7689 | 1.0139 | 1.1352 | 1.0238 | 220.9552 | 51.7956 | 66080.9874 | 64185.333 | 1279.7067 | 1179.7067 | 947.4511
3801 | 109.4352 | 1.0413 | 1.108 | 0.99807 | 209.4422 | 53.2615 | 65180.2857 | 65944.8982 | 1314.898 | 1214.898 | 947.4855
3802 | 95.1817 | 0.92678 | 1.3581 | 1.0239 | 183.951 | 49.8445 | 73890.8506 | 72978.2787 | 1455.5656 | 1355.5656 | 947.5651
3803 | 96.2457 | 1.017 | 1.2315 | 0.99116 | 206.0481 | 49.3321 | 65249.2376 | 65131.9658 | 1298.6393 | 1198.6393 | 947.6433
3804 | 123.7266 | 0.95164 | 0.80736 | 1.1062 | 195.6608 | 52.0025 | 52763.3386 | 53346.2179 | 1062.9244 | 962.9244 | 947.6826
3805 | 120.7971 | 1.085 | 1.3358 | 0.99544 | 191.7267 | 50.8721 | 74503.8921 | 75915.9316 | 1514.3186 | 1414.3186 | 947.6977
3806 | 101.4602 | 1.0181 | 0.90343 | 0.90928 | 204.01 | 52.4369 | 48695.3905 | 50118.338 | 998.3668 | 898.3668 | 947.7035
3807 | 109.3763 | 0.97852 | 0.79167 | 1.0474 | 206.3175 | 46.5229 | 43762.4738 | 44051.4274 | 877.0285 | 777.0285 | 947.7113
3808 | 99.5354 | 0.95199 | 0.98557 | 1.0657 | 192.9231 | 47.5337 | 54622.4359 | 55373.767 | 1103.4753 | 1003.4753 | 947.8169
3809 | 105.8514 | 1.0003 | 1.2041 | 0.92582 | 201.2361 | 48.2738 | 59129.0862 | 58768.8151 | 1171.3763 | 1071.3763 | 947.8288
3810 | 87.9691 | 1.0104 | 0.9892 | 1.0202 | 184.5291 | 48.4843 | 53423.3918 | 52445.7329 | 1044.9147 | 944.9147 | 947.8352
3811 | 93.6251 | 0.95054 | 0.75019 | 1.0255 | 214.2096 | 45.0398 | 38872.5308 | 38516.3536 | 766.3271 | 666.3271 | 947.9149
3812 | 83.66 | 0.96882 | 0.95907 | 1.022 | 205.9159 | 51.482 | 54837.682 | 52600.5249 | 1048.0105 | 948.0105 | 947.9151
3813 | 98.5412 | 0.91006 | 1.0658 | 1.0411 | 208.0653 | 49.2868 | 59317.1219 | 57586.8408 | 1147.7368 | 1047.7368 | 947.9221
3814 | 109.2688 | 1.0468 | 1.0574 | 1.0357 | 200.1731 | 52.9122 | 64199.8645 | 65473.0074 | 1305.4601 | 1205.4601 | 948.0093
3815 | 62.1797 | 1.0031 | 0.78672 | 0.85254 | 200.5762 | 50.8311 | 37464.0952 | 36814.915 | 732.2983 | 632.2983 | 948.0105
3816 | 83.3279 | 1.0222 | 1.2418 | 1.0299 | 225.8547 | 49.6104 | 67900.1804 | 64674.1942 | 1289.4839 | 1189.4839 | 948.0305
3817 | 105.1431 | 1.0631 | 0.98666 | 1.028 | 198.9971 | 52.1978 | 58976.7186 | 57362.8425 | 1143.2568 | 1043.2568 | 948.121
3818 | 103.148 | 0.96356 | 0.9096 | 1.0106 | 197.311 | 52.9102 | 54092.5472 | 54995.3798 | 1095.9076 | 995.9076 | 948.21
3819 | 107.9652 | 1.0608 | 1.0404 | 1.0038 | 206.3339 | 50.0903 | 58252.0584 | 57571.2492 | 1147.425 | 1047.425 | 948.2836
3820 | 86.0744 | 0.94363 | 0.9925 | 1.0063 | 222.419 | 52.2966 | 56703.7755 | 57145.0059 | 1138.9001 | 1038.9001 | 948.2983
3821 | 111.3364 | 0.957 | 0.69051 | 0.98138 | 189.3036 | 50.1351 | 39505.3931 | 38170.044 | 759.4009 | 659.4009 | 948.3184
3822 | 90.7408 | 1.0045 | 0.91161 | 0.95698 | 205.4655 | 52.1654 | 50468.5825 | 49142.3386 | 978.8468 | 878.8468 | 948.3424
3823 | 102.3261 | 0.89075 | 0.89148 | 1.0316 | 198.4428 | 51.7323 | 52491.2548 | 54458.013 | 1085.1603 | 985.1603 | 948.3643
3824 | 96.5931 | 0.93868 | 0.96139 | 1.0332 | 204.6321 | 55.5746 | 60444.7023 | 60710.8548 | 1210.2171 | 1110.2171 | 948.3897
3825 | 106.3023 | 0.91181 | 1.0713 | 0.98711 | 195.3577 | 50.1265 | 58060.4993 | 59332.7504 | 1182.655 | 1082.655 | 948.4026
3826 | 114.1221 | 0.992 | 1.1028 | 1.0222 | 187.9498 | 52.4766 | 65280.422 | 65226.8912 | 1300.5378 | 1200.5378 | 948.4604
3827 | 100.8674 | 0.97665 | 0.98694 | 1.0133 | 189.1858 | 47.6666 | 52552.4362 | 52127.7231 | 1038.5545 | 938.5545 | 948.4787
3828 | 95.5629 | 0.90422 | 0.66546 | 0.94299 | 174.7758 | 47.607 | 34163.1909 | 33348.6866 | 662.9737 | 562.9737 | 948.4804
3829 | 123.9614 | 0.97638 | 1.1912 | 0.99576 | 195.2393 | 47.8103 | 62690.2814 | 63259.5548 | 1261.1911 | 1161.1911 | 948.5191
3830 | 84.7274 | 1.0621 | 1.0295 | 1.066 | 193.2946 | 49.6857 | 59188.7558 | 59856.9975 | 1193.1399 | 1093.1399 | 948.6892
3831 | 96.989 | 0.97529 | 0.917 | 0.98624 | 201.472 | 53.4619 | 53608.6164 | 52080.8794 | 1037.6176 | 937.6176 | 948.7231
3832 | 97.7079 | 0.95063 | 1.2262 | 0.99922 | 211.601 | 52.0843 | 68866.3973 | 69015.9559 | 1376.3191 | 1276.3191 | 948.8814
3833 | 90.8174 | 0.92047 | 1.0508 | 1.0187 | 200.3872 | 51.006 | 59066.2197 | 58208.2546 | 1160.1651 | 1060.1651 | 948.9004
3834 | 105.8926 | 0.8531 | 0.97186 | 1.0373 | 199.9067 | 49.6063 | 54691.7902 | 54198.5146 | 1079.9703 | 979.9703 | 948.9325
3835 | 100.2599 | 0.89955 | 0.80436 | 0.97249 | 193.5912 | 50.7116 | 44435.5892 | 43682.5288 | 869.6506 | 769.6506 | 948.9687
3836 | 100.6798 | 0.99717 | 1.0423 | 1.0119 | 182.8622 | 54.2795 | 62877.6184 | 61912.8996 | 1234.258 | 1134.258 | 948.9789
3837 | 88.6341 | 1.0042 | 1.1065 | 0.95062 | 195.0066 | 52.7497 | 60375.5544 | 61400.7617 | 1224.0152 | 1124.0152 | 949.0593
3838 | 116.4779 | 1.027 | 0.8049 | 0.95548 | 205.5284 | 45.4022 | 40553.8889 | 39603.8891 | 788.0778 | 688.0778 | 949.1898
3839 | 92.6739 | 1.0212 | 0.88735 | 0.90445 | 214.5769 | 52.7127 | 47508.5126 | 46133.3702 | 918.6674 | 818.6674 | 949.235
3840 | 89.5336 | 1.0582 | 0.85143 | 1.0862 | 196.6541 | 51.1086 | 52304.1639 | 51939.5326 | 1034.7907 | 934.7907 | 949.3002
3841 | 87.4005 | 1.0151 | 1.1498 | 1.0199 | 197.7174 | 48.2149 | 61017.0107 | 60613.1351 | 1208.2627 | 1108.2627 | 949.3959
3842 | 102.0769 | 1.0938 | 0.95126 | 0.96392 | 191.608 | 52.881 | 54584.6217 | 57037.5531 | 1136.7511 | 1036.7511 | 949.4119
3843 | 91.7888 | 0.95787 | 1.0823 | 0.98732 | 206.3294 | 48.9486 | 56813.8378 | 55953.8281 | 1115.0766 | 1015.0766 | 949.4191
3844 | 92.8616 | 1.0774 | 0.96885 | 1.0605 | 204.2548 | 50.5668 | 57219.5039 | 56842.2951 | 1132.8459 | 1032.8459 | 949.4506
3845 | 111.8636 | 1.1091 | 0.90869 | 0.97891 | 195.8415 | 50.7426 | 51628.1506 | 51133.621 | 1018.6724 | 918.6724 | 949.4675
3846 | 107.5096 | 0.96453 | 1.0507 | 1.0519 | 210.4529 | 52.9102 | 64178.3339 | 64711.864 | 1290.2373 | 1190.2373 | 949.4721
3847 | 109.0191 | 0.9682 | 1.0095 | 0.9926 | 230.7123 | 47.9638 | 53355.3494 | 53607.0498 | 1068.141 | 968.141 | 949.5207
3848 | 100.0646 | 1.0959 | 1.0655 | 1.0524 | 200.4595 | 53.9357 | 66591.7716 | 68619.0861 | 1368.3817 | 1268.3817 | 949.6075
3849 | 105.1148 | 0.99521 | 0.77443 | 1.0282 | 199.0316 | 51.2327 | 46354.6138 | 46901.292 | 934.0258 | 834.0258 | 949.617
3850 | 97.3863 | 1.1411 | 0.9199 | 1.0703 | 186.7772 | 50.6288 | 55662.403 | 55362.063 | 1103.2413 | 1003.2413 | 949.6767
3851 | 105.1024 | 0.97441 | 1.0335 | 0.98854 | 195.724 | 47.0167 | 53046.6213 | 52327.798 | 1042.556 | 942.556 | 949.7962
3852 | 98.8486 | 1.1069 | 0.95109 | 1.0558 | 192.0512 | 52.5463 | 58707.6369 | 58992.3612 | 1175.8472 | 1075.8472 | 949.8389
3853 | 111.6218 | 1.0318 | 1.2339 | 0.9871 | 200.3278 | 47.2179 | 63147.0504 | 62158.4699 | 1239.1694 | 1139.1694 | 949.9209
3854 | 109.2483 | 0.97633 | 0.92506 | 0.96371 | 213.2333 | 49.4906 | 49612.4093 | 49348.7107 | 982.9742 | 882.9742 | 949.9458
3855 | 100.9522 | 1.007 | 1.0822 | 0.99153 | 198.6175 | 50.7464 | 59808.2378 | 61050.4866 | 1217.0097 | 1117.0097 | 949.9604
3856 | 98.3727 | 1.0076 | 0.84905 | 0.89816 | 217.4964 | 52.8001 | 45715.5695 | 46957.4795 | 935.1496 | 835.1496 | 949.9838
3857 | 106.2731 | 1.0715 | 0.77953 | 1.0408 | 206.1749 | 52.1279 | 48435.5804 | 48923.6823 | 974.4736 | 874.4736 | 950.0015
3858 | 99.693 | 0.99273 | 1.0329 | 1.0708 | 180.3852 | 49.8042 | 60195.838 | 61210.7297 | 1220.2146 | 1120.2146 | 950.0261
3859 | 95.081 | 0.98173 | 1.112 | 1.0454 | 201.7478 | 53.3678 | 67221.468 | 66396.9591 | 1323.9392 | 1223.9392 | 950.0615
3860 | 110.7618 | 0.93573 | 1.059 | 1.0572 | 204.4689 | 51.2472 | 62890.887 | 63469.7408 | 1265.3948 | 1165.3948 | 950.0655
3861 | 91.7362 | 0.94281 | 0.97335 | 0.94611 | 193.5167 | 48.6454 | 49198.4774 | 47659.8732 | 949.1975 | 849.1975 | 950.1355
3862 | 106.071 | 1.0634 | 1.0385 | 0.89814 | 204.6233 | 51.5143 | 54064.4078 | 54489.2051 | 1085.7841 | 985.7841 | 950.1768
3863 | 98.026 | 1.0728 | 0.89915 | 0.95236 | 205.0123 | 47.5716 | 45944.2842 | 45723.824 | 910.4765 | 810.4765 | 950.1823
3864 | 84.2261 | 1.0027 | 1.1339 | 0.96021 | 192.6863 | 49.2841 | 58013.1252 | 59476.5406 | 1185.5308 | 1085.5308 | 950.2957
3865 | 105.6782 | 0.99034 | 1.0875 | 0.93308 | 206.5865 | 53.2102 | 59767.062 | 57778.3498 | 1151.567 | 1051.567 | 950.3216
3866 | 104.0597 | 0.99409 | 0.68335 | 0.92483 | 205.2367 | 50.514 | 37354.8267 | 38208.402 | 760.168 | 660.168 | 950.4467
3867 | 97.4513 | 1.0218 | 1.0711 | 0.92117 | 221.7055 | 54.4776 | 59396.6501 | 60346.7696 | 1202.9354 | 1102.9354 | 950.4582
3868 | 109.4183 | 1.0539 | 0.83643 | 1.0524 | 208.5048 | 50.885 | 50869.6178 | 51011.6084 | 1016.2322 | 916.2322 | 950.4886
3869 | 80.11 | 1.0334 | 0.93243 | 0.99758 | 205.1035 | 48.1176 | 48946.1822 | 49254.1436 | 981.0829 | 881.0829 | 950.4965
3870 | 93.1326 | 1.021 | 1.0009 | 1.0159 | 190.3452 | 49.439 | 55163.0628 | 53058.8971 | 1057.1779 | 957.1779 | 950.5271
3871 | 88.7212 | 1.0284 | 0.88919 | 0.96291 | 210.5939 | 47.8272 | 45524.2618 | 45748.5484 | 910.971 | 810.971 | 950.5804
3872 | 104.6341 | 1.0293 | 1.0951 | 1.0186 | 217.1245 | 53.3809 | 65508.5766 | 64447.6083 | 1284.9522 | 1184.9522 | 950.646
3873 | 96.3403 | 1.0664 | 0.90427 | 1.021 | 187.1615 | 50.1686 | 51661.1013 | 51928.0296 | 1034.5606 | 934.5606 | 950.6587
3874 | 122.6035 | 0.97914 | 0.63884 | 0.9769 | 202.7867 | 51.4547 | 38491.6133 | 38158.4444 | 759.1689 | 659.1689 | 950.7644
3875 | 84.0856 | 1.005 | 0.98059 | 1.0023 | 191.3625 | 52.9548 | 56712.6178 | 56575.7181 | 1127.5144 | 1027.5144 | 950.8922
3876 | 93.572 | 1.0499 | 1.065 | 0.96318 | 201.591 | 52.8387 | 59595.8909 | 59537.6732 | 1186.7535 | 1086.7535 | 950.9313
3877 | 97.09 | 0.94944 | 0.97803 | 1.009 | 192.2014 | 47.9057 | 51882.4859 | 53724.1334 | 1070.4827 | 970.4827 | 950.9657
3878 | 102.8749 | 1.0973 | 0.96653 | 0.94639 | 203.3803 | 46.3776 | 47861.5549 | 48022.6764 | 956.4535 | 856.4535 | 950.989
3879 | 115.7023 | 0.99087 | 1.0947 | 0.94701 | 210.5503 | 49.0937 | 56735.2419 | 56119.9147 | 1118.3983 | 1018.3983 | 950.9972
3880 | 105.8057 | 1.0059 | 1.0487 | 1.0643 | 192.3314 | 55.0255 | 67460.5487 | 71011.8944 | 1416.2379 | 1316.2379 | 951.0076
3881 | 106.5188 | 1.057 | 0.95137 | 1.008 | 210.0363 | 50.3681 | 54181.6047 | 53204.9008 | 1060.098 | 960.098 | 951.0529
3882 | 113.5771 | 0.89578 | 0.79927 | 0.98731 | 187.0354 | 48.3988 | 43303.7984 | 43542.4463 | 866.8489 | 766.8489 | 951.0663
3883 | 108.4619 | 0.98034 | 1.0274 | 0.98955 | 201.3352 | 48.3381 | 54485.588 | 53382.3232 | 1063.6465 | 963.6465 | 951.1018
3884 | 128.0593 | 1.0226 | 0.96549 | 1.1021 | 213.0995 | 48.527 | 58205.7757 | 57730.9214 | 1150.6184 | 1050.6184 | 951.1383
3885 | 92.5285 | 0.99426 | 1.0661 | 1.0087 | 216.1115 | 46.2377 | 54195.2574 | 54436.9815 | 1084.7396 | 984.7396 | 951.2367
3886 | 97.5141 | 0.92269 | 0.93015 | 0.95278 | 203.6891 | 53.9097 | 52830.6742 | 54519.7099 | 1086.3942 | 986.3942 | 951.3452
3887 | 107.3412 | 1.0033 | 1.031 | 0.96282 | 201.309 | 51.3619 | 56719.3916 | 57479.4943 | 1145.5899 | 1045.5899 | 951.375
3888 | 93.1283 | 1.0045 | 1.0112 | 1.0196 | 209.0544 | 49.9185 | 56343.1653 | 55657.2453 | 1109.1449 | 1009.1449 | 951.4129
3889 | 107.7819 | 1.0614 | 0.9 | 0.97827 | 177.6954 | 45.9824 | 45923.1404 | 46731.4043 | 930.6281 | 830.6281 | 951.4259
3890 | 97.9541 | 0.99056 | 0.92336 | 1.1093 | 184.0544 | 54.4991 | 61296.7567 | 61567.5428 | 1227.3509 | 1127.3509 | 951.4926
3891 | 108.6706 | 1.0277 | 1.1464 | 0.99587 | 201.3878 | 49.8697 | 62705.8869 | 59775.0162 | 1191.5003 | 1091.5003 | 951.6873
3892 | 98.0146 | 0.98997 | 1.004 | 0.97936 | 201.5404 | 50.9842 | 55280.2774 | 55975.9588 | 1115.5192 | 1015.5192 | 951.7043
3893 | 95.3047 | 0.95852 | 0.8475 | 1.0561 | 188.6284 | 51.356 | 50844.1426 | 49385.9731 | 983.7195 | 883.7195 | 951.7087
3894 | 83.3508 | 0.99439 | 1.0974 | 0.98438 | 193.8519 | 47.9377 | 55951.2015 | 55513.9741 | 1106.2795 | 1006.2795 | 951.8037
3895 | 103.8999 | 1.0483 | 0.92094 | 0.94817 | 191.4752 | 51.3836 | 50656.4193 | 50836.8932 | 1012.7379 | 912.7379 | 951.8837
3896 | 110.7001 | 1.0719 | 0.9735 | 0.96347 | 199.3911 | 49.2381 | 52224.1858 | 52636.1551 | 1048.7231 | 948.7231 | 952.0164
3897 | 96.1011 | 0.97346 | 1.1949 | 0.92484 | 197.4102 | 49.2983 | 59286.958 | 59124.951 | 1178.499 | 1078.499 | 952.0298
3898 | 97.8959 | 0.99902 | 0.77755 | 0.93415 | 220.7719 | 49.3367 | 40881.45 | 41240.9599 | 820.8192 | 720.8192 | 952.0514
3899 | 99.1761 | 0.96011 | 1.1976 | 0.96427 | 190.7419 | 52.9828 | 66418.2644 | 65840.9739 | 1312.8195 | 1212.8195 | 952.0544
3900 | 105.0841 | 0.98175 | 1.1563 | 1.0774 | 204.1873 | 55.8963 | 75603.465 | 76515.6023 | 1526.312 | 1426.312 | 952.0956
3901 | 108.6164 | 1.0234 | 1.0646 | 1.0171 | 192.4415 | 46.4553 | 55661.031 | 58671.3692 | 1169.4274 | 1069.4274 | 952.1868
3902 | 106.7738 | 1.0156 | 1.036 | 0.98377 | 191.0363 | 48.7744 | 55189.9767 | 55998.8153 | 1115.9763 | 1015.9763 | 952.2086
3903 | 83.9445 | 0.9665 | 0.97667 | 1.0348 | 207.8502 | 47.5245 | 52094.3729 | 52163.6055 | 1039.2721 | 939.2721 | 952.2097
3904 | 100.9563 | 1.0799 | 1.1236 | 1.0421 | 192.7007 | 50.4361 | 64748.3603 | 62413.9857 | 1244.2797 | 1144.2797 | 952.2097
3905 | 92.5067 | 1.0152 | 1.1057 | 1.018 | 205.0204 | 46.8422 | 57333.3954 | 57191.3055 | 1139.8261 | 1039.8261 | 952.2179
3906 | 104.9068 | 0.99749 | 0.83479 | 1.078 | 199.0898 | 46.5725 | 46983.6576 | 46780.3679 | 931.6074 | 831.6074 | 952.2713
3907 | 96.3229 | 1.0106 | 0.89008 | 1.0214 | 186.5122 | 49.0264 | 49530.8043 | 49160.1036 | 979.2021 | 879.2021 | 952.3687
3908 | 105.0659 | 1.0303 | 1.149 | 0.96377 | 211.2884 | 49.6571 | 60573.795 | 60818.6211 | 1212.3724 | 1112.3724 | 952.4233
3909 | 102.8342 | 1.0042 | 0.79198 | 1.0097 | 194.7499 | 46.5352 | 42213.0799 | 42351.7532 | 843.0351 | 743.0351 | 952.6026
3910 | 104.5169 | 1.0073 | 1.0963 | 0.96261 | 194.7669 | 54.9405 | 63956.7622 | 63874.2565 | 1273.4851 | 1173.4851 | 952.7261
3911 | 118.7971 | 1.0238 | 1.0513 | 1 | 189.3618 | 51.1105 | 60138.0805 | 62857.4768 | 1253.1495 | 1153.1495 | 952.7516
3912 | 99.5865 | 1.0362 | 0.98447 | 1.0064 | 198.3497 | 48.0335 | 52745.1058 | 51953.3335 | 1035.0667 | 935.0667 | 952.8087
3913 | 113.1776 | 1.0548 | 0.99496 | 0.99235 | 190.0434 | 51.0006 | 56633.6357 | 55932.7061 | 1114.6541 | 1014.6541 | 952.8448
3914 | 93.0042 | 1.0091 | 1.1581 | 0.97628 | 184.4857 | 49.1536 | 60369.8341 | 59150.8725 | 1179.0175 | 1079.0175 | 952.8817
3915 | 113.7047 | 0.90137 | 1.1357 | 1.0365 | 215.6449 | 48.7748 | 62633.2863 | 64271.358 | 1281.4272 | 1181.4272 | 952.9096
3916 | 108.9382 | 0.9338 | 0.99511 | 0.96991 | 195.7267 | 50.3985 | 53965.4113 | 53749.5077 | 1070.9902 | 970.9902 | 952.9377
3917 | 84.7944 | 0.96311 | 1.1876 | 1.0837 | 183.6055 | 48.9018 | 67115.3144 | 68355.3435 | 1363.1069 | 1263.1069 | 952.964
3918 | 90.2257 | 0.94412 | 0.79302 | 1.0298 | 211.2659 | 47.0706 | 42661.2653 | 42110.5151 | 838.2103 | 738.2103 | 952.9696
3919 | 112.0303 | 0.97266 | 1.1149 | 1.0115 | 213.2345 | 51.2816 | 63632.4759 | 61554.4532 | 1227.0891 | 1127.0891 | 952.9702
3920 | 76.3511 | 1.0366 | 0.99252 | 0.9745 | 202.3622 | 54.1492 | 56861.8826 | 57777.1527 | 1151.5431 | 1051.5431 | 953.0987
3921 | 125.7228 | 1.0169 | 1.0014 | 0.95018 | 202.1674 | 47.8208 | 51817.9224 | 51632.6114 | 1028.6522 | 928.6522 | 953.1146
3922 | 109.1941 | 1.0001 | 0.88115 | 0.91382 | 209.4423 | 49.8643 | 45806.4127 | 45082.4128 | 897.6483 | 797.6483 | 953.1637
3923 | 100.1433 | 0.98571 | 1.0521 | 0.94682 | 187.1858 | 54.3175 | 59658.5772 | 60250.052 | 1201.001 | 1101.001 | 953.1814
3924 | 96.9566 | 0.99801 | 0.87741 | 0.99172 | 191.1952 | 46.6832 | 45329.4488 | 44007.9811 | 876.1596 | 776.1596 | 953.2151
3925 | 84.5782 | 0.96236 | 0.97519 | 1.111 | 203.2856 | 51.9919 | 60765.2839 | 62110.7949 | 1238.2159 | 1138.2159 | 953.284
3926 | 98.4916 | 1.101 | 1.0598 | 0.98742 | 200.398 | 51.3689 | 59525.1916 | 59423.7 | 1184.474 | 1084.474 | 953.3308
3927 | 95.1278 | 0.9619 | 0.99442 | 0.93858 | 215.0328 | 51.2422 | 52730.527 | 52844.0869 | 1052.8817 | 952.8817 | 953.3617
3928 | 99.5784 | 1.0223 | 1.0662 | 1.0492 | 200.2126 | 51.1319 | 62602.9977 | 60961.5856 | 1215.2317 | 1115.2317 | 953.374
3929 | 105.6022 | 1.0669 | 1.0455 | 0.98589 | 201.5768 | 50.1783 | 57574.4994 | 57660.7787 | 1149.2156 | 1049.2156 | 953.3757
3930 | 104.3054 | 1.0313 | 1.0497 | 1.0079 | 204.8741 | 52.7822 | 61723.9871 | 62268.0126 | 1241.3603 | 1141.3603 | 953.4358
3931 | 113.8538 | 1.0579 | 0.93148 | 1.0249 | 194.83 | 49.3636 | 53264.7944 | 52388.0509 | 1043.761 | 943.761 | 953.4672
3932 | 88.862 | 0.96981 | 0.80427 | 1.0274 | 204.5399 | 46.9938 | 43085.3886 | 44031.4742 | 876.6295 | 776.6295 | 953.4838
3933 | 95.6552 | 0.93347 | 1.1638 | 1.035 | 184.1354 | 46.5525 | 60417.9645 | 59582.2279 | 1187.6446 | 1087.6446 | 953.5301
3934 | 112.8009 | 1.0474 | 0.82229 | 1.0488 | 195.0244 | 50.5074 | 49720.4116 | 48820.3432 | 972.4069 | 872.4069 | 953.5303
3935 | 103.4782 | 1.0435 | 0.96521 | 0.91671 | 201.6207 | 49.6888 | 49532.4374 | 50452.4356 | 1005.0487 | 905.0487 | 953.5357
3936 | 98.2706 | 1.0114 | 0.87793 | 1.0096 | 199.4115 | 51.5167 | 50983.0758 | 51204.5871 | 1020.0917 | 920.0917 | 953.5734
3937 | 104.2978 | 0.9815 | 0.66057 | 0.92329 | 186.0274 | 52.8985 | 37863.7722 | 38269.8938 | 761.3979 | 661.3979 | 953.5831
3938 | 98.7441 | 0.83687 | 0.9186 | 1.0236 | 207.8191 | 51.2311 | 52612.2536 | 51469.4968 | 1025.3899 | 925.3899 | 953.7689
3939 | 98.9501 | 1.0478 | 1.091 | 0.97446 | 183.8261 | 53.2691 | 62336.5927 | 62979.8371 | 1255.5967 | 1155.5967 | 953.8017
3940 | 106.3435 | 0.96953 | 1.1086 | 0.95752 | 201.6194 | 49.0239 | 57294.1015 | 58426.8286 | 1164.5366 | 1064.5366 | 953.8735
3941 | 82.3497 | 0.99945 | 1.0299 | 0.93652 | 207.0209 | 52.8209 | 55503.3608 | 56120.9188 | 1118.4184 | 1018.4184 | 953.8764
3942 | 99.1793 | 1.0063 | 1.1378 | 1.07 | 186.3331 | 48.8039 | 64474.7293 | 62876.1489 | 1253.523 | 1153.523 | 953.8984
3943 | 86.2552 | 0.97365 | 1.0232 | 1.0482 | 199.139 | 50.2398 | 58299.7089 | 58896.4805 | 1173.9296 | 1073.9296 | 953.9261
3944 | 98.6359 | 1.0025 | 0.85584 | 0.98955 | 197.6953 | 48.0994 | 45688.8227 | 45565.5268 | 907.3105 | 807.3105 | 954.026
3945 | 81.8307 | 0.92457 | 0.96072 | 0.92395 | 210.9809 | 50.1184 | 48490.991 | 49752.6909 | 991.0538 | 891.0538 | 954.0429
3946 | 108.3162 | 0.98384 | 0.86218 | 1.0584 | 208.26 | 49.0272 | 50172.7625 | 49541.249 | 986.825 | 886.825 | 954.0543
3947 | 77.5426 | 0.95847 | 1.1749 | 0.9775 | 205.642 | 49.1711 | 60331.589 | 60496.627 | 1205.9325 | 1105.9325 | 954.0551
3948 | 107.6162 | 1.0592 | 0.92183 | 1.0027 | 196.3555 | 49.6758 | 51774.9944 | 51938.5133 | 1034.7703 | 934.7703 | 954.0938
3949 | 93.3557 | 0.97169 | 1.0185 | 1.0572 | 210.8474 | 48.976 | 57388.6098 | 55020.4415 | 1096.4088 | 996.4088 | 954.1011
3950 | 83.9141 | 1.0155 | 1.1563 | 1.0231 | 182.7364 | 53.9963 | 68663.3787 | 69403.9436 | 1384.0789 | 1284.0789 | 954.1415
3951 | 89.38 | 1.0404 | 0.97271 | 1.0689 | 207.3646 | 55.5534 | 63132.7519 | 66145.0382 | 1318.9008 | 1218.9008 | 954.2001
3952 | 115.7658 | 1.0025 | 1.1441 | 0.99301 | 205.3532 | 44.9723 | 56518.7993 | 56707.767 | 1130.1553 | 1030.1553 | 954.2855
3953 | 91.6676 | 0.91091 | 1.0489 | 0.95661 | 209.4185 | 45.4831 | 49646.2125 | 50123.8897 | 998.4778 | 898.4778 | 954.313
3954 | 103.4778 | 1.0521 | 1.1215 | 0.95162 | 208.3434 | 48.5778 | 57342.1148 | 55820.7803 | 1112.4156 | 1012.4156 | 954.3508
3955 | 125.9814 | 1.0041 | 1.007 | 0.95845 | 196.4434 | 50.1049 | 54893.1904 | 55678.5051 | 1109.5701 | 1009.5701 | 954.4063
3956 | 109.0334 | 0.95324 | 1.0432 | 0.98903 | 200.1903 | 51.0296 | 58155.3198 | 59965.6732 | 1195.3135 | 1095.3135 | 954.5986
3957 | 94.6156 | 1.0621 | 0.97847 | 1.0496 | 198.6489 | 52.0097 | 58839.2084 | 59941.7705 | 1194.8354 | 1094.8354 | 954.6003
3958 | 95.1181 | 0.9208 | 0.97672 | 0.93928 | 210.9166 | 50.7233 | 51187.7685 | 51210.1592 | 1020.2032 | 920.2032 | 954.6407
3959 | 111.7697 | 0.9647 | 0.98186 | 1.0547 | 194.1237 | 50.6744 | 58136.8861 | 56804.979 | 1132.0996 | 1032.0996 | 954.7518
3960 | 72.887 | 1.0743 | 1.0845 | 0.93407 | 197.5325 | 46.942 | 51427.4613 | 52714.785 | 1050.2957 | 950.2957 | 954.8083
3961 | 90.2403 | 1.077 | 0.97185 | 1.0144 | 200.4216 | 46.1128 | 50141.1762 | 50853.7265 | 1013.0745 | 913.0745 | 954.8472
3962 | 85.5995 | 1.0273 | 0.87668 | 0.95071 | 206.99 | 53.6 | 49594.3422 | 49436.7577 | 984.7352 | 884.7352 | 954.9339
3963 | 111.6522 | 1.0296 | 0.92876 | 1.0135 | 193.5411 | 49.9378 | 52942.9009 | 53752.3429 | 1071.0469 | 971.0469 | 954.9371
3964 | 114.8257 | 0.97085 | 1.1571 | 1.0417 | 185.7044 | 49.7915 | 65756.9017 | 66132.0607 | 1318.6412 | 1218.6412 | 955.0658
3965 | 87.7273 | 1.0608 | 1.22 | 1.0323 | 205.9909 | 47.8194 | 64881.9144 | 66402.1377 | 1324.0428 | 1224.0428 | 955.0671
3966 | 94.2663 | 1.0125 | 0.97705 | 1.0239 | 202.1695 | 52.1632 | 57366.2695 | 55397.6175 | 1103.9524 | 1003.9524 | 955.0797
3967 | 121.0687 | 1.0563 | 1.1481 | 1.0166 | 206.4401 | 46.1956 | 60032.9344 | 60126.2722 | 1198.5254 | 1098.5254 | 955.1146
3968 | 89.7823 | 1.0295 | 1.0983 | 1.1171 | 193.87 | 48.9623 | 64791.5716 | 65334.9558 | 1302.6991 | 1202.6991 | 955.1303
3969 | 96.1691 | 1.0405 | 0.81659 | 0.98604 | 192.3052 | 51.2024 | 46543.8789 | 46804.168 | 932.0834 | 832.0834 | 955.176
3970 | 89.9798 | 1.0196 | 0.85696 | 0.97959 | 190.8331 | 46.2875 | 43294.3564 | 43173.25 | 859.465 | 759.465 | 955.2044
3971 | 87.1016 | 1.0026 | 0.89976 | 1.0295 | 220.9631 | 50.5277 | 51436.3418 | 50413.3004 | 1004.266 | 904.266 | 955.2868
3972 | 119.8169 | 0.97203 | 0.59523 | 0.93861 | 207.0644 | 49.0158 | 33300.0121 | 32120.8574 | 638.4171 | 538.4171 | 955.2961
3973 | 117.8279 | 1.0411 | 0.92924 | 1.033 | 201.6087 | 52.3887 | 56918.3218 | 57950.6763 | 1155.0135 | 1055.0135 | 955.3321
3974 | 108.6766 | 0.97334 | 1.0323 | 1.0254 | 219.2249 | 51.9685 | 60728.1052 | 59510.7136 | 1186.2143 | 1086.2143 | 955.349
3975 | 108.3253 | 1.0627 | 1.064 | 0.94562 | 207.1434 | 48.5394 | 54630.0706 | 55735.1464 | 1110.7029 | 1010.7029 | 955.3942
3976 | 90.9701 | 0.98995 | 0.92926 | 0.94351 | 205.737 | 48.2117 | 46817.6938 | 48372.4654 | 963.4493 | 863.4493 | 955.4108
3977 | 79.5014 | 0.97706 | 0.98315 | 1.0433 | 221.3994 | 47.4411 | 52569.1067 | 52110.0183 | 1038.2004 | 938.2004 | 955.4276
3978 | 103.4501 | 0.98297 | 1.1445 | 1.0249 | 210.3656 | 46.3452 | 59286.8379 | 59967.4217 | 1195.3484 | 1095.3484 | 955.4645
3979 | 102.8904 | 1.0773 | 0.95157 | 1.0155 | 196.7039 | 54.8509 | 59280.4461 | 59074.6379 | 1177.4928 | 1077.4928 | 955.5034
3980 | 86.5076 | 0.99013 | 1.0397 | 0.98701 | 213.7524 | 47.7832 | 53340.2781 | 53331.599 | 1062.632 | 962.632 | 955.5738
3981 | 103.9737 | 1.0401 | 0.97622 | 0.95202 | 198.9518 | 44.9994 | 46886.7288 | 46604.03 | 928.0806 | 828.0806 | 955.6134
3982 | 97.9654 | 0.95048 | 1.0529 | 1.013 | 194.3792 | 48.6728 | 56636.9969 | 56325.3584 | 1122.5072 | 1022.5072 | 955.6535
3983 | 86.3269 | 0.94563 | 0.91666 | 0.96934 | 204.828 | 53.4779 | 52088.7064 | 51481.0783 | 1025.6216 | 925.6216 | 955.8354
3984 | 72.5827 | 0.99221 | 0.8768 | 1.0274 | 204.2195 | 48.5356 | 47419.8687 | 47540.3424 | 946.8068 | 846.8068 | 955.8727
3985 | 106.6393 | 1.0004 | 1.2213 | 1.0867 | 198.9927 | 53.5479 | 76976.0633 | 75855.5901 | 1513.1118 | 1413.1118 | 955.8932
3986 | 108.7082 | 0.99675 | 1.1273 | 0.88385 | 210.0137 | 51.0434 | 56596.5606 | 57110.4738 | 1138.2095 | 1038.2095 | 955.9178
3987 | 113.3429 | 0.96941 | 0.90439 | 1.0873 | 198.3157 | 49.8661 | 54712.7056 | 55559.2177 | 1107.1844 | 1007.1844 | 956.0372
3988 | 84.4022 | 1.0493 | 0.99195 | 1.0479 | 187.0043 | 50.8547 | 57554.0369 | 58065.6401 | 1157.3128 | 1057.3128 | 956.0967
3989 | 108.1995 | 0.86413 | 1.2891 | 1.0106 | 205.8005 | 52.1514 | 73024.8119 | 70492.0127 | 1405.8403 | 1305.8403 | 956.1465
3990 | 109.9252 | 1.0319 | 1.0116 | 0.97195 | 203.4106 | 51.0924 | 56232.7708 | 56860.0154 | 1133.2003 | 1033.2003 | 956.2529
3991 | 97.0793 | 1.0624 | 1.2235 | 0.92436 | 191.6308 | 53.8158 | 66604.2076 | 66077.8724 | 1317.5574 | 1217.5574 | 956.2545
3992 | 89.0189 | 0.95803 | 1.1268 | 0.88561 | 196.8007 | 45.257 | 49220.0438 | 49480.0735 | 985.6015 | 885.6015 | 956.2763
3993 | 97.84 | 0.98757 | 0.83915 | 1.0941 | 197.0864 | 51.9 | 52861.5697 | 53910.4462 | 1074.2089 | 974.2089 | 956.2958
3994 | 99.9386 | 0.99003 | 0.95811 | 0.93647 | 218.2024 | 47.0332 | 47071.5212 | 46904.1146 | 934.0823 | 834.0823 | 956.313
3995 | 88.0549 | 0.97864 | 0.98011 | 1.1041 | 206.6601 | 48.6165 | 57006.3169 | 56453.6336 | 1125.0727 | 1025.0727 | 956.3637
3996 | 88.9998 | 0.97363 | 1.0935 | 0.97346 | 186.4413 | 49.2948 | 56931.6255 | 58427.1659 | 1164.5433 | 1064.5433 | 956.5169
3997 | 96.612 | 1.0824 | 0.86461 | 1.0302 | 202.5658 | 52.263 | 52221.7586 | 51165.2827 | 1019.3057 | 919.3057 | 956.5629
3998 | 81.4165 | 0.98541 | 1.1413 | 0.93248 | 210.8435 | 52.4223 | 60206.2402 | 59215.1634 | 1180.3033 | 1080.3033 | 956.6174
3999 | 90.6513 | 0.94696 | 0.97933 | 1.0078 | 211.107 | 49.1535 | 52945.6429 | 52669.794 | 1049.3959 | 949.3959 | 956.6629
4000 | 103.9824 | 0.93804 | 1.1327 | 0.977 | 199.7624 | 48.1092 | 58130.9969 | 60197.9238 | 1199.9585 | 1099.9585 | 956.7319
4001 | 93.2284 | 0.95454 | 1.2804 | 0.93431 | 191.2641 | 53.9743 | 69561.1085 | 69180.6787 | 1379.6136 | 1279.6136 | 956.7889
4002 | 104.9276 | 0.95846 | 0.9668 | 0.85893 | 188.4783 | 46.0983 | 43105.1874 | 44343.4598 | 882.8692 | 782.8692 | 956.7895
4003 | 108.5147 | 1.0402 | 0.92137 | 1.0479 | 206.264 | 47.8527 | 51809.9432 | 51965.6935 | 1035.3139 | 935.3139 | 956.821
4004 | 112.0408 | 1.002 | 1.1236 | 1.0685 | 211.3112 | 45.8165 | 60361.2483 | 60675.5532 | 1209.5111 | 1109.5111 | 956.8395
4005 | 99.0935 | 1.052 | 0.97654 | 0.96084 | 209.8645 | 48.1444 | 50402.4073 | 52303.1303 | 1042.0626 | 942.0626 | 956.8506
4006 | 105.3558 | 0.95779 | 1.0658 | 0.95136 | 198.827 | 47.2126 | 52832.7115 | 52481.4785 | 1045.6296 | 945.6296 | 956.8532
4007 | 99.6758 | 0.84718 | 1.0914 | 0.95361 | 197.8657 | 52.2746 | 59019.9521 | 59752.0407 | 1191.0408 | 1091.0408 | 956.9344
4008 | 117.1618 | 1.0292 | 1.0461 | 1.021 | 186.6697 | 53.6794 | 63992.1554 | 62018.8873 | 1236.3777 | 1136.3777 | 956.946
4009 | 99.7741 | 1.0407 | 1.2527 | 1.0004 | 212.1935 | 53.6069 | 72955.9131 | 75048.1115 | 1496.9622 | 1396.9622 | 957.0199
4010 | 79.5897 | 0.97675 | 0.98659 | 1.0461 | 189.2545 | 52.3096 | 58243.7874 | 58905.5305 | 1174.1106 | 1074.1106 | 957.0978
4011 | 121.7056 | 1.0771 | 0.89743 | 0.96991 | 198.0612 | 54.651 | 54932.472 | 54813.7622 | 1092.2752 | 992.2752 | 957.1008
4012 | 97.2085 | 1.1381 | 0.91206 | 0.98356 | 199.2273 | 48.7445 | 49319.4489 | 49210.3765 | 980.2075 | 880.2075 | 957.1108
4013 | 96.9758 | 0.95616 | 1.4111 | 1.0224 | 209.3466 | 49.3013 | 75911.3209 | 78317.65 | 1562.353 | 1462.353 | 957.1244
4014 | 96.2601 | 1.1202 | 0.83035 | 0.95733 | 195.6037 | 52.7719 | 47835.3966 | 47908.4504 | 954.169 | 854.169 | 957.1358
4015 | 103.1278 | 1.1115 | 1.1259 | 1.0506 | 196.1026 | 53.7091 | 69882.6858 | 69468.6384 | 1385.3728 | 1285.3728 | 957.1745
4016 | 92.352 | 0.98042 | 0.89351 | 0.98176 | 210.2052 | 50.7362 | 49310.7309 | 48632.9347 | 968.6587 | 868.6587 | 957.1779
4017 | 122.6715 | 0.97077 | 0.76793 | 0.96259 | 183.4957 | 54.0728 | 46593.2465 | 47236.6329 | 940.7327 | 840.7327 | 957.3207
4018 | 107.2058 | 0.99094 | 0.96919 | 1.0038 | 215.5472 | 50.1349 | 54314.8095 | 54673.5576 | 1089.4712 | 989.4712 | 957.3746
4019 | 99.4425 | 1.0227 | 1.1228 | 1.0377 | 218.8626 | 51.4545 | 65399.0481 | 67508.7802 | 1346.1756 | 1246.1756 | 957.4184
4020 | 95.6468 | 1.0606 | 0.71516 | 0.95419 | 183.208 | 50.1213 | 39470.3483 | 40106.6752 | 798.1335 | 698.1335 | 957.451
4021 | 106.3762 | 0.91669 | 0.96052 | 1.0893 | 214.4727 | 48.9465 | 56200.9385 | 56481.2828 | 1125.6257 | 1025.6257 | 957.4815
4022 | 105.8435 | 1.0008 | 1.0362 | 0.98517 | 206.2366 | 50.5979 | 57219.532 | 56441.1137 | 1124.8223 | 1024.8223 | 957.5082
4023 | 105.7858 | 0.99737 | 1.1454 | 0.9343 | 200.438 | 50.7533 | 59868.7613 | 58787.6103 | 1171.7522 | 1071.7522 | 957.6542
4024 | 97.7115 | 0.9797 | 0.75281 | 1.08 | 193.4771 | 47.8324 | 43661.4466 | 43059.2407 | 857.1848 | 757.1848 | 957.6764
4025 | 90.3265 | 0.93171 | 0.89639 | 1.0494 | 181.4883 | 48.3399 | 49723.1775 | 50507.498 | 1006.15 | 906.15 | 957.6968
4026 | 105.4798 | 1.0163 | 0.82982 | 0.93675 | 183.8071 | 49.9805 | 44393.0213 | 45211.4007 | 900.228 | 800.228 | 957.7338
4027 | 101.0336 | 1.032 | 0.88379 | 1.0454 | 190.675 | 48.9148 | 50484.0597 | 48994.746 | 975.8949 | 875.8949 | 957.7708
4028 | 96.669 | 1.0572 | 1.2047 | 1.0368 | 189.7455 | 49.813 | 67494.463 | 66640.8324 | 1328.8166 | 1228.8166 | 957.8824
4029 | 84.3948 | 0.91512 | 0.98518 | 1.0561 | 184.7599 | 46.4948 | 52153.272 | 51738.3674 | 1030.7673 | 930.7673 | 957.955
4030 | 99.8363 | 0.97754 | 1.0873 | 0.95458 | 203.4216 | 49.4799 | 56390.4535 | 55845.4864 | 1112.9097 | 1012.9097 | 957.9599
4031 | 101.1318 | 1.0128 | 1.069 | 0.97056 | 195.3402 | 48.0478 | 54967.7112 | 54729.342 | 1090.5868 | 990.5868 | 957.9755
4032 | 91.0494 | 1.0108 | 1.0712 | 1.0005 | 193.2238 | 52.3101 | 61070.6591 | 62114.43 | 1238.2886 | 1138.2886 | 957.9835
4033 | 95.6312 | 0.9208 | 1.1698 | 0.96336 | 206.6192 | 49.7074 | 60600.8707 | 58723.7069 | 1170.4741 | 1070.4741 | 958.0051
4034 | 101.7607 | 1.0882 | 1.2868 | 1.0451 | 199.0999 | 48.6925 | 71073.4479 | 72223.528 | 1440.4706 | 1340.4706 | 958.033
4035 | 98.4933 | 1.0189 | 0.66202 | 1.0179 | 204.5389 | 50.9107 | 39620.6546 | 38857.5769 | 773.1515 | 673.1515 | 958.0525
4036 | 97.6585 | 0.96809 | 1.0903 | 1.0002 | 191.9441 | 44.6289 | 53079.6208 | 53499.7524 | 1065.995 | 965.995 | 958.0728
4037 | 91.8937 | 0.98434 | 1.0403 | 0.90847 | 197.9186 | 50.361 | 52347.2615 | 51982.1949 | 1035.6439 | 935.6439 | 958.118
4038 | 103.17 | 1.0072 | 1.0018 | 0.99485 | 201.1585 | 52.6091 | 58097.99 | 57971.1034 | 1155.4221 | 1055.4221 | 958.1785
4039 | 97.2037 | 0.96735 | 0.81857 | 1.0115 | 209.9394 | 52.5166 | 48632.8987 | 50271.3795 | 1001.4276 | 901.4276 | 958.297
4040 | 107.8528 | 1.0419 | 0.83407 | 0.95608 | 213.3847 | 52.2964 | 47793.306 | 47514.2013 | 946.284 | 846.284 | 958.4097
4041 | 98.8658 | 0.96082 | 1.2769 | 0.89981 | 190.304 | 52.8641 | 65948.976 | 70412.3226 | 1404.2465 | 1304.2465 | 958.4354
4042 | 103.9474 | 1.0094 | 1.1987 | 1.0614 | 203.3676 | 52.6358 | 72694.0309 | 76079.8176 | 1517.5964 | 1417.5964 | 958.5185
4043 | 100.7149 | 0.97189 | 1.039 | 0.95564 | 193.1905 | 50.7047 | 55500.7456 | 57679.1316 | 1149.5826 | 1049.5826 | 958.5208
4044 | 89.535 | 0.97873 | 1.1131 | 1.0046 | 203.0818 | 51.521 | 62332.4798 | 63638.0794 | 1268.7616 | 1168.7616 | 958.5887
4045 | 121.5218 | 1.0334 | 0.83059 | 1.0612 | 194.573 | 47.5398 | 48067.8027 | 49718.6373 | 990.3727 | 890.3727 | 958.5928
4046 | 95.3387 | 0.95161 | 0.94718 | 1.034 | 203.9324 | 50.2847 | 54015.0167 | 53542.1007 | 1066.842 | 966.842 | 958.5997
4047 | 101.0846 | 1.0395 | 1.2358 | 0.98186 | 196.8822 | 48.9493 | 64733.1739 | 64224.6671 | 1280.4933 | 1180.4933 | 958.692
4048 | 104.4085 | 1.0025 | 1.1061 | 0.99407 | 206.8597 | 49.5357 | 59859.257 | 63871.6382 | 1273.4328 | 1173.4328 | 958.7143
4049 | 78.2868 | 0.95669 | 1.0427 | 0.85871 | 209.0894 | 52.7223 | 51364.2988 | 51492.7602 | 1025.8552 | 925.8552 | 958.7459
4050 | 107.0703 | 1.0344 | 1.1989 | 1.0219 | 193.7241 | 48.7127 | 65267.3949 | 65160.2323 | 1299.2046 | 1199.2046 | 958.762
4051 | 104.5029 | 0.9947 | 1.1999 | 1.0274 | 216.5342 | 53.3453 | 71524.2526 | 73939.142 | 1474.7828 | 1374.7828 | 958.7886
4052 | 119.4244 | 1.0072 | 1.1585 | 1.0761 | 196.7239 | 49.7931 | 68261.3083 | 69541.4027 | 1386.8281 | 1286.8281 | 958.8023
4053 | 95.0474 | 1.0994 | 1.4155 | 0.91477 | 188.624 | 50.3519 | 70648.28 | 70574.7068 | 1407.4941 | 1307.4941 | 958.8283
4054 | 87.9885 | 0.92379 | 1.0034 | 0.90944 | 205.3284 | 52.4815 | 52364.2068 | 52840.4373 | 1052.8087 | 952.8087 | 958.8509
4055 | 104.8661 | 1.0274 | 1.0959 | 0.94086 | 220.0278 | 51.0033 | 58305.8544 | 57569.466 | 1147.3893 | 1047.3893 | 958.9078
4056 | 94.448 | 0.96434 | 1.1528 | 1.018 | 199.8318 | 49.9533 | 63371.6736 | 63344.0942 | 1262.8819 | 1162.8819 | 958.9476
4057 | 94.1715 | 0.9936 | 0.95202 | 0.99127 | 189.3995 | 48.3151 | 50305.8139 | 52070.8658 | 1037.4173 | 937.4173 | 958.9928
4058 | 86.5957 | 1.021 | 0.87453 | 1.0574 | 203.2716 | 54.8926 | 55818.5192 | 54127.8929 | 1078.5579 | 978.5579 | 959.0195
4059 | 109.8721 | 1.043 | 1.0892 | 1.0406 | 196.943 | 48.6497 | 60916.6882 | 60948.394 | 1214.9679 | 1114.9679 | 959.0557
4060 | 105.4235 | 0.93283 | 1.0274 | 0.98947 | 188.6193 | 48.4904 | 54252.601 | 52724.8261 | 1050.4965 | 950.4965 | 959.0973
4061 | 98.383 | 1.0396 | 1.0231 | 1.0277 | 206.059 | 43.3423 | 50214.6495 | 51179.8474 | 1019.5969 | 919.5969 | 959.1152
4062 | 100.1266 | 1.1021 | 1.0111 | 1.0761 | 197.7061 | 47.6156 | 57257.6444 | 57347.0785 | 1142.9416 | 1042.9416 | 959.1363
4063 | 106.9546 | 0.97752 | 1.1905 | 1.0514 | 171.167 | 51.653 | 70226.7623 | 70683.942 | 1409.6788 | 1309.6788 | 959.1701
4064 | 99.9011 | 0.9851 | 0.90069 | 1.0161 | 194.5391 | 50.0303 | 50905.1716 | 51526.8803 | 1026.5376 | 926.5376 | 959.2692
4065 | 108.8921 | 0.98035 | 1.2039 | 0.96684 | 200.6136 | 51.0222 | 65037.1203 | 64398.6903 | 1283.9738 | 1183.9738 | 959.2746
4066 | 112.7933 | 0.9545 | 1.0026 | 0.93796 | 181.1251 | 51.3445 | 53992.6595 | 52970.5407 | 1055.4108 | 955.4108 | 959.3156
4067 | 103.7118 | 0.94148 | 1.2153 | 1.0081 | 205.4514 | 50.5565 | 67078.2508 | 64941.4051 | 1294.8281 | 1194.8281 | 959.3254
4068 | 65.5819 | 0.98352 | 1.003 | 0.87815 | 210.4663 | 52.8884 | 50206.386 | 51558.114 | 1027.1623 | 927.1623 | 959.4485
4069 | 81.3795 | 0.98865 | 0.85399 | 0.94078 | 189.5549 | 50.6461 | 44954.0735 | 44470.2556 | 885.4051 | 785.4051 | 959.4979
4070 | 100.8302 | 0.9914 | 0.9325 | 1.0092 | 189.8551 | 50.3429 | 52599.2806 | 52532.3394 | 1046.6468 | 946.6468 | 959.5356
4071 | 113.7132 | 1.0889 | 0.87022 | 0.93759 | 193.1062 | 51.6442 | 48725.2733 | 49165.792 | 979.3158 | 879.3158 | 959.5527
4072 | 97.2224 | 0.96793 | 1.1227 | 0.99911 | 222.2082 | 54.0969 | 65993.7885 | 69104.3054 | 1378.0861 | 1278.0861 | 959.5901
4073 | 83.6292 | 1.0134 | 0.81949 | 0.94982 | 199.7288 | 48.3643 | 41943.9268 | 39152.173 | 779.0435 | 679.0435 | 959.6529
4074 | 99.5809 | 1.0018 | 1.0658 | 1.1364 | 198.5235 | 45.8745 | 60336.7493 | 60514.3501 | 1206.287 | 1106.287 | 959.681
4075 | 90.2656 | 0.98091 | 0.94047 | 0.98662 | 198.4918 | 45.9971 | 46951.2146 | 48919.9828 | 974.3997 | 874.3997 | 959.7277
4076 | 82.9378 | 0.98578 | 0.87593 | 0.96397 | 198.9583 | 49.177 | 45743.0143 | 46452.0129 | 925.0403 | 825.0403 | 959.7726
4077 | 95.3769 | 0.99071 | 1.2122 | 0.96578 | 201.9421 | 47.4419 | 60224.0838 | 58804.2882 | 1172.0858 | 1072.0858 | 959.7914
4078 | 94.9071 | 1.0157 | 0.96343 | 1.007 | 187.246 | 51.4524 | 55064.6977 | 56138.6723 | 1118.7734 | 1018.7734 | 959.7978
4079 | 95.7502 | 1.0228 | 1.0702 | 0.98147 | 201.7824 | 56.8707 | 65503.9273 | 64275.1536 | 1281.5031 | 1181.5031 | 959.8162
4080 | 105.5365 | 0.89851 | 1.2352 | 0.98462 | 200.5678 | 54.36 | 71467.8852 | 69214.2179 | 1380.2844 | 1280.2844 | 959.833
4081 | 90.4411 | 0.95667 | 0.83789 | 1.0274 | 195.3663 | 50.0909 | 47649.0025 | 48456.1665 | 965.1233 | 865.1233 | 959.8465
4082 | 104.067 | 1.029 | 1.1569 | 1.0221 | 206.1525 | 52.2816 | 67625.907 | 66829.4972 | 1332.5899 | 1232.5899 | 959.9287
4083 | 109.308 | 0.9602 | 0.78081 | 1.0813 | 203.4503 | 54.1827 | 51637.2346 | 51173.4751 | 1019.4695 | 919.4695 | 959.9386
4084 | 87.4351 | 0.90105 | 0.82445 | 1.0497 | 217.088 | 50.3166 | 47727.6195 | 48680.0141 | 969.6003 | 869.6003 | 960.0388
4085 | 95.8078 | 1.114 | 1.2941 | 0.90725 | 183.6339 | 50.3253 | 64639.4476 | 64995.6469 | 1295.9129 | 1195.9129 | 960.0482
4086 | 96.5659 | 0.96834 | 1.171 | 1.0005 | 212.7784 | 49.2642 | 62537.3494 | 62380.0421 | 1243.6008 | 1143.6008 | 960.057
4087 | 106.5373 | 1.0562 | 1.1523 | 1.0588 | 195.5159 | 47.9394 | 64073.5033 | 60970.0199 | 1215.4004 | 1115.4004 | 960.098
4088 | 118.1701 | 1.0474 | 1.0187 | 1.0338 | 197.5363 | 50.817 | 60005.8444 | 59956.87 | 1195.1374 | 1095.1374 | 960.1103
4089 | 91.1076 | 0.92921 | 1.0947 | 1.0053 | 206.5858 | 50.3634 | 59897.4746 | 58094.9331 | 1157.8987 | 1057.8987 | 960.1644
4090 | 105.189 | 1.0598 | 1.1401 | 1.0293 | 221.2372 | 44.7246 | 57690.8821 | 55465.1437 | 1105.3029 | 1005.3029 | 960.1693
4091 | 93.362 | 0.94352 | 1.0867 | 1.0138 | 204.2865 | 46.5645 | 55607.545 | 54880.4272 | 1093.6085 | 993.6085 | 960.248
4092 | 95.2488 | 0.96748 | 0.92167 | 1.0408 | 206.2463 | 50.422 | 53222.1416 | 54178.7601 | 1079.5752 | 979.5752 | 960.2703
4093 | 87.6441 | 1.0304 | 0.99916 | 1.0266 | 177.7254 | 55.2051 | 61789.6453 | 61029.2377 | 1216.5848 | 1116.5848 | 960.2758
4094 | 116.3834 | 0.99983 | 1.0269 | 0.99423 | 191.0728 | 52.2971 | 59669.7247 | 58155.0479 | 1159.101 | 1059.101 | 960.3066
4095 | 109.0496 | 0.91239 | 1.1113 | 1.0523 | 197.5195 | 51.1647 | 65120.8311 | 65063.708 | 1297.2742 | 1197.2742 | 960.3075
4096 | 113.6871 | 0.98221 | 0.99557 | 0.94026 | 187.2667 | 52.3418 | 55028.9373 | 55268.3823 | 1101.3676 | 1001.3676 | 960.3271
4097 | 87.6794 | 0.96145 | 0.79577 | 0.95217 | 186.2046 | 54.0899 | 45730.1906 | 45915.0188 | 914.3004 | 814.3004 | 960.3541
4098 | 118.4743 | 1.0211 | 1.1101 | 0.97987 | 209.922 | 47.1675 | 57224.5746 | 58184.1273 | 1159.6825 | 1059.6825 | 960.3611
4099 | 89.8507 | 0.91147 | 0.91179 | 1.0603 | 203.1666 | 53.03 | 55815.7923 | 54690.127 | 1089.8025 | 989.8025 | 960.3986
4100 | 94.7451 | 0.98919 | 0.84369 | 0.97941 | 202.942 | 50.2785 | 46461.3018 | 46558.944 | 927.1789 | 827.1789 | 960.4273
4101 | 94.1773 | 0.99603 | 1.0495 | 1.0421 | 201.5145 | 49.4533 | 58923.0146 | 59610.6299 | 1188.2126 | 1088.2126 | 960.4658
4102 | 118.5581 | 1.0526 | 0.86004 | 1.0511 | 180.6615 | 48.873 | 50460.285 | 49641.1854 | 988.8237 | 888.8237 | 960.4668
4103 | 94.3399 | 1.0658 | 1.1123 | 1.048 | 202.0073 | 47.9473 | 60917.5995 | 61037.5037 | 1216.7501 | 1116.7501 | 960.467
4104 | 109.2603 | 1.0528 | 1.1642 | 0.98738 | 196.8252 | 56.6862 | 71881.4771 | 70304.6735 | 1402.0935 | 1302.0935 | 960.5052
4105 | 95.1034 | 0.98474 | 1.341 | 1.0145 | 193.5182 | 48.6935 | 71000.2623 | 70865.5965 | 1413.3119 | 1313.3119 | 960.5165
4106 | 95.5239 | 1.0264 | 1.1402 | 1.0472 | 205.9122 | 54.5302 | 70661.4216 | 69868.0269 | 1393.3605 | 1293.3605 | 960.5283
4107 | 98.9568 | 0.98647 | 0.96967 | 1.0282 | 204.5334 | 48.4607 | 53252.3364 | 51940.5194 | 1034.8104 | 934.8104 | 960.5703
4108 | 114.3361 | 1.0044 | 1.0533 | 1.0749 | 203.1277 | 47.7569 | 59759.177 | 58001.9995 | 1156.04 | 1056.04 | 960.5772
4109 | 99.6395 | 0.92929 | 1.0562 | 1.0242 | 198.7998 | 48.6287 | 57305.5398 | 58356.9494 | 1163.139 | 1063.139 | 960.6299
4110 | 87.469 | 0.92956 | 0.9492 | 1.0011 | 194.9174 | 53.3071 | 55182.4583 | 54881.8445 | 1093.6369 | 993.6369 | 960.6528
4111 | 102.129 | 1.0266 | 0.83544 | 0.96762 | 215.7588 | 50.4673 | 46304.1316 | 47690.1765 | 949.8035 | 849.8035 | 960.7193
4112 | 109.4114 | 0.98965 | 0.95462 | 1.0492 | 198.9295 | 50.4919 | 56237.7236 | 55920.2655 | 1114.4053 | 1014.4053 | 960.8502
4113 | 105.7951 | 0.96439 | 0.93616 | 0.95477 | 216.8531 | 46.0526 | 46078.1246 | 47212.3934 | 940.2479 | 840.2479 | 960.943
4114 | 96.0855 | 1.0331 | 0.93754 | 0.98497 | 200.6794 | 48.4128 | 49712.7438 | 50358.6929 | 1003.1739 | 903.1739 | 960.9457
4115 | 96.9326 | 1.0292 | 1.2789 | 0.98104 | 199.3283 | 52.9186 | 71875.1503 | 69878.9472 | 1393.5789 | 1293.5789 | 960.9783
4116 | 108.1538 | 1.0548 | 0.85624 | 0.99391 | 191.1515 | 48.9979 | 47479.5575 | 47185.9544 | 939.7191 | 839.7191 | 961.0226
4117 | 108.3635 | 1.0085 | 0.85156 | 1.0033 | 205.1976 | 44.3188 | 42911.7093 | 42260.6447 | 841.2129 | 741.2129 | 961.0245
4118 | 99.4805 | 1.0085 | 1.2152 | 0.85964 | 197.5101 | 51.0046 | 58597.3399 | 60613.4154 | 1208.2683 | 1108.2683 | 961.0838
4119 | 104.7349 | 0.93624 | 1.0174 | 0.9923 | 206.6941 | 46.8367 | 52084.0337 | 53560.5933 | 1067.2119 | 967.2119 | 961.0896
4120 | 108.7106 | 1.0125 | 0.97238 | 0.97419 | 191.897 | 52.4799 | 55681.4772 | 56947.3873 | 1134.9477 | 1034.9477 | 961.1307
4121 | 92.5306 | 0.97506 | 1.0092 | 0.98326 | 207.1994 | 53.2311 | 57833.5172 | 56583.4324 | 1127.6686 | 1027.6686 | 961.2694
4122 | 98.6947 | 0.9712 | 0.92249 | 1.0027 | 219.2292 | 48.249 | 49472.722 | 50450.2779 | 1005.0056 | 905.0056 | 961.2887
4123 | 115.2008 | 0.95439 | 1.2154 | 1.0239 | 201.301 | 48.051 | 65281.812 | 64831.7359 | 1292.6347 | 1192.6347 | 961.3528
4124 | 104.9275 | 0.9768 | 0.9297 | 0.92941 | 207.5097 | 45.8533 | 44527.7935 | 44267.8131 | 881.3563 | 781.3563 | 961.3587
4125 | 94.298 | 1.0492 | 0.79623 | 0.92274 | 198.5439 | 48.39 | 40539.1089 | 39868.1391 | 793.3628 | 693.3628 | 961.3912
4126 | 96.4286 | 1.0278 | 0.85796 | 1.0085 | 213.0595 | 50.5907 | 48999.2704 | 49699.2463 | 989.9849 | 889.9849 | 961.3965
4127 | 100.5174 | 0.88321 | 1.0711 | 0.95332 | 202.6147 | 53.0622 | 59095.4467 | 56448.0948 | 1124.9619 | 1024.9619 | 961.4024
4128 | 93.636 | 1.058 | 0.92157 | 0.9546 | 194.1725 | 50.5701 | 49692.4878 | 49075.6839 | 977.5137 | 877.5137 | 961.4176
4129 | 110.0424 | 0.94615 | 1.1456 | 1.0317 | 207.093 | 53.0142 | 68388.4819 | 69139.4856 | 1378.7897 | 1278.7897 | 961.4407
4130 | 109.8638 | 1.01 | 0.92787 | 0.99318 | 205.0479 | 48.8319 | 50624.2224 | 49970.3424 | 995.4068 | 895.4068 | 961.4436
4131 | 112.3863 | 0.99442 | 0.83584 | 1.0104 | 184.9659 | 54.2848 | 52096.0717 | 51216.4529 | 1020.3291 | 920.3291 | 961.4745
4132 | 115.9991 | 1.027 | 1.3764 | 0.96766 | 193.3388 | 45.7748 | 66613.1107 | 64864.1807 | 1293.2836 | 1193.2836 | 961.5609
4133 | 99.7693 | 0.83295 | 0.95037 | 1.1239 | 195.7524 | 52.0609 | 60127.9847 | 61109.4873 | 1218.1897 | 1118.1897 | 961.6114
4134 | 99.5019 | 0.92347 | 1.1887 | 1.0271 | 201.2961 | 48.7961 | 64261.9341 | 64196.427 | 1279.9285 | 1179.9285 | 961.6349
4135 | 111.3138 | 0.93863 | 0.89876 | 1.1201 | 200.1858 | 51.4017 | 57318.3411 | 57710.6848 | 1150.2137 | 1050.2137 | 961.6379
4136 | 102.2045 | 0.94297 | 0.82562 | 1.0192 | 214.69 | 49.5527 | 46688.0839 | 46622.9607 | 928.4592 | 828.4592 | 961.6798
4137 | 106.979 | 0.92519 | 0.82973 | 0.99546 | 208.3743 | 50.822 | 47215.7141 | 46385.9132 | 923.7183 | 823.7183 | 961.6906
4138 | 108.1285 | 0.98777 | 1.128 | 0.95393 | 188.3199 | 45.5326 | 54045.3957 | 53154.8663 | 1059.0973 | 959.0973 | 961.7151
4139 | 108.9401 | 1.1127 | 0.91367 | 1.0227 | 195.5115 | 47.5772 | 50417.5357 | 50573.123 | 1007.4625 | 907.4625 | 961.729
4140 | 84.1643 | 0.97501 | 0.8873 | 1.0509 | 206.3242 | 48.5725 | 49483.4032 | 51447.4555 | 1024.9491 | 924.9491 | 961.7481
4141 | 114.8271 | 1.0194 | 0.72077 | 1.0133 | 189.4069 | 53.0888 | 45177.7815 | 47057.6113 | 937.1522 | 837.1522 | 961.7675
4142 | 104.2565 | 0.95475 | 1.0017 | 0.97991 | 181.0431 | 49.9193 | 54150.7026 | 55291.0172 | 1101.8203 | 1001.8203 | 961.8532
4143 | 108.3023 | 1.0007 | 0.94429 | 1.0636 | 209.8266 | 49.3424 | 55114.3275 | 55428.1166 | 1104.5623 | 1004.5623 | 961.9018
4144 | 104.6099 | 0.94947 | 0.96592 | 0.93978 | 198.2974 | 50.7264 | 51283.4953 | 50283.3414 | 1001.6668 | 901.6668 | 961.9263
4145 | 94.7463 | 0.97684 | 1.1185 | 0.99182 | 197.2801 | 44.9878 | 54270.2339 | 54429.3123 | 1084.5862 | 984.5862 | 961.993
4146 | 97.2709 | 1.0703 | 1.0782 | 1.0082 | 204.3378 | 46.7644 | 55905.5119 | 55331.9674 | 1102.6393 | 1002.6393 | 962.0159
4147 | 87.1481 | 1.0617 | 0.86546 | 0.95885 | 213.2623 | 50.5549 | 46843.6052 | 46214.5595 | 920.2912 | 820.2912 | 962.1812
4148 | 109.7261 | 0.98094 | 0.84637 | 0.95907 | 210.731 | 51.7746 | 47810.5898 | 46619.0412 | 928.3808 | 828.3808 | 962.2024
4149 | 94.1137 | 0.98051 | 0.88728 | 1.0792 | 185.3647 | 52.0493 | 54830.1422 | 53734.6553 | 1070.6931 | 970.6931 | 962.2104
4150 | 105.4421 | 0.98979 | 1.072 | 0.91805 | 196.0881 | 47.8163 | 52244.9261 | 51187.0463 | 1019.7409 | 919.7409 | 962.2359
4151 | 109.1331 | 0.95227 | 1.0374 | 0.95521 | 205.7626 | 44.3882 | 48803.4821 | 48779.7392 | 971.5948 | 871.5948 | 962.3035
4152 | 91.2927 | 0.99813 | 1.0429 | 1.0526 | 200.6733 | 51.6549 | 61609.1736 | 61149.3375 | 1218.9868 | 1118.9868 | 962.311
4153 | 126.9973 | 1.0078 | 1.1935 | 1.0698 | 193.3298 | 50.0709 | 70533.9408 | 70888.5652 | 1413.7713 | 1313.7713 | 962.3835
4154 | 108.484 | 0.9545 | 1.0441 | 1.049 | 206.5121 | 48.0076 | 57754.7188 | 57606.6258 | 1148.1325 | 1048.1325 | 962.409
4155 | 105.5936 | 0.98768 | 1.0715 | 0.96736 | 203.7084 | 47.8125 | 54748.4229 | 55440.2166 | 1104.8043 | 1004.8043 | 962.4708
4156 | 109.1641 | 1.1346 | 0.65896 | 1.0183 | 197.2665 | 47.1524 | 37677.1208 | 37178.0938 | 739.5619 | 639.5619 | 962.489
4157 | 106.6088 | 0.92395 | 1.1098 | 1.0387 | 202.7445 | 48.9231 | 61417.45 | 60808.0629 | 1212.1613 | 1112.1613 | 962.5066
4158 | 83.5168 | 1.0193 | 1.0671 | 1.0277 | 200.032 | 48.5555 | 57579.2083 | 59507.2098 | 1186.1442 | 1086.1442 | 962.546
4159 | 113.4253 | 0.95826 | 1.1805 | 0.97253 | 199.8868 | 50.4518 | 63607.4755 | 66866.5732 | 1333.3315 | 1233.3315 | 962.6213
4160 | 103.4087 | 0.98481 | 1.253 | 1.0698 | 200.2832 | 52.9882 | 76627.1951 | 79515.0391 | 1586.3008 | 1486.3008 | 962.6309
4161 | 103.0424 | 0.96127 | 1.1611 | 0.99576 | 202.9916 | 49.9891 | 62949.9243 | 60134.0413 | 1198.6808 | 1098.6808 | 962.632
4162 | 117.6608 | 1.0257 | 0.68422 | 1.0458 | 214.2281 | 52.2493 | 43908.0815 | 43262.2153 | 861.2443 | 761.2443 | 962.6357
4163 | 113.3302 | 1.0685 | 0.95 | 1.0641 | 202.0128 | 47.8391 | 54357.1627 | 54515.7307 | 1086.3146 | 986.3146 | 962.7269
4164 | 104.9553 | 1.0247 | 0.86502 | 1.027 | 197.3055 | 48.7091 | 48706.1711 | 47803.6591 | 952.0732 | 852.0732 | 962.7325
4165 | 106.6459 | 0.99356 | 0.8878 | 0.94887 | 195.767 | 51.4271 | 48967.0259 | 49011.966 | 976.2393 | 876.2393 | 962.7466
4166 | 84.347 | 0.96758 | 1.0417 | 1.0127 | 200.2028 | 50.5659 | 57672.8294 | 58545.0357 | 1166.9007 | 1066.9007 | 962.8228
4167 | 92.053 | 1.0601 | 1.0208 | 1.0219 | 193.208 | 53.7481 | 61503.9983 | 60115.2878 | 1198.3058 | 1098.3058 | 962.8538
4168 | 100.5362 | 0.99261 | 1.096 | 0.97975 | 201.7512 | 52.5135 | 61831.527 | 62664.2597 | 1249.2852 | 1149.2852 | 962.9088
4169 | 105.3274 | 1.019 | 0.98363 | 1.0096 | 199.117 | 51.4374 | 56800.014 | 57663.7777 | 1149.2756 | 1049.2756 | 962.9244
4170 | 104.6624 | 1.068 | 0.90314 | 0.94877 | 192.933 | 51.4613 | 50041.0994 | 50239.8042 | 1000.7961 | 900.7961 | 963.1774
4171 | 96.3068 | 0.95308 | 1.2246 | 1.1126 | 215.4843 | 51.4803 | 75077.7593 | 72857.8139 | 1453.1563 | 1353.1563 | 963.2051
4172 | 96.7684 | 0.9156 | 0.9831 | 0.9626 | 198.5208 | 51.6839 | 53688.002 | 53267.9363 | 1061.3587 | 961.3587 | 963.2243
4173 | 91.7706 | 0.98478 | 0.98575 | 1.0362 | 185.2376 | 51.9596 | 57953.6328 | 59984.5449 | 1195.6909 | 1095.6909 | 963.2781
4174 | 96.0327 | 0.9742 | 1.0351 | 0.9475 | 198.5199 | 54.0519 | 58265.4505 | 57991.4149 | 1155.8283 | 1055.8283 | 963.3009
4175 | 90.006 | 0.94734 | 1.1355 | 0.94993 | 209.7967 | 52.1143 | 60868.3007 | 59402.6361 | 1184.0527 | 1084.0527 | 963.3174
4176 | 85.2146 | 0.89972 | 0.94604 | 0.92252 | 206.7452 | 50.5361 | 48185.9089 | 47515.9343 | 946.3187 | 846.3187 | 963.3473
4177 | 107.9769 | 0.97541 | 0.9374 | 1.0227 | 192.1051 | 52.5526 | 56107.1971 | 56270.6718 | 1121.4134 | 1021.4134 | 963.384
4178 | 116.0796 | 1.0466 | 0.89131 | 1.0104 | 217.1398 | 50.9473 | 52291.0702 | 53728.27 | 1070.5654 | 970.5654 | 963.4178
4179 | 107.3224 | 1.0309 | 0.93466 | 1.0159 | 195.0935 | 55.1401 | 58654.4527 | 55648.2595 | 1108.9652 | 1008.9652 | 963.4406
4180 | 100.0049 | 1.0658 | 0.98441 | 1.0505 | 209.4986 | 48.2388 | 55234.7994 | 56451.679 | 1125.0336 | 1025.0336 | 963.4519
4181 | 115.313 | 1.0394 | 1.0833 | 1.0432 | 197.7238 | 49.4921 | 62063.8991 | 63360.6788 | 1263.2136 | 1163.2136 | 963.4688
4182 | 100.9102 | 1.0372 | 0.99447 | 1.0472 | 196.7809 | 52.3776 | 60224.5216 | 59514.352 | 1186.287 | 1086.287 | 963.5066
4183 | 88.8025 | 0.97936 | 0.91477 | 0.96226 | 209.4121 | 57.4905 | 55815.5767 | 54892.0434 | 1093.8409 | 993.8409 | 963.5166
4184 | 99.0413 | 0.8557 | 1.06 | 0.98868 | 212.8515 | 49.0787 | 55805.2086 | 55522.0709 | 1106.4414 | 1006.4414 | 963.524
4185 | 104.0431 | 0.94244 | 1.0674 | 1.0417 | 207.7435 | 48.0858 | 58388.7681 | 57209.454 | 1140.1891 | 1040.1891 | 963.5741
4186 | 95.6587 | 0.97146 | 1.2238 | 1.0209 | 200.8455 | 49.1147 | 66128.3706 | 64047.8202 | 1276.9564 | 1176.9564 | 963.5963
4187 | 110.6296 | 0.90332 | 0.99699 | 1.0199 | 201.8133 | 44.9003 | 50343.6908 | 50143.394 | 998.8679 | 898.8679 | 963.6007
4188 | 102.373 | 1.0251 | 1.0002 | 1.1066 | 203.5356 | 48.9685 | 59538.3513 | 59189.2246 | 1179.7845 | 1079.7845 | 963.6182
4189 | 89.2317 | 1.0754 | 0.94527 | 0.92092 | 208.6762 | 51.9536 | 50420.8874 | 51528.6134 | 1026.5723 | 926.5723 | 963.6465
4190 | 86.7391 | 1.0678 | 0.75826 | 1.0435 | 211.1618 | 50.8663 | 45170.2419 | 44677.8886 | 889.5578 | 789.5578 | 963.6575
4191 | 98.924 | 1.0091 | 1.055 | 0.98686 | 218.6388 | 51.5478 | 59032.9514 | 58192.8205 | 1159.8564 | 1059.8564 | 963.6698
4192 | 90.8553 | 0.98215 | 0.88386 | 1.0255 | 203.7997 | 52.8575 | 52831.6636 | 52468.9672 | 1045.3793 | 945.3793 | 963.6881
4193 | 103.9853 | 0.9191 | 1.1491 | 1.0397 | 210.5092 | 50.1512 | 64919.8448 | 64498.7749 | 1285.9755 | 1185.9755 | 963.7005
4194 | 101.8786 | 1.0024 | 1.0776 | 1.0091 | 210.4038 | 45.2422 | 54028.5498 | 52866.5398 | 1053.3308 | 953.3308 | 963.7155
4195 | 103.8946 | 1.0188 | 0.96015 | 1.0458 | 210.3153 | 52.4457 | 58424.3459 | 58119.9285 | 1158.3986 | 1058.3986 | 963.7365
4196 | 84.9431 | 1.047 | 1.1285 | 1.0392 | 186.7876 | 52.6786 | 66648.3669 | 64377.668 | 1283.5534 | 1183.5534 | 963.7462
4197 | 102.4169 | 1.0372 | 1.2842 | 0.94215 | 201.177 | 53.645 | 70805.7414 | 69219.2846 | 1380.3857 | 1280.3857 | 963.7657
4198 | 106.0748 | 1.0593 | 1.0801 | 0.95714 | 198.4557 | 49.6084 | 57055.5928 | 58999.5931 | 1175.9919 | 1075.9919 | 963.8452
4199 | 100.9696 | 0.99136 | 0.89487 | 1.0518 | 198.9614 | 52.0362 | 54387.6502 | 55460.0113 | 1105.2002 | 1005.2002 | 963.8762
4200 | 115.8076 | 0.92159 | 1.1695 | 0.95972 | 203.4868 | 51.5094 | 63514.9007 | 63449.1108 | 1264.9822 | 1164.9822 | 963.9042
4201 | 104.0547 | 0.86998 | 1.1512 | 1.0557 | 194.7933 | 49.0435 | 64239.9722 | 63656.5215 | 1269.1304 | 1169.1304 | 963.9242
4202 | 97.4118 | 0.9714 | 0.72351 | 0.90549 | 206.9699 | 48.2728 | 36399.9108 | 36136.1752 | 718.7235 | 618.7235 | 963.9586
4203 | 102.7758 | 0.94453 | 1.093 | 0.96646 | 202.3973 | 54.4362 | 62988.6771 | 62816.5682 | 1252.3314 | 1152.3314 | 963.9723
4204 | 98.1744 | 0.98227 | 1.2006 | 1.0555 | 209.5549 | 50.1898 | 68655.9066 | 70624.8804 | 1408.4976 | 1308.4976 | 964.0193
4205 | 109.8466 | 1.0393 | 1.321 | 1.0286 | 212.5003 | 52.0825 | 76926.4497 | 75039.5214 | 1496.7904 | 1396.7904 | 964.04
4206 | 109.0794 | 0.98381 | 1.3187 | 1.0414 | 200.5437 | 48.5018 | 72014.0111 | 71426.4125 | 1424.5283 | 1324.5283 | 964.064
4207 | 86.4069 | 1.0268 | 1.0979 | 1.0103 | 177.945 | 51.5928 | 61980.972 | 60881.582 | 1213.6316 | 1113.6316 | 964.1499
4208 | 105.1936 | 0.9827 | 0.99363 | 1.0764 | 216.9817 | 50.4314 | 59369.7009 | 59482.5207 | 1185.6504 | 1085.6504 | 964.1759
4209 | 99.9393 | 1.0466 | 0.92392 | 0.9863 | 201.5293 | 50.7802 | 51786.5523 | 53341.1414 | 1062.8228 | 962.8228 | 964.1988
4210 | 131.8011 | 1.12 | 1.0731 | 0.99408 | 203.948 | 49.1224 | 59854.5814 | 59962.4134 | 1195.2483 | 1095.2483 | 964.2475
4211 | 105.8108 | 0.98783 | 0.93163 | 1.0531 | 198.7387 | 54.444 | 59302.8705 | 61709.9832 | 1230.1997 | 1130.1997 | 964.2973
4212 | 86.2157 | 0.94431 | 0.95154 | 0.91575 | 196.2154 | 49.0705 | 46950.2577 | 46877.6065 | 933.5521 | 833.5521 | 964.3291
4213 | 98.4933 | 0.94572 | 1.0559 | 0.95553 | 226.1892 | 54.8995 | 60730.87 | 58883.4072 | 1173.6681 | 1073.6681 | 964.3381
4214 | 96.96 | 0.99383 | 0.93381 | 0.93756 | 211.0183 | 47.4095 | 46286.681 | 46562.9483 | 927.259 | 827.259 | 964.3578
4215 | 89.5059 | 1.0376 | 0.92246 | 0.97932 | 194.9427 | 50.3065 | 50313.0665 | 48666.0208 | 969.3204 | 869.3204 | 964.3797
4216 | 111.6779 | 1.0572 | 1.0272 | 1.0384 | 197.7611 | 52.0071 | 61814.1985 | 62816.5505 | 1252.331 | 1152.331 | 964.3826
4217 | 102.8219 | 1.0157 | 0.94074 | 0.98034 | 204.232 | 47.5081 | 48980.0172 | 48295.5194 | 961.9104 | 861.9104 | 964.3998
4218 | 105.7932 | 0.94972 | 1.265 | 0.99763 | 203.9377 | 52.5976 | 71865.1796 | 69058.4771 | 1377.1695 | 1277.1695 | 964.4039
4219 | 105.8602 | 0.97353 | 1.2099 | 1.0368 | 186.5905 | 52.5958 | 71583.8951 | 72289.0547 | 1441.7811 | 1341.7811 | 964.4684
4220 | 109.9021 | 0.9939 | 1.1028 | 1.0489 | 209.9141 | 53.5037 | 67941.093 | 67901.0988 | 1354.022 | 1254.022 | 964.5486
4221 | 78.4672 | 0.97399 | 0.9512 | 1.0332 | 207.1546 | 50.9112 | 54131.9105 | 53705.9694 | 1070.1194 | 970.1194 | 964.5702
4222 | 94.4078 | 1.0197 | 1.019 | 1.0798 | 202.3764 | 50.4088 | 60522.5859 | 59483.1694 | 1185.6634 | 1085.6634 | 964.6056
4223 | 113.0983 | 1.0586 | 1.0675 | 0.99616 | 187.3325 | 48.5562 | 57633.3934 | 58689.3067 | 1169.7861 | 1069.7861 | 964.6788
4224 | 112.6574 | 0.92544 | 0.8401 | 1.0315 | 197.0261 | 49.8339 | 48576.2652 | 49535.304 | 986.7061 | 886.7061 | 964.7184
4225 | 100.6961 | 0.89561 | 0.85886 | 1.0086 | 204.8132 | 51.0965 | 49077.2424 | 51234.4479 | 1020.689 | 920.689 | 964.8285
4226 | 81.302 | 0.95137 | 1.056 | 0.92165 | 218.8101 | 48.8247 | 51516.0549 | 50390.6489 | 1003.813 | 903.813 | 964.8362
4227 | 100.1607 | 1.0302 | 1.1942 | 0.95733 | 195.3885 | 50.3571 | 62962.5892 | 65850.4312 | 1313.0086 | 1213.0086 | 964.861
4228 | 98.9094 | 1.0198 | 0.94679 | 0.94824 | 192.8865 | 50.4078 | 50532.9033 | 50743.9704 | 1010.8794 | 910.8794 | 964.9203
4229 | 101.0931 | 1.0418 | 1.1226 | 0.94864 | 186.9353 | 47.0625 | 55262.0968 | 55162.5228 | 1099.2505 | 999.2505 | 964.927
4230 | 109.7124 | 0.96533 | 0.92556 | 1.0084 | 195.1668 | 48.979 | 51096.1031 | 51285.4028 | 1021.7081 | 921.7081 | 964.9542
4231 | 101.4496 | 0.94979 | 1.0596 | 1.0032 | 204.7545 | 53.5958 | 62336.8938 | 62737.732 | 1250.7546 | 1150.7546 | 964.9629
4232 | 97.0063 | 1.0013 | 1.0512 | 1.0399 | 191.4436 | 48.4001 | 57796.6416 | 58800.1797 | 1172.0036 | 1072.0036 | 964.9784
4233 | 86.283 | 1.0435 | 1.1304 | 1.0217 | 198.4703 | 53.2286 | 66463.6971 | 66626.8018 | 1328.536 | 1228.536 | 965.0101
4234 | 105.8996 | 1.0828 | 0.81312 | 1.0213 | 199.8399 | 51.6135 | 48981.2165 | 49400.6733 | 984.0135 | 884.0135 | 965.0172
4235 | 92.6838 | 1.0069 | 1.0419 | 0.99753 | 197.9698 | 46.5483 | 52922.8071 | 52358.8227 | 1043.1765 | 943.1765 | 965.1008
4236 | 91.0358 | 0.98887 | 0.84275 | 1.0565 | 194.9874 | 49.056 | 48290.2447 | 50670.4683 | 1009.4094 | 909.4094 | 965.1825
4237 | 111.0773 | 1.0342 | 0.88058 | 1.0429 | 196.0005 | 49.1914 | 51022.1984 | 51731.041 | 1030.6208 | 930.6208 | 965.2046
4238 | 88.715 | 0.96777 | 1.2711 | 1.0592 | 208.8511 | 51.0858 | 73371.7098 | 73855.6517 | 1473.113 | 1373.113 | 965.244
4239 | 116.0879 | 0.94422 | 0.9647 | 1.0092 | 202.4055 | 51.9142 | 56436.7832 | 54916.735 | 1094.3347 | 994.3347 | 965.3381
4240 | 100.3769 | 0.91282 | 0.8522 | 0.96431 | 204.2342 | 46.2757 | 42472.8799 | 43270.4063 | 861.4081 | 761.4081 | 965.3394
4241 | 103.3455 | 0.9759 | 0.96757 | 1.0177 | 221.778 | 49.7599 | 54239.7694 | 54582.7887 | 1087.6558 | 987.6558 | 965.399
4242 | 93.1239 | 0.92435 | 0.99678 | 1.0661 | 205.3585 | 48.9365 | 56419.7612 | 56742.9043 | 1130.8581 | 1030.8581 | 965.4164
4243 | 91.5005 | 0.9913 | 0.89278 | 0.96753 | 196.492 | 53.3935 | 51160.3281 | 50863.4425 | 1013.2688 | 913.2688 | 965.4459
4244 | 92.5727 | 0.9539 | 1.1443 | 0.98275 | 201.1345 | 50.0722 | 60932.3016 | 60270.3553 | 1201.4071 | 1101.4071 | 965.543
4245 | 109.6955 | 1.0867 | 1.0527 | 1.0303 | 202.3269 | 49.8164 | 60171.0531 | 59730.5924 | 1190.6118 | 1090.6118 | 965.5576
4246 | 96.2153 | 1.0798 | 1.0814 | 1.0399 | 207.5044 | 49.8832 | 61484.637 | 62604.6456 | 1248.0929 | 1148.0929 | 965.6086
4247 | 99.0045 | 1.0375 | 0.57568 | 1.0133 | 193.2963 | 49.8011 | 34359.3717 | 35289.3063 | 701.7861 | 601.7861 | 965.6374
4248 | 103.9406 | 0.99355 | 1.2127 | 1.0336 | 201.2135 | 47.7534 | 64990.5281 | 66260.324 | 1321.2065 | 1221.2065 | 965.7041
4249 | 111.1021 | 1.0039 | 0.88334 | 0.9723 | 198.162 | 50.9728 | 49662.1266 | 49217.172 | 980.3434 | 880.3434 | 965.7199
4250 | 86.1627 | 1.101 | 0.98446 | 1.057 | 197.8849 | 49.5952 | 56508.8175 | 55049.5059 | 1096.9901 | 996.9901 | 965.8199
4251 | 86.1185 | 0.98228 | 1.2025 | 0.99201 | 199.8074 | 49.9926 | 64066.7267 | 63305.0556 | 1262.1011 | 1162.1011 | 965.8455
4252 | 80.9577 | 1.0562 | 1.0044 | 0.99634 | 188.6489 | 51.3301 | 55946.8024 | 57747.4148 | 1150.9483 | 1050.9483 | 965.8587
4253 | 101.14 | 0.97766 | 0.65476 | 1.0882 | 186.403 | 51.6257 | 42074.4118 | 41414.9574 | 824.2991 | 724.2991 | 965.8617
4254 | 115.7345 | 1.0792 | 0.90357 | 0.99548 | 210.923 | 54.3218 | 55857.2816 | 57240.3848 | 1140.8077 | 1040.8077 | 965.8803
4255 | 86.7231 | 1.016 | 1.0591 | 0.99309 | 199.9562 | 49.5111 | 56638.024 | 56916.0048 | 1134.3201 | 1034.3201 | 965.9078
4256 | 94.5458 | 1.0548 | 1.0477 | 1.055 | 199.6878 | 42.7444 | 51711.6699 | 51776.4789 | 1031.5296 | 931.5296 | 965.9368
4257 | 84.3299 | 1.0193 | 1.008 | 1.0402 | 205.0337 | 45.9752 | 52362.4703 | 52738.219 | 1050.7644 | 950.7644 | 965.9864
4258 | 82.2547 | 0.99754 | 0.92922 | 0.97076 | 207.0963 | 49.7313 | 49148.0072 | 49155.5305 | 979.1106 | 879.1106 | 965.995
4259 | 104.1935 | 0.91568 | 0.87986 | 1.0238 | 217.6749 | 55.0673 | 55076.4496 | 54953.14 | 1095.0628 | 995.0628 | 966.3163
4260 | 118.1327 | 0.91398 | 1.1057 | 1.0494 | 208.197 | 49.7315 | 63283.0009 | 61514.362 | 1226.2872 | 1126.2872 | 966.3182
4261 | 101.4408 | 1.0101 | 0.863 | 0.86411 | 205.6669 | 50.3966 | 42952.1767 | 41624.9124 | 828.4982 | 728.4982 | 966.4491
4262 | 118.8033 | 1.0718 | 0.90949 | 1.0376 | 206.6513 | 49.5329 | 53257.7906 | 52501.7387 | 1046.0348 | 946.0348 | 966.4665
4263 | 93.7712 | 0.97491 | 0.94119 | 1.0114 | 219.5081 | 48.3131 | 50625.6918 | 49161.3422 | 979.2268 | 879.2268 | 966.5441
4264 | 88.9425 | 0.8847 | 0.89952 | 0.95863 | 184.6209 | 46.9778 | 44390.7403 | 45113.5244 | 898.2705 | 798.2705 | 966.6226
4265 | 88.7375 | 0.97413 | 1.1292 | 0.91843 | 202.61 | 55.7885 | 62880.927 | 63654.7868 | 1269.0957 | 1169.0957 | 966.6911
4266 | 94.3956 | 0.99921 | 0.89168 | 0.93683 | 198.1407 | 48.7584 | 45527.4826 | 45154.6197 | 899.0924 | 799.0924 | 966.7397
4267 | 99.1679 | 1.0606 | 1.1716 | 0.97891 | 209.0157 | 52.0326 | 65359.4178 | 63487.4041 | 1265.7481 | 1165.7481 | 966.8148
4268 | 91.1302 | 1.0133 | 0.87619 | 1.0499 | 182.0123 | 51.193 | 52001.7879 | 51168.8772 | 1019.3775 | 919.3775 | 966.8299
4269 | 82.9953 | 1.13 | 1.4033 | 1.0522 | 187.8767 | 50.1324 | 78913.2788 | 81266.2246 | 1621.3245 | 1521.3245 | 966.8388
4270 | 103.1675 | 0.97979 | 1.1691 | 1.0439 | 208.1836 | 53.3538 | 70712.8519 | 72120.2522 | 1438.405 | 1338.405 | 966.842
4271 | 109.0678 | 1.0757 | 0.83607 | 0.94419 | 191.2218 | 45.4843 | 41433.3009 | 40920.2701 | 814.4054 | 714.4054 | 967.0852
4272 | 86.0501 | 1.0268 | 1.094 | 0.97795 | 195.7254 | 50.5466 | 58739.1233 | 58789.7339 | 1171.7947 | 1071.7947 | 967.1396
4273 | 103.8398 | 1.0306 | 1.0252 | 1.0035 | 186.0245 | 50.7259 | 57801.8211 | 60320.0918 | 1202.4018 | 1102.4018 | 967.1554
4274 | 103.2039 | 1.0502 | 0.80452 | 1.0406 | 191.4989 | 49.5591 | 47052.2109 | 48434.5807 | 964.6916 | 864.6916 | 967.1841
4275 | 95.6279 | 0.97176 | 1.0681 | 0.98406 | 185.0742 | 50.0915 | 57490.6612 | 57338.4827 | 1142.7697 | 1042.7697 | 967.1844
4276 | 83.6468 | 1.0611 | 0.93444 | 1.0253 | 196.1118 | 45.9734 | 48322.1744 | 47096.95 | 937.939 | 837.939 | 967.211
4277 | 81.2226 | 0.96166 | 1.1171 | 1.0187 | 209.8914 | 49.4485 | 60347.0558 | 59707.9805 | 1190.1596 | 1090.1596 | 967.2119
4278 | 90.3922 | 0.98403 | 1.1603 | 0.98877 | 216.2526 | 48.8004 | 60542.9778 | 61021.8272 | 1216.4365 | 1116.4365 | 967.2131
4279 | 100.6672 | 1.011 | 0.95815 | 1.0231 | 201.9601 | 49.7447 | 54030.6369 | 53829.3051 | 1072.5861 | 972.5861 | 967.2618
4280 | 88.8855 | 0.99524 | 0.86479 | 1.0786 | 202.1436 | 48.2586 | 49486.7527 | 50737.3793 | 1010.7476 | 910.7476 | 967.3123
4281 | 108.057 | 1.0873 | 0.59505 | 1.0045 | 216.2262 | 50.7286 | 36497.4986 | 36247.4938 | 720.9499 | 620.9499 | 967.3555
4282 | 102.1353 | 0.97648 | 0.86426 | 1.0478 | 196.1812 | 48.7286 | 49183.3459 | 50025.3476 | 996.507 | 896.507 | 967.3584
4283 | 106.0064 | 1.0056 | 0.87331 | 1.0024 | 188.6545 | 47.7237 | 47053.7737 | 48252.1932 | 961.0439 | 861.0439 | 967.4107
4284 | 100.5309 | 0.98991 | 1.1227 | 0.94669 | 200.8266 | 48.8217 | 56948.8829 | 56797.1892 | 1131.9438 | 1031.9438 | 967.4374
4285 | 106.1635 | 1.0226 | 0.90704 | 0.94512 | 203.8003 | 51.0068 | 49467.4387 | 48840.2036 | 972.8041 | 872.8041 | 967.4751
4286 | 93.364 | 0.98626 | 0.81317 | 0.91943 | 198.9193 | 48.6829 | 41079.7912 | 41985.3255 | 835.7065 | 735.7065 | 967.4827
4287 | 89.0576 | 1.0042 | 0.96329 | 1.0308 | 211.1755 | 45.0349 | 48956.075 | 49471.2038 | 985.4241 | 885.4241 | 967.5341
4288 | 101.5665 | 1.0158 | 1.2922 | 0.94974 | 189.0984 | 47.8813 | 63892.4607 | 64848.8165 | 1292.9763 | 1192.9763 | 967.5481
4289 | 91.0818 | 1.0072 | 0.89624 | 0.99112 | 205.5688 | 50.1047 | 49309.0854 | 48978.6894 | 975.5738 | 875.5738 | 967.5902
4290 | 101.5237 | 0.94542 | 1.0893 | 0.9353 | 190.5324 | 51.0182 | 57063.5572 | 56284.0106 | 1121.6802 | 1021.6802 | 967.6184
4291 | 104.5778 | 1.0071 | 0.79342 | 0.98461 | 200.0548 | 47.3516 | 42178.8652 | 40882.8867 | 813.6577 | 713.6577 | 967.6437
4292 | 117.1978 | 1.0222 | 0.93249 | 1.0353 | 197.9623 | 48.9333 | 53299.6325 | 52648.945 | 1048.9789 | 948.9789 | 967.6768
4293 | 93.5 | 1.0241 | 1.0363 | 0.98147 | 204.3904 | 43.8027 | 48952.1941 | 49006.5185 | 976.1304 | 876.1304 | 967.7344
4294 | 82.7831 | 0.93932 | 1.2769 | 1.0505 | 194.6234 | 49.6287 | 70624.1262 | 70920.8045 | 1414.4161 | 1314.4161 | 967.7511
4295 | 83.1401 | 0.91259 | 1.0576 | 1.0815 | 205.7822 | 46.2769 | 56647.9388 | 56785.8075 | 1131.7161 | 1031.7161 | 967.7525
4296 | 95.7855 | 0.96615 | 0.64335 | 1.0336 | 201.526 | 48.1686 | 36690.2609 | 35703.8388 | 710.0768 | 610.0768 | 967.7726
4297 | 93.588 | 0.96502 | 1.049 | 0.93637 | 191.021 | 50.1731 | 54006.9277 | 54655.3381 | 1089.1068 | 989.1068 | 967.8201
4298 | 106.4279 | 1.0696 | 0.94193 | 0.98095 | 196.7067 | 52.2097 | 54381.3005 | 55457.21 | 1105.1442 | 1005.1442 | 967.8279
4299 | 99.2806 | 0.99858 | 0.78349 | 0.99909 | 196.072 | 51.5496 | 45658.4447 | 45510.3013 | 906.206 | 806.206 | 967.8496
4300 | 94.7747 | 1.0335 | 1.0301 | 1.0261 | 213.5756 | 52.2622 | 60572.217 | 58197.9869 | 1159.9597 | 1059.9597 | 967.9345
4301 | 92.0672 | 0.97015 | 0.93975 | 1.0475 | 205.0964 | 49.1214 | 52947.9324 | 53331.7828 | 1062.6357 | 962.6357 | 967.9865
4302 | 97.2323 | 0.99904 | 1.0276 | 0.9733 | 206.2222 | 54.5969 | 60112.9701 | 62490.1241 | 1245.8025 | 1145.8025 | 968.0075
4303 | 106.4858 | 1.069 | 0.97092 | 0.9491 | 191.8582 | 49.0936 | 51020.3486 | 50399.4005 | 1003.988 | 903.988 | 968.0217
4304 | 122.7311 | 1.0151 | 1.1507 | 1.0152 | 214.9949 | 51.398 | 66657.3853 | 68124.9043 | 1358.4981 | 1258.4981 | 968.0911
4305 | 95.0223 | 1.0083 | 1.0066 | 0.93214 | 209.0664 | 53.2322 | 55255.0739 | 54784.0014 | 1091.68 | 991.68 | 968.1233
4306 | 101.8921 | 0.9817 | 0.75138 | 0.95848 | 181.4842 | 49.7768 | 41008.9121 | 40628.0121 | 808.5602 | 708.5602 | 968.141
4307 | 96.1111 | 0.96058 | 1.0629 | 1.0607 | 184.5684 | 49.4998 | 60561.6242 | 62568.7324 | 1247.3746 | 1147.3746 | 968.2212
4308 | 119.5294 | 1.0497 | 0.90069 | 1.0489 | 202.3053 | 47.6249 | 51172.2875 | 52476.3715 | 1045.5274 | 945.5274 | 968.2841
4309 | 97.4987 | 0.94766 | 1.0211 | 1.0732 | 184.2733 | 52.6032 | 62689.2842 | 63922.4901 | 1274.4498 | 1174.4498 | 968.3402
4310 | 92.9693 | 1.0151 | 1.2663 | 1.0705 | 199.0551 | 47.5754 | 69177.8022 | 68509.7744 | 1366.1955 | 1266.1955 | 968.3559
4311 | 101.3202 | 0.97693 | 1.2534 | 0.94147 | 199.0725 | 49.7349 | 63812.076 | 64157.8209 | 1279.1564 | 1179.1564 | 968.4539
4312 | 110.9938 | 1.0395 | 1.0521 | 0.99357 | 201.7722 | 51.3238 | 59775.5145 | 61644.4535 | 1228.8891 | 1128.8891 | 968.4876
4313 | 112.5354 | 0.9933 | 0.78438 | 0.9288 | 202.0792 | 51.7226 | 43665.0239 | 43904.2873 | 874.0857 | 774.0857 | 968.5385
4314 | 83.0299 | 1.0329 | 0.92835 | 0.99489 | 205.4181 | 47.4753 | 48125.644 | 48481.5342 | 965.6307 | 865.6307 | 968.5505
4315 | 91.3266 | 0.9496 | 0.9894 | 1.0605 | 189.8311 | 55.2708 | 62978.7649 | 65086.9987 | 1297.74 | 1197.74 | 968.5676
4316 | 75.0667 | 0.97269 | 0.56744 | 0.96891 | 184.491 | 52.051 | 32602.6889 | 31502.3918 | 626.0478 | 526.0478 | 968.5701
4317 | 97.7959 | 1.1037 | 0.99406 | 0.95095 | 197.88 | 52.4652 | 55456.4588 | 56366.3455 | 1123.3269 | 1023.3269 | 968.5809
4318 | 97.3494 | 0.89515 | 1.3269 | 1.0367 | 200.7033 | 48.0108 | 70429.3021 | 70149.0444 | 1398.9809 | 1298.9809 | 968.5898
4319 | 96.6084 | 1.0706 | 1.0438 | 0.94599 | 190.19 | 45.5044 | 49826.9821 | 48691.6977 | 969.834 | 869.834 | 968.5955
4320 | 97.0087 | 1.0322 | 0.98127 | 0.91992 | 206.1425 | 43.8801 | 44209.8806 | 44173.2511 | 879.465 | 779.465 | 968.6865
4321 | 106.4744 | 1.0137 | 0.86704 | 0.96314 | 206.6553 | 53.2886 | 50458.548 | 50583.3151 | 1007.6663 | 907.6663 | 968.6963
4322 | 93.0073 | 1.0686 | 0.8452 | 0.9635 | 211.7174 | 56.0108 | 51390.948 | 51287.2244 | 1021.7445 | 921.7445 | 968.7353
4323 | 106.4959 | 0.97024 | 0.90146 | 0.98232 | 209.6015 | 48.9141 | 48578.0743 | 49361.9133 | 983.2383 | 883.2383 | 968.769
4324 | 95.3269 | 1.0158 | 1.1073 | 0.98985 | 196.0532 | 47.6681 | 57060.4262 | 57769.2489 | 1151.385 | 1051.385 | 968.9142
4325 | 117.3439 | 1.0065 | 1.1156 | 0.99421 | 199.0305 | 49.9063 | 61445.2311 | 61561.051 | 1227.221 | 1127.221 | 968.9408
4326 | 99.7293 | 0.94683 | 0.98722 | 1.0301 | 202.9912 | 52.272 | 58294.1096 | 57007.3701 | 1136.1474 | 1036.1474 | 969.0513
4327 | 99.2398 | 1.0222 | 1.2545 | 1.0116 | 203.1715 | 48.9106 | 67233.3394 | 66379.6384 | 1323.5928 | 1223.5928 | 969.0517
4328 | 92.1187 | 1.0234 | 1.2298 | 1.0081 | 205.9582 | 50.6555 | 67780.9251 | 66871.731 | 1333.4346 | 1233.4346 | 969.1732
4329 | 70.3751 | 1.0566 | 1.3792 | 1.0294 | 190.0988 | 47.137 | 70619.3981 | 70831.3846 | 1412.6277 | 1312.6277 | 969.2469
4330 | 93.6242 | 1.0276 | 1.1116 | 1.0337 | 195.2516 | 46.4476 | 58034.6987 | 58543.5673 | 1166.8713 | 1066.8713 | 969.281
4331 | 83.8853 | 0.98632 | 0.85315 | 0.93362 | 213.5858 | 48.3577 | 42732.6948 | 41358.7853 | 823.1757 | 723.1757 | 969.2884
4332 | 111.7942 | 1.0484 | 0.80838 | 1.0479 | 210.3079 | 52.2667 | 50609.5265 | 50332.8542 | 1002.6571 | 902.6571 | 969.3209
4333 | 111.4923 | 0.97171 | 1.2977 | 1.0853 | 213.7302 | 49.0526 | 74615.7341 | 75170.597 | 1499.4119 | 1399.4119 | 969.3436
4334 | 92.7944 | 1.0211 | 1.0149 | 1.0168 | 195.8746 | 46.9103 | 53052.0865 | 54312.9166 | 1082.2583 | 982.2583 | 969.3719
4335 | 112.8484 | 0.99579 | 1.1083 | 0.98522 | 213.9285 | 53.4952 | 64637.2708 | 62575.9555 | 1247.5191 | 1147.5191 | 969.3819
4336 | 108.1957 | 0.86979 | 0.97138 | 0.96453 | 198.1928 | 53.8658 | 55735.2698 | 56257.9887 | 1121.1598 | 1021.1598 | 969.4462
4337 | 93.7951 | 0.97864 | 0.92276 | 0.96156 | 197.3257 | 47.6952 | 46894.61 | 46742.2882 | 930.8458 | 830.8458 | 969.4801
4338 | 106.7247 | 1.0168 | 0.97677 | 1.032 | 206.5027 | 54.4926 | 61050.6044 | 61155.0405 | 1219.1008 | 1119.1008 | 969.5403
4339 | 102.7444 | 0.9729 | 1.0253 | 1.0288 | 211.0749 | 51.0009 | 59101.8155 | 58675.0522 | 1169.501 | 1069.501 | 969.6152
4340 | 101.7082 | 1.0651 | 0.95706 | 1.0768 | 191.0899 | 54.3709 | 62112.5681 | 60189.5982 | 1199.792 | 1099.792 | 969.7258
4341 | 98.7858 | 1.007 | 1.1257 | 0.99023 | 195.2846 | 47.4046 | 57755.3809 | 56257.1688 | 1121.1434 | 1021.1434 | 969.7647
4342 | 99.8042 | 1.0336 | 1.2131 | 1.0177 | 205.8523 | 46.8035 | 62813.4467 | 62810.8801 | 1252.2176 | 1152.2176 | 969.8101
4343 | 100.7573 | 0.99953 | 1.2059 | 0.95664 | 205.3783 | 51.3655 | 64633.0713 | 64007.7347 | 1276.1547 | 1176.1547 | 969.8826
4344 | 89.987 | 0.99878 | 1.0522 | 0.9886 | 206.747 | 51.1561 | 58016.3992 | 56351.501 | 1123.03 | 1023.03 | 969.9548
4345 | 82.3734 | 1.0605 | 0.98318 | 0.98711 | 206.5491 | 49.6816 | 52763.071 | 50897.0426 | 1013.9409 | 913.9409 | 969.9744
4346 | 99.1277 | 0.96755 | 1.0241 | 1.0488 | 213.3617 | 57.0281 | 66938.2609 | 67734.4735 | 1350.6895 | 1250.6895 | 969.9792
4347 | 103.1565 | 1.0874 | 1.1899 | 1.0954 | 211.3198 | 53.47 | 75904.7823 | 78599.3248 | 1567.9865 | 1467.9865 | 969.9962
4348 | 104.3655 | 0.98101 | 1.0031 | 1.0449 | 211.1101 | 52.2896 | 60368.5046 | 61360.1296 | 1223.2026 | 1123.2026 | 970.0007
4349 | 102.1411 | 1.0441 | 0.94718 | 1.02 | 190.2065 | 45.9118 | 49442.0389 | 48266.4626 | 961.3293 | 861.3293 | 970.0027
4350 | 85.2999 | 0.9927 | 0.91971 | 0.98712 | 203.5469 | 44.871 | 44739.9046 | 44688.1995 | 889.764 | 789.764 | 970.1162
4351 | 108.0339 | 1.0938 | 1.0949 | 0.98769 | 192.7087 | 50.9528 | 61316.5665 | 61163.2795 | 1219.2656 | 1119.2656 | 970.1194
4352 | 98.2635 | 1.0166 | 0.80644 | 1.0405 | 204.0181 | 49.5775 | 46758.9129 | 47332.2429 | 942.6449 | 842.6449 | 970.1711
4353 | 84.8921 | 1.0102 | 1.0232 | 0.94317 | 205.5204 | 53.4711 | 56395.3933 | 53559.2036 | 1067.1841 | 967.1841 | 970.1978
4354 | 102.0181 | 0.955 | 1.2041 | 1.0341 | 199.5195 | 53.5855 | 72144.9239 | 73200.7908 | 1460.0158 | 1360.0158 | 970.2129
4355 | 94.9956 | 0.98518 | 1.2167 | 1.0438 | 210.5566 | 45.7348 | 62570.0365 | 62101.2 | 1238.024 | 1138.024 | 970.2215
4356 | 91.077 | 0.99318 | 1.0538 | 1.0565 | 199.4222 | 47.0968 | 56894.983 | 56075.1487 | 1117.503 | 1017.503 | 970.2999
4357 | 97.1444 | 1.0092 | 0.7771 | 0.99143 | 200.4702 | 50.1349 | 43742.0121 | 43925.9963 | 874.5199 | 774.5199 | 970.341
4358 | 100.7246 | 1.1121 | 1.1854 | 0.97381 | 201.0626 | 47.0301 | 59759.4651 | 58147.349 | 1158.947 | 1058.947 | 970.3663
4359 | 97.7013 | 0.94742 | 0.86731 | 0.97179 | 208.6365 | 51.1618 | 48066.135 | 48185.1148 | 959.7023 | 859.7023 | 970.3781
4360 | 122.6375 | 0.93994 | 1.34 | 0.99863 | 200.8561 | 49.7691 | 72536.7687 | 72249.7338 | 1440.9947 | 1340.9947 | 970.4827
4361 | 108.2256 | 0.9672 | 0.91961 | 0.9279 | 208.4247 | 50.5674 | 48651.4402 | 46722.7187 | 930.4544 | 830.4544 | 970.5387
4362 | 114.391 | 1.0325 | 0.90752 | 0.95652 | 190.943 | 50.6384 | 50128.6996 | 48715.3482 | 970.307 | 870.307 | 970.5654
4363 | 96.1638 | 0.91631 | 1.0415 | 1.0184 | 179.6767 | 51.1384 | 58925.2413 | 59503.7181 | 1186.0744 | 1086.0744 | 970.5782
4364 | 98.8138 | 1.0184 | 1.1367 | 1.0658 | 199.4504 | 49.813 | 65558.4261 | 66345.0603 | 1322.9012 | 1222.9012 | 970.5825
4365 | 109.7467 | 1.0198 | 1.007 | 1.0281 | 195.3741 | 52.1041 | 59969.5234 | 59041.1899 | 1176.8238 | 1076.8238 | 970.6139
4366 | 98.2232 | 1.1099 | 1.1357 | 1.0419 | 200.4853 | 50.5296 | 65503.4058 | 66006.0474 | 1316.1209 | 1216.1209 | 970.6309
4367 | 98.6133 | 1.0229 | 0.8555 | 0.96661 | 209.6634 | 51.0535 | 47577.4002 | 47262.5829 | 941.2517 | 841.2517 | 970.665
4368 | 107.1407 | 1.0254 | 0.80192 | 0.96407 | 219.8235 | 48.1484 | 42733.457 | 44414.4448 | 884.2889 | 784.2889 | 970.6813
4369 | 100.3648 | 1.0129 | 0.86745 | 0.97843 | 202.5734 | 54.3363 | 51843.8205 | 51955.6774 | 1035.1135 | 935.1135 | 970.6822
4370 | 116.9116 | 1.0284 | 0.76797 | 1.021 | 211.4698 | 47.2194 | 42913.5899 | 44178.2627 | 879.5653 | 779.5653 | 970.6931
4371 | 96.425 | 1.0111 | 1.076 | 0.9364 | 212.9345 | 51.4213 | 57037.7282 | 56740.2101 | 1130.8042 | 1030.8042 | 970.7777
4372 | 97.754 | 0.96832 | 1.3313 | 1.1104 | 200.5959 | 51.117 | 80605.5047 | 82027.7185 | 1636.5544 | 1536.5544 | 970.7898
4373 | 100.9442 | 0.98138 | 1.0467 | 1.1174 | 205.5322 | 48.4146 | 61627.0986 | 61625.529 | 1228.5106 | 1128.5106 | 970.8699
4374 | 96.7523 | 0.99468 | 0.93167 | 0.90918 | 218.4258 | 49.9009 | 47289.8202 | 48785.453 | 971.7091 | 871.7091 | 970.9138
4375 | 75.6134 | 0.99706 | 0.92911 | 1.0102 | 213.589 | 47.5491 | 48426.6878 | 48682.2952 | 969.6459 | 869.6459 | 970.9508
4376 | 86.2006 | 1.0134 | 1.2361 | 0.96113 | 191.1131 | 50.7276 | 64891.6336 | 64028.5481 | 1276.571 | 1176.571 | 970.9902
4377 | 95.2272 | 1.0133 | 0.96283 | 1.0427 | 194.0333 | 47.258 | 52197.9869 | 52645.0178 | 1048.9004 | 948.9004 | 971.0041
4378 | 112.7442 | 1.057 | 0.94166 | 1.0714 | 211.9704 | 52.4883 | 59422.4929 | 60693.0135 | 1209.8603 | 1109.8603 | 971.0096
4379 | 87.4119 | 1.0401 | 0.91748 | 0.98775 | 209.9873 | 46.3535 | 46431.7731 | 46673.6267 | 929.4725 | 829.4725 | 971.0099
4380 | 94.6406 | 1.0369 | 0.86971 | 0.99672 | 200.0281 | 52.2954 | 50664.7948 | 51210.0361 | 1020.2007 | 920.2007 | 971.0469
4381 | 83.8705 | 0.92997 | 1.0126 | 0.9998 | 187.3481 | 50.8599 | 55646.5526 | 56437.6106 | 1124.7522 | 1024.7522 | 971.0631
4382 | 113.4904 | 1.0867 | 1.1102 | 1.0316 | 211.9829 | 48.5227 | 61766.6979 | 61269.9938 | 1221.3999 | 1121.3999 | 971.0917
4383 | 97.1028 | 1.014 | 0.6938 | 0.97025 | 184.4021 | 50.0867 | 38832.6941 | 38520.4082 | 766.4082 | 666.4082 | 971.0932
4384 | 108.243 | 1.0136 | 0.98719 | 1.0275 | 185.3635 | 53.4181 | 60228.1999 | 60732.4741 | 1210.6495 | 1110.6495 | 971.1
4385 | 98.6899 | 0.99028 | 0.76377 | 1.0125 | 196.7675 | 49.9609 | 43716.5938 | 44177.9851 | 879.5597 | 779.5597 | 971.1064
4386 | 111.6778 | 1.074 | 0.99681 | 1.0004 | 192.2092 | 51.2001 | 57388.7971 | 58218.7876 | 1160.3758 | 1060.3758 | 971.1435
4387 | 95.6465 | 0.95319 | 1.1405 | 1.0403 | 207.9903 | 53.4647 | 68515.6884 | 70645.0532 | 1408.9011 | 1308.9011 | 971.1738
4388 | 95.1738 | 1.0327 | 1.0787 | 0.9506 | 200.3903 | 49.8618 | 56227.9637 | 55484.6146 | 1105.6923 | 1005.6923 | 971.2053
4389 | 85.0691 | 0.90528 | 0.83135 | 1.058 | 207.2169 | 48.6056 | 46703.9232 | 47957.7246 | 955.1545 | 855.1545 | 971.2225
4390 | 96.3893 | 1.0375 | 1.2432 | 0.99684 | 214.555 | 50.9557 | 68456.2904 | 67115.4084 | 1338.3082 | 1238.3082 | 971.2352
4391 | 103.5219 | 0.94583 | 0.86119 | 0.99542 | 192.0969 | 48.7545 | 46760.2537 | 46796.7227 | 931.9345 | 831.9345 | 971.2758
4392 | 92.7335 | 1.0052 | 0.83012 | 0.95149 | 190.8732 | 49.1808 | 43620.7658 | 42981.4139 | 855.6283 | 755.6283 | 971.2872
4393 | 107.5254 | 1.0161 | 1.0669 | 0.9851 | 193.3175 | 52.2475 | 60814.504 | 60934.9121 | 1214.6982 | 1114.6982 | 971.328
4394 | 93.5324 | 1.0375 | 1.348 | 1.0241 | 195.3618 | 48.8534 | 72382.3668 | 72169.4798 | 1439.3896 | 1339.3896 | 971.3393
4395 | 96.0304 | 1.0371 | 1.0823 | 1.0151 | 190.3184 | 47.4027 | 56988.5055 | 55494.994 | 1105.8999 | 1005.8999 | 971.3421
4396 | 92.6672 | 0.97222 | 0.96565 | 0.98195 | 213.5025 | 51.7252 | 53920.4252 | 53446.015 | 1064.9203 | 964.9203 | 971.3431
4397 | 92.7121 | 1.046 | 0.98319 | 0.94249 | 210.4694 | 51.3985 | 52823.0761 | 52871.7908 | 1053.4358 | 953.4358 | 971.3531
4398 | 99.042 | 0.97564 | 1.0746 | 0.94676 | 187.9188 | 56.6771 | 63328.9709 | 61719.3907 | 1230.3878 | 1130.3878 | 971.4972
4399 | 100.358 | 1.0406 | 1.0792 | 1.014 | 201.1856 | 52.5893 | 63239.6553 | 63074.7335 | 1257.4947 | 1157.4947 | 971.5215
4400 | 117.4473 | 1.0656 | 1.0355 | 0.97972 | 190.5422 | 51.1678 | 58504.2751 | 59601.1888 | 1188.0238 | 1088.0238 | 971.5773
4401 | 106.5719 | 0.99722 | 0.96252 | 0.93503 | 209.1596 | 52.6244 | 53162.9272 | 54401.5338 | 1084.0307 | 984.0307 | 971.6193
4402 | 120.6235 | 0.87282 | 1.131 | 0.94163 | 203.8247 | 52.1678 | 61254.2369 | 60602.4663 | 1208.0493 | 1108.0493 | 971.6354
4403 | 107.6977 | 0.93749 | 0.9338 | 0.96333 | 200.5875 | 41.7288 | 41951.4611 | 42385.8072 | 843.7161 | 743.7161 | 971.6679
4404 | 117.2393 | 1.0061 | 0.99997 | 0.93995 | 203.3705 | 43.4138 | 46129.7633 | 44481.1969 | 885.6239 | 785.6239 | 971.6982
4405 | 93.8465 | 0.9004 | 0.91612 | 1.0321 | 213.7524 | 52.8289 | 54629.7709 | 53379.8167 | 1063.5963 | 963.5963 | 971.7318
4406 | 96.6553 | 1.0021 | 1.0173 | 0.99239 | 187.8789 | 49.5137 | 54970.7445 | 55606.5712 | 1108.1314 | 1008.1314 | 971.7439
4407 | 91.6303 | 0.99794 | 0.99775 | 0.9684 | 201.8 | 50.9628 | 54103.016 | 52804.7824 | 1052.0956 | 952.0956 | 971.8332
4408 | 117.8095 | 1.1412 | 0.87927 | 0.98053 | 193.7473 | 52.9539 | 52967.1145 | 52342.7255 | 1042.8545 | 942.8545 | 971.8346
4409 | 122.5529 | 0.99614 | 0.80667 | 0.99049 | 200.0797 | 45.2527 | 41881.5677 | 41784.6562 | 831.6931 | 731.6931 | 971.9074
4410 | 124.1294 | 1.0016 | 0.62875 | 0.90932 | 187.9567 | 49.7651 | 34827.2953 | 35842.2359 | 712.8447 | 612.8447 | 971.9156
4411 | 97.2285 | 0.96112 | 1.0908 | 1.0173 | 182.9471 | 48.7789 | 58869.6033 | 58645.0023 | 1168.9 | 1068.9 | 972.076
4412 | 109.1974 | 1.0871 | 0.99762 | 1.0722 | 202.7774 | 50.0526 | 59684.8049 | 60658.892 | 1209.1778 | 1109.1778 | 972.162
4413 | 96.8005 | 0.97878 | 1.0159 | 1.0687 | 211.3541 | 52.2414 | 61878.7034 | 63507.2852 | 1266.1457 | 1166.1457 | 972.2121
4414 | 93.3291 | 1.075 | 0.93172 | 0.9717 | 199.1288 | 54.1735 | 54680.6506 | 54381.0758 | 1083.6215 | 983.6215 | 972.309
4415 | 114.1138 | 1.0056 | 1.0602 | 1.0575 | 214.2322 | 51.5554 | 63929.9976 | 62792.2442 | 1251.8449 | 1151.8449 | 972.3227
4416 | 88.8297 | 1.1584 | 0.97942 | 0.97983 | 217.3301 | 49.0463 | 52332.2816 | 51013.2068 | 1016.2641 | 916.2641 | 972.3624
4417 | 89.3305 | 1.0763 | 1.2033 | 0.99674 | 205.4962 | 49.9629 | 64933.4756 | 65802.5876 | 1312.0518 | 1212.0518 | 972.385
4418 | 122.8941 | 0.99286 | 0.76648 | 0.9979 | 204.3143 | 48.1712 | 42926.4281 | 40691.8657 | 809.8373 | 709.8373 | 972.3917
4419 | 98.0601 | 0.99869 | 0.97006 | 1.0368 | 196.8963 | 53.6958 | 59459.7471 | 59977.3539 | 1195.5471 | 1095.5471 | 972.4076
4420 | 103.5772 | 1.0664 | 1.1029 | 0.98507 | 198.0546 | 49.2918 | 59196.7632 | 58544.6105 | 1166.8922 | 1066.8922 | 972.5861
4421 | 105.4754 | 0.98398 | 0.9074 | 0.96653 | 191.9963 | 52.7489 | 51928.4197 | 52964.8059 | 1055.2961 | 955.2961 | 972.7259
4422 | 111.2418 | 0.99813 | 1.2001 | 0.9512 | 193.8788 | 50.2287 | 63106.8657 | 62876.8603 | 1253.5372 | 1153.5372 | 972.7765
4423 | 85.8476 | 0.99137 | 0.88999 | 1.0339 | 194.4259 | 47.582 | 48028.4333 | 49382.2822 | 983.6456 | 883.6456 | 972.7857
4424 | 118.7327 | 1.0597 | 1.1803 | 1.0016 | 213.5509 | 56.0798 | 73567.645 | 71969.0913 | 1435.3818 | 1335.3818 | 972.7979
4425 | 117.8213 | 0.93365 | 1.0144 | 1.0328 | 190.4833 | 48.1572 | 55944.5304 | 55273.0283 | 1101.4606 | 1001.4606 | 972.8617
4426 | 99.9115 | 0.96922 | 0.91039 | 0.97469 | 219.8334 | 47.9831 | 47443.8674 | 46404.9489 | 924.099 | 824.099 | 972.8922
4427 | 110.0941 | 1.049 | 1.0575 | 1.045 | 196.284 | 49.1664 | 60208.513 | 61759.2547 | 1231.1851 | 1131.1851 | 972.9321
4428 | 82.6958 | 1.0119 | 1.0069 | 1.0043 | 194.5607 | 51.6678 | 56767.8299 | 56579.3633 | 1127.5873 | 1027.5873 | 972.962
4429 | 90.229 | 1.0007 | 0.98881 | 1.0252 | 217.6665 | 48.2155 | 53446.387 | 50786.6916 | 1011.7338 | 911.7338 | 973.0003
4430 | 99.0873 | 1.0147 | 1.0313 | 1.0847 | 219.3066 | 48.2766 | 59075.1908 | 59041.194 | 1176.8239 | 1076.8239 | 973.1982
4431 | 99.5385 | 1.0629 | 1.3872 | 1.1245 | 193.4107 | 49.8974 | 83308.472 | 83590.2559 | 1667.8051 | 1567.8051 | 973.2049
4432 | 71.6764 | 0.98063 | 1.0449 | 1.0043 | 186.3594 | 49.1873 | 55258.258 | 54428.2133 | 1084.5643 | 984.5643 | 973.254
4433 | 116.4884 | 0.94173 | 1.4102 | 1.0399 | 204.8507 | 50.0454 | 79083.2247 | 78702.7737 | 1570.0555 | 1470.0555 | 973.3858
4434 | 107.6542 | 1.0376 | 0.9434 | 1.0234 | 210.4772 | 47.9474 | 51857.1682 | 51608.0063 | 1028.1601 | 928.1601 | 973.5433
4435 | 124.9234 | 1.0801 | 1.1448 | 1.0052 | 192.3518 | 51.1308 | 65929.8866 | 67579.4403 | 1347.5888 | 1247.5888 | 973.5748
4436 | 90.0365 | 1.0394 | 0.82027 | 1.0153 | 200.8913 | 47.8407 | 44521.2582 | 43437.1693 | 864.7434 | 764.7434 | 973.6631
4437 | 106.0709 | 0.94538 | 0.96724 | 1.0761 | 192.3953 | 53.6084 | 61363.7479 | 61052.3738 | 1217.0475 | 1117.0475 | 973.7399
4438 | 94.1954 | 0.9853 | 0.96138 | 0.95421 | 213.3382 | 49.9604 | 50681.5996 | 51364.475 | 1023.2895 | 923.2895 | 973.7549
4439 | 96.7565 | 1.1025 | 1.1474 | 1.0377 | 212.0937 | 52.5868 | 68430.9005 | 70282.222 | 1401.6444 | 1301.6444 | 973.7778
4440 | 87.0742 | 1.009 | 0.7045 | 1.0309 | 195.6267 | 48.6477 | 39802.2637 | 40090.3989 | 797.808 | 697.808 | 973.8106
4441 | 100.1433 | 0.95007 | 1.1051 | 1.0022 | 192.9092 | 47.8695 | 57764.6402 | 58189.7006 | 1159.794 | 1059.794 | 973.8184
4442 | 107.0419 | 0.96836 | 1.1842 | 0.92324 | 200.2285 | 49.9313 | 59966.5513 | 57908.1857 | 1154.1637 | 1054.1637 | 973.8191
4443 | 113.803 | 0.98536 | 0.81788 | 0.94104 | 206.6417 | 47.3099 | 41924.1151 | 41994.8338 | 835.8967 | 735.8967 | 973.827
4444 | 95.7945 | 0.93752 | 0.92123 | 1.0112 | 195.6554 | 44.3412 | 45485.7797 | 45599.0379 | 907.9808 | 807.9808 | 973.8689
4445 | 102.0425 | 0.98627 | 0.92552 | 0.9349 | 208.6575 | 49.5796 | 48098.1 | 49938.5499 | 994.771 | 894.771 | 973.8929
4446 | 106.4557 | 1.0076 | 0.98895 | 0.97762 | 203.9265 | 50.0067 | 53914.9297 | 52591.7577 | 1047.8352 | 947.8352 | 973.9093
4447 | 111.363 | 1 | 1.0964 | 0.98741 | 211.9564 | 46.4607 | 55685.3735 | 57485.1361 | 1145.7027 | 1045.7027 | 973.9181
4448 | 101.0691 | 0.98792 | 0.66716 | 1.0235 | 190.8165 | 53.3563 | 41953.7296 | 40436.451 | 804.729 | 704.729 | 973.9367
4449 | 114.7151 | 0.92905 | 1.0482 | 0.98497 | 201.8974 | 50.168 | 57342.783 | 57548.6325 | 1146.9727 | 1046.9727 | 973.9722
4450 | 97.8255 | 0.98753 | 1.2218 | 0.93738 | 198.6051 | 51.0482 | 63594.5809 | 63146.3929 | 1258.9279 | 1158.9279 | 974.0866
4451 | 107.4247 | 1.0217 | 1.0124 | 1.0859 | 205.7342 | 48.0226 | 58268.3207 | 59416.391 | 1184.3278 | 1084.3278 | 974.1259
4452 | 100.6643 | 1.0269 | 1.1876 | 1.0073 | 201.9093 | 47.9997 | 62580.7306 | 64224.619 | 1280.4924 | 1180.4924 | 974.1652
4453 | 105.8227 | 0.98307 | 0.83208 | 0.92442 | 209.71 | 52.2421 | 45828.2995 | 46095.2239 | 917.9045 | 817.9045 | 974.2089
4454 | 82.5797 | 0.95016 | 0.88448 | 1.0563 | 201.838 | 50.2077 | 51049.9237 | 50435.4541 | 1004.7091 | 904.7091 | 974.2177
4455 | 102.1004 | 1.0106 | 1.0325 | 1.0055 | 192.1277 | 52.6609 | 60301.5928 | 60906.3318 | 1214.1266 | 1114.1266 | 974.2223
4456 | 97.3089 | 1.0574 | 1.304 | 0.93654 | 190.3005 | 50.1252 | 66563.5535 | 66562.9699 | 1327.2594 | 1227.2594 | 974.2662
4457 | 113.5262 | 1.0678 | 1.141 | 1.0216 | 204.7887 | 52.085 | 67226.8703 | 64695.1793 | 1289.9036 | 1189.9036 | 974.2873
4458 | 112.6519 | 1.0571 | 0.94199 | 0.97203 | 200.6161 | 52.9225 | 54960.7155 | 55129.7532 | 1098.5951 | 998.5951 | 974.2979
4459 | 114.1679 | 1.0757 | 0.92933 | 1.0248 | 193.9944 | 47.3555 | 51110.9811 | 50965.6813 | 1015.3136 | 915.3136 | 974.3026
4460 | 82.0553 | 1.0645 | 0.91355 | 1.0778 | 203.5061 | 48.8103 | 52526.2491 | 54403.8429 | 1084.0769 | 984.0769 | 974.3265
4461 | 88.2727 | 1.0174 | 0.77799 | 1.0452 | 204.1775 | 49.9206 | 45280.2354 | 43560.1033 | 867.2021 | 767.2021 | 974.3297
4462 | 88.7472 | 1.0388 | 0.91191 | 1.0416 | 199.1757 | 45.8953 | 48024.6464 | 45034.3567 | 896.6871 | 796.6871 | 974.3746
4463 | 113.5343 | 1.0155 | 1.153 | 1.0174 | 201.1798 | 51.3849 | 66407.3519 | 67638.6272 | 1348.7725 | 1248.7725 | 974.3847
4464 | 98.388 | 1.0139 | 1.105 | 0.98773 | 196.0551 | 50.3889 | 60219.0974 | 62123.705 | 1238.4741 | 1138.4741 | 974.3933
4465 | 103.7097 | 1.0231 | 1.2181 | 1.0625 | 174.9849 | 46.3135 | 65026.531 | 65528.2223 | 1306.5644 | 1206.5644 | 974.4135
4466 | 106.0413 | 0.94581 | 1.1183 | 1.0096 | 180.115 | 49.0084 | 60425.7146 | 59662.9519 | 1189.259 | 1089.259 | 974.5232
4467 | 94.7038 | 1.0189 | 1.1324 | 1.0654 | 202.4451 | 51.6243 | 67466.0768 | 68115.145 | 1358.3029 | 1258.3029 | 974.6651
4468 | 66.8233 | 1.0212 | 1.0819 | 0.91956 | 192.1502 | 52.2581 | 55747.6792 | 55412.9898 | 1104.2598 | 1004.2598 | 974.6848
4469 | 115.6849 | 0.97958 | 1.281 | 0.94539 | 209.0336 | 48.686 | 64688.0996 | 63647.5883 | 1268.9518 | 1168.9518 | 974.7524
4470 | 100.8079 | 1.0231 | 1.1348 | 0.92269 | 206.5948 | 48.0869 | 55515.1502 | 56115.0062 | 1118.3001 | 1018.3001 | 974.8
4471 | 89.1044 | 0.95811 | 1.1053 | 1.0216 | 206.9411 | 50.3339 | 61343.5476 | 60526.6463 | 1206.5329 | 1106.5329 | 974.8112
4472 | 96.1759 | 1.0638 | 0.78935 | 0.97195 | 202.0787 | 51.9629 | 45384.7279 | 45377.5205 | 903.5504 | 803.5504 | 974.8684
4473 | 106.161 | 0.96705 | 0.94256 | 1.0304 | 195.2171 | 50.448 | 54371.1376 | 53836.2958 | 1072.7259 | 972.7259 | 975.1011
4474 | 103.8557 | 1.0881 | 0.92402 | 0.89656 | 173.7108 | 49.209 | 46501.3837 | 48741.9425 | 970.8389 | 870.8389 | 975.1432
4475 | 69.4648 | 1.0351 | 0.97727 | 1.0417 | 198.9583 | 49.9663 | 54658.5319 | 54507.0673 | 1086.1413 | 986.1413 | 975.1914
4476 | 95.4876 | 1.0696 | 0.83877 | 1.0317 | 204.6601 | 52.4919 | 50989.6277 | 51658.0775 | 1029.1615 | 929.1615 | 975.2152
4477 | 102.2095 | 0.98894 | 0.88523 | 1.0264 | 212.6605 | 51.1921 | 51898.7872 | 52902.7165 | 1054.0543 | 954.0543 | 975.2247
4478 | 90.9611 | 0.98502 | 1.2821 | 1.0459 | 203.1821 | 53.7199 | 77049.6047 | 74571.3348 | 1487.4267 | 1387.4267 | 975.2592
4479 | 89.5641 | 0.99633 | 0.78504 | 0.9474 | 209.117 | 50.9757 | 42671.0813 | 42803.6076 | 852.0722 | 752.0722 | 975.2824
4480 | 84.147 | 1.0756 | 1.0118 | 1.0411 | 193.4957 | 51.2479 | 58815.6476 | 60542.3164 | 1206.8463 | 1106.8463 | 975.409
4481 | 109.9717 | 0.95796 | 1.0606 | 0.97785 | 177.201 | 53.3169 | 61091.1972 | 61313.6951 | 1222.2739 | 1122.2739 | 975.4503
4482 | 97.4791 | 1.0152 | 1.038 | 1.0339 | 209.452 | 51.1969 | 60223.6688 | 59421.8654 | 1184.4373 | 1084.4373 | 975.5233
4483 | 83.7889 | 0.97745 | 1.1822 | 1.0105 | 179.8856 | 51.9688 | 66517.3766 | 66740.8054 | 1330.8161 | 1230.8161 | 975.628
4484 | 121.0513 | 0.93626 | 1.0152 | 1.0389 | 213.2703 | 52.1263 | 61098.2786 | 60565.0157 | 1207.3003 | 1107.3003 | 975.6799
4485 | 99.9337 | 1.003 | 0.92564 | 0.96798 | 201.2761 | 50.8275 | 50837.6502 | 50118.677 | 998.3735 | 898.3735 | 976.1062
4486 | 113.2711 | 1.0152 | 1.012 | 1.0982 | 218.6352 | 51.6472 | 63556.1273 | 66143.4763 | 1318.8695 | 1218.8695 | 976.1114
4487 | 87.0487 | 0.92566 | 0.72778 | 0.95747 | 193.6271 | 49.7876 | 38898.881 | 39202.6451 | 780.0529 | 680.0529 | 976.1353
4488 | 107.0845 | 1.0608 | 1.0205 | 1.0221 | 198.8161 | 53.3984 | 61963.9919 | 60200.3839 | 1200.0077 | 1100.0077 | 976.1485
4489 | 92.1742 | 0.99134 | 0.99535 | 1.0387 | 216.947 | 50.5853 | 57136.0091 | 57623.0487 | 1148.461 | 1048.461 | 976.162
4490 | 102.5951 | 1.0393 | 1.0118 | 1.0108 | 208.21 | 47.9683 | 54384.3852 | 55576.7951 | 1107.5359 | 1007.5359 | 976.2797
4491 | 111.4498 | 0.95906 | 0.87783 | 0.92281 | 205.5433 | 47.5271 | 43785.773 | 44444.6218 | 884.8924 | 784.8924 | 976.3184
4492 | 96.6782 | 1.0275 | 0.9428 | 0.97759 | 188.4668 | 52.1281 | 53411.5205 | 53645.7103 | 1068.9142 | 968.9142 | 976.4298
4493 | 105.8935 | 0.95068 | 1.2138 | 1.0037 | 195.4883 | 48.6528 | 64367.9883 | 64404.474 | 1284.0895 | 1184.0895 | 976.5282
4494 | 119.6494 | 0.94957 | 1.0928 | 1.0098 | 196.963 | 52.2192 | 63751.4886 | 64340.9937 | 1282.8199 | 1182.8199 | 976.5351
4495 | 93.6252 | 0.90878 | 0.80247 | 0.97707 | 207.5172 | 45.3489 | 39622.7387 | 39036.6963 | 776.7339 | 676.7339 | 976.5596
4496 | 90.1402 | 1.0143 | 0.99953 | 0.98019 | 198.166 | 51.3826 | 55236.7833 | 55965.8596 | 1115.3172 | 1015.3172 | 976.6286
4497 | 104.4283 | 0.92581 | 1.1503 | 1.0071 | 194.3709 | 48.5253 | 61101.2634 | 60203.3829 | 1200.0677 | 1100.0677 | 976.6296
4498 | 103.7622 | 1.0298 | 0.75724 | 1.1136 | 200.9652 | 45.6977 | 43619.058 | 42828.9364 | 852.5787 | 752.5787 | 976.6498
4499 | 96.6352 | 0.93308 | 0.85191 | 0.91442 | 212.5514 | 47.6229 | 41605.1699 | 42734.1072 | 850.6821 | 750.6821 | 976.6849
4500 | 107.7447 | 1.0409 | 0.89862 | 0.9497 | 211.8066 | 48.1843 | 46736.6889 | 46643.5157 | 928.8703 | 828.8703 | 976.7137
4501 | 89.5937 | 1.0131 | 1.26 | 0.95141 | 194.0966 | 50.6053 | 65453.3818 | 64939.0057 | 1294.7801 | 1194.7801 | 976.7752
4502 | 93.4422 | 0.89952 | 1.3359 | 1.0336 | 191.2036 | 52.2389 | 76711.7089 | 75630.7355 | 1508.6147 | 1408.6147 | 976.7976
4503 | 97.7698 | 1.0039 | 1.0874 | 0.98981 | 232.1163 | 49.9944 | 58949.6639 | 59541.4191 | 1186.8284 | 1086.8284 | 976.805
4504 | 114.5652 | 0.97583 | 0.58428 | 1.0976 | 201.8714 | 52.0529 | 39403.2056 | 38439.902 | 764.798 | 664.798 | 976.8207
4505 | 85.2884 | 1.0255 | 0.94462 | 0.96944 | 200.7995 | 52.1085 | 52476.2915 | 52390.5525 | 1043.8111 | 943.8111 | 976.9148
4506 | 98.1125 | 0.944 | 0.91313 | 1.011 | 198.2452 | 49.0842 | 50057.5101 | 50026.3987 | 996.528 | 896.528 | 976.9731
4507 | 106.1713 | 1.0261 | 0.95526 | 0.93222 | 201.2162 | 57.7499 | 57919.6225 | 57898.1602 | 1153.9632 | 1053.9632 | 977.0111
4508 | 77.2534 | 0.97861 | 0.75918 | 0.96378 | 200.5788 | 52.3296 | 42445.4324 | 42703.8444 | 850.0769 | 750.0769 | 977.0909
4509 | 77.7005 | 0.95804 | 1.0158 | 0.97024 | 205.1435 | 48.9195 | 52059.7045 | 52278.0523 | 1041.561 | 941.561 | 977.1621
4510 | 101.6626 | 1.0164 | 0.87359 | 1.0164 | 208.4368 | 53.6677 | 53407.4392 | 54351.7788 | 1083.0356 | 983.0356 | 977.1867
4511 | 98.2099 | 0.95077 | 1.1486 | 0.97408 | 208.6932 | 49.6627 | 60411.3625 | 59910.228 | 1194.2046 | 1094.2046 | 977.2293
4512 | 100.3299 | 0.9987 | 0.97478 | 1.0336 | 207.7916 | 42.788 | 47603.6465 | 49107.3972 | 978.1479 | 878.1479 | 977.2645
4513 | 96.424 | 0.976 | 1.0876 | 0.96859 | 203.7139 | 55.3263 | 63695.5308 | 64055.0245 | 1277.1005 | 1177.1005 | 977.2949
4514 | 94.0592 | 1.0154 | 1.1282 | 1.1269 | 208.9997 | 47.7917 | 65536.5178 | 64337.0887 | 1282.7418 | 1182.7418 | 977.3105
4515 | 100.4191 | 0.97171 | 0.79292 | 0.95572 | 197.6561 | 51.1864 | 43981.8744 | 44116.4543 | 878.3291 | 778.3291 | 977.4161
4516 | 102.8369 | 1.0725 | 1.0267 | 1.0088 | 194.4408 | 50.5307 | 58104.6625 | 56081.9482 | 1117.639 | 1017.639 | 977.4342
4517 | 94.2016 | 0.91589 | 1.4413 | 1.0713 | 218.4904 | 46.0171 | 75243.6594 | 76282.249 | 1521.645 | 1421.645 | 977.5505
4518 | 87.6395 | 0.96763 | 1.0337 | 0.96889 | 199.0513 | 48.775 | 53183.5362 | 55957.296 | 1115.1459 | 1015.1459 | 977.5906
4519 | 100.4828 | 0.95436 | 0.9798 | 0.98708 | 206.9005 | 49.2 | 52508.3174 | 51232.4463 | 1020.6489 | 920.6489 | 977.6055
4520 | 97.2076 | 1.0155 | 0.9595 | 1.0365 | 205.5603 | 47.7993 | 52460.8912 | 52004.0427 | 1036.0809 | 936.0809 | 977.7261
4521 | 91.106 | 1.0493 | 1.1568 | 0.91315 | 194.4778 | 49.4508 | 57159.4586 | 56277.6103 | 1121.5522 | 1021.5522 | 977.9233
4522 | 108.7215 | 0.98768 | 1.1706 | 0.9677 | 189.0588 | 47.2656 | 58804.4046 | 59796.6865 | 1191.9337 | 1091.9337 | 977.9389
4523 | 91.8607 | 1.0485 | 0.88663 | 1.0348 | 201.2722 | 49.7224 | 50610.3103 | 48877.3229 | 973.5465 | 873.5465 | 977.9456
4524 | 99.0086 | 1.02 | 0.94544 | 1.0562 | 195.4675 | 49.6878 | 54832.3452 | 56667.1154 | 1129.3423 | 1029.3423 | 977.9521
4525 | 109.8775 | 0.98351 | 1.2412 | 1.0324 | 203.0987 | 47.5415 | 66263.5377 | 63772.9197 | 1271.4584 | 1171.4584 | 977.9542
4526 | 104.4917 | 0.89847 | 0.67898 | 0.98166 | 195.7725 | 51.1426 | 39085.1303 | 38711.9169 | 770.2383 | 670.2383 | 978.104
4527 | 105.5839 | 1.007 | 0.8885 | 1.0481 | 208.3513 | 48.9864 | 51034.4586 | 52169.205 | 1039.3841 | 939.3841 | 978.1781
4528 | 103.692 | 1.0459 | 0.97989 | 0.92118 | 175.5532 | 47.1427 | 47842.166 | 47591.3548 | 947.8271 | 847.8271 | 978.2792
4529 | 98.6976 | 0.93202 | 0.88059 | 1.0529 | 198.008 | 51.0174 | 52192.7337 | 53577.4072 | 1067.5481 | 967.5481 | 978.347
4530 | 104.9036 | 0.99241 | 1.0159 | 1.0196 | 199.4252 | 51.5672 | 58981.1671 | 59358.5355 | 1183.1707 | 1083.1707 | 978.3742
4531 | 100.0545 | 1.0212 | 1.0066 | 1.0855 | 184.8679 | 50.5039 | 60528.1245 | 61152.6769 | 1219.0535 | 1119.0535 | 978.3958
4532 | 91.7183 | 0.96284 | 0.94772 | 1.0701 | 181.811 | 50.7579 | 56139.2151 | 55752.0171 | 1111.0403 | 1011.0403 | 978.4393
4533 | 97.0393 | 1.0108 | 1.1039 | 1.013 | 187.2408 | 53.444 | 65198.7612 | 66468.1466 | 1325.3629 | 1225.3629 | 978.4541
4534 | 116.1014 | 0.92549 | 1.0103 | 0.99539 | 193.7671 | 46.3136 | 51744.062 | 51150.1666 | 1019.0033 | 919.0033 | 978.4636
4535 | 90.546 | 1.0811 | 0.92215 | 1.0215 | 202.7284 | 50.3157 | 52524.3672 | 51126.8464 | 1018.5369 | 918.5369 | 978.521
4536 | 94.97 | 0.99098 | 1.0604 | 1.0221 | 195.3433 | 54.5205 | 64415.3424 | 65611.7041 | 1308.2341 | 1208.2341 | 978.5579
4537 | 100.5599 | 1.0335 | 0.94514 | 0.99757 | 212.8643 | 54.8028 | 57578.8372 | 58876.3445 | 1173.5269 | 1073.5269 | 978.6964
4538 | 100.8889 | 1.0568 | 0.96851 | 0.99349 | 187.8314 | 50.8706 | 54559.883 | 54199.6617 | 1079.9932 | 979.9932 | 978.8092
4539 | 106.8286 | 1.0238 | 0.83159 | 0.96912 | 197.4353 | 48.2861 | 44392.5577 | 43769.2721 | 871.3854 | 771.3854 | 978.8262
4540 | 90.097 | 1.0106 | 0.98027 | 1.0437 | 188.4534 | 47.6316 | 53259.9594 | 53766.9654 | 1071.3393 | 971.3393 | 978.8989
4541 | 99.9938 | 1.0059 | 0.53976 | 1.0426 | 190.4606 | 49.9325 | 33314.0743 | 33553.8144 | 667.0763 | 567.0763 | 978.9274
4542 | 93.9721 | 0.92209 | 1.0056 | 1.0299 | 203.6812 | 50.7143 | 57121.6664 | 57654.6498 | 1149.093 | 1049.093 | 978.9304
4543 | 77.849 | 1.0025 | 0.99417 | 0.97209 | 201.0865 | 47.2629 | 49565.0959 | 48322.2097 | 962.4442 | 862.4442 | 978.9344
4544 | 100.9269 | 0.94882 | 1.0708 | 0.93746 | 198.4877 | 50.3063 | 55513.2275 | 55852.6761 | 1113.0535 | 1013.0535 | 978.9895
4545 | 96.4948 | 1.0416 | 0.86539 | 1.0652 | 209.2348 | 48.9887 | 50291.0487 | 50024.9816 | 996.4996 | 896.4996 | 978.9922
4546 | 110.6764 | 0.93152 | 0.83825 | 1.0152 | 204.2128 | 48.7167 | 46685.1958 | 47725.9521 | 950.519 | 850.519 | 979.0154
4547 | 84.8596 | 1.0245 | 1.1472 | 0.96785 | 219.5758 | 50.3753 | 60531.3653 | 60167.3853 | 1199.3477 | 1099.3477 | 979.0419
4548 | 91.2745 | 0.97296 | 1.0365 | 1.0558 | 206.1578 | 50.9128 | 60445.3811 | 61352.9312 | 1223.0586 | 1123.0586 | 979.1689
4549 | 96.5507 | 0.93773 | 0.99781 | 0.89823 | 211.7636 | 45.4708 | 45082.6246 | 45578.1291 | 907.5626 | 807.5626 | 979.2407
4550 | 82.5082 | 0.97837 | 1.039 | 1.036 | 206.4171 | 52.8404 | 61346.5976 | 61069.7141 | 1217.3943 | 1117.3943 | 979.2932
4551 | 97.8283 | 1.0716 | 1.2476 | 1.0049 | 202.936 | 49.0048 | 66776.981 | 68220.763 | 1360.4153 | 1260.4153 | 979.3042
4552 | 123.5899 | 1.0419 | 0.98275 | 1.0534 | 186.8356 | 49.0891 | 57325.505 | 58949.6466 | 1174.9929 | 1074.9929 | 979.3363
4553 | 79.8676 | 1.0498 | 1.3101 | 0.98612 | 186.9582 | 52.1309 | 71905.6386 | 72949.4109 | 1454.9882 | 1354.9882 | 979.3897
4554 | 101.9518 | 1.0336 | 1.1306 | 0.99129 | 213.2866 | 49.2754 | 60632.0826 | 58931.0361 | 1174.6207 | 1074.6207 | 979.4454
4555 | 108.9519 | 1.0028 | 1.1284 | 0.9576 | 206.1255 | 53.107 | 63393.5498 | 62162.2109 | 1239.2442 | 1139.2442 | 979.5187
4556 | 101.1344 | 1.0179 | 0.76206 | 0.98379 | 186.2687 | 54.8061 | 46917.139 | 45715.576 | 910.3115 | 810.3115 | 979.5281
4557 | 116.9771 | 1.0545 | 0.98631 | 1.1314 | 190.2188 | 50.8859 | 63250.156 | 66361.1884 | 1323.2238 | 1223.2238 | 979.5735
4558 | 110.7797 | 0.99119 | 0.88549 | 1.0831 | 197.8536 | 45.8866 | 49245.386 | 47321.7725 | 942.4355 | 842.4355 | 979.5752
4559 | 105.1418 | 1.0343 | 1.0482 | 1.0073 | 190.3223 | 50.6973 | 59230.2762 | 61253.5699 | 1221.0714 | 1121.0714 | 979.8168
4560 | 105.1288 | 1.0789 | 1.2307 | 0.97437 | 198.0592 | 48.7537 | 64191.8308 | 65609.5392 | 1308.1908 | 1208.1908 | 979.8528
4561 | 89.6962 | 1.0067 | 0.98558 | 1.0402 | 190.5465 | 49.0643 | 54923.1999 | 54522.622 | 1086.4524 | 986.4524 | 979.8987
4562 | 105.3222 | 0.95834 | 1.2213 | 0.99393 | 202.8586 | 48.6869 | 64216.2979 | 62181.555 | 1239.6311 | 1139.6311 | 979.9622
4563 | 111.8455 | 0.90403 | 0.87754 | 0.95638 | 209.2542 | 51.7828 | 48904.5378 | 48932.7432 | 974.6549 | 874.6549 | 979.9698
4564 | 103.9023 | 1.0097 | 0.91635 | 1.0268 | 188.9494 | 46.8088 | 49144.7748 | 49707.8197 | 990.1564 | 890.1564 | 979.9703
4565 | 87.1937 | 0.97559 | 0.97166 | 0.93876 | 200.9837 | 49.0622 | 49126.9538 | 49894.9659 | 993.8993 | 893.8993 | 979.9745
4566 | 92.5535 | 1.0172 | 1.1749 | 1.0722 | 191.9118 | 49.4431 | 67127.82 | 68347.9031 | 1362.9581 | 1262.9581 | 979.9932
4567 | 96.5533 | 0.94785 | 1.0051 | 1.0191 | 210.5632 | 51.1691 | 57303.4542 | 57210.2461 | 1140.2049 | 1040.2049 | 980.145
4568 | 93.2787 | 0.96826 | 0.90333 | 1.0665 | 196.5427 | 55.0036 | 58155.6688 | 59159.2444 | 1179.1849 | 1079.1849 | 980.1504
4569 | 91.6843 | 0.97693 | 1.1503 | 1.0282 | 203.7703 | 57.9421 | 73923.5115 | 76224.5124 | 1520.4902 | 1420.4902 | 980.166
4570 | 105.9205 | 1.0081 | 0.96914 | 0.96708 | 214.5283 | 50.614 | 53056.6023 | 52724.4312 | 1050.4886 | 950.4886 | 980.1941
4571 | 96.9081 | 0.99504 | 1.0837 | 0.99546 | 219.0443 | 50.788 | 59902.9091 | 61292.8953 | 1221.8579 | 1121.8579 | 980.205
4572 | 105.296 | 0.95347 | 0.99714 | 1.0159 | 200.6166 | 49.2044 | 54984.8018 | 55091.2251 | 1097.8245 | 997.8245 | 980.3559
4573 | 99.2085 | 1.0422 | 1.2275 | 1.0305 | 194.2952 | 54.1069 | 74229.4215 | 74398.8301 | 1483.9766 | 1383.9766 | 980.4411
4574 | 95.2497 | 0.97421 | 1.0029 | 1.0642 | 197.9596 | 47.9928 | 55875.1782 | 57368.9903 | 1143.3798 | 1043.3798 | 980.4593
4575 | 96.8497 | 1.0259 | 1.1474 | 1.0003 | 209.8593 | 48.9065 | 61201.6861 | 61612.5239 | 1228.2505 | 1128.2505 | 980.5208
4576 | 108.9773 | 0.99939 | 0.92232 | 0.95951 | 193.0664 | 50.1408 | 50027.2906 | 49562.7829 | 987.2557 | 887.2557 | 980.5437
4577 | 108.2472 | 1.0136 | 0.99888 | 0.95534 | 191.2832 | 50.6693 | 54102.5293 | 52971.3784 | 1055.4276 | 955.4276 | 980.6038
4578 | 109.9745 | 1.029 | 0.93923 | 0.99925 | 208.3687 | 46.288 | 48889.0429 | 47744.7917 | 950.8958 | 850.8958 | 980.6204
4579 | 98.4023 | 1.0814 | 1.1504 | 1.0284 | 210.7299 | 52.7975 | 68293.3505 | 69905.7842 | 1394.1157 | 1294.1157 | 980.6257
4580 | 114.1314 | 1.0033 | 0.85801 | 0.99247 | 213.6987 | 44.9775 | 43664.6022 | 43306.6168 | 862.1323 | 762.1323 | 980.6674
4581 | 109.9746 | 0.99915 | 0.97595 | 0.92377 | 201.8441 | 49.2139 | 49978.1754 | 49806.6559 | 992.1331 | 892.1331 | 980.7167
4582 | 91.8559 | 1.0344 | 0.75828 | 0.8736 | 200.3366 | 48.3787 | 36844.7765 | 35513.7131 | 706.2743 | 606.2743 | 980.7584
4583 | 109.98 | 1.0667 | 0.87159 | 1.0124 | 209.5028 | 47.5282 | 47724.2671 | 48254.6017 | 961.092 | 861.092 | 980.875
4584 | 102.6971 | 0.97632 | 1.2019 | 0.98658 | 203.8326 | 48.135 | 62109.1884 | 62153.2128 | 1239.0643 | 1139.0643 | 980.8994
4585 | 104.9005 | 0.95688 | 0.91592 | 0.93599 | 209.995 | 46.748 | 44979.0663 | 44687.2517 | 889.745 | 789.745 | 980.9007
4586 | 104.1583 | 0.91001 | 1.1544 | 0.97316 | 212.6104 | 51.5642 | 63029.7911 | 62568.9933 | 1247.3799 | 1147.3799 | 980.9288
4587 | 96.4509 | 1.074 | 0.96315 | 1.018 | 194.2813 | 45.9711 | 50031.8438 | 50387.1482 | 1003.743 | 903.743 | 980.9375
4588 | 122.1522 | 1.0067 | 1.2096 | 1.063 | 208.8039 | 51.1951 | 72333.9021 | 70125.5286 | 1398.5106 | 1298.5106 | 980.9458
4589 | 102.9703 | 1.0535 | 0.95649 | 0.98178 | 206.923 | 47.928 | 50413.9941 | 49581.9829 | 987.6397 | 887.6397 | 981.0362
4590 | 114.67 | 0.89827 | 1.0553 | 1.0624 | 199.1104 | 50.8763 | 62477.3572 | 61353.8584 | 1223.0772 | 1123.0772 | 981.0679
4591 | 105.4944 | 0.97882 | 1.0956 | 1.0106 | 201.1146 | 48.7207 | 59176.7975 | 58022.6879 | 1156.4538 | 1056.4538 | 981.0907
4592 | 95.6602 | 1.0306 | 1.0599 | 0.99333 | 211.2331 | 51.0724 | 59018.7485 | 60032.7563 | 1196.6551 | 1096.6551 | 981.103
4593 | 109.1538 | 1.0451 | 1.1686 | 1.0406 | 194.802 | 50.297 | 67092.6857 | 69207.9138 | 1380.1583 | 1280.1583 | 981.187
4594 | 109.3922 | 0.99509 | 1.1655 | 1.118 | 188.8886 | 51.6406 | 73102.3989 | 75102.5417 | 1498.0508 | 1398.0508 | 981.2452
4595 | 99.9722 | 1.0054 | 1.0601 | 1.0258 | 195.2776 | 53.9089 | 64238.6524 | 64775.2717 | 1291.5054 | 1191.5054 | 981.2496
4596 | 99.2477 | 1.0018 | 1.0308 | 1.0218 | 200.9206 | 49.7975 | 57606.5369 | 57023.2175 | 1136.4644 | 1036.4644 | 981.2526
4597 | 107.2156 | 1.0075 | 0.95901 | 0.97001 | 194.5917 | 54.3387 | 56613.0222 | 57932.6156 | 1154.6523 | 1054.6523 | 981.3092
4598 | 94.3482 | 0.92226 | 0.83863 | 0.94515 | 200.5641 | 50.7749 | 44864.5536 | 45502.879 | 906.0576 | 806.0576 | 981.3173
4599 | 108.0946 | 1.0239 | 1.0831 | 1.0023 | 193.8737 | 47.5944 | 57132.1344 | 57014.208 | 1136.2842 | 1036.2842 | 981.3355
4600 | 84.1083 | 0.91756 | 0.93251 | 0.93937 | 191.5225 | 51.6306 | 49402.8739 | 49701.8855 | 990.0377 | 890.0377 | 981.3738
4601 | 77.0038 | 1.0437 | 0.99169 | 0.9881 | 208.0177 | 52.4005 | 55766.0483 | 54544.4403 | 1086.8888 | 986.8888 | 981.4444
4602 | 99.2137 | 0.9888 | 0.98036 | 0.99966 | 198.5287 | 46.2567 | 50069.4892 | 49430.4578 | 984.6092 | 884.6092 | 981.4498
4603 | 95.8002 | 0.95903 | 0.84709 | 1.0043 | 200.2363 | 49.6159 | 46967.2054 | 45817.4228 | 912.3485 | 812.3485 | 981.4678
4604 | 98.7832 | 0.97703 | 1.0113 | 0.96011 | 202.0061 | 51.9967 | 55708.3622 | 55163.3261 | 1099.2665 | 999.2665 | 981.568
4605 | 92.2805 | 0.92421 | 0.85756 | 0.9635 | 187.6581 | 49.3253 | 45149.6833 | 46879.6972 | 933.5939 | 833.5939 | 981.5835
4606 | 96.3742 | 0.95354 | 1.2624 | 1.1056 | 228.4094 | 51.1145 | 76261.4536 | 75420.131 | 1504.4026 | 1404.4026 | 981.6555
4607 | 89.905 | 1.0179 | 0.91158 | 0.9394 | 202.4592 | 47.4851 | 45211.2299 | 45131.683 | 898.6337 | 798.6337 | 981.6558
4608 | 90.0299 | 1.0596 | 1.1149 | 1.0453 | 199.1265 | 49.2484 | 62289.0874 | 63221.8455 | 1260.4369 | 1160.4369 | 981.6902
4609 | 109.7429 | 0.98992 | 0.84343 | 1.066 | 187.4022 | 46.6794 | 47229.3766 | 46274.9917 | 921.4998 | 821.4998 | 981.7038
4610 | 110.8774 | 1.0209 | 0.84082 | 0.92232 | 199.5926 | 48.4247 | 43235.1311 | 42483.8616 | 845.6772 | 745.6772 | 981.7872
4611 | 98.327 | 0.98511 | 1.1652 | 0.99931 | 198.1634 | 51.4472 | 65084.5673 | 64612.5381 | 1288.2508 | 1188.2508 | 981.8428
4612 | 92.598 | 1.025 | 0.80829 | 0.98738 | 213.5276 | 54.4219 | 48812.8488 | 49471.2512 | 985.425 | 885.425 | 981.9216
4613 | 106.6572 | 0.95702 | 1.0412 | 1.0223 | 188.8328 | 53.0471 | 62066.9449 | 62286.357 | 1241.7271 | 1141.7271 | 981.9226
4614 | 112.8381 | 0.91015 | 1.094 | 1.0528 | 202.7227 | 48.7447 | 61350.9692 | 60968.1017 | 1215.362 | 1115.362 | 981.9243
4615 | 112.9101 | 0.95336 | 1.2479 | 0.98865 | 216.7734 | 52.3579 | 70447.1082 | 71988.5926 | 1435.7719 | 1335.7719 | 982.1756
4616 | 104.6711 | 1.0161 | 0.95379 | 0.99393 | 203.449 | 52.3992 | 55450.8933 | 56062.2362 | 1117.2447 | 1017.2447 | 982.2251
4617 | 102.0576 | 1.0314 | 0.98508 | 0.93453 | 199.9487 | 47.2872 | 48709.4384 | 51033.3763 | 1016.6675 | 916.6675 | 982.2274
4618 | 86.9998 | 1.0383 | 1.3525 | 1.0071 | 195.729 | 52.603 | 76600.3423 | 77972.6519 | 1555.453 | 1455.453 | 982.2298
4619 | 111.1914 | 0.98597 | 0.75537 | 0.95839 | 205.9766 | 52.549 | 44008.9419 | 46002.2324 | 916.0446 | 816.0446 | 982.2583
4620 | 104.8988 | 1.0408 | 1.1704 | 0.96607 | 188.9584 | 55.0709 | 68470.8092 | 67904.4837 | 1354.0897 | 1254.0897 | 982.293
4621 | 93.4989 | 1.0158 | 0.98719 | 1.0298 | 182.1002 | 53.8091 | 59996.4447 | 60813.1072 | 1212.2621 | 1112.2621 | 982.3076
4622 | 98.1687 | 0.9694 | 0.91743 | 0.88991 | 189.7062 | 47.8767 | 43834.1032 | 44017.0089 | 876.3402 | 776.3402 | 982.3632
4623 | 90.8478 | 0.89366 | 1.0197 | 1.0489 | 210.1567 | 46.9472 | 54233.0419 | 53677.013 | 1069.5403 | 969.5403 | 982.4106
4624 | 82.5633 | 0.98255 | 1.0572 | 1.0495 | 197.7014 | 53.8857 | 64359.282 | 64618.1304 | 1288.3626 | 1188.3626 | 982.4402
4625 | 90.4664 | 1.0364 | 0.97101 | 1.0707 | 205.4962 | 52.4331 | 59633.456 | 58752.6752 | 1171.0535 | 1071.0535 | 982.4458
4626 | 104.7693 | 1.0544 | 1.0736 | 1.0603 | 194.9062 | 54.0263 | 67662.8005 | 67480.4781 | 1345.6096 | 1245.6096 | 982.4591
4627 | 110.8633 | 1.0205 | 0.78045 | 0.91305 | 186.9358 | 51.589 | 42785.1258 | 41539.2275 | 826.7845 | 726.7845 | 982.6436
4628 | 105.3185 | 0.96197 | 1.0957 | 0.94439 | 208.5509 | 48.1567 | 54917.4577 | 54039.882 | 1076.7976 | 976.7976 | 982.6555
4629 | 102.8641 | 1.0768 | 0.97797 | 0.98496 | 200.2934 | 48.2671 | 52040.3995 | 52689.8085 | 1049.7962 | 949.7962 | 982.6692
4630 | 109.3653 | 0.93012 | 1.2528 | 1.0108 | 200.485 | 50.7982 | 69695.4829 | 69843.1749 | 1392.8635 | 1292.8635 | 982.6956
4631 | 99.4707 | 1.0608 | 1.3468 | 0.98093 | 197.0333 | 44.3468 | 63465.761 | 63988.3079 | 1275.7662 | 1175.7662 | 982.7793
4632 | 90.0013 | 0.99387 | 0.94414 | 1.0034 | 193.3623 | 51.6627 | 53757.7509 | 56792.8761 | 1131.8575 | 1031.8575 | 982.8043
4633 | 100.3394 | 1.0149 | 1.0659 | 0.95643 | 188.9495 | 52.6332 | 59205.0354 | 56465.487 | 1125.3097 | 1025.3097 | 982.8465
4634 | 112.3875 | 1.0085 | 0.87357 | 1.0251 | 184.403 | 50.2248 | 50854.6743 | 50028.6595 | 996.5732 | 896.5732 | 982.9029
4635 | 93.9146 | 1.0244 | 1.2456 | 0.98344 | 202.3151 | 49.2363 | 65251.86 | 66630.6256 | 1328.6125 | 1228.6125 | 982.9115
4636 | 99.5468 | 0.96291 | 0.74315 | 0.98709 | 191.5639 | 47.7732 | 39815.0161 | 39368.819 | 783.3764 | 683.3764 | 982.9347
4637 | 97.189 | 1.0412 | 1.116 | 0.9988 | 197.6843 | 51.1782 | 62425.0911 | 62446.7654 | 1244.9353 | 1144.9353 | 983.0275
4638 | 85.713 | 1.1436 | 0.7714 | 1.0004 | 207.0888 | 48.5981 | 42473.2036 | 43009.9233 | 856.1985 | 756.1985 | 983.0356
4639 | 81.3479 | 0.98076 | 1.2715 | 1.0251 | 200.2544 | 46.6539 | 64731.7201 | 65885.1904 | 1313.7038 | 1213.7038 | 983.1197
4640 | 106.4723 | 1.053 | 1.0818 | 1.0031 | 194.0079 | 52.1181 | 62595.1406 | 62430.546 | 1244.6109 | 1144.6109 | 983.1547
4641 | 95.8967 | 0.95568 | 0.99812 | 0.97708 | 188.8349 | 55.6204 | 59529.8129 | 58491.9363 | 1165.8387 | 1065.8387 | 983.2622
4642 | 91.0916 | 0.94185 | 0.90414 | 1.0064 | 188.2119 | 48.7156 | 48694.7156 | 48484.5325 | 965.6906 | 865.6906 | 983.2928
4643 | 99.8284 | 1.0498 | 0.84351 | 0.98913 | 194.5385 | 53.0479 | 50013.7621 | 51597.0414 | 1027.9408 | 927.9408 | 983.3175
4644 | 84.0849 | 0.98822 | 0.9097 | 0.93178 | 201.3411 | 49.6345 | 46397.5806 | 47106.1823 | 938.1236 | 838.1236 | 983.3481
4645 | 96.3443 | 0.95779 | 0.92537 | 0.97889 | 192.4573 | 47.6277 | 47730.2048 | 47015.2388 | 936.3048 | 836.3048 | 983.374
4646 | 89.8275 | 1.0332 | 1.1706 | 1.062 | 198.8592 | 51.5358 | 69046.966 | 68890.9892 | 1373.8198 | 1273.8198 | 983.377
4647 | 113.6314 | 1.0153 | 1.2516 | 1.0274 | 212.7466 | 50.3764 | 70804.7378 | 70854.9058 | 1413.0981 | 1313.0981 | 983.3908
4648 | 97.4095 | 0.97527 | 0.87986 | 1.0036 | 198.9023 | 45.6368 | 44833.9543 | 44987.2641 | 895.7453 | 795.7453 | 983.4057
4649 | 96.2415 | 1.0074 | 1.0789 | 1.0298 | 189.9235 | 49.837 | 60394.5473 | 59520.449 | 1186.409 | 1086.409 | 983.4547
4650 | 101.5252 | 1.0163 | 1.0979 | 1.0508 | 213.4954 | 51.534 | 64980.9876 | 64472.4797 | 1285.4496 | 1185.4496 | 983.5213
4651 | 71.3365 | 1.0471 | 0.91579 | 1.0438 | 214.6116 | 52.1374 | 53948.4407 | 55394.8584 | 1103.8972 | 1003.8972 | 983.5413
4652 | 111.7782 | 1.0092 | 0.91688 | 1.0516 | 213.8433 | 48.3782 | 52315.061 | 50791.5898 | 1011.8318 | 911.8318 | 983.5541
4653 | 100.6912 | 1.0859 | 0.85678 | 0.99692 | 196.0397 | 52.5122 | 50790.2763 | 50635.2198 | 1008.7044 | 908.7044 | 983.5555
4654 | 98.3633 | 0.95978 | 1.0456 | 1.0058 | 203.8725 | 54.1465 | 62260.857 | 61816.2777 | 1232.3256 | 1132.3256 | 983.5828
4655 | 106.7813 | 1.0662 | 0.90433 | 1.0461 | 217.1932 | 45.8661 | 48827.5685 | 49502.4477 | 986.049 | 886.049 | 983.6215
4656 | 117.4446 | 1.0279 | 0.85746 | 1.0127 | 199.6546 | 50.0624 | 49716.3914 | 48972.4392 | 975.4488 | 875.4488 | 983.76
4657 | 90.4016 | 1.0001 | 0.98461 | 0.95411 | 194.0073 | 50.5427 | 52244.5067 | 52874.1876 | 1053.4838 | 953.4838 | 983.7698
4658 | 102.2004 | 0.98971 | 0.93954 | 0.87448 | 187.4279 | 51.3534 | 47574.003 | 47976.9487 | 955.539 | 855.539 | 983.7817
4659 | 98.0543 | 1.0482 | 1.2909 | 1.0088 | 193.5293 | 50.9986 | 71849.5909 | 73822.7327 | 1472.4547 | 1372.4547 | 983.7821
4660 | 97.0656 | 1.1018 | 1.1042 | 1.0058 | 221.5873 | 51.7715 | 63257.74 | 60582.9982 | 1207.66 | 1107.66 | 983.8686
4661 | 87.3986 | 0.95726 | 0.95357 | 1.0073 | 204.1421 | 54.5904 | 57207.2301 | 58261.8153 | 1161.2363 | 1061.2363 | 983.8795
4662 | 104.9143 | 0.98335 | 0.69692 | 1.1025 | 223.177 | 50.3481 | 44101.0275 | 45529.9607 | 906.5992 | 806.5992 | 983.8827
4663 | 110.1055 | 1.0909 | 1.0199 | 0.93929 | 190.7891 | 51.09 | 55268.607 | 55327.5332 | 1102.5507 | 1002.5507 | 983.9066
4664 | 114.7501 | 1.0373 | 0.73008 | 1.0329 | 185.2436 | 54.6563 | 47905.7531 | 47919.5477 | 954.391 | 854.391 | 983.9778
4665 | 97.989 | 0.91837 | 1.2029 | 0.99556 | 200.551 | 50.5137 | 65241.7441 | 65943.5513 | 1314.871 | 1214.871 | 983.9839
4666 | 112.5989 | 1.0141 | 0.89767 | 0.98933 | 192.3173 | 52.7819 | 53094.5094 | 51597.224 | 1027.9445 | 927.9445 | 984.0052
4667 | 102.7465 | 1.0044 | 1.1534 | 0.92505 | 215.9843 | 47.7243 | 56062.092 | 56771.3432 | 1131.4269 | 1031.4269 | 984.0307
4668 | 104.0539 | 1.0609 | 1.0076 | 0.99762 | 214.7381 | 53.8477 | 60285.1888 | 57150.1396 | 1139.0028 | 1039.0028 | 984.0769
4669 | 106.2873 | 0.97091 | 0.78625 | 0.95613 | 190.0783 | 47.7815 | 41040.8228 | 42187.696 | 839.7539 | 739.7539 | 984.159
4670 | 103.4745 | 0.95925 | 1.1737 | 0.93907 | 213.2262 | 53.7862 | 64834.7328 | 65340.9781 | 1302.8196 | 1202.8196 | 984.2258
4671 | 121.0537 | 0.89276 | 0.90894 | 0.97915 | 207.5749 | 53.1872 | 53291.4742 | 53097.9952 | 1057.9599 | 957.9599 | 984.232
4672 | 103.7867 | 1.0327 | 1.1906 | 1.0233 | 216.0495 | 48.9834 | 65148.2576 | 66381.3415 | 1323.6268 | 1223.6268 | 984.247
4673 | 99.1241 | 1.0561 | 1.0126 | 0.922 | 178.2696 | 53.2553 | 55473.9034 | 56910.2317 | 1134.2046 | 1034.2046 | 984.4291
4674 | 107.2089 | 1.0601 | 0.86339 | 0.99259 | 172.7991 | 51.0198 | 49694.2942 | 49051.0819 | 977.0216 | 877.0216 | 984.4634
4675 | 104.6706 | 0.98278 | 1.0247 | 1.0461 | 208.3985 | 51.7558 | 61013.5932 | 62070.8412 | 1237.4168 | 1137.4168 | 984.4787
4676 | 104.336 | 0.98335 | 1.0116 | 1.0214 | 190.532 | 47.256 | 53864.9555 | 54595.4224 | 1087.9084 | 987.9084 | 984.5643
4677 | 107.5002 | 1.1001 | 0.70989 | 1.0252 | 190.5658 | 54.4401 | 46248.2167 | 46174.4814 | 919.4896 | 819.4896 | 984.5862
4678 | 96.7236 | 1.0649 | 1.0825 | 0.88091 | 228.6344 | 49.5682 | 52599.9682 | 52904.6918 | 1054.0938 | 954.0938 | 984.6312
4679 | 115.1109 | 1.0141 | 1.263 | 0.97005 | 205.2926 | 48.7341 | 65601.703 | 66747.2931 | 1330.9459 | 1230.9459 | 984.6389
4680 | 103.0445 | 0.98909 | 0.87678 | 0.9705 | 200.3084 | 52.0518 | 49797.0864 | 50648.5207 | 1008.9704 | 908.9704 | 984.6456
4681 | 99.0208 | 0.93582 | 0.93736 | 1.0259 | 198.9274 | 53.4319 | 56530.3676 | 55483.5944 | 1105.6719 | 1005.6719 | 984.6961
4682 | 108.5981 | 1.1591 | 1.0415 | 1.0301 | 206.5293 | 49.6503 | 59723.2825 | 59649.4271 | 1188.9885 | 1088.9885 | 984.7219
4683 | 102.7745 | 0.92968 | 1.1169 | 1.0383 | 230.1022 | 49.3917 | 62228.3704 | 62684.1694 | 1249.6834 | 1149.6834 | 984.7396
4684 | 92.1394 | 0.91496 | 1.0463 | 0.97498 | 193.667 | 50.3562 | 55805.9487 | 54086.3074 | 1077.7261 | 977.7261 | 984.7968
4685 | 93.4466 | 1.0157 | 0.87518 | 1.0587 | 188.3397 | 48.9573 | 50197.008 | 52171.5698 | 1039.4314 | 939.4314 | 984.9759
4686 | 99.0016 | 1.0329 | 0.97293 | 1.0107 | 185.6628 | 49.5693 | 53996.8658 | 54377.7059 | 1083.5541 | 983.5541 | 985.0161
4687 | 93.12 | 1.0661 | 1.087 | 1.0185 | 188.4311 | 47.6813 | 57709.6454 | 58371.8381 | 1163.4368 | 1063.4368 | 985.0665
4688 | 108.2407 | 0.93144 | 1.1155 | 1.0074 | 202.4499 | 49.2668 | 60535.7542 | 62334.8685 | 1242.6974 | 1142.6974 | 985.0689
4689 | 83.8494 | 1.0156 | 0.93912 | 0.94955 | 200.7812 | 48.402 | 47484.3776 | 45845.822 | 912.9164 | 812.9164 | 985.0973
4690 | 108.278 | 1.0383 | 0.93406 | 1.0215 | 207.4285 | 54.4033 | 58233.127 | 57587.8703 | 1147.7574 | 1047.7574 | 985.0991
4691 | 99.5239 | 1.0523 | 1.0455 | 0.91719 | 205.1198 | 46.9166 | 50109.8476 | 49621.8813 | 988.4376 | 888.4376 | 985.1503
4692 | 110.3763 | 0.87292 | 1.1971 | 1.0402 | 210.6517 | 50.4679 | 67913.4836 | 69646.7579 | 1388.9352 | 1288.9352 | 985.1603
4693 | 108.4882 | 1.024 | 1.0766 | 0.97417 | 193.3505 | 52.8973 | 61546.9846 | 62696.2507 | 1249.925 | 1149.925 | 985.3416
4694 | 96.3974 | 1.0858 | 0.94891 | 0.98258 | 180.498 | 48.2848 | 50254.0657 | 50753.3871 | 1011.0677 | 911.0677 | 985.3527
4695 | 71.5637 | 1.0125 | 1.0666 | 1.0071 | 186.5916 | 50.03 | 57550.6302 | 57313.0651 | 1142.2613 | 1042.2613 | 985.4275
4696 | 104.6286 | 1.0319 | 0.84593 | 0.98179 | 183.3403 | 49.6687 | 46797.1985 | 47049.5173 | 936.9903 | 836.9903 | 985.4738
4697 | 102.1062 | 1.0514 | 1.021 | 1.043 | 205.8534 | 52.0453 | 61215.3703 | 62133.93 | 1238.6786 | 1138.6786 | 985.4906
4698 | 110.094 | 0.99648 | 1.1407 | 1.0683 | 195.6548 | 47.6016 | 63423.9717 | 66335.15 | 1322.703 | 1222.703 | 985.582
4699 | 88.1488 | 0.9329 | 1.2667 | 1.0829 | 203.3095 | 52.1563 | 76039.3513 | 74236.6042 | 1480.7321 | 1380.7321 | 985.588
4700 | 82.8198 | 0.89497 | 0.98124 | 1.0114 | 203.1348 | 50.0207 | 53551.6354 | 52894.9199 | 1053.8984 | 953.8984 | 985.7841
4701 | 105.9884 | 1.0407 | 0.93444 | 0.93795 | 197.6062 | 52.9963 | 52492.2028 | 52801.4911 | 1052.0298 | 952.0298 | 985.8116
4702 | 89.767 | 1.1342 | 1.1774 | 0.98059 | 198.5219 | 48.6323 | 61296.6162 | 62109.4865 | 1238.1897 | 1138.1897 | 985.8147
4703 | 97.3836 | 0.94795 | 0.8606 | 1.0457 | 200.7735 | 50.6658 | 50471.7202 | 50742.8466 | 1010.8569 | 910.8569 | 986.0557
4704 | 101.8011 | 1.0233 | 1.207 | 1.023 | 205.4311 | 48.986 | 65796.721 | 65592.295 | 1307.8459 | 1207.8459 | 986.1024
4705 | 103.7771 | 0.96432 | 1.1356 | 0.92725 | 198.1713 | 50.5304 | 58462.9082 | 58343.8298 | 1162.8766 | 1062.8766 | 986.1413
4706 | 97.8955 | 1.044 | 1.1857 | 1.0037 | 206.5864 | 48.3268 | 62660.569 | 61677.1998 | 1229.544 | 1129.544 | 986.1617
4707 | 91.5036 | 0.98671 | 1.0929 | 0.98039 | 220.6924 | 49.9944 | 58300.6998 | 59076.3497 | 1177.527 | 1077.527 | 986.2129
4708 | 104.5137 | 0.99712 | 1.1295 | 1.0174 | 218.4153 | 49.6633 | 62466.2735 | 63473.0793 | 1265.4616 | 1165.4616 | 986.2386
4709 | 98.9771 | 0.9431 | 1.0888 | 1.0375 | 202.1557 | 45.6409 | 56016.7027 | 58286.5662 | 1161.7313 | 1061.7313 | 986.3146
4710 | 90.3155 | 1.059 | 1.1692 | 1.0099 | 200.2358 | 44.5311 | 57043.0577 | 58219.8946 | 1160.3979 | 1060.3979 | 986.3942
4711 | 99.1837 | 1.0075 | 1.1011 | 1.0555 | 194.7769 | 49.4998 | 62668.7615 | 62234.9463 | 1240.6989 | 1140.6989 | 986.4524
4712 | 96.6981 | 0.9714 | 1.0106 | 0.98223 | 207.37 | 52.3718 | 57114.2041 | 58525.2864 | 1166.5057 | 1066.5057 | 986.4553
4713 | 84.9336 | 1.0335 | 1.1401 | 0.98161 | 194.5306 | 51.1239 | 61896.801 | 61745.3989 | 1230.908 | 1130.908 | 986.4581
4714 | 111.7245 | 1.0797 | 0.79484 | 0.99386 | 196.918 | 45.9707 | 42056.9706 | 41342.3144 | 822.8463 | 722.8463 | 986.4715
4715 | 83.7801 | 1.0285 | 1.1154 | 1.0338 | 178.349 | 51.9034 | 64502.8712 | 65629.7611 | 1308.5952 | 1208.5952 | 986.4998
4716 | 98.2038 | 0.95753 | 1.0747 | 0.97685 | 210.6391 | 51.1799 | 58753.3928 | 57517.2507 | 1146.345 | 1046.345 | 986.6117
4717 | 107.0868 | 0.98302 | 0.77811 | 1.0491 | 218.6193 | 49.5086 | 45844.7647 | 47339.5331 | 942.7907 | 842.7907 | 986.6725
4718 | 102.2628 | 0.9394 | 1.1294 | 1.1016 | 192.8247 | 48.714 | 65478.4242 | 66233.4629 | 1320.6693 | 1220.6693 | 986.7107
4719 | 104.57 | 1.0063 | 0.96796 | 1.0116 | 195.3734 | 49.442 | 53811.9967 | 53888.8906 | 1073.7778 | 973.7778 | 986.8174
4720 | 99.2205 | 0.95715 | 1.2304 | 0.92721 | 193.2116 | 50.9644 | 63177.8158 | 63074.9444 | 1257.4989 | 1157.4989 | 986.8487
4721 | 100.7709 | 1.0215 | 0.86659 | 1.1229 | 207.6482 | 52.4046 | 56595.4347 | 56354.8449 | 1123.0969 | 1023.0969 | 986.8879
4722 | 95.8716 | 0.99046 | 0.96526 | 1.0456 | 205.3402 | 48.5701 | 53836.1311 | 56295.1986 | 1121.904 | 1021.904 | 986.8888
4723 | 98.2467 | 0.95981 | 1.1097 | 1.0141 | 203.3373 | 49.4436 | 60502.53 | 61066.7285 | 1217.3346 | 1117.3346 | 986.8917
4724 | 97.564 | 0.90602 | 1.1858 | 0.9512 | 200.1546 | 50.8072 | 61999.1608 | 61510.732 | 1226.2146 | 1126.2146 | 986.8982
4725 | 92.4234 | 1.0218 | 0.98911 | 0.96572 | 226.3717 | 52.4904 | 55322.5142 | 55281.7449 | 1101.6349 | 1001.6349 | 986.971
4726 | 84.8815 | 0.95831 | 1.0608 | 0.98532 | 195.8831 | 50.3932 | 56965.4145 | 58773.2926 | 1171.4659 | 1071.4659 | 987.0343
4727 | 110.2356 | 0.94393 | 0.91563 | 1.0335 | 194.16 | 45.2066 | 47676.4479 | 48046.4162 | 956.9283 | 856.9283 | 987.0902
4728 | 98.9753 | 0.97109 | 1.1972 | 0.97263 | 213.5755 | 49.606 | 62742.5152 | 61036.7159 | 1216.7343 | 1116.7343 | 987.1114
4729 | 102.0246 | 0.94402 | 0.88847 | 0.9564 | 205.6806 | 48.8561 | 46425.9908 | 48213.1209 | 960.2624 | 860.2624 | 987.1293
4730 | 97.7451 | 1.0216 | 0.90882 | 1.0767 | 181.6986 | 54.4787 | 58931.6764 | 56901.8814 | 1134.0376 | 1034.0376 | 987.2021
4731 | 105.4125 | 0.99037 | 0.99874 | 1.0327 | 192.6634 | 54.5625 | 62163.7458 | 60845.6195 | 1212.9124 | 1112.9124 | 987.2886
4732 | 92.4756 | 1.0531 | 0.98172 | 1.0067 | 189.2168 | 51.1174 | 55685 | 55953.2432 | 1115.0649 | 1015.0649 | 987.2966
4733 | 109.6161 | 0.92315 | 0.92523 | 1.1642 | 199.9818 | 45.9336 | 54326.1191 | 55445.3255 | 1104.9065 | 1004.9065 | 987.3201
4734 | 86.1251 | 1.0303 | 1.0118 | 0.94041 | 193.6421 | 52.3682 | 54670.1579 | 55121.9456 | 1098.4389 | 998.4389 | 987.3217
4735 | 107.166 | 1.0239 | 0.91395 | 0.98891 | 208.8213 | 51.7247 | 52633.8713 | 56499.9019 | 1125.998 | 1025.998 | 987.3656
4736 | 98.4419 | 0.97352 | 0.91204 | 0.92373 | 187.3989 | 51.4014 | 48418.1282 | 47146.3155 | 938.9263 | 838.9263 | 987.406
4737 | 104.8693 | 1.0325 | 1.027 | 1.069 | 205.3267 | 47.7028 | 57739.8175 | 57038.7458 | 1136.7749 | 1036.7749 | 987.4068
4738 | 108.8164 | 0.97014 | 0.75667 | 0.97493 | 199.1852 | 48.0695 | 40734.5855 | 42069.4708 | 837.3894 | 737.3894 | 987.448
4739 | 100.6273 | 1.0326 | 1.0349 | 0.95345 | 203.164 | 52.6623 | 57636.3597 | 57444.2678 | 1144.8854 | 1044.8854 | 987.4551
4740 | 95.0098 | 0.99483 | 0.90054 | 0.93667 | 198.4091 | 48.8552 | 46025.7886 | 45318.0715 | 902.3614 | 802.3614 | 987.5082
4741 | 110.9513 | 1.0002 | 0.99374 | 1.0327 | 209.5455 | 49.1668 | 56120.7108 | 55336.1702 | 1102.7234 | 1002.7234 | 987.5686
4742 | 108.4162 | 0.97645 | 1.0652 | 1.0303 | 202.0314 | 51.9219 | 62682.8554 | 61830.5737 | 1232.6115 | 1132.6115 | 987.6147
4743 | 98.9759 | 0.96817 | 1.0907 | 1.0015 | 189.3115 | 48.3219 | 57601.6083 | 58852.2144 | 1173.0443 | 1073.0443 | 987.617
4744 | 110.1783 | 0.95003 | 1.1016 | 1.005 | 208.0764 | 49.5072 | 60204.0444 | 60464.6854 | 1205.2937 | 1105.2937 | 987.6511
4745 | 112.9822 | 1.0141 | 0.90249 | 0.97789 | 212.3407 | 50.5614 | 50627.8814 | 50355.0412 | 1003.1008 | 903.1008 | 987.6558
4746 | 95.8856 | 1.0503 | 0.8553 | 1.0289 | 196.1539 | 49.1482 | 48395.5034 | 48397.5401 | 963.9508 | 863.9508 | 987.656
4747 | 96.8311 | 0.97327 | 0.95286 | 0.94641 | 205.6509 | 49.3906 | 49400.2937 | 52038.1694 | 1036.7634 | 936.7634 | 987.6593
4748 | 93.511 | 1.0418 | 0.98235 | 0.89225 | 196.0031 | 50.3435 | 49226.8935 | 47260.5752 | 941.2115 | 841.2115 | 987.6604
4749 | 88.7987 | 0.99893 | 0.92214 | 1.025 | 201.3308 | 55.1722 | 57243.5948 | 57032.906 | 1136.6581 | 1036.6581 | 987.7428
4750 | 102.6144 | 1.0288 | 0.98965 | 1.0145 | 211.8231 | 50.8511 | 56634.8868 | 56261.3807 | 1121.2276 | 1021.2276 | 987.7477
4751 | 82.9283 | 0.95237 | 1.0964 | 0.96424 | 204.7614 | 55.8657 | 63678.9815 | 65457.5391 | 1305.1508 | 1205.1508 | 987.9028
4752 | 89.0142 | 0.91351 | 0.85063 | 1.0777 | 208.1175 | 51.17 | 51276.6635 | 52560.1948 | 1047.2039 | 947.2039 | 987.9084
4753 | 94.6972 | 1.0006 | 0.86994 | 1.0318 | 191.85 | 46.2298 | 46067.2138 | 47157.5436 | 939.1509 | 839.1509 | 987.9447
4754 | 113.9648 | 1.0249 | 0.88876 | 0.9996 | 202.6003 | 50.2447 | 50709.0814 | 50543.2606 | 1006.8652 | 906.8652 | 987.9633
4755 | 91.5762 | 0.99825 | 1.137 | 0.92219 | 168.5812 | 50.8795 | 58169.1689 | 57013.0705 | 1136.2614 | 1036.2614 | 987.9703
4756 | 99.8108 | 1.0049 | 0.92948 | 0.92456 | 183.8573 | 47.4862 | 45754.843 | 44851.2215 | 893.0244 | 793.0244 | 988.0733
4757 | 92.078 | 0.9911 | 1.0411 | 0.93946 | 209.8755 | 53.4792 | 57394.9119 | 57266.5496 | 1141.331 | 1041.331 | 988.0954
4758 | 97.0053 | 1.1013 | 0.96089 | 1.0123 | 197.0301 | 49.7421 | 53894.1636 | 55053.7477 | 1097.075 | 997.075 | 988.1061
4759 | 105.0015 | 0.98918 | 0.84828 | 0.9382 | 203.948 | 47.6557 | 43080.7034 | 43767.3815 | 871.3476 | 771.3476 | 988.2213
4760 | 104.3688 | 0.97075 | 0.82814 | 0.98729 | 192.009 | 51.0621 | 47114.7096 | 46256.03 | 921.1206 | 821.1206 | 988.2411
4761 | 108.5111 | 1.0232 | 0.87069 | 1.0585 | 202.5912 | 51.5629 | 53449.4288 | 52661.7484 | 1049.235 | 949.235 | 988.2586
4762 | 90.0322 | 0.99512 | 0.9657 | 0.92859 | 191.7425 | 48.6568 | 48183.3783 | 49189.056 | 979.7811 | 879.7811 | 988.2877
4763 | 108.5391 | 1.0029 | 1.0793 | 0.96006 | 190.2525 | 47.3387 | 54394.4211 | 53396.2094 | 1063.9242 | 963.9242 | 988.4079
4764 | 80.3902 | 1.0359 | 1.1324 | 1.047 | 207.8363 | 46.4274 | 59120.5921 | 59054.7145 | 1177.0943 | 1077.0943 | 988.4103
4765 | 113.6755 | 0.95683 | 0.75504 | 1.0274 | 190.9499 | 49.8924 | 44319.6123 | 43786.9883 | 871.7398 | 771.7398 | 988.4598
4766 | 109.4845 | 1.0719 | 0.95052 | 0.9757 | 207.4531 | 54.2935 | 56931.5963 | 56102.7619 | 1118.0552 | 1018.0552 | 988.4907
4767 | 91.0252 | 1.0262 | 0.9263 | 0.95553 | 229.6555 | 49.5452 | 48710.6028 | 49696.9503 | 989.939 | 889.939 | 988.5203
4768 | 96.4071 | 0.90628 | 1.1356 | 1.0152 | 216.7551 | 53.5596 | 66643.3099 | 65453.5917 | 1305.0718 | 1205.0718 | 988.6195
4769 | 109.3446 | 0.99605 | 0.90192 | 1.0011 | 190.0702 | 49.1217 | 49894.7419 | 47166.4387 | 939.3288 | 839.3288 | 988.6556
4770 | 106.2569 | 1.0394 | 0.88683 | 0.96994 | 186.6208 | 51.0897 | 49774.7659 | 50433.8053 | 1004.6761 | 904.6761 | 988.6801
4771 | 99.1508 | 1.0941 | 0.97701 | 0.94691 | 200.9062 | 44.1467 | 45831.4215 | 46071.969 | 917.4394 | 817.4394 | 988.6953
4772 | 88.3947 | 1.0416 | 0.84514 | 0.96611 | 202.8641 | 49.9547 | 45590.2085 | 46628.8811 | 928.5776 | 828.5776 | 988.7866
4773 | 113.1245 | 0.96587 | 1.2634 | 0.98368 | 206.898 | 48.9463 | 66383.198 | 67480.0753 | 1345.6015 | 1245.6015 | 988.8076
4774 | 93.5529 | 0.99011 | 1.1102 | 1.0244 | 207.4886 | 53.9081 | 66515.0406 | 64300.8529 | 1282.0171 | 1182.0171 | 988.8913
4775 | 113.4601 | 0.9536 | 1.0634 | 0.99113 | 201.4987 | 48.1216 | 56127.6456 | 56175.9343 | 1119.5187 | 1019.5187 | 988.9021
4776 | 114.3704 | 1.0852 | 0.97668 | 1.1223 | 194.7414 | 46.0435 | 56379.999 | 54393.9737 | 1083.8795 | 983.8795 | 988.9451
4777 | 96.4568 | 0.99153 | 0.95553 | 1.0236 | 208.0985 | 52.797 | 56895.4211 | 55911.5243 | 1114.2305 | 1014.2305 | 988.9457
4778 | 100.9952 | 0.92602 | 0.91414 | 1.0564 | 194.0802 | 50.106 | 53267.1741 | 51809.824 | 1032.1965 | 932.1965 | 988.9479
4779 | 98.5663 | 1.0038 | 1.1989 | 0.89104 | 189.0429 | 55.1213 | 64529.0242 | 61839.5189 | 1232.7904 | 1132.7904 | 988.9705
4780 | 113.3404 | 0.96921 | 1.2501 | 0.94297 | 193.9421 | 49.5839 | 64092.7834 | 63355.5884 | 1263.1118 | 1163.1118 | 988.9875
4781 | 101.861 | 1.0115 | 0.97445 | 0.92149 | 208.7701 | 46.6973 | 46951.5197 | 46525.0762 | 926.5015 | 826.5015 | 989.0642
4782 | 92.5171 | 1.0555 | 0.84531 | 1.0094 | 201.9055 | 47.2696 | 45151.1404 | 45236.2806 | 900.7256 | 800.7256 | 989.1068
4783 | 82.628 | 1.0723 | 1.0845 | 0.97702 | 224.9304 | 50.2999 | 57976.3024 | 56377.7236 | 1123.5545 | 1023.5545 | 989.1095
4784 | 110.7139 | 0.93108 | 0.97933 | 1.0893 | 193.1033 | 50.4828 | 59249.9601 | 60795.7333 | 1211.9147 | 1111.9147 | 989.1121
4785 | 103.004 | 0.99522 | 1.1873 | 0.98294 | 192.8407 | 47.9705 | 61092.1383 | 61106.6759 | 1218.1335 | 1118.1335 | 989.4712
4786 | 99.131 | 1.0425 | 0.71531 | 1.0532 | 208.5926 | 47.6151 | 40999.8179 | 41410.9132 | 824.2183 | 724.2183 | 989.5811
4787 | 104.7862 | 1.0113 | 0.81242 | 1.0124 | 211.4028 | 54.4096 | 50728.8219 | 50546.3854 | 1006.9277 | 906.9277 | 989.6466
4788 | 102.2113 | 1.0314 | 0.9086 | 1.0276 | 193.7438 | 47.5047 | 49556.2592 | 49686.0958 | 989.7219 | 889.7219 | 989.6852
4789 | 97.8986 | 0.96852 | 1.0173 | 1.0262 | 193.651 | 47.2049 | 53946.5873 | 54984.8746 | 1095.6975 | 995.6975 | 989.7271
4790 | 107.134 | 1.0055 | 0.97924 | 0.99506 | 199.5294 | 51.4396 | 55863.8547 | 56191.8108 | 1119.8362 | 1019.8362 | 989.8025
4791 | 119.5019 | 0.99903 | 0.81557 | 0.94622 | 206.694 | 51.7128 | 46287.6991 | 47013.2473 | 936.2649 | 836.2649 | 989.8111
4792 | 96.3372 | 0.96451 | 0.70114 | 0.89873 | 217.4983 | 53.1389 | 38639.657 | 40359.5913 | 803.1918 | 703.1918 | 989.8189
4793 | 102.3434 | 1.0512 | 1.1741 | 0.95929 | 205.3003 | 51.1076 | 63265.3946 | 63222.7558 | 1260.4551 | 1160.4551 | 989.8732
4794 | 112.8936 | 1.033 | 1.2458 | 1.0194 | 204.0238 | 50.6079 | 70373.2944 | 72703.0475 | 1450.0609 | 1350.0609 | 989.907
4795 | 81.2824 | 0.9691 | 1.0908 | 1.0154 | 187.1619 | 55.7433 | 66322.5412 | 66633.201 | 1328.664 | 1228.664 | 990.1209
4796 | 87.3084 | 0.95212 | 1.2264 | 0.96147 | 193.8836 | 49.4828 | 62654.295 | 61788.0468 | 1231.7609 | 1131.7609 | 990.198
4797 | 98.3195 | 0.96783 | 1.0746 | 0.99323 | 203.3642 | 48.6298 | 56734.9915 | 57142.4624 | 1138.8492 | 1038.8492 | 990.232
4798 | 114.2564 | 1.0698 | 0.98928 | 1.0554 | 203.8663 | 50.5024 | 59106.3787 | 59169.6064 | 1179.3921 | 1079.3921 | 990.2565
4799 | 89.5587 | 1.0015 | 0.90979 | 0.95954 | 173.098 | 53.8552 | 52017.8816 | 52264.5157 | 1041.2903 | 941.2903 | 990.3149
4800 | 84.4293 | 1.0184 | 0.82739 | 1.0474 | 214.4923 | 47.8661 | 45811.6087 | 46200.2732 | 920.0055 | 820.0055 | 990.3256
4801 | 111.1439 | 0.92202 | 1.2208 | 1.0006 | 194.5186 | 48.7697 | 64765.2225 | 64414.2339 | 1284.2847 | 1184.2847 | 990.3351
4802 | 93.1059 | 1.0305 | 1.1475 | 0.91455 | 200.9059 | 46.2046 | 53122.9289 | 52410.7087 | 1044.2142 | 944.2142 | 990.3377
4803 | 103.5508 | 0.93783 | 0.75054 | 1.0399 | 191.8375 | 51.8673 | 45711.913 | 44482.7107 | 885.6542 | 785.6542 | 990.3512
4804 | 84.2353 | 1.0092 | 0.98997 | 0.99771 | 186.8839 | 50.7284 | 54603.8698 | 54771.064 | 1091.4213 | 991.4213 | 990.364
4805 | 112.7603 | 0.93183 | 1.0216 | 0.94624 | 210.3208 | 49.1625 | 52898.481 | 53587.6241 | 1067.7525 | 967.7525 | 990.3712
4806 | 108.7094 | 1.0696 | 0.98841 | 1.0036 | 199.5058 | 50.3846 | 56037.3975 | 53791.7288 | 1071.8346 | 971.8346 | 990.4326
4807 | 101.5329 | 0.96965 | 0.80735 | 1.0239 | 209.8499 | 49.8871 | 46358.6808 | 47175.8261 | 939.5165 | 839.5165 | 990.5204
4808 | 92.0315 | 1.0642 | 1.1 | 0.93637 | 184.9342 | 50.7467 | 57422.5664 | 57648.5826 | 1148.9717 | 1048.9717 | 990.569
4809 | 105.9209 | 1.0083 | 0.76974 | 1.0258 | 188.5639 | 53.4544 | 48102.9693 | 46592.2719 | 927.8454 | 827.8454 | 990.5868
4810 | 96.2127 | 1.0325 | 0.73463 | 0.99811 | 192.1145 | 51.0981 | 42735.4163 | 42244.9779 | 840.8996 | 740.8996 | 990.6062
4811 | 99.4504 | 0.90202 | 1.3307 | 0.90914 | 210.5044 | 48.0855 | 62696.7256 | 64050.2471 | 1277.0049 | 1177.0049 | 990.6323
4812 | 108.9762 | 1.0065 | 0.94964 | 1.0275 | 199.7597 | 52.7424 | 57448.3698 | 57162.5233 | 1139.2505 | 1039.2505 | 990.6426
4813 | 100.8533 | 1.0903 | 0.84033 | 1.0322 | 219.167 | 50.9699 | 50035.8542 | 49464.1826 | 985.2837 | 885.2837 | 990.6475
4814 | 99.5486 | 0.98856 | 0.87459 | 0.9285 | 194.5191 | 52.9234 | 48379.5386 | 46826.805 | 932.5361 | 832.5361 | 990.6938
4815 | 89.739 | 0.95257 | 1.1486 | 0.97398 | 205.9163 | 53.6581 | 64822.7635 | 65812.5599 | 1312.2512 | 1212.2512 | 990.7337
4816 | 77.1608 | 0.96454 | 0.96883 | 0.92744 | 221.0961 | 48.5403 | 47448.7842 | 48710.6976 | 970.214 | 870.214 | 990.7612
4817 | 101.7613 | 1.0175 | 0.97496 | 0.99166 | 200.0191 | 49.1342 | 52791.4545 | 51153.984 | 1019.0797 | 919.0797 | 990.7748
4818 | 99.8558 | 0.96233 | 1.2391 | 1.0273 | 203.2003 | 46.6707 | 64098.2321 | 64547.9102 | 1286.9582 | 1186.9582 | 990.7888
4819 | 104.1195 | 0.92556 | 1.0237 | 0.92833 | 199.0182 | 47.6575 | 50080.2542 | 50334.0502 | 1002.681 | 902.681 | 990.8692
4820 | 78.4289 | 1.0386 | 1.3218 | 1.0552 | 195.4776 | 47.3659 | 70116.4144 | 68174.2305 | 1359.4846 | 1259.4846 | 990.8727
4821 | 88.3235 | 1.0152 | 1.2345 | 0.98475 | 206.3847 | 48.7692 | 63864.9835 | 64915.5398 | 1294.3108 | 1194.3108 | 990.8885
4822 | 95.0427 | 1.0031 | 1.1487 | 0.97403 | 198.0939 | 50.6193 | 61659.2912 | 60928.0305 | 1214.5606 | 1114.5606 | 990.8974
4823 | 111.6008 | 0.99555 | 1.1707 | 1.0438 | 207.7139 | 49.191 | 65779.7191 | 63111.1534 | 1258.2231 | 1158.2231 | 990.9797
4824 | 98.1397 | 1.013 | 0.77007 | 0.98343 | 188.5243 | 50.5741 | 43516.7311 | 43167.0736 | 859.3415 | 759.3415 | 991.0214
4825 | 81.2132 | 0.98652 | 0.91337 | 0.9776 | 190.9546 | 52.9236 | 51687.3534 | 51803.4936 | 1032.0699 | 932.0699 | 991.0416
4826 | 99.857 | 0.99364 | 0.95438 | 0.92998 | 210.8845 | 48.7869 | 48352.5736 | 48188.3844 | 959.7677 | 859.7677 | 991.0556
4827 | 110.4968 | 0.99368 | 0.94153 | 1.0713 | 204.0182 | 50.1076 | 56248.4964 | 56421.8895 | 1124.4378 | 1024.4378 | 991.1458
4828 | 101.0607 | 0.99345 | 1.0474 | 0.94222 | 214.2553 | 52.7208 | 57536.9105 | 58515.8433 | 1166.3169 | 1066.3169 | 991.2513
4829 | 96.0033 | 0.99198 | 1.2709 | 1.0024 | 222.8819 | 46.3793 | 63720.3455 | 63048.7416 | 1256.9748 | 1156.9748 | 991.2855
4830 | 125.6034 | 1.023 | 0.84872 | 1.006 | 202.8253 | 47.8557 | 47212.5497 | 46836.5979 | 932.732 | 832.732 | 991.314
4831 | 100.2885 | 0.87908 | 1.1329 | 1.0347 | 200.8186 | 44.4196 | 56187.889 | 55022.4435 | 1096.4489 | 996.4489 | 991.3196
4832 | 95.8397 | 1.0582 | 1.1776 | 1.077 | 208.9757 | 45.677 | 62772.7849 | 62879.3756 | 1253.5875 | 1153.5875 | 991.3406
4833 | 106.4661 | 0.94175 | 1.0065 | 1.0024 | 203.3089 | 48.525 | 54024.0717 | 54448.7937 | 1084.9759 | 984.9759 | 991.3469
4834 | 80.25 | 1.0887 | 0.9906 | 1.0071 | 215.2205 | 49.948 | 54410.7686 | 55275.6272 | 1101.5125 | 1001.5125 | 991.3479
4835 | 84.9621 | 0.95705 | 1.1857 | 0.92635 | 184.8279 | 49.7314 | 58851.6828 | 57215.2177 | 1140.3044 | 1040.3044 | 991.3548
4836 | 105.4716 | 1.0128 | 0.88845 | 0.99556 | 210.8565 | 47.0607 | 46863.5073 | 47206.7551 | 940.1351 | 840.1351 | 991.4213
4837 | 92.5916 | 0.94455 | 0.97601 | 0.91702 | 210.5869 | 51.0572 | 50373.4429 | 50631.0592 | 1008.6212 | 908.6212 | 991.4506
4838 | 104.3276 | 1.043 | 0.85295 | 1.0079 | 206.3438 | 52.8361 | 51379.4815 | 52600.4638 | 1048.0093 | 948.0093 | 991.5283
4839 | 91.3437 | 1.0214 | 0.92442 | 0.98171 | 197.4673 | 44.9267 | 45160.39 | 46244.8178 | 920.8964 | 820.8964 | 991.68
4840 | 93.7902 | 1.0404 | 1.2278 | 1.0073 | 191.154 | 47.1473 | 63102.201 | 62859.4776 | 1253.1896 | 1153.1896 | 991.7284
4841 | 102.3403 | 0.99872 | 1.1147 | 1.0518 | 205.5744 | 50.508 | 64588.5741 | 66384.1004 | 1323.682 | 1223.682 | 991.7648
4842 | 103.5549 | 1.0002 | 1.2706 | 1.001 | 191.9661 | 52.2666 | 72081.1979 | 72477.7013 | 1445.554 | 1345.554 | 991.7697
4843 | 120.634 | 0.99115 | 0.76726 | 1.0406 | 192.735 | 47.6184 | 43905.6562 | 43913.7037 | 874.2741 | 774.2741 | 991.7996
4844 | 115.1835 | 1.0342 | 1.1047 | 1.0925 | 186.1576 | 53.2696 | 70823.9564 | 72128.0844 | 1438.5617 | 1338.5617 | 991.9671
4845 | 103.6055 | 1.0321 | 1.0234 | 1.0608 | 194.1823 | 51.5468 | 61667.2778 | 63343.4202 | 1262.8684 | 1162.8684 | 992.008
4846 | 103.312 | 1.054 | 1.0472 | 0.97911 | 209.4201 | 54.8515 | 62425.56 | 62820.5143 | 1252.4103 | 1152.4103 | 992.068
4847 | 98.2838 | 1.0756 | 0.80092 | 0.94483 | 195.4364 | 47.9568 | 41555.6362 | 41420.131 | 824.4026 | 724.4026 | 992.1039
4848 | 98.2463 | 0.97431 | 0.95874 | 1.077 | 184.2771 | 49.1986 | 55695.1201 | 56762.073 | 1131.2415 | 1031.2415 | 992.161
4849 | 101.3106 | 1.0208 | 0.98799 | 1.0283 | 184.1996 | 56.1286 | 63011.5062 | 64080.632 | 1277.6126 | 1177.6126 | 992.1899
4850 | 106.1344 | 0.9773 | 1.0135 | 0.95803 | 193.9238 | 49.7426 | 53652.3988 | 53102.6275 | 1058.0525 | 958.0525 | 992.1957
4851 | 99.5044 | 0.94849 | 1.0633 | 1.0747 | 215.7496 | 53.691 | 66635.474 | 66573.3918 | 1327.4678 | 1227.4678 | 992.1988
4852 | 115.2457 | 0.97174 | 1.3338 | 1.0369 | 196.072 | 48.2766 | 72370.4087 | 69395.6215 | 1383.9124 | 1283.9124 | 992.2752
4853 | 111.7061 | 1.0512 | 1.0159 | 1.0053 | 191.1689 | 51.8508 | 59236.3506 | 60374.8647 | 1203.4973 | 1103.4973 | 992.3262
4854 | 86.7074 | 0.99546 | 1.0759 | 0.98463 | 220.4759 | 46.199 | 53149.5534 | 52842.2416 | 1052.8448 | 952.8448 | 992.6142
4855 | 85.2329 | 1.0027 | 1.3579 | 1.0057 | 210.2065 | 48.6122 | 70748.9261 | 70432.2726 | 1404.6455 | 1304.6455 | 992.6501
4856 | 85.6928 | 1.0473 | 1.0271 | 0.98501 | 201.5827 | 50.8463 | 56205.9019 | 56151.4403 | 1119.0288 | 1019.0288 | 992.723
4857 | 100.3917 | 1.0386 | 0.92583 | 1.0054 | 196.623 | 53.1867 | 55251.2437 | 56074.7455 | 1117.4949 | 1017.4949 | 992.8913
4858 | 103.7841 | 1.0242 | 0.9065 | 0.98684 | 196.701 | 50.2676 | 50507.7984 | 52276.13 | 1041.5226 | 941.5226 | 992.9163
4859 | 102.5816 | 0.98501 | 1.1171 | 1.0384 | 201.2226 | 50.2185 | 63527.3672 | 64854.5796 | 1293.0916 | 1193.0916 | 992.9222
4860 | 103.4242 | 0.93976 | 0.89787 | 1.0703 | 214.1098 | 53.0282 | 56327.0801 | 56270.2828 | 1121.4057 | 1021.4057 | 992.9307
4861 | 104.6002 | 0.98616 | 0.75786 | 0.9704 | 213.9511 | 49.4356 | 41669.9 | 42403.5869 | 844.0717 | 744.0717 | 992.9325
4862 | 90.4111 | 0.95011 | 0.93161 | 1.0621 | 186.5918 | 51.1807 | 55222.5282 | 55878.1774 | 1113.5635 | 1013.5635 | 992.9545
4863 | 107.8795 | 0.98246 | 1.1327 | 0.96211 | 194.6025 | 49.3033 | 59151.3898 | 58837.7538 | 1172.7551 | 1072.7551 | 993.0198
4864 | 99.244 | 0.98691 | 1.076 | 1.0292 | 182.1336 | 52.8484 | 63885.8905 | 66049.4469 | 1316.9889 | 1216.9889 | 993.0332
4865 | 104.4437 | 0.95688 | 0.93841 | 0.99923 | 205.4687 | 49.5785 | 51649.6679 | 52440.873 | 1044.8175 | 944.8175 | 993.1173
4866 | 105.6134 | 1.0499 | 0.78047 | 0.98555 | 208.46 | 48.6989 | 43067.3503 | 42379.1032 | 843.5821 | 743.5821 | 993.2051
4867 | 127.3661 | 1.0567 | 1.0303 | 0.94942 | 221.7042 | 50.22 | 56104.1061 | 55830.269 | 1112.6054 | 1012.6054 | 993.2177
4868 | 89.837 | 0.96542 | 0.72891 | 1.0846 | 200.4879 | 50.5161 | 44518.2208 | 43696.2742 | 869.9255 | 769.9255 | 993.2591
4869 | 113.6424 | 0.97696 | 1.0658 | 0.97593 | 201.2588 | 46.1579 | 53337.4598 | 53647.0396 | 1068.9408 | 968.9408 | 993.3023
4870 | 91.7557 | 1.0675 | 0.97032 | 1.041 | 206.1802 | 54.348 | 60425.8322 | 59378.6618 | 1183.5732 | 1083.5732 | 993.3445
4871 | 95.8596 | 0.97527 | 1.0071 | 1.0059 | 204.7774 | 48.5559 | 53936.1405 | 54063.2247 | 1077.2645 | 977.2645 | 993.3927
4872 | 98.5353 | 0.96168 | 1.0969 | 1.0113 | 212.8787 | 44.6485 | 53973.8233 | 53891.3498 | 1073.827 | 973.827 | 993.4189
4873 | 100.7701 | 1.0501 | 1.0377 | 0.97845 | 193.2033 | 48.9075 | 55025.3808 | 53714.9961 | 1070.2999 | 970.2999 | 993.4589
4874 | 97.6779 | 0.93015 | 0.66918 | 1.0365 | 190.6389 | 51.0703 | 40252.1846 | 39923.9908 | 794.4798 | 694.4798 | 993.4944
4875 | 99.2556 | 1.0477 | 1.2306 | 0.91117 | 207.3855 | 53.0355 | 65189.9046 | 65803.2582 | 1312.0652 | 1212.0652 | 993.5076
4876 | 113.3714 | 0.99561 | 1.0138 | 1.0234 | 203.9535 | 49.8858 | 57589.6904 | 56159.7546 | 1119.1951 | 1019.1951 | 993.5093
4877 | 104.8786 | 0.98927 | 1.1886 | 1.0716 | 192.6417 | 51.6783 | 71377.407 | 72660.4997 | 1449.21 | 1349.21 | 993.5355
4878 | 93.47 | 0.97817 | 0.94682 | 0.98739 | 211.6036 | 50.344 | 51880.1525 | 51829.4531 | 1032.5891 | 932.5891 | 993.6085
4879 | 74.7787 | 1.074 | 0.82463 | 0.9467 | 209.944 | 49.6526 | 42960.2225 | 43177.2508 | 859.545 | 759.545 | 993.6369
4880 | 110.4962 | 1.0671 | 1.101 | 1.0968 | 183.4681 | 49.606 | 65935.7187 | 64926.4293 | 1294.5286 | 1194.5286 | 993.6438
4881 | 106.4715 | 1.0289 | 1.2439 | 0.96179 | 210.9802 | 47.6799 | 62478.0702 | 63131.3033 | 1258.6261 | 1158.6261 | 993.6881
4882 | 90.1412 | 0.96913 | 0.87457 | 1.0603 | 199.6068 | 47.3801 | 48274.3309 | 48884.822 | 973.6964 | 873.6964 | 993.7063
4883 | 101.9859 | 0.95088 | 1.0574 | 0.98446 | 215.1441 | 51.5909 | 58923.3146 | 58406.3201 | 1164.1264 | 1064.1264 | 993.8409
4884 | 95.4851 | 0.88112 | 0.68692 | 0.98956 | 195.983 | 49.1596 | 37748.0897 | 37069.4993 | 737.39 | 637.39 | 993.9326
4885 | 110.5568 | 1.002 | 1.0797 | 0.94187 | 193.2702 | 53.464 | 60486.9799 | 59365.372 | 1183.3074 | 1083.3074 | 993.9511
4886 | 99.6097 | 0.98872 | 1.1789 | 0.97513 | 207.6212 | 51.1029 | 63985.4109 | 64361.5978 | 1283.232 | 1183.232 | 993.9676
4887 | 101.1787 | 1.0006 | 1.0388 | 0.98257 | 174.0943 | 51.6431 | 58112.1789 | 57270.8051 | 1141.4161 | 1041.4161 | 993.9745
4888 | 105.5754 | 1.0297 | 0.8369 | 1.0504 | 204.4164 | 53.9275 | 53473.3364 | 52700.0775 | 1050.0015 | 950.0015 | 993.9749
4889 | 100.2189 | 1.077 | 0.99929 | 0.94382 | 207.332 | 46.522 | 49106.1547 | 49998.7681 | 995.9754 | 895.9754 | 994.1041
4890 | 78.8793 | 0.9892 | 1.12 | 0.98131 | 191.3446 | 47.0785 | 55604.9254 | 53176.7783 | 1059.5356 | 959.5356 | 994.1084
4891 | 101.4349 | 1.064 | 0.91484 | 1.0531 | 206.8517 | 46.9336 | 50487.9643 | 50775.7669 | 1011.5153 | 911.5153 | 994.1315
4892 | 94.9176 | 0.95998 | 1.1396 | 0.98702 | 217.1517 | 51.6808 | 63056.2633 | 64237.1018 | 1280.742 | 1180.742 | 994.1622
4893 | 94.4698 | 0.96032 | 0.88013 | 0.98272 | 193.3177 | 50.3223 | 48283.3849 | 49172.5352 | 979.4507 | 879.4507 | 994.164
4894 | 92.8783 | 0.96753 | 1.0657 | 0.95839 | 193.0503 | 47.773 | 53279.4866 | 53264.4331 | 1061.2887 | 961.2887 | 994.2442
4895 | 97.2418 | 0.9983 | 1.1969 | 1.0182 | 198.4804 | 50.8197 | 67063.8066 | 66793.4674 | 1331.8693 | 1231.8693 | 994.3347
4896 | 112.5143 | 1.0584 | 0.90402 | 1.0388 | 208.9571 | 50.1898 | 53316.4522 | 53098.7755 | 1057.9755 | 957.9755 | 994.4425
4897 | 83.7089 | 0.96873 | 1.0075 | 1.0578 | 199.2764 | 54.9584 | 63231.7838 | 64589.9366 | 1287.7987 | 1187.7987 | 994.4732
4898 | 101.8541 | 0.98399 | 1.1315 | 1.1064 | 210.1099 | 50.6393 | 68680.064 | 65839.7404 | 1312.7948 | 1212.7948 | 994.6381
4899 | 115.2378 | 0.95773 | 0.9799 | 0.97631 | 198.7071 | 46.365 | 49672.9109 | 49290.668 | 981.8134 | 881.8134 | 994.75
4900 | 122.7976 | 0.93885 | 0.91093 | 0.95396 | 196.1045 | 46.6758 | 46138.4203 | 46180.9242 | 919.6185 | 819.6185 | 994.7549
4901 | 102.1479 | 0.90384 | 1.0468 | 1.0493 | 213.847 | 53.2728 | 63647.5851 | 62994.3828 | 1255.8877 | 1155.8877 | 994.7594
4902 | 108.0707 | 0.97707 | 0.95411 | 1.0154 | 201.0204 | 46.7459 | 50426.539 | 47816.0442 | 952.3209 | 852.3209 | 994.848
4903 | 105.8779 | 0.98118 | 0.96779 | 1.0265 | 193.0874 | 52.0235 | 57279.966 | 58078.6058 | 1157.5721 | 1057.5721 | 994.8658
4904 | 111.6384 | 0.95408 | 0.88627 | 1.0178 | 203.7475 | 49.7224 | 50350.8279 | 50227.8493 | 1000.557 | 900.557 | 994.8833
4905 | 99.9475 | 1.0088 | 0.66638 | 0.92567 | 192.2144 | 47.6701 | 34403.9671 | 34687.8405 | 689.7568 | 589.7568 | 994.8921
4906 | 113.89 | 1.0088 | 1.1117 | 1.0004 | 216.3232 | 48.5466 | 59781.6657 | 59341.5277 | 1182.8306 | 1082.8306 | 994.8992
4907 | 99.2984 | 0.97849 | 1.1255 | 1.0183 | 204.427 | 49.2105 | 61386.0262 | 60516.1595 | 1206.3232 | 1106.3232 | 994.9112
4908 | 102.0946 | 0.94971 | 0.98948 | 0.99347 | 218.7824 | 52.1351 | 56523.5267 | 58381.7867 | 1163.6357 | 1063.6357 | 994.9237
4909 | 93.0929 | 0.99967 | 1.1297 | 0.92971 | 214.2975 | 47.9969 | 55091.526 | 55205.2593 | 1100.1052 | 1000.1052 | 994.9397
4910 | 99.1362 | 0.94153 | 1.2586 | 0.95368 | 202.2296 | 46.2518 | 60034.697 | 60453.815 | 1205.0763 | 1105.0763 | 994.9464
4911 | 109.0817 | 0.97196 | 0.84235 | 0.99739 | 193.3031 | 51.2391 | 48674.6929 | 48000.7572 | 956.0151 | 856.0151 | 995.0241
4912 | 86.0201 | 0.90749 | 0.96284 | 1.0076 | 189.4903 | 48.7629 | 51303.4698 | 51254.2045 | 1021.0841 | 921.0841 | 995.0375
4913 | 106.2616 | 1.0362 | 0.93297 | 1.0559 | 211.8609 | 50.4015 | 55413.7205 | 55310.2945 | 1102.2059 | 1002.2059 | 995.0429
4914 | 111.1345 | 0.93144 | 0.87626 | 0.94038 | 211.5751 | 49.4439 | 46072.3904 | 47398.7137 | 943.9743 | 843.9743 | 995.0628
4915 | 102.0993 | 0.96116 | 1.1219 | 0.97773 | 197.0685 | 50.0438 | 60000.0715 | 59830.8099 | 1192.6162 | 1092.6162 | 995.0788
4916 | 98.2833 | 1.0431 | 1.107 | 1.0936 | 188.7629 | 46.9028 | 61779.148 | 60050.5205 | 1197.0104 | 1097.0104 | 995.1033
4917 | 105.373 | 0.97997 | 0.92871 | 1.0248 | 185.7603 | 48.6438 | 51505.7232 | 51953.5037 | 1035.0701 | 935.0701 | 995.1475
4918 | 78.3613 | 1.0051 | 0.95713 | 0.87006 | 206.512 | 50.6727 | 46395.9234 | 46965.0096 | 935.3002 | 835.3002 | 995.1604
4919 | 103.7083 | 1.0235 | 0.63405 | 1.0615 | 184.6973 | 52.7622 | 41295.8105 | 41186.8164 | 819.7363 | 719.7363 | 995.192
4920 | 101.6115 | 0.91691 | 1.1176 | 1.0442 | 198.7738 | 50.2947 | 63580.0089 | 64968.2771 | 1295.3655 | 1195.3655 | 995.2452
4921 | 98.3559 | 0.98717 | 1.1502 | 1.0513 | 189.3329 | 52.6115 | 68918.8933 | 68641.4241 | 1368.8285 | 1268.8285 | 995.2689
4922 | 97.2138 | 0.9698 | 1.1702 | 1.0311 | 205.2371 | 51.2681 | 66898.8379 | 64063.6832 | 1277.2737 | 1177.2737 | 995.3332
4923 | 105.2708 | 0.99548 | 1.0528 | 0.94774 | 198.8923 | 50.4671 | 55844.857 | 54554.5076 | 1087.0902 | 987.0902 | 995.3854
4924 | 89.6532 | 0.95623 | 0.96228 | 1.0154 | 187.1415 | 50.4431 | 53799.8203 | 53270.1204 | 1061.4024 | 961.4024 | 995.5535
4925 | 89.7271 | 0.95413 | 0.69902 | 1.0679 | 188.5065 | 48.5186 | 40560.7993 | 40509.6859 | 806.1937 | 706.1937 | 995.5891
4926 | 103.984 | 1.0084 | 0.65406 | 1.0621 | 190.5761 | 51.3081 | 41211.6057 | 41207.2083 | 820.1442 | 720.1442 | 995.6031
4927 | 94.9055 | 0.98369 | 1.0093 | 1.0358 | 193.479 | 49.8928 | 57007.6128 | 56684.8677 | 1129.6974 | 1029.6974 | 995.6764
4928 | 94.1468 | 0.99811 | 0.81509 | 0.94794 | 207.5075 | 51.3809 | 44735.592 | 46133.523 | 918.6705 | 818.6705 | 995.6975
4929 | 93.8616 | 1.0103 | 0.81748 | 0.96117 | 189.3771 | 44.4957 | 39370.6001 | 40298.7015 | 801.974 | 701.974 | 995.7544
4930 | 101.8694 | 0.96418 | 0.95031 | 1.0359 | 206.5227 | 45.3346 | 49287.3559 | 50056.4549 | 997.1291 | 897.1291 | 995.7975
4931 | 88.1298 | 1.0225 | 1.0413 | 0.99952 | 196.9228 | 50.1854 | 56950.6374 | 55226.1832 | 1100.5237 | 1000.5237 | 995.8213
4932 | 95.1226 | 1.0434 | 1.0157 | 0.99784 | 195.0667 | 43.9195 | 49068.2808 | 48378.6662 | 963.5733 | 863.5733 | 995.8475
4933 | 112.3601 | 0.907 | 1.2401 | 1.0588 | 202.487 | 50.4938 | 71647.4395 | 69938.2533 | 1394.7651 | 1294.7651 | 995.8476
4934 | 93.4274 | 1.0165 | 0.99466 | 1.0695 | 219.6377 | 51.0718 | 59399.5788 | 57863.1719 | 1153.2634 | 1053.2634 | 995.8869
4935 | 99.1956 | 1.0182 | 1.1438 | 0.99526 | 191.5191 | 53.2949 | 66243.6549 | 65432.9155 | 1304.6583 | 1204.6583 | 995.9076
4936 | 104.6617 | 0.94841 | 1.1992 | 1.0852 | 186.3548 | 51.1947 | 71889.402 | 73706.2185 | 1470.1244 | 1370.1244 | 995.9543
4937 | 103.9097 | 0.98878 | 1.1285 | 0.9306 | 208.0694 | 51.9547 | 60107.5431 | 59789.6834 | 1191.7937 | 1091.7937 | 996
4938 | 111.4908 | 0.98363 | 0.78774 | 0.9666 | 190.1896 | 49.6916 | 43476.27 | 41866.9574 | 833.3391 | 733.3391 | 996.0618
4939 | 99.2809 | 1.005 | 1.2226 | 0.94388 | 180.2059 | 48.7529 | 61305.8853 | 60066.4927 | 1197.3299 | 1097.3299 | 996.2013
4940 | 108.7959 | 1.0407 | 0.82615 | 0.97525 | 217.1962 | 52.0594 | 48056.2331 | 48837.0964 | 972.7419 | 872.7419 | 996.2034
4941 | 95.6443 | 1.0053 | 1.0393 | 1.0677 | 200.3598 | 51.6145 | 62435.6582 | 61824.7336 | 1232.4947 | 1132.4947 | 996.4088
4942 | 83.8023 | 1.037 | 0.92938 | 0.99505 | 215.0641 | 49.6915 | 50487.0409 | 49967.4899 | 995.3498 | 895.3498 | 996.4347
4943 | 102.6617 | 1.0472 | 1.0047 | 1.0137 | 192.4457 | 51.9308 | 58668.4252 | 59759.3862 | 1191.1877 | 1091.1877 | 996.4489
4944 | 113.2215 | 1.0647 | 0.69186 | 0.99226 | 199.5205 | 49.0222 | 39763.345 | 38772.6942 | 771.4539 | 671.4539 | 996.4698
4945 | 95.1856 | 1.0746 | 0.99786 | 1.0194 | 195.4776 | 50.7815 | 57045.6002 | 56729.7437 | 1130.5949 | 1030.5949 | 996.4973
4946 | 100.4679 | 0.95125 | 0.75799 | 0.95392 | 207.9198 | 48.9926 | 40314.766 | 40693.7043 | 809.8741 | 709.8741 | 996.5445
4947 | 92.9741 | 0.99652 | 0.94701 | 0.96903 | 222.8127 | 50.6757 | 51421.6915 | 51578.0045 | 1027.5601 | 927.5601 | 996.5708
4948 | 100.5785 | 0.97999 | 1.0284 | 0.98942 | 194.7641 | 47.3144 | 53000.9909 | 53196.9314 | 1059.9386 | 959.9386 | 996.6738
4949 | 103.6513 | 0.94247 | 1.214 | 0.94933 | 219.4537 | 47.5021 | 59606.2415 | 60191.581 | 1199.8316 | 1099.8316 | 996.7329
4950 | 107.6406 | 0.98052 | 0.95732 | 1.026 | 207.2689 | 49.5903 | 54147.8759 | 56499.3189 | 1125.9864 | 1025.9864 | 996.7749
4951 | 97.1537 | 1.0042 | 1.0062 | 1.0246 | 212.9287 | 50.7578 | 57491.4781 | 59731.4515 | 1190.629 | 1090.629 | 996.7757
4952 | 93.758 | 0.96005 | 1.0172 | 0.99593 | 197.9057 | 51.4164 | 56914.5009 | 57219.1862 | 1140.3837 | 1040.3837 | 996.7884
4953 | 101.245 | 0.95375 | 1.084 | 0.99607 | 190.5382 | 54.3401 | 64113.1082 | 67164.1252 | 1339.2825 | 1239.2825 | 996.8036
4954 | 99.6641 | 1.0183 | 1.1421 | 0.98254 | 191.5268 | 46.5221 | 57118.4811 | 54870.943 | 1093.4189 | 993.4189 | 996.8053
4955 | 96.9122 | 1.0586 | 0.80357 | 0.92329 | 205.4571 | 48.7322 | 41360.6571 | 42185.0176 | 839.7004 | 739.7004 | 996.8076
4956 | 100.3649 | 0.99511 | 1.2313 | 1.0282 | 195.4294 | 50.715 | 69467.9581 | 69460.2589 | 1385.2052 | 1285.2052 | 996.8897
4957 | 96.6136 | 1.0008 | 1.0751 | 0.93776 | 196.2813 | 49.2014 | 54556.3127 | 55474.4867 | 1105.4897 | 1005.4897 | 996.9282
4958 | 104.614 | 0.87214 | 0.83507 | 1.0047 | 198.5794 | 54.6307 | 51020.0271 | 52624.0182 | 1048.4804 | 948.4804 | 996.9641
4959 | 86.337 | 0.94409 | 1.0815 | 0.94311 | 200.0617 | 55.5582 | 61396.1908 | 62637.3164 | 1248.7463 | 1148.7463 | 996.9901
4960 | 116.0103 | 0.97735 | 1.1291 | 0.98926 | 191.8121 | 48.7701 | 60196.4149 | 59448.1639 | 1184.9633 | 1084.9633 | 997.0077
4961 | 106.6509 | 1.0713 | 1.0743 | 0.98781 | 196.9665 | 50.837 | 59954.1068 | 59077.9455 | 1177.5589 | 1077.5589 | 997.075
4962 | 119.2976 | 1.0016 | 1.0849 | 0.97019 | 203.6038 | 46.7185 | 54959.2539 | 54647.2829 | 1088.9457 | 988.9457 | 997.2603
4963 | 91.2369 | 0.9061 | 1.1335 | 0.98032 | 194.7842 | 50.1774 | 60102.3671 | 57694.5377 | 1149.8908 | 1049.8908 | 997.2822
4964 | 97.3062 | 0.99947 | 1.3795 | 0.94503 | 207.3206 | 47.4347 | 66657.3396 | 67663.0877 | 1349.2618 | 1249.2618 | 997.2924
4965 | 109.274 | 1.0013 | 0.98424 | 1.0675 | 201.988 | 48.8695 | 56893.1454 | 56421.273 | 1124.4255 | 1024.4255 | 997.2961
4966 | 96.961 | 0.98854 | 0.80681 | 1.0132 | 185.4191 | 50.3395 | 46160.7594 | 46537.5045 | 926.7501 | 826.7501 | 997.3703
4967 | 101.6195 | 1.0224 | 0.9627 | 0.99094 | 181.392 | 50.2748 | 53365.8686 | 56726.1414 | 1130.5228 | 1030.5228 | 997.3789
4968 | 91.5858 | 1.1149 | 1.1562 | 1.0096 | 224.9972 | 49.8916 | 63556.0959 | 62514.331 | 1246.2866 | 1146.2866 | 997.3977
4969 | 99.1745 | 0.96033 | 1.3345 | 1.1315 | 198.4949 | 48.4934 | 78039.5172 | 76621.2145 | 1528.4243 | 1428.4243 | 997.4524
4970 | 101.2471 | 0.84576 | 1.1629 | 1.0966 | 218.5707 | 44.8824 | 61296.374 | 60856.8223 | 1213.1364 | 1113.1364 | 997.4726
4971 | 94.8469 | 0.98059 | 1.0522 | 1.0734 | 195.429 | 53.2383 | 65270.4903 | 66439.7213 | 1324.7944 | 1224.7944 | 997.5137
4972 | 117.1074 | 0.9508 | 0.92668 | 1.0989 | 209.8367 | 50.447 | 57197.0367 | 57001.7507 | 1136.035 | 1036.035 | 997.5851
4973 | 109.0998 | 0.91884 | 0.86567 | 1.0494 | 215.4365 | 49.1293 | 49770.5935 | 49777.4926 | 991.5499 | 891.5499 | 997.5888
4974 | 95.3688 | 1.0626 | 0.99045 | 1.0342 | 215.3694 | 50.1291 | 56643.7934 | 57980.7272 | 1155.6145 | 1055.6145 | 997.6368
4975 | 85.69 | 0.90615 | 1.0194 | 1.0282 | 199.8056 | 51.7352 | 58443.1929 | 59268.6225 | 1181.3725 | 1081.3725 | 997.6895
4976 | 93.1642 | 0.90649 | 1.0873 | 1.0017 | 199.6142 | 54.6968 | 64392.4993 | 67189.7911 | 1339.7958 | 1239.7958 | 997.7073
4977 | 89.6256 | 0.93271 | 1.0918 | 1.128 | 192.036 | 49.1748 | 64863.8016 | 64046.9802 | 1276.9396 | 1176.9396 | 997.7353
4978 | 92.2938 | 0.96605 | 0.84125 | 0.93876 | 204.4191 | 50.3605 | 44465.8671 | 44480.6783 | 885.6136 | 785.6136 | 997.7714
4979 | 104.8945 | 1.0231 | 1.1459 | 0.97354 | 190.2607 | 55.3404 | 67863.8315 | 70382.9691 | 1403.6594 | 1303.6594 | 997.7738
4980 | 109.1672 | 1.0447 | 0.99759 | 1.029 | 198.9747 | 48.9813 | 56065.3555 | 57044.8236 | 1136.8965 | 1036.8965 | 997.813
4981 | 96.1171 | 0.99331 | 1.2332 | 0.9881 | 198.2735 | 52.042 | 68581.2554 | 68734.4099 | 1370.6882 | 1270.6882 | 997.8181
4982 | 100.4872 | 1.0123 | 1.0935 | 1.0189 | 211.3936 | 46.3307 | 56543.4813 | 53934.2424 | 1074.6848 | 974.6848 | 997.8245
4983 | 107.3263 | 0.94601 | 1.2923 | 0.957 | 193.8921 | 50.6625 | 67991.2558 | 67387.4047 | 1343.7481 | 1243.7481 | 997.862
4984 | 111.1813 | 1.0082 | 1.0795 | 0.95958 | 198.6206 | 46.7469 | 53862.1577 | 54767.3967 | 1091.3479 | 991.3479 | 997.8818
4985 | 112.1894 | 0.97545 | 0.85642 | 0.99871 | 204.576 | 47.2136 | 45753.8336 | 45855.8965 | 913.1179 | 813.1179 | 997.9536
4986 | 102.0927 | 1.0193 | 1.0282 | 0.95549 | 191.0917 | 51.6766 | 56336.6665 | 55845.3036 | 1112.9061 | 1012.9061 | 997.9616
4987 | 90.6909 | 1.0369 | 1.0849 | 0.96005 | 206.4743 | 52.4819 | 59802.1775 | 60348.6351 | 1202.9727 | 1102.9727 | 997.9778
4988 | 95.2533 | 0.98331 | 0.9966 | 0.99797 | 198.2553 | 48.6436 | 53134.4513 | 54262.4809 | 1081.2496 | 981.2496 | 997.9819
4989 | 97.0426 | 0.99086 | 1.1131 | 0.91982 | 193.1052 | 45.7049 | 51382.6833 | 51212.3708 | 1020.2474 | 920.2474 | 998.0194
4990 | 105.8369 | 0.98901 | 1.0637 | 0.9551 | 209.2632 | 44.6257 | 50218.301 | 49401.1391 | 984.0228 | 884.0228 | 998.0343
4991 | 119.7057 | 1.0218 | 0.96924 | 1.0238 | 177.088 | 46.5383 | 52049.4386 | 51732.5952 | 1030.6519 | 930.6519 | 998.0391
4992 | 96.8884 | 1.058 | 0.85511 | 0.96056 | 205.0798 | 51.0391 | 47359.8046 | 46948.5734 | 934.9715 | 834.9715 | 998.078
4993 | 99.6304 | 0.99152 | 0.8694 | 1.0012 | 176.9885 | 47.823 | 46529.9936 | 46049.6312 | 916.9926 | 816.9926 | 998.095
4994 | 91.387 | 1.0262 | 0.81349 | 1.0296 | 199.8091 | 51.1746 | 47859.8924 | 48465.6238 | 965.3125 | 865.3125 | 998.1098
4995 | 99.3113 | 1.0315 | 0.84758 | 1.0103 | 194.3129 | 44.6731 | 43024.2727 | 44556.7103 | 887.1342 | 787.1342 | 998.2716
4996 | 97.8599 | 0.96961 | 1.0786 | 0.99125 | 191.3046 | 48.883 | 57095.2651 | 56928.8364 | 1134.5767 | 1034.5767 | 998.278
4997 | 99.8275 | 0.95878 | 0.74977 | 1.0167 | 197.6944 | 51.7214 | 44573.9877 | 45593.0027 | 907.8601 | 807.8601 | 998.296
4998 | 107.9047 | 0.95133 | 1.1764 | 1.0769 | 203.497 | 50.4172 | 69249.2623 | 69685.7548 | 1389.7151 | 1289.7151 | 998.3053
4999 | 102.0838 | 0.96185 | 0.90846 | 1.0202 | 191.9174 | 48.57 | 49977.5448 | 49805.6571 | 992.1131 | 892.1131 | 998.3301
5000 | 100.8925 | 1.0107 | 0.92719 | 0.97753 | 209.6744 | 52.3633 | 53009.1575 | 51582.1073 | 1027.6421 | 927.6421 | 998.3302
5001 | 95.8297 | 0.99139 | 0.77437 | 1.0355 | 197.8014 | 48.5532 | 43743.2135 | 43806.883 | 872.1377 | 772.1377 | 998.3393
5002 | 85.9646 | 1.0101 | 1.0565 | 1.0605 | 188.8006 | 52.6559 | 63758.0781 | 63029.216 | 1256.5843 | 1156.5843 | 998.4389
5003 | 126.5518 | 0.95867 | 0.73969 | 1.0856 | 200.8834 | 49.8893 | 46316.2468 | 45704.185 | 910.0837 | 810.0837 | 998.5262
5004 | 87.995 | 1.0132 | 0.66122 | 1.0351 | 191.4639 | 49.6353 | 38589.9083 | 38836.2759 | 772.7255 | 672.7255 | 998.5622
5005 | 116.4865 | 1.0542 | 1.0246 | 1.0524 | 183.785 | 49.772 | 59965.4619 | 60407.3616 | 1204.1472 | 1104.1472 | 998.5951
5006 | 116.1727 | 0.944 | 1.0746 | 0.9884 | 200.8496 | 45.5129 | 53533.4817 | 54630.9752 | 1088.6195 | 988.6195 | 998.6027
5007 | 93.111 | 0.96637 | 1.1077 | 0.96671 | 186.6198 | 50.3178 | 58595.3046 | 57541.2343 | 1146.8247 | 1046.8247 | 998.6507
5008 | 93.7214 | 0.99362 | 1.31 | 0.91522 | 200.3115 | 51.2848 | 66463.0274 | 65463.9367 | 1305.2787 | 1205.2787 | 998.6867
5009 | 102.8786 | 1.1009 | 1.0593 | 1.0098 | 211.8728 | 51.4588 | 61083.9186 | 59096.5394 | 1177.9308 | 1077.9308 | 998.6942
5010 | 109.6356 | 0.97194 | 1.064 | 0.99609 | 212.5726 | 54.9855 | 64349.7266 | 65120.5162 | 1298.4103 | 1198.4103 | 998.7367
5011 | 91.6743 | 0.99575 | 0.96449 | 1.0881 | 216.8416 | 53.0786 | 60767.6097 | 63219.7885 | 1260.3958 | 1160.3958 | 998.7753
5012 | 115.4694 | 1.0163 | 1.1614 | 1.0911 | 191.3235 | 51.2157 | 71101.202 | 73050.2339 | 1457.0047 | 1357.0047 | 998.8495
5013 | 103.9351 | 1.0531 | 1.3205 | 1.0407 | 214.6689 | 49.5892 | 73791.4086 | 73775.2905 | 1471.5058 | 1371.5058 | 998.8744
5014 | 108.676 | 0.96412 | 1.1232 | 0.97844 | 191.0889 | 49.1466 | 59353.8925 | 58112.2541 | 1158.2451 | 1058.2451 | 998.8891
5015 | 116.3137 | 1.001 | 0.92615 | 1.0415 | 204.2426 | 49.2514 | 53445.6201 | 53767.6526 | 1071.3531 | 971.3531 | 998.9346
5016 | 104.9802 | 0.91212 | 1.09 | 0.9944 | 187.0225 | 54.3939 | 64352.7884 | 61014.2874 | 1216.2857 | 1116.2857 | 998.951
5017 | 93.4917 | 1.0092 | 0.96608 | 1.044 | 179.8331 | 47.6254 | 52707.8854 | 52920.314 | 1054.4063 | 954.4063 | 999.014
5018 | 79.2759 | 1.05 | 0.89708 | 1.0707 | 205.3449 | 50.7553 | 53181.0004 | 53409.942 | 1064.1988 | 964.1988 | 999.0299
5019 | 98.9823 | 0.97853 | 0.91332 | 1.0638 | 216.9892 | 49.3546 | 52950.6293 | 52164.3925 | 1039.2879 | 939.2879 | 999.0571
5020 | 93.8789 | 1.0535 | 0.9696 | 0.94935 | 197.4799 | 56.5617 | 57856.3026 | 57085.5289 | 1137.7106 | 1037.7106 | 999.1193
5021 | 89.2277 | 1.0072 | 0.84338 | 1.0015 | 190.9697 | 51.1934 | 48033.5448 | 46321.238 | 922.4248 | 822.4248 | 999.1715
5022 | 107.1584 | 0.97234 | 0.79444 | 1.0678 | 200.1052 | 47.8757 | 45801.3137 | 45467.5162 | 905.3503 | 805.3503 | 999.2505
5023 | 97.1135 | 0.93525 | 1.1177 | 0.94893 | 195.5555 | 51.9103 | 59966.6869 | 61900.8116 | 1234.0162 | 1134.0162 | 999.2665
5024 | 104.6205 | 1.0663 | 1.0024 | 0.97229 | 188.8425 | 48.8546 | 53255.7848 | 54582.9651 | 1087.6593 | 987.6593 | 999.2888
5025 | 94.9353 | 1.0184 | 1.0035 | 0.96576 | 193.6753 | 49.1232 | 52552.3866 | 51495.4383 | 1025.9088 | 925.9088 | 999.3338
5026 | 76.3642 | 1.0453 | 1.036 | 1.0021 | 197.8263 | 48.9635 | 54938.7092 | 56414.8808 | 1124.2976 | 1024.2976 | 999.372
5027 | 89.3353 | 1.0051 | 1.25 | 1.059 | 218.7664 | 47.9358 | 67973.7721 | 67151.4389 | 1339.0288 | 1239.0288 | 999.4191
5028 | 107.2619 | 1.0691 | 1.0773 | 0.98394 | 207.6664 | 49.5142 | 58370.9606 | 57078.114 | 1137.5623 | 1037.5623 | 999.5699
5029 | 103.6502 | 1.0575 | 1.1079 | 0.89934 | 188.5282 | 48.4303 | 53752.9567 | 55098.0811 | 1097.9616 | 997.9616 | 999.5997
5030 | 98.4429 | 1.0707 | 1.1517 | 1.01 | 208.847 | 48.4114 | 61629.0347 | 61527.2506 | 1226.545 | 1126.545 | 999.6188
5031 | 104.1064 | 0.97526 | 1.1338 | 0.96188 | 209.6181 | 51.0302 | 61042.6184 | 61909.6812 | 1234.1936 | 1134.1936 | 999.6583
5032 | 97.6278 | 0.98054 | 1.2139 | 1.0016 | 204.2318 | 51.268 | 67447.1317 | 65520.7781 | 1306.4156 | 1206.4156 | 999.8123
5033 | 86.1042 | 0.99641 | 0.79476 | 0.99419 | 203.2558 | 49.195 | 43294.9378 | 42550.2116 | 847.0042 | 747.0042 | 999.9438
5034 | 107.1914 | 0.99859 | 0.80127 | 1.0107 | 210.2616 | 52.2659 | 48132.7753 | 46672.3305 | 929.4466 | 829.4466 | 1000.083
5035 | 97.528 | 0.99642 | 0.92517 | 1.0249 | 195.4428 | 52.0728 | 54633.597 | 54739.4423 | 1090.7888 | 990.7888 | 1000.1052
5036 | 95.044 | 0.97709 | 0.8308 | 1.051 | 193.1062 | 41.5316 | 40315.0195 | 39676.0243 | 789.5205 | 689.5205 | 1000.115
5037 | 108.1909 | 1.0247 | 0.90849 | 1.0055 | 189.9684 | 47.8947 | 49248.9086 | 50651.9529 | 1009.0391 | 909.0391 | 1000.1369
5038 | 72.9171 | 1.042 | 1.1823 | 0.99242 | 214.8292 | 45.9343 | 57601.2869 | 57413.5025 | 1144.2701 | 1044.2701 | 1000.2041
5039 | 99.7041 | 0.98198 | 1.0576 | 1.0268 | 207.494 | 52.3305 | 62158.535 | 63012.3664 | 1256.2473 | 1156.2473 | 1000.214
5040 | 93.6862 | 0.99781 | 0.65055 | 0.93153 | 202.5558 | 51.7278 | 36385.1638 | 36548.4038 | 726.9681 | 626.9681 | 1000.2536
5041 | 107.4028 | 1.0088 | 0.87845 | 0.96712 | 186.3848 | 47.2959 | 45491.6693 | 46592.1539 | 927.8431 | 827.8431 | 1000.3354
5042 | 98.1204 | 0.9251 | 1.2469 | 0.95152 | 190.8347 | 50.3382 | 64484.4614 | 65591.8751 | 1307.8375 | 1207.8375 | 1000.3408
5043 | 93.0106 | 1.0938 | 0.92536 | 1.0723 | 208.2365 | 52.1777 | 57292.7146 | 56597.174 | 1127.9435 | 1027.9435 | 1000.3431
5044 | 102.1736 | 1.0036 | 0.97775 | 1.0159 | 192.5753 | 51.2616 | 56365.087 | 55360.8521 | 1103.217 | 1003.217 | 1000.3569
5045 | 95.6174 | 0.9804 | 1.0699 | 1.0269 | 194.8043 | 50.7103 | 60666.215 | 59745.6086 | 1190.9122 | 1090.9122 | 1000.361
5046 | 117.2475 | 0.98846 | 1.0697 | 0.97768 | 218.2183 | 49.8048 | 58076.399 | 56840.1776 | 1132.8036 | 1032.8036 | 1000.4136
5047 | 98.1283 | 1.0191 | 0.9887 | 0.97418 | 211.639 | 51.7083 | 55186.6198 | 53386.823 | 1063.7365 | 963.7365 | 1000.4433
5048 | 102.6287 | 1.0179 | 0.97523 | 0.98882 | 190.6926 | 53.7367 | 57623.9878 | 59298.4407 | 1181.9688 | 1081.9688 | 1000.4474
5049 | 105.6866 | 0.95345 | 1.0099 | 0.932 | 179.1656 | 56.3895 | 58937.3777 | 57627.1387 | 1148.5428 | 1048.5428 | 1000.5237
5050 | 92.2329 | 1.0269 | 1.0167 | 0.99279 | 190.8886 | 48.2626 | 53478.0854 | 53082.7085 | 1057.6542 | 957.6542 | 1000.5485
5051 | 99.6872 | 1.0729 | 0.89585 | 0.95639 | 203.9948 | 51.2999 | 49643.3723 | 47796.9349 | 951.9387 | 851.9387 | 1000.6079
5052 | 108.2855 | 0.99324 | 1.147 | 1.0385 | 191.2242 | 51.2897 | 66799.4253 | 68269.6738 | 1361.3935 | 1261.3935 | 1000.6186
5053 | 96.3408 | 1.0329 | 0.91498 | 1.0077 | 216.148 | 51.5821 | 52908.7965 | 54620.3931 | 1088.4079 | 988.4079 | 1000.7213
5054 | 99.1911 | 0.97677 | 0.96631 | 1.0388 | 214.2044 | 50.7758 | 56100.4792 | 53242.5112 | 1060.8502 | 960.8502 | 1000.8101
5055 | 105.3694 | 0.99085 | 1.0486 | 1.0408 | 193.167 | 52.434 | 62891.3546 | 61858.4128 | 1233.1683 | 1133.1683 | 1000.8315
5056 | 93.9369 | 0.90421 | 0.84426 | 0.90865 | 210.5683 | 44.1308 | 37813.4528 | 38575.3596 | 767.5072 | 667.5072 | 1000.8436
5057 | 89.5955 | 1.0556 | 1.0269 | 0.93465 | 206.5184 | 51.8125 | 54835.422 | 55942.8311 | 1114.8566 | 1014.8566 | 1000.8507
5058 | 103.1534 | 1.0315 | 1.1141 | 1.003 | 208.8364 | 51.2236 | 62900.4562 | 63510.305 | 1266.2061 | 1166.2061 | 1000.8703
5059 | 116.1275 | 0.96004 | 0.78377 | 0.99434 | 203.0996 | 52.9102 | 47336.6746 | 46354.7099 | 923.0942 | 823.0942 | 1000.9041
5060 | 87.1288 | 0.92413 | 1.0349 | 0.89176 | 185.4597 | 49.1529 | 49507.1773 | 50353.8568 | 1003.0771 | 903.0771 | 1000.9129
5061 | 87.6018 | 1.0565 | 1.0433 | 1.0035 | 210.126 | 52.1607 | 59649.9602 | 60479.1252 | 1205.5825 | 1105.5825 | 1000.9314
5062 | 73.7503 | 0.99325 | 0.883 | 1.0508 | 191.7655 | 45.004 | 45246.5185 | 45368.2554 | 903.3651 | 803.3651 | 1000.9852
5063 | 103.718 | 1.0329 | 1.2163 | 1.0258 | 203.1286 | 51.1629 | 69520.8156 | 66418.4494 | 1324.369 | 1224.369 | 1001.0966
5064 | 102.4298 | 1.1036 | 0.94496 | 1.0065 | 184.6868 | 51.3199 | 54797.2195 | 55553.1501 | 1107.063 | 1007.063 | 1001.1146
5065 | 103.72 | 1.023 | 0.89348 | 0.94501 | 199.8285 | 51.7835 | 49417.8288 | 49081.9315 | 977.6386 | 877.6386 | 1001.1229
5066 | 114.3141 | 1.0783 | 0.92253 | 0.97931 | 205.7481 | 46.792 | 48247.0696 | 48776.0553 | 971.5211 | 871.5211 | 1001.1248
5067 | 82.5258 | 0.95156 | 1.0447 | 0.95221 | 204.9677 | 50.3139 | 54205.4368 | 55432.6585 | 1104.6532 | 1004.6532 | 1001.1745
5068 | 124.408 | 1.0069 | 0.9495 | 1.011 | 205.3405 | 47.2681 | 51500.3314 | 50414.9932 | 1004.2999 | 904.2999 | 1001.3093
5069 | 102.2064 | 0.9988 | 1.0179 | 1.0003 | 196.3169 | 47.8836 | 53842.1674 | 55064.8029 | 1097.2961 | 997.2961 | 1001.3676
5070 | 113.5312 | 1.044 | 1.2791 | 0.96003 | 181.4943 | 50.4172 | 68070.5603 | 67229.1688 | 1340.5834 | 1240.5834 | 1001.3834
5071 | 92.5809 | 1.0199 | 1.023 | 1.0857 | 187.1836 | 53.2463 | 64354.3758 | 64257.9492 | 1281.159 | 1181.159 | 1001.4198
5072 | 87.3179 | 1.0753 | 1.011 | 0.99241 | 206.8927 | 52.4426 | 57747.761 | 58381.9781 | 1163.6396 | 1063.6396 | 1001.4606
5073 | 112.0743 | 0.95738 | 0.99705 | 1.0046 | 198.0331 | 43.9131 | 48892.5831 | 48574.2905 | 967.4858 | 867.4858 | 1001.5053
5074 | 82.5261 | 0.99211 | 0.72992 | 0.98895 | 202.1154 | 48.805 | 39428.1496 | 39926.5688 | 794.5314 | 694.5314 | 1001.5125
5075 | 99.0776 | 1.0194 | 0.93486 | 1.0295 | 216.5029 | 51.8662 | 55374.2944 | 55033.6878 | 1096.6738 | 996.6738 | 1001.6286
5076 | 121.3145 | 0.97163 | 0.94372 | 1.0372 | 190.4602 | 55.0079 | 60516.3023 | 59373.5513 | 1183.471 | 1083.471 | 1001.6349
5077 | 96.9752 | 0.94428 | 1.0304 | 0.99639 | 196.0168 | 49.6775 | 55748.7591 | 57129.3336 | 1138.5867 | 1038.5867 | 1001.6588
5078 | 88.4434 | 0.94951 | 1.3367 | 0.91326 | 198.0628 | 53.2429 | 69667.6588 | 68064.0235 | 1357.2805 | 1257.2805 | 1001.661
5079 | 112.314 | 0.90491 | 1.1388 | 1.0615 | 184.6361 | 52.5547 | 69053.7798 | 71257.0125 | 1421.1402 | 1321.1402 | 1001.6823
5080 | 93.0507 | 0.98877 | 1.1618 | 1.0161 | 196.4736 | 48.0069 | 61287.3442 | 63224.6966 | 1260.4939 | 1160.4939 | 1001.8203
5081 | 107.1518 | 1.0515 | 0.93467 | 0.99833 | 220.4618 | 48.9797 | 51442.574 | 50847.0318 | 1012.9406 | 912.9406 | 1001.8494
5082 | 117.2583 | 0.9562 | 0.93717 | 0.97202 | 179.3707 | 50.992 | 52347.5929 | 53007.323 | 1056.1465 | 956.1465 | 1001.9076
5083 | 102.6912 | 1.052 | 0.66534 | 1.0159 | 201.9243 | 47.2842 | 37271.828 | 36935.8298 | 734.7166 | 634.7166 | 1001.9177
5084 | 103.2203 | 0.88514 | 1.0429 | 1.0874 | 190.1596 | 49.7211 | 61116.7044 | 62751.0754 | 1251.0215 | 1151.0215 | 1001.9436
5085 | 109.7234 | 1.0708 | 1.3122 | 1.012 | 205.614 | 50.6636 | 73437.9086 | 72559.9471 | 1447.1989 | 1347.1989 | 1001.9848
5086 | 96.2146 | 1.0202 | 0.7836 | 0.92059 | 203.3726 | 51.3661 | 42299.624 | 41512.0857 | 826.2417 | 726.2417 | 1002.0086
5087 | 92.9312 | 1.038 | 0.75478 | 0.96856 | 193.5225 | 46.4053 | 38594.6367 | 39778.8239 | 791.5765 | 691.5765 | 1002.0382
5088 | 107.8232 | 1.0205 | 1.1559 | 1.087 | 189.1414 | 48.3572 | 66270.6157 | 66142.4458 | 1318.8489 | 1218.8489 | 1002.0692
5089 | 97.8335 | 0.9481 | 0.94934 | 0.92895 | 198.4513 | 49.9449 | 48877.1474 | 49970.5279 | 995.4106 | 895.4106 | 1002.2059
5090 | 96.8415 | 1.0426 | 1.0672 | 0.96834 | 198.1144 | 47.6837 | 54288.7575 | 54867.2232 | 1093.3445 | 993.3445 | 1002.2154
5091 | 106.0137 | 0.99266 | 0.87776 | 0.88826 | 177.7355 | 50.3774 | 44757.5301 | 45004.7763 | 896.0955 | 796.0955 | 1002.2336
5092 | 97.5195 | 0.99321 | 0.96297 | 1.0486 | 186.5705 | 53.2691 | 59136.1441 | 59501.0702 | 1186.0214 | 1086.0214 | 1002.4615
5093 | 80.4658 | 1.0355 | 0.9683 | 0.97911 | 213.9272 | 49.2773 | 51038.2301 | 51101.9103 | 1018.0382 | 918.0382 | 1002.4642
5094 | 95.6194 | 0.9879 | 0.96672 | 1.0965 | 194.4412 | 55.774 | 64583.7313 | 66957.8328 | 1335.1567 | 1235.1567 | 1002.4699
5095 | 104.3664 | 0.93251 | 1.0716 | 0.97873 | 219.9559 | 52.864 | 60807.0813 | 60212.1821 | 1200.2436 | 1100.2436 | 1002.4926
5096 | 92.21 | 1.0579 | 0.99132 | 1.0575 | 214.4452 | 48.5607 | 55858.7838 | 56825.5479 | 1132.511 | 1032.511 | 1002.4982
5097 | 102.5662 | 1.0638 | 0.95969 | 0.99373 | 192.1698 | 51.8876 | 55337.557 | 58190.7277 | 1159.8146 | 1059.8146 | 1002.5097
5098 | 106.5965 | 0.94126 | 0.88517 | 1.0619 | 202.6597 | 53.4174 | 55773.0872 | 57471.1598 | 1145.4232 | 1045.4232 | 1002.5218
5099 | 90.855 | 0.9824 | 0.7914 | 0.99274 | 193.949 | 52.3454 | 45991.266 | 46151.0094 | 919.0202 | 819.0202 | 1002.5507
5100 | 109.2237 | 1.031 | 1.1326 | 0.89013 | 184.465 | 53.6966 | 60367.7141 | 59958.2826 | 1195.1657 | 1095.1657 | 1002.6393
5101 | 125.1133 | 1.1025 | 1.1536 | 0.99959 | 216.9381 | 52.1964 | 67605.2642 | 65758.7094 | 1311.1742 | 1211.1742 | 1002.7234
5102 | 90.4168 | 1.0329 | 1.291 | 0.96119 | 199.2903 | 49.1335 | 65758.8227 | 64495.3413 | 1285.9068 | 1185.9068 | 1002.7475
5103 | 109.1728 | 1.0546 | 0.86418 | 1.0388 | 196.932 | 52.6169 | 53488.9371 | 54897.5541 | 1093.9511 | 993.9511 | 1002.808
5104 | 96.8563 | 0.983 | 0.83246 | 1.0622 | 183.5353 | 52.078 | 51192.2156 | 49905.0969 | 994.1019 | 894.1019 | 1002.8742
5105 | 82.0128 | 1.0615 | 1.0319 | 1.0275 | 205.4723 | 51.8176 | 59656.39 | 58256.3133 | 1161.1263 | 1061.1263 | 1002.9036
5106 | 107.7769 | 1.0197 | 1.0945 | 0.99666 | 206.8634 | 49.7735 | 59974.3131 | 58489.5155 | 1165.7903 | 1065.7903 | 1002.9059
5107 | 97.7023 | 0.93426 | 0.87575 | 1.0575 | 199.7721 | 46.7507 | 47762.0457 | 48057.3357 | 957.1467 | 857.1467 | 1002.9533
5108 | 102.4098 | 1.0359 | 0.85344 | 0.98439 | 218.6783 | 49.9396 | 47471.6216 | 46935.3084 | 934.7062 | 834.7062 | 1003.0951
5109 | 108.0413 | 1.023 | 1.0389 | 0.99417 | 195.0492 | 51.6756 | 59281.0927 | 59983.7034 | 1195.6741 | 1095.6741 | 1003.1092
5110 | 96.4405 | 1.0811 | 1.0923 | 1.0285 | 201.407 | 55.6011 | 68462.6709 | 70697.8261 | 1409.9565 | 1309.9565 | 1003.217
5111 | 91.3853 | 0.88477 | 0.84647 | 0.90165 | 188.4011 | 46.1438 | 39137.0926 | 41344.7035 | 822.8941 | 722.8941 | 1003.2413
5112 | 104.4503 | 1.0807 | 1.0444 | 1.0531 | 195.0966 | 50.4588 | 61388.51 | 59715.0627 | 1190.3013 | 1090.3013 | 1003.2621
5113 | 96.2501 | 1.0701 | 1.0733 | 0.96433 | 191.7088 | 51.5558 | 58861.6612 | 60590.3993 | 1207.808 | 1107.808 | 1003.288
5114 | 86.0534 | 1.093 | 0.98221 | 0.8993 | 225.3779 | 51.1815 | 50247.8441 | 49915.2205 | 994.3044 | 894.3044 | 1003.3064
5115 | 93.4987 | 1.0787 | 0.94152 | 1.0376 | 213.0498 | 48.5274 | 52516.0795 | 52017.1887 | 1036.3438 | 936.3438 | 1003.4127
5116 | 85.631 | 0.98256 | 0.96453 | 1.0224 | 205.765 | 49.432 | 53112.0021 | 54262.6298 | 1081.2526 | 981.2526 | 1003.4435
5117 | 100.6344 | 0.88818 | 0.75385 | 1.0647 | 222.0133 | 47.6472 | 42723.9044 | 41065.6398 | 817.3128 | 717.3128 | 1003.4464
5118 | 101.7453 | 1.0613 | 0.96983 | 0.96362 | 200.3596 | 54.6306 | 57154.2149 | 59461.6024 | 1185.232 | 1085.232 | 1003.4753
5119 | 88.87 | 1.048 | 0.96371 | 1.0258 | 207.0525 | 48.8038 | 52998.5184 | 53149.6392 | 1058.9928 | 958.9928 | 1003.5415
5120 | 110.4372 | 1.082 | 1.0096 | 0.90804 | 186.9818 | 47.9499 | 49873.4545 | 51056.8487 | 1017.137 | 917.137 | 1003.5584
5121 | 96.3236 | 0.98462 | 1.0616 | 1.0534 | 199.4704 | 51.0044 | 62073.9188 | 63471.6196 | 1265.4324 | 1165.4324 | 1003.6145
5122 | 88.9089 | 1.0464 | 1.1583 | 1.0478 | 193.4399 | 53.3507 | 69907.0007 | 69069.6117 | 1377.3922 | 1277.3922 | 1003.6631
5123 | 105.8005 | 1.0367 | 1.1349 | 0.97911 | 205.5803 | 49.4046 | 60523.781 | 59209.9581 | 1180.1992 | 1080.1992 | 1003.6953
5124 | 101.8735 | 0.95976 | 0.80441 | 0.96104 | 209.1879 | 52.8735 | 46253.8094 | 47108.5815 | 938.1716 | 838.1716 | 1003.8972
5125 | 100.369 | 0.97099 | 1.1403 | 0.96132 | 194.1117 | 50.3593 | 60307.1979 | 60589.9568 | 1207.7991 | 1107.7991 | 1003.9524
5126 | 96.2273 | 1.0274 | 1.0087 | 0.97252 | 195.8041 | 53.7749 | 58265.0043 | 57829.0503 | 1152.581 | 1052.581 | 1003.9562
5127 | 106.7991 | 1.0394 | 1.0927 | 1.056 | 206.1197 | 48.6693 | 61767.2254 | 61500.0664 | 1226.0013 | 1126.0013 | 1004.0374
5128 | 98.763 | 0.8866 | 0.74041 | 1.017 | 206.9427 | 46.3613 | 39175.7886 | 39479.9384 | 785.5988 | 685.5988 | 1004.2317
5129 | 95.771 | 1.0661 | 0.89734 | 1.0177 | 191.9485 | 50.0454 | 51001.731 | 51551.3751 | 1027.0275 | 927.0275 | 1004.2598
5130 | 117.1363 | 1.0526 | 1.064 | 0.94103 | 191.9522 | 50.9834 | 57523.6701 | 59029.4481 | 1176.589 | 1076.589 | 1004.2771
5131 | 128.0817 | 1.0238 | 0.79562 | 0.94703 | 190.783 | 50.3578 | 44737.4395 | 43663.0349 | 869.2607 | 769.2607 | 1004.306
5132 | 104.9279 | 0.93538 | 1.0273 | 1.0551 | 183.4167 | 53.4046 | 63309.5658 | 63706.7522 | 1270.135 | 1170.135 | 1004.3067
5133 | 113.8233 | 1.0109 | 1.1396 | 1.0143 | 193.4385 | 47.6608 | 60766.7995 | 59788.6108 | 1191.7722 | 1091.7722 | 1004.3353
5134 | 99.6662 | 0.94218 | 0.8468 | 0.98044 | 200.4496 | 47.2829 | 43896.6199 | 44224.1001 | 880.482 | 780.482 | 1004.348
5135 | 91.4206 | 1.0093 | 1.1062 | 0.95985 | 211.7513 | 50.7182 | 58742.7942 | 57658.5131 | 1149.1703 | 1049.1703 | 1004.3553
5136 | 94.3665 | 1.0281 | 0.77885 | 1.0424 | 194.957 | 45.873 | 41887.7911 | 40897.8678 | 813.9574 | 713.9574 | 1004.42
5137 | 103.3217 | 1.0029 | 0.98928 | 1.11 | 209.2578 | 52.7714 | 63625.6362 | 63133.4214 | 1258.6684 | 1158.6684 | 1004.4695
5138 | 103.8987 | 0.91567 | 0.9497 | 1.0129 | 193.7418 | 50.7158 | 53806.569 | 53286.4488 | 1061.729 | 961.729 | 1004.5485
5139 | 99.3048 | 1.0566 | 1.106 | 1.0022 | 188.0972 | 50.2398 | 61143.6589 | 61099.43 | 1217.9886 | 1117.9886 | 1004.5623
5140 | 113.7843 | 0.94139 | 0.94699 | 1.0406 | 185.0197 | 48.8159 | 53519.6815 | 53441.4267 | 1064.8285 | 964.8285 | 1004.6093
5141 | 89.7511 | 1.0297 | 1.1677 | 1.0718 | 200.8917 | 54.5344 | 73487.6997 | 74521.3036 | 1486.4261 | 1386.4261 | 1004.6532
5142 | 97.3583 | 1.0193 | 1.0881 | 1.0654 | 228.9429 | 49.3572 | 62344.7983 | 62844.0342 | 1252.8807 | 1152.8807 | 1004.7347
5143 | 91.4988 | 0.94323 | 0.60543 | 1.093 | 200.9082 | 50.9829 | 38336.7164 | 38511.6891 | 766.2338 | 666.2338 | 1004.8043
5144 | 92.15 | 1.0323 | 1.028 | 1.0413 | 196.0958 | 49.887 | 58345.3918 | 58919.395 | 1174.3879 | 1074.3879 | 1004.8423
5145 | 99.6541 | 1.064 | 1.3183 | 1.0466 | 194.2123 | 51.1456 | 76181.0862 | 74549.9703 | 1486.9994 | 1386.9994 | 1004.9007
5146 | 105.4175 | 0.97579 | 0.96908 | 1.0755 | 211.8235 | 47.1152 | 54165.2917 | 53047.299 | 1056.946 | 956.946 | 1004.903
5147 | 90.9668 | 0.9653 | 1.1866 | 0.99944 | 203.0163 | 50.7299 | 64817.9949 | 65332.1278 | 1302.6426 | 1202.6426 | 1004.9065
5148 | 101.422 | 0.98962 | 1.2285 | 1.1205 | 213.0986 | 49.9032 | 73914.4382 | 73110.9285 | 1458.2186 | 1358.2186 | 1004.9658
5149 | 94.5148 | 0.92939 | 0.96165 | 1.0544 | 209.8863 | 47.7305 | 52797.591 | 50188.7256 | 999.7745 | 899.7745 | 1004.9834
5150 | 74.1211 | 0.99503 | 0.93161 | 1.0301 | 191.8446 | 48.5527 | 50368.0421 | 51203.3081 | 1020.0662 | 920.0662 | 1004.9942
5151 | 110.0381 | 0.96001 | 0.94314 | 0.92142 | 219.2455 | 49.1026 | 48077.7953 | 47368.6127 | 943.3723 | 843.3723 | 1005.0015
5152 | 90.8526 | 0.97 | 0.92957 | 1.0008 | 194.1197 | 48.8864 | 49983.472 | 50392.2351 | 1003.8447 | 903.8447 | 1005.0853
5153 | 100.5326 | 0.98184 | 1.1034 | 1.0425 | 195.6852 | 48.8063 | 61150.8771 | 61917.719 | 1234.3544 | 1134.3544 | 1005.1442
5154 | 95.5193 | 0.97146 | 1.3076 | 1.0156 | 190.5727 | 49.5642 | 70611.5642 | 70353.9393 | 1403.0788 | 1303.0788 | 1005.1462
5155 | 108.9501 | 1.0352 | 0.83101 | 1.0217 | 194.2953 | 49.4433 | 47750.3208 | 49022.5977 | 976.452 | 876.452 | 1005.1537
5156 | 111.0184 | 1.0238 | 0.79443 | 0.96049 | 200.6902 | 56.9005 | 50085.6352 | 50961.557 | 1015.2311 | 915.2311 | 1005.1986
5157 | 97.6042 | 1.0018 | 0.85335 | 1.0764 | 187.3942 | 52.2603 | 53302.7084 | 51667.7503 | 1029.355 | 929.355 | 1005.2002
5158 | 89.619 | 1.1029 | 0.9969 | 0.93525 | 202.7443 | 50.7594 | 52545.2156 | 50635.0318 | 1008.7006 | 908.7006 | 1005.3029
5159 | 106.9783 | 1.1171 | 1.1631 | 0.97863 | 198.4042 | 48.081 | 60672.7147 | 60086.5745 | 1197.7315 | 1097.7315 | 1005.387
5160 | 100.7979 | 0.94456 | 0.82892 | 1.0444 | 214.3814 | 45.4061 | 43845.6128 | 44699.3736 | 889.9875 | 789.9875 | 1005.4201
5161 | 82.4799 | 1.0327 | 0.92087 | 1.0268 | 188.5972 | 50.1341 | 51861.3215 | 52953.3572 | 1055.0671 | 955.0671 | 1005.4718
5162 | 113.3649 | 0.97757 | 0.95185 | 0.99903 | 208.5046 | 48.8817 | 52108.6777 | 52562.6125 | 1047.2523 | 947.2523 | 1005.4805
5163 | 113.6528 | 0.91361 | 0.8977 | 1.0443 | 217.9745 | 54.0543 | 56505.8718 | 56986.68 | 1135.7336 | 1035.7336 | 1005.4897
5164 | 112.5525 | 0.93571 | 0.80041 | 0.96451 | 190.0307 | 51.8884 | 45712.8572 | 46477.7022 | 925.554 | 825.554 | 1005.5045
5165 | 94.6259 | 1.0274 | 1.0602 | 0.94661 | 204.6595 | 47.7633 | 52782.5935 | 53428.5124 | 1064.5702 | 964.5702 | 1005.6192
5166 | 96.5577 | 1.0475 | 0.85175 | 1.0405 | 198.0126 | 49.0661 | 48645.1194 | 48874.1372 | 973.4827 | 873.4827 | 1005.6499
5167 | 108.6608 | 1.031 | 1.1365 | 0.99972 | 192.033 | 50.2117 | 62866.6884 | 62218.2939 | 1240.3659 | 1140.3659 | 1005.6719
5168 | 110.0903 | 0.93341 | 1.0659 | 1.0314 | 208.4114 | 52.6467 | 63501.4668 | 64984.6778 | 1295.6936 | 1195.6936 | 1005.6923
5169 | 105.8713 | 0.99773 | 1.2127 | 0.98926 | 188.827 | 51.5522 | 67482.9912 | 67994.6326 | 1355.8927 | 1255.8927 | 1005.6988
5170 | 119.0602 | 1.0135 | 0.78254 | 1.0215 | 202.1021 | 49.4529 | 45699.5006 | 44961.7246 | 895.2345 | 795.2345 | 1005.7321
5171 | 110.2264 | 1.0187 | 1.1282 | 0.9562 | 209.3937 | 50.9059 | 60840.844 | 60544.472 | 1206.8894 | 1106.8894 | 1005.7405
5172 | 96.9029 | 0.96626 | 1.24 | 1.0458 | 191.4653 | 46.4208 | 64736.1002 | 66258.3107 | 1321.1662 | 1221.1662 | 1005.7722
5173 | 113.1338 | 0.96222 | 0.85548 | 0.98344 | 194.5216 | 54.4273 | 51909.7651 | 51904.5965 | 1034.0919 | 934.0919 | 1005.8615
5174 | 107.3429 | 0.98841 | 0.78247 | 0.99428 | 216.3741 | 53.9291 | 47894.5632 | 46842.3978 | 932.848 | 832.848 | 1005.8698
5175 | 97.4945 | 0.97239 | 0.7612 | 0.92673 | 211.9479 | 52.1858 | 41972.6573 | 43329.4687 | 862.5894 | 762.5894 | 1005.8953
5176 | 121.754 | 0.97456 | 0.99354 | 0.93639 | 188.5239 | 48.6034 | 51173.3414 | 52541.522 | 1046.8304 | 946.8304 | 1005.8999
5177 | 97.9618 | 0.9894 | 1.0838 | 1.0191 | 197.7689 | 56.6462 | 68251.4078 | 67507.5921 | 1346.1518 | 1246.1518 | 1005.9128
5178 | 102.969 | 0.95613 | 1.0766 | 0.88986 | 198.8865 | 52.7746 | 55953.0174 | 56117.8413 | 1118.3568 | 1018.3568 | 1005.9226
5179 | 84.3858 | 0.95513 | 0.73239 | 1.0238 | 182.313 | 46.9594 | 39179.2428 | 39320.2835 | 782.4057 | 682.4057 | 1005.9428
5180 | 97.9214 | 1.0835 | 0.97706 | 1.0304 | 211.7529 | 46.3392 | 51778.6231 | 52140.1157 | 1038.8023 | 938.8023 | 1006.0132
5181 | 99.4789 | 1.0287 | 0.89679 | 1.1061 | 192.9576 | 52.0315 | 57131.3435 | 57287.0417 | 1141.7408 | 1041.7408 | 1006.0265
5182 | 92.9116 | 1.0216 | 1.0708 | 0.90472 | 217.9622 | 49.3517 | 52713.7546 | 51586.1356 | 1027.7227 | 927.7227 | 1006.0913
5183 | 104.1812 | 0.99736 | 0.85549 | 0.96189 | 221.949 | 56.3512 | 52447.7981 | 53447.7117 | 1064.9542 | 964.9542 | 1006.2181
5184 | 105.5415 | 1.0286 | 1.006 | 1.0201 | 200.3966 | 47.8367 | 54485.6259 | 54149.6103 | 1078.9922 | 978.9922 | 1006.2553
5185 | 108.7652 | 0.96563 | 1.1298 | 1.0681 | 204.2096 | 48.996 | 64475.5987 | 64425.0773 | 1284.5015 | 1184.5015 | 1006.2795
5186 | 93.9615 | 1.0042 | 1.1344 | 1.0049 | 206.6665 | 51.586 | 63879.9536 | 64637.9085 | 1288.7582 | 1188.7582 | 1006.3624
5187 | 96.0404 | 0.92987 | 1.1079 | 1.0423 | 179.9503 | 57.1504 | 71282.7819 | 69236.9691 | 1380.7394 | 1280.7394 | 1006.3893
5188 | 111.3958 | 1.0306 | 1.1355 | 1.1013 | 199.0394 | 49.3923 | 67633.7149 | 68580.9991 | 1367.62 | 1267.62 | 1006.4211
5189 | 101.5674 | 0.98972 | 1.4433 | 0.99132 | 198.2452 | 50.4394 | 77433.7879 | 77362.2941 | 1543.2459 | 1443.2459 | 1006.4414
5190 | 116.3975 | 0.94498 | 1.0956 | 1.1069 | 191.0935 | 46.8501 | 62161.7351 | 61450.2744 | 1225.0055 | 1125.0055 | 1006.5048
5191 | 104.7659 | 1.004 | 0.9753 | 0.98502 | 194.0112 | 50.9593 | 54510.4016 | 53898.612 | 1073.9722 | 973.9722 | 1006.57
5192 | 109.3979 | 1.0051 | 1.267 | 1.0805 | 215.6546 | 50.7606 | 75288.7527 | 75218.9091 | 1500.3782 | 1400.3782 | 1006.6553
5193 | 101.689 | 1.0981 | 1.062 | 1.0161 | 201.155 | 50.7098 | 60585.4184 | 62832.907 | 1252.6581 | 1152.6581 | 1006.744
5194 | 71.3194 | 0.99826 | 0.87047 | 1.0615 | 187.4445 | 49.5453 | 49496.2108 | 48986.3225 | 975.7264 | 875.7264 | 1006.8701
5195 | 91.911 | 0.98748 | 0.95058 | 0.99672 | 192.437 | 50.1136 | 52221.6047 | 52112.2298 | 1038.2446 | 938.2446 | 1006.8806
5196 | 105.9367 | 0.99322 | 0.89934 | 1.0135 | 190.8921 | 48.7732 | 49778.8467 | 52190.3086 | 1039.8062 | 939.8062 | 1006.895
5197 | 109.3846 | 1.0023 | 0.84844 | 1.0696 | 206.1493 | 47.9592 | 48987.6629 | 47692.4765 | 949.8495 | 849.8495 | 1006.9109
5198 | 92.9674 | 1.0206 | 0.88087 | 1.0575 | 181.9074 | 52.4185 | 53985.2878 | 52697.291 | 1049.9458 | 949.9458 | 1006.9303
5199 | 107.7104 | 1.0174 | 1.0883 | 0.99479 | 193.6012 | 51.2162 | 61254.9044 | 59678.8298 | 1189.5766 | 1089.5766 | 1007.0561
5200 | 102.6645 | 0.91731 | 0.72922 | 0.97413 | 191.3663 | 50.1027 | 40500.526 | 40004.925 | 796.0985 | 696.0985 | 1007.0578
5201 | 104.1521 | 0.93914 | 1.1427 | 1.1052 | 196.9548 | 52.0319 | 70996.077 | 72272.5976 | 1441.452 | 1341.452 | 1007.063
5202 | 92.4798 | 0.9971 | 0.84582 | 0.96182 | 203.7509 | 44.8234 | 40801.6666 | 40957.7288 | 815.1546 | 715.1546 | 1007.113
5203 | 94.5738 | 0.98214 | 1.0776 | 1.0212 | 209.0945 | 49.2679 | 58999.6828 | 59175.052 | 1179.501 | 1079.501 | 1007.1844
5204 | 104.578 | 1.0167 | 0.73887 | 0.9824 | 196.8345 | 48.6401 | 40674.8074 | 39842.1748 | 792.8435 | 692.8435 | 1007.2157
5205 | 105.5804 | 1.0052 | 0.98645 | 0.93763 | 204.9525 | 50.8236 | 52606.5309 | 53213.515 | 1060.2703 | 960.2703 | 1007.2183
5206 | 101.3067 | 1.0108 | 1.0198 | 1.109 | 200.344 | 50.8045 | 62856.4647 | 61141.1776 | 1218.8236 | 1118.8236 | 1007.2213
5207 | 84.6288 | 0.90444 | 0.95294 | 1.057 | 196.2471 | 47.3243 | 51485.8534 | 53223.2903 | 1060.4658 | 960.4658 | 1007.2679
5208 | 101.0303 | 1.005 | 0.76995 | 1.0169 | 212.8539 | 49.6829 | 44155.4849 | 42534.8476 | 846.697 | 746.697 | 1007.2929
5209 | 77.6624 | 1.0777 | 1.0703 | 1.1117 | 197.3915 | 51.6919 | 66030.2568 | 63995.219 | 1275.9044 | 1175.9044 | 1007.3594
5210 | 87.2584 | 1.0833 | 0.94522 | 0.97699 | 189.0414 | 48.273 | 49330.9448 | 48604.1544 | 968.0831 | 868.0831 | 1007.3786
5211 | 109.4928 | 1.0208 | 1.1328 | 1.0799 | 199.8077 | 47.7463 | 63943.5255 | 65358.6374 | 1303.1727 | 1203.1727 | 1007.3918
5212 | 92.2355 | 0.99003 | 0.84245 | 0.97171 | 211.5255 | 50.9113 | 46537.3963 | 45789.3502 | 911.787 | 811.787 | 1007.512
5213 | 99.4397 | 0.94172 | 1.1092 | 1.0295 | 204.4804 | 47.842 | 59318.5654 | 60042.0272 | 1196.8405 | 1096.8405 | 1007.5359
5214 | 83.8144 | 1.0095 | 1.0178 | 1.0248 | 201.2723 | 51.1499 | 57882.9714 | 57836.6713 | 1152.7334 | 1052.7334 | 1007.5577
5215 | 112.3158 | 0.97902 | 0.98227 | 1.0262 | 198.6237 | 48.277 | 54172.3001 | 53388.2842 | 1063.7657 | 963.7657 | 1007.5888
5216 | 113.6932 | 0.94655 | 1.0217 | 0.91309 | 221.5121 | 50.8689 | 53151.6736 | 53331.5468 | 1062.6309 | 962.6309 | 1007.5894
5217 | 88.4226 | 0.98561 | 0.93595 | 0.97634 | 186.99 | 46.9377 | 47169.1603 | 46716.4565 | 930.3291 | 830.3291 | 1007.6225
5218 | 96.0837 | 1.0857 | 1.2009 | 0.93318 | 195.4403 | 46.7294 | 57436.9125 | 55971.2526 | 1115.4251 | 1015.4251 | 1007.654
5219 | 106.0822 | 0.9824 | 0.83605 | 0.99889 | 211.4861 | 49.2949 | 46515.8494 | 47308.2298 | 942.1646 | 842.1646 | 1007.6949
5220 | 100.0008 | 0.96822 | 1.344 | 0.97543 | 198.6972 | 54.043 | 76281.2222 | 75403.3199 | 1504.0664 | 1404.0664 | 1007.7687
5221 | 96.1843 | 1.0745 | 1.0932 | 0.96974 | 192.7506 | 52.133 | 60846.4579 | 62086.2244 | 1237.7245 | 1137.7245 | 1007.7896
5222 | 80.6997 | 1.0181 | 0.94413 | 1.0994 | 203.2143 | 47.8263 | 53776.8197 | 54548.5506 | 1086.971 | 986.971 | 1007.8475
5223 | 96.677 | 0.99595 | 0.95733 | 1.0335 | 200.4105 | 51.9286 | 56576.4648 | 57065.5463 | 1137.3109 | 1037.3109 | 1007.8483
5224 | 101.1202 | 1.0264 | 0.86341 | 0.98893 | 185.3641 | 48.756 | 46875.9653 | 45891.8008 | 913.836 | 813.836 | 1007.9595
5225 | 89.8444 | 1.0481 | 0.9092 | 0.98848 | 197.1781 | 51.6444 | 51474.2053 | 50573.7475 | 1007.475 | 907.475 | 1007.9856
5226 | 105.2876 | 0.99275 | 1.1425 | 0.96299 | 202.5975 | 49.0099 | 59247.2038 | 58272.0533 | 1161.4411 | 1061.4411 | 1007.9915
5227 | 93.1506 | 1.0648 | 1.1564 | 1.0372 | 198.3967 | 49.002 | 63833.5392 | 62692.7737 | 1249.8555 | 1149.8555 | 1008.0051
5228 | 115.0629 | 0.95063 | 1.0486 | 0.93902 | 186.7719 | 49.7938 | 54664.2051 | 53325.3318 | 1062.5066 | 962.5066 | 1008.0054
5229 | 128.4908 | 1.0509 | 0.85202 | 0.97271 | 195.8434 | 49.653 | 48051.4941 | 47127.5014 | 938.55 | 838.55 | 1008.0884
5230 | 119.2788 | 0.95702 | 0.88714 | 0.95921 | 204.5969 | 45.279 | 43903.7124 | 42626.3513 | 848.527 | 748.527 | 1008.1055
5231 | 108.5044 | 1.0412 | 0.9858 | 1.043 | 202.9938 | 49.5077 | 56698.5971 | 55488.6112 | 1105.7722 | 1005.7722 | 1008.1257
5232 | 113.6046 | 0.9081 | 0.71853 | 0.97456 | 186.7943 | 49.7514 | 40157.4414 | 41180.3053 | 819.6061 | 719.6061 | 1008.1265
5233 | 107.7266 | 0.89859 | 1.0811 | 1.0783 | 195.7318 | 48.8917 | 61926.3894 | 62253.3839 | 1241.0677 | 1141.0677 | 1008.1314
5234 | 98.8445 | 0.97803 | 0.97305 | 0.90843 | 185.0626 | 50.6657 | 49869.1427 | 49154.075 | 979.0815 | 879.0815 | 1008.2648
5235 | 109.1659 | 1.0201 | 0.87038 | 0.96729 | 217.4879 | 48.6463 | 46590.8655 | 47963.9043 | 955.2781 | 855.2781 | 1008.3345
5236 | 106.235 | 0.91115 | 0.99031 | 1.0167 | 189.7531 | 47.0641 | 52132.7414 | 51555.2441 | 1027.1049 | 927.1049 | 1008.4096
5237 | 81.659 | 0.99044 | 0.96321 | 0.96109 | 213.0819 | 48.6856 | 49220.6928 | 48249.1082 | 960.9822 | 860.9822 | 1008.4512
5238 | 94.1417 | 1.0387 | 1.04 | 0.87948 | 195.8271 | 49.3569 | 50168.7487 | 50202.9694 | 1000.0594 | 900.0594 | 1008.4941
5239 | 122.6278 | 0.98793 | 1.0859 | 1.0229 | 191.3303 | 48.569 | 60023.1994 | 61282.134 | 1221.6427 | 1121.6427 | 1008.5898
5240 | 86.5619 | 0.95351 | 0.87196 | 0.98582 | 182.7261 | 52.9659 | 50083.6802 | 50279.1226 | 1001.5825 | 901.5825 | 1008.5908
5241 | 100.0234 | 1.0545 | 1.0458 | 0.99745 | 180.9061 | 47.1186 | 54299.787 | 57120.0827 | 1138.4017 | 1038.4017 | 1008.6001
5242 | 105.9305 | 0.96135 | 1.1379 | 1.1284 | 178.6878 | 47.856 | 66498.1088 | 67073.9242 | 1337.4785 | 1237.4785 | 1008.6268
5243 | 104.5962 | 0.96879 | 0.87699 | 1.0338 | 197.2315 | 47.0395 | 47610.4331 | 47545.2502 | 946.905 | 846.905 | 1008.6482
5244 | 100.1963 | 0.92693 | 1.0452 | 1.0736 | 192.3276 | 43.7757 | 53383.434 | 53560.5523 | 1067.211 | 967.211 | 1008.8218
5245 | 105.062 | 0.9375 | 0.99436 | 0.90118 | 200.8307 | 50.0392 | 49969.2153 | 50112.3469 | 998.2469 | 898.2469 | 1008.8637
5246 | 79.3449 | 1.0434 | 1.2325 | 0.94006 | 187.7968 | 49.1846 | 61247.3213 | 62261.6902 | 1241.2338 | 1141.2338 | 1008.9652
5247 | 100.666 | 1.0233 | 1.1363 | 0.9904 | 219.6015 | 49.5527 | 61089.7467 | 62157.5821 | 1239.1516 | 1139.1516 | 1008.9759
5248 | 93.3153 | 0.95889 | 0.79004 | 1.0182 | 210.3086 | 52.8429 | 47446.2627 | 46197.5611 | 919.9512 | 819.9512 | 1008.9849
5249 | 112.1779 | 0.92413 | 0.71379 | 1.007 | 198.2995 | 49.2914 | 40739.5131 | 40963.3336 | 815.2667 | 715.2667 | 1009.0674
5250 | 106.8789 | 0.95163 | 0.57805 | 0.98607 | 207.9125 | 55.2084 | 37291.691 | 36867.7636 | 733.3553 | 633.3553 | 1009.0837
5251 | 101.1754 | 1.0287 | 0.92697 | 1.0069 | 199.6374 | 47.3574 | 49328.3474 | 50286.2482 | 1001.725 | 901.725 | 1009.1382
5252 | 93.0426 | 0.98009 | 1.1231 | 0.94854 | 203.3005 | 52.3212 | 60713.3109 | 59927.6522 | 1194.553 | 1094.553 | 1009.1449
5253 | 94.2991 | 1.0451 | 1.1588 | 0.95086 | 201.3005 | 52.4158 | 63123.0174 | 62197.8936 | 1239.9579 | 1139.9579 | 1009.1877
5254 | 111.7994 | 0.96499 | 0.83918 | 0.95453 | 202.533 | 49.032 | 44768.0479 | 46372.0643 | 923.4413 | 823.4413 | 1009.1921
5255 | 103.4995 | 0.97469 | 1.3288 | 0.92347 | 186.7323 | 48.3203 | 64355.8667 | 64374.7367 | 1283.4947 | 1183.4947 | 1009.2057
5256 | 85.7097 | 0.99113 | 1.0549 | 0.95516 | 201.7155 | 52.1306 | 57158.7925 | 56126.1132 | 1118.5223 | 1018.5223 | 1009.2221
5257 | 108.991 | 0.92329 | 1.0735 | 0.91968 | 207.9225 | 54.0273 | 58984.3028 | 56626.0026 | 1128.5201 | 1028.5201 | 1009.2758
5258 | 100.2726 | 0.95383 | 1.0689 | 1.0087 | 189.9035 | 43.6701 | 51450.3885 | 50185.2308 | 999.7046 | 899.7046 | 1009.2822
5259 | 96.4889 | 1.0077 | 1.144 | 0.97255 | 203.6694 | 48.9178 | 59384.7311 | 58223.0941 | 1160.4619 | 1060.4619 | 1009.2935
5260 | 87.3421 | 1.0402 | 0.91739 | 1.0742 | 209.3691 | 54.0123 | 58344.6591 | 60321.8844 | 1202.4377 | 1102.4377 | 1009.3808
5261 | 92.0705 | 1.0262 | 0.94247 | 1.0578 | 206.4313 | 50.9984 | 55868.9869 | 56404.4389 | 1124.0888 | 1024.0888 | 1009.4111
5262 | 99.8953 | 1.0534 | 0.83204 | 0.99087 | 195.3233 | 48.1975 | 45003.393 | 46298.2269 | 921.9645 | 821.9645 | 1009.4145
5263 | 109.3065 | 1.0444 | 1.0466 | 1.022 | 180.1733 | 51.6978 | 61379.9914 | 60784.1434 | 1211.6829 | 1111.6829 | 1009.5701
5264 | 107.0858 | 0.92049 | 1.0537 | 0.97396 | 191.8431 | 50.9603 | 57513.0609 | 57231.2276 | 1140.6246 | 1040.6246 | 1009.6012
5265 | 92.9151 | 1.0293 | 1.1724 | 0.8969 | 203.5251 | 47.584 | 54790.1829 | 56506.6487 | 1126.133 | 1026.133 | 1009.6056
5266 | 109.793 | 0.98395 | 0.99701 | 1.0079 | 183.4398 | 45.6996 | 51044.082 | 50772.3485 | 1011.447 | 911.447 | 1009.6354
5267 | 109.0911 | 0.89195 | 1.0117 | 0.9418 | 195.4395 | 50.2876 | 53002.9019 | 52408.9573 | 1044.1791 | 944.1791 | 1009.7769
5268 | 111.4988 | 0.97998 | 1.1211 | 0.9401 | 198.5549 | 53.4946 | 62424.114 | 62913.9858 | 1254.2797 | 1154.2797 | 1009.8803
5269 | 92.319 | 1.0106 | 0.97403 | 1.0382 | 192.9256 | 51.0018 | 56528.0541 | 57191.8188 | 1139.8364 | 1039.8364 | 1009.909
5270 | 109.5983 | 0.98242 | 1.1083 | 0.95205 | 201.6937 | 53.402 | 62299.2473 | 63443.3575 | 1264.8672 | 1164.8672 | 1009.9388
5271 | 99.5979 | 0.93709 | 0.91788 | 1.0132 | 213.9699 | 48.0859 | 49420.2175 | 50295.5099 | 1001.9102 | 901.9102 | 1010.0241
5272 | 87.3151 | 0.9358 | 1.0218 | 0.96185 | 189.5974 | 51.2567 | 54751.4847 | 54865.1145 | 1093.3023 | 993.3023 | 1010.0415
5273 | 106.3398 | 1.0699 | 0.79718 | 0.98141 | 201.194 | 57.2471 | 51502.7448 | 52248.5755 | 1040.9715 | 940.9715 | 1010.0623
5274 | 95.5093 | 0.94176 | 0.93843 | 1.0107 | 184.3192 | 49.4092 | 51490.8342 | 53055.0423 | 1057.1008 | 957.1008 | 1010.0801
5275 | 110.9673 | 1.0182 | 1.2656 | 0.93006 | 203.3942 | 46.1261 | 59709.287 | 58728.8326 | 1170.5767 | 1070.5767 | 1010.1074
5276 | 92.3203 | 0.98623 | 1.2157 | 0.99923 | 187.5252 | 49.6898 | 65073.4913 | 64224.9168 | 1280.4983 | 1180.4983 | 1010.1299
5277 | 100.7658 | 0.98416 | 1.1066 | 1.0166 | 182.5124 | 46.2614 | 56812.3128 | 57994.3686 | 1155.8874 | 1055.8874 | 1010.1502
5278 | 98.3177 | 0.90151 | 0.95104 | 1.0864 | 197.2006 | 48.207 | 54275.9619 | 55143.7195 | 1098.8744 | 998.8744 | 1010.4767
5279 | 80.3494 | 0.98411 | 1.1268 | 0.97146 | 189.1443 | 49.1905 | 57925.0016 | 55510.905 | 1106.2181 | 1006.2181 | 1010.5754
5280 | 116.4136 | 1.0192 | 1.2734 | 1.0342 | 183.4519 | 48.3191 | 69552.1325 | 69208.7247 | 1380.1745 | 1280.1745 | 1010.5803
5281 | 118.8318 | 1.0198 | 1.1374 | 0.95101 | 197.8061 | 49.793 | 60089.9582 | 62909.4838 | 1254.1897 | 1154.1897 | 1010.6674
5282 | 100.4253 | 0.9706 | 0.87361 | 0.95353 | 199.8217 | 50.1448 | 46858.8255 | 45983.731 | 915.6746 | 815.6746 | 1010.6734
5283 | 94.9936 | 1.0335 | 0.88452 | 1.0416 | 205.6892 | 49.3143 | 50482.1215 | 50791.8118 | 1011.8362 | 911.8362 | 1010.7029
5284 | 85.7078 | 1.0829 | 0.91565 | 1.042 | 218.7194 | 47.2451 | 49682.3564 | 50488.4796 | 1005.7696 | 905.7696 | 1010.7066
5285 | 99.1225 | 1.0579 | 1.2444 | 1.0479 | 207.8428 | 49.6187 | 70112.8423 | 69807.1348 | 1392.1427 | 1292.1427 | 1010.7187
5286 | 92.7962 | 0.99208 | 0.95821 | 1.0209 | 206.276 | 51.0406 | 54834.2889 | 54311.3713 | 1082.2274 | 982.2274 | 1010.7383
5287 | 84.8669 | 1.0066 | 0.66881 | 1.0842 | 205.1924 | 50.3982 | 41055.9381 | 40127.1656 | 798.5433 | 698.5433 | 1010.8065
5288 | 106.4574 | 1.0678 | 0.93466 | 1.0364 | 196.3237 | 48.6729 | 52879.1249 | 53754.6599 | 1071.0932 | 971.0932 | 1010.8865
5289 | 84.4963 | 0.90936 | 1.0698 | 1.0244 | 215.1263 | 51.8186 | 60988.589 | 61230.6127 | 1220.6123 | 1120.6123 | 1010.909
5290 | 81.4545 | 0.99724 | 0.92231 | 0.97605 | 202.3458 | 46.8113 | 46145.3726 | 46095.7889 | 917.9158 | 817.9158 | 1010.9861
5291 | 104.6795 | 0.94671 | 1.098 | 0.98207 | 187.5037 | 50.3077 | 59422.0967 | 58745.3607 | 1170.9072 | 1070.9072 | 1011.0403
5292 | 103.4541 | 0.97538 | 0.73904 | 1.0172 | 211.8558 | 49.7444 | 42627.0878 | 42502.1362 | 846.0427 | 746.0427 | 1011.077
5293 | 106.3965 | 1.022 | 0.73933 | 0.88443 | 201.9324 | 53.0296 | 40643.228 | 40658.7279 | 809.1746 | 709.1746 | 1011.1214
5294 | 79.9968 | 0.93964 | 1.0626 | 1.0349 | 192.3824 | 47.0026 | 55413.5241 | 55908.3963 | 1114.1679 | 1014.1679 | 1011.1796
5295 | 102.4335 | 0.98387 | 0.83977 | 0.96125 | 196.5036 | 45.8762 | 41852.5685 | 41295.1369 | 821.9027 | 721.9027 | 1011.2074
5296 | 89.3607 | 1.0119 | 1.2921 | 1.0418 | 205.197 | 51.7694 | 74570.8689 | 73333.8438 | 1462.6769 | 1362.6769 | 1011.2728
5297 | 93.3596 | 0.93581 | 0.78627 | 1.05 | 205.1805 | 50.976 | 46744.1322 | 47927.9236 | 954.5585 | 854.5585 | 1011.2818
5298 | 107.6766 | 1.0663 | 1.0765 | 1.0026 | 187.9492 | 50.9013 | 60966.8233 | 62570.9165 | 1247.4183 | 1147.4183 | 1011.3086
5299 | 86.245 | 0.98384 | 0.95869 | 0.90469 | 210.9362 | 52.1233 | 49840.9811 | 50587.0165 | 1007.7403 | 907.7403 | 1011.4705
5300 | 110.9784 | 1.0103 | 0.8574 | 1.025 | 206.705 | 48.8637 | 48627.0424 | 48076.4309 | 957.5286 | 857.5286 | 1011.5396
5301 | 94.9472 | 0.97712 | 1.2645 | 0.9852 | 208.5637 | 49.9258 | 67036.9218 | 66575.3509 | 1327.507 | 1227.507 | 1011.6066
5302 | 116.8837 | 0.96932 | 0.81715 | 1.0041 | 199.9093 | 48.3078 | 45310.5338 | 46316.0847 | 922.3217 | 822.3217 | 1011.6246
5303 | 99.0831 | 1.1201 | 0.83717 | 0.99791 | 196.4455 | 52.6708 | 50044.3553 | 50212.7556 | 1000.2551 | 900.2551 | 1011.6938
5304 | 98.0224 | 0.98933 | 0.93661 | 0.9911 | 202.5637 | 48.3215 | 49744.2646 | 48001.2799 | 956.0256 | 856.0256 | 1011.6978
5305 | 89.3291 | 1.0037 | 1.0598 | 1.0072 | 203.0083 | 50.1512 | 58235.54 | 58754.82 | 1171.0964 | 1071.0964 | 1011.744
5306 | 100.6731 | 0.95965 | 1.1223 | 0.95949 | 204.9992 | 51.4576 | 60587.3797 | 61522.4971 | 1226.4499 | 1126.4499 | 1011.7457
5307 | 100.1876 | 1.0054 | 1.1085 | 1.0155 | 199.131 | 50.3568 | 61956.1455 | 62936.8134 | 1254.7363 | 1154.7363 | 1011.7692
5308 | 97.6965 | 1.0321 | 1.0032 | 0.98343 | 216.1971 | 48.2753 | 52711.1576 | 52517.6447 | 1046.3529 | 946.3529 | 1011.7772
5309 | 95.9868 | 0.95549 | 1.2581 | 1.0158 | 202.7176 | 53.5253 | 73514.8835 | 71660.0624 | 1429.2012 | 1329.2012 | 1011.9102
5310 | 104.9257 | 0.97697 | 1.0315 | 0.94251 | 205.4354 | 53.0156 | 57181.1765 | 56341.2035 | 1122.8241 | 1022.8241 | 1011.9487
5311 | 98.6935 | 0.98208 | 0.83669 | 1.01 | 206.3893 | 49.6411 | 46967.6266 | 47471.9152 | 945.4383 | 845.4383 | 1012.0066
5312 | 94.321 | 1.0107 | 0.96737 | 0.97897 | 206.6071 | 49.7721 | 52086.9644 | 50607.5111 | 1008.1502 | 908.1502 | 1012.0764
5313 | 91.392 | 0.93174 | 1.0912 | 0.92427 | 208.8059 | 51.9584 | 57035.6159 | 54946.1856 | 1094.9237 | 994.9237 | 1012.1339
5314 | 123.4696 | 1.0459 | 0.87539 | 0.90205 | 205.1365 | 48.2275 | 44515.4349 | 45234.5844 | 900.6917 | 800.6917 | 1012.2739
5315 | 101.3903 | 1.0111 | 1.0989 | 1.0197 | 209.7089 | 52.8969 | 64902.8709 | 66012.1563 | 1316.2431 | 1216.2431 | 1012.3929
5316 | 114.978 | 0.98866 | 1.0712 | 0.98417 | 212.851 | 43.6774 | 51225.0354 | 52848.198 | 1052.964 | 952.964 | 1012.4156
5317 | 105.8222 | 0.96195 | 0.92668 | 1.053 | 205.6932 | 51.4968 | 55700.3681 | 54624.5341 | 1088.4907 | 988.4907 | 1012.4778
5318 | 97.7028 | 0.94701 | 0.82124 | 0.91739 | 186.4595 | 49.8512 | 42356.7651 | 42077.0632 | 837.5413 | 737.5413 | 1012.5242
5319 | 98.647 | 0.97891 | 0.7352 | 1.087 | 203.7888 | 47.0822 | 42375.8751 | 40719.5212 | 810.3904 | 710.3904 | 1012.5719
5320 | 91.2502 | 1.0375 | 1.1623 | 0.97291 | 202.4966 | 48.6436 | 59811.9176 | 59418.2298 | 1184.3646 | 1084.3646 | 1012.6054
5321 | 108.1413 | 1.0358 | 1.1419 | 0.93839 | 204.4978 | 50.232 | 59656.8685 | 58834.1569 | 1172.6831 | 1072.6831 | 1012.6552
5322 | 101.6233 | 1.0028 | 0.98974 | 0.92153 | 190.6181 | 48.5905 | 49460.2798 | 50910.9794 | 1014.2196 | 914.2196 | 1012.723
5323 | 93.9316 | 1.0298 | 0.82493 | 1.0023 | 196.8601 | 50.3572 | 46706.8835 | 48130.5315 | 958.6106 | 858.6106 | 1012.7305
5324 | 103.9754 | 1.0029 | 1.1576 | 1.0356 | 195.958 | 55.0617 | 71946.3037 | 72731.428 | 1450.6286 | 1350.6286 | 1012.832
5325 | 93.8178 | 0.95205 | 0.8647 | 0.92742 | 211.9108 | 53.1309 | 47565.2189 | 47765.327 | 951.3065 | 851.3065 | 1012.8567
5326 | 110.317 | 1.0138 | 0.80823 | 0.98443 | 200.6841 | 50.0623 | 45631.6774 | 44896.5366 | 893.9307 | 793.9307 | 1012.857
5327 | 108.8948 | 1.0449 | 0.78173 | 0.99985 | 197.178 | 52.564 | 47262.3987 | 46633.7797 | 928.6756 | 828.6756 | 1012.864
5328 | 120.5332 | 1.0832 | 0.67963 | 0.98122 | 198.7575 | 53.8871 | 43169.6369 | 43381.4583 | 863.6292 | 763.6292 | 1012.9061
5329 | 116.0806 | 1.1228 | 0.97237 | 1.0291 | 192.1344 | 47.5672 | 53990.1889 | 53784.9091 | 1071.6982 | 971.6982 | 1012.9097
5330 | 76.6437 | 1.0491 | 1.1107 | 1.0394 | 197.2537 | 47.6303 | 59018.1519 | 58298.6042 | 1161.9721 | 1061.9721 | 1012.9234
5331 | 81.6988 | 0.96026 | 1.1502 | 0.95808 | 207.8285 | 51.555 | 61065.7299 | 58801.1639 | 1172.0233 | 1072.0233 | 1012.9352
5332 | 87.0042 | 1.0091 | 1.1202 | 0.98441 | 205.0312 | 48.5268 | 57975.574 | 55606.3235 | 1108.1265 | 1008.1265 | 1013.0334
5333 | 105.6677 | 1.0012 | 1.1722 | 0.97121 | 211.9935 | 49.6011 | 61929.7558 | 63690.7855 | 1269.8157 | 1169.8157 | 1013.0535
5334 | 108.4338 | 0.98352 | 0.7904 | 1.0298 | 208.7303 | 48.1968 | 44578.6069 | 44596.2205 | 887.9244 | 787.9244 | 1013.1321
5335 | 98.3021 | 1.0004 | 0.88028 | 1.0153 | 203.0688 | 49.7548 | 49563.93 | 49186.8969 | 979.7379 | 879.7379 | 1013.1914
5336 | 93.7296 | 1.0511 | 0.83762 | 0.93487 | 208.0502 | 51.4397 | 45556.6845 | 46418.3315 | 924.3666 | 824.3666 | 1013.1953
5337 | 98.0774 | 1.0497 | 1.1625 | 0.99726 | 192.6412 | 53.4016 | 67599.6993 | 65943.0935 | 1314.8619 | 1214.8619 | 1013.2605
5338 | 86.9829 | 1.0323 | 1.0325 | 1.0313 | 211.7375 | 51.8181 | 60041.6917 | 62002.0897 | 1236.0418 | 1136.0418 | 1013.2657
5339 | 84.1864 | 1.0038 | 1.0448 | 0.99281 | 205.0686 | 49.479 | 55711.4649 | 57073.4997 | 1137.47 | 1037.47 | 1013.292
5340 | 104.9976 | 1.0296 | 1.0424 | 0.94504 | 201.7171 | 50.9715 | 55922.1241 | 53733.2478 | 1070.665 | 970.665 | 1013.3425
5341 | 93.4457 | 1.0082 | 1.2074 | 0.90574 | 201.683 | 49.6826 | 59214.3797 | 59000.5187 | 1176.0104 | 1076.0104 | 1013.3803
5342 | 103.373 | 0.92974 | 1.1218 | 1.0461 | 212.8715 | 52.1011 | 66362.1141 | 66508.3771 | 1326.1675 | 1226.1675 | 1013.4184
5343 | 92.8628 | 0.96879 | 1.0084 | 0.94746 | 207.0606 | 52.8576 | 55465.454 | 58031.5337 | 1156.6307 | 1056.6307 | 1013.487
5344 | 104.8759 | 0.99287 | 1.0435 | 1.0201 | 202.3915 | 48.7848 | 57208.2664 | 58228.9347 | 1160.5787 | 1060.5787 | 1013.5448
5345 | 112.0364 | 1.0532 | 0.97739 | 1.0713 | 191.7094 | 53.6726 | 62724.1057 | 64300.1811 | 1282.0036 | 1182.0036 | 1013.5535
5346 | 99.9122 | 1.0154 | 1.0002 | 0.97315 | 206.7625 | 43.2246 | 46662.4289 | 46950.0622 | 935.0012 | 835.0012 | 1013.5635
5347 | 111.0918 | 1.0567 | 1.0083 | 1.1012 | 204.3301 | 54.1964 | 66743.8638 | 66925.8021 | 1334.516 | 1234.516 | 1013.5863
5348 | 96.6991 | 0.97467 | 0.8771 | 0.95481 | 200.858 | 48.1135 | 45028.6662 | 46034.2393 | 916.6848 | 816.6848 | 1013.5883
5349 | 106.1039 | 1.0516 | 1.0159 | 1.0261 | 184.7176 | 49.1491 | 56904.4867 | 56216.5783 | 1120.3316 | 1020.3316 | 1013.6001
5350 | 83.0661 | 1.0527 | 0.79962 | 1.0106 | 195.1418 | 44.3099 | 39875.3339 | 39551.0856 | 787.0217 | 687.0217 | 1013.717
5351 | 88.6797 | 0.983 | 1.4403 | 1.0422 | 199.9484 | 45.8047 | 72948.0156 | 72478.1601 | 1445.5632 | 1345.5632 | 1013.8005
5352 | 90.51 | 0.97403 | 1.0892 | 0.92789 | 185.9673 | 51.8792 | 57191.839 | 57512.387 | 1146.2477 | 1046.2477 | 1013.8156
5353 | 89.1887 | 1.0544 | 1.1997 | 0.95631 | 200.1382 | 49.1134 | 61165.7022 | 63222.5088 | 1260.4502 | 1160.4502 | 1013.9195
5354 | 96.8231 | 0.93003 | 0.963 | 0.97526 | 195.3476 | 52.1111 | 53829.3217 | 52277.1332 | 1041.5427 | 941.5427 | 1013.9469
5355 | 88.8293 | 1.0527 | 0.93531 | 0.93878 | 215.9885 | 46.6241 | 45513.6515 | 45848.2212 | 912.9644 | 812.9644 | 1013.9731
5356 | 96.0765 | 0.99386 | 1.2762 | 1.0037 | 195.822 | 53.8311 | 74291.9677 | 75332.7972 | 1502.6559 | 1402.6559 | 1014.1679
5357 | 100.2008 | 0.95583 | 1.2394 | 0.93097 | 214.5823 | 50.2339 | 62990.3234 | 62353.0685 | 1243.0614 | 1143.0614 | 1014.2195
5358 | 97.5426 | 0.9729 | 0.93582 | 0.93997 | 212.9785 | 52.5129 | 51389.0774 | 51054.7973 | 1017.0959 | 917.0959 | 1014.2208
5359 | 109.0582 | 1.0376 | 0.9076 | 0.9021 | 207.7924 | 53.7606 | 50307.1842 | 50513.4953 | 1006.2699 | 906.2699 | 1014.2305
5360 | 97.6199 | 0.93929 | 0.93235 | 0.92683 | 200.0223 | 49.4731 | 47487.6673 | 47888.5083 | 953.7702 | 853.7702 | 1014.2726
5361 | 99.1232 | 0.91841 | 0.80009 | 0.98276 | 205.7164 | 50.9165 | 44876.0271 | 44605.5411 | 888.1108 | 788.1108 | 1014.2846
5362 | 93.8842 | 0.98497 | 1.2783 | 1.0689 | 202.5478 | 53.1411 | 77731.6342 | 77533.0932 | 1546.6619 | 1446.6619 | 1014.3065
5363 | 110.0207 | 0.92897 | 0.90423 | 1.069 | 207.3106 | 53.4928 | 57380.8808 | 58821.6858 | 1172.4337 | 1072.4337 | 1014.3137
5364 | 106.2458 | 1.0101 | 0.99878 | 1.1088 | 196.8681 | 49.6249 | 60478.1517 | 61622.1932 | 1228.4439 | 1128.4439 | 1014.3216
5365 | 100.4338 | 1.0166 | 1.145 | 0.96797 | 203.262 | 50.6004 | 61451.1286 | 61734.1735 | 1230.6835 | 1130.6835 | 1014.3292
5366 | 110.1026 | 0.99824 | 0.9841 | 1.0016 | 187.4998 | 51.6839 | 56809.2379 | 56463.4913 | 1125.2698 | 1025.2698 | 1014.4053
5367 | 99.6554 | 1.0224 | 0.74543 | 1.0045 | 212.9019 | 46.5586 | 39819.9077 | 39378.0621 | 783.5612 | 683.5612 | 1014.4368
5368 | 103.5964 | 0.95619 | 1.1399 | 0.96805 | 206.1876 | 50.2786 | 60667.0913 | 60949.4549 | 1214.9891 | 1114.9891 | 1014.4379
5369 | 113.6517 | 0.97407 | 0.93958 | 0.99536 | 204.1339 | 47.8735 | 50276.3106 | 51548.0293 | 1026.9606 | 926.9606 | 1014.4519
5370 | 90.4005 | 1.0074 | 1.0093 | 1.0136 | 201.2621 | 51.1567 | 57194.9337 | 57102.4452 | 1138.0489 | 1038.0489 | 1014.4673
5371 | 107.896 | 0.92216 | 0.99076 | 0.93795 | 199.3888 | 52.4357 | 54144.6615 | 53540.7377 | 1066.8148 | 966.8148 | 1014.5539
5372 | 96.7627 | 1.0251 | 0.99498 | 0.98288 | 200.0015 | 47.2953 | 51143.9331 | 51610.5046 | 1028.2101 | 928.2101 | 1014.5661
5373 | 100.7773 | 0.92036 | 1.1553 | 1.0005 | 193.9715 | 48.4123 | 60645.6166 | 59790.9989 | 1191.82 | 1091.82 | 1014.5679
5374 | 103.6533 | 0.93544 | 0.73972 | 1.1029 | 187.5557 | 47.9213 | 43929.7518 | 43356.9461 | 863.1389 | 763.1389 | 1014.5916
5375 | 90.3662 | 0.98143 | 0.88504 | 0.97541 | 201.2688 | 48.8396 | 46694.901 | 47272.3546 | 941.4471 | 841.4471 | 1014.6541
5376 | 92.2471 | 1.0525 | 0.95772 | 0.99737 | 185.9599 | 54.655 | 57698.9063 | 57634.4063 | 1148.6881 | 1048.6881 | 1014.715
5377 | 113.4725 | 1.016 | 1.0754 | 1.0857 | 190.6183 | 50.852 | 65424.2378 | 64570.3524 | 1287.407 | 1187.407 | 1014.7403
5378 | 75.8436 | 0.99527 | 1.0369 | 0.98305 | 215.5674 | 50.8503 | 55887.0981 | 55758.9804 | 1111.1796 | 1011.1796 | 1014.8284
5379 | 92.3586 | 1.0073 | 0.97002 | 0.89541 | 184.5619 | 54.3876 | 52483.7601 | 52285.827 | 1041.7165 | 941.7165 | 1014.8566
5380 | 87.0182 | 0.97629 | 1.0776 | 1.0434 | 212.59 | 45.0075 | 54642.6541 | 55170.9546 | 1099.4191 | 999.4191 | 1014.9089
5381 | 98.4509 | 0.97977 | 0.88443 | 0.94071 | 199.9659 | 51.4558 | 47973.9937 | 47976.5187 | 955.5304 | 855.5304 | 1014.9618
5382 | 105.7681 | 0.93317 | 0.98503 | 1.0319 | 208.5483 | 49.2099 | 55085.0241 | 56632.5474 | 1128.6509 | 1028.6509 | 1014.9636
5383 | 101.116 | 1.0055 | 0.92669 | 1.0616 | 211.7191 | 48.8534 | 53238.7921 | 52355.6692 | 1043.1134 | 943.1134 | 1015.0448
5384 | 96.4785 | 0.94496 | 0.93375 | 0.93339 | 197.4989 | 46.3944 | 44862.6415 | 45909.5988 | 914.192 | 814.192 | 1015.0649
5385 | 86.2725 | 0.94778 | 1.3751 | 0.99548 | 193.752 | 48.0482 | 69893.2362 | 71443.8091 | 1424.8762 | 1324.8762 | 1015.0766
5386 | 116.3197 | 0.94178 | 0.76375 | 0.97702 | 202.2714 | 44.8555 | 38587.4291 | 40536.9971 | 806.7399 | 706.7399 | 1015.0813
5387 | 116.4597 | 1.034 | 1.0629 | 0.98245 | 191.4116 | 48.0871 | 56195.8846 | 56968.7162 | 1135.3743 | 1035.3743 | 1015.1312
5388 | 109.3567 | 0.97603 | 0.85776 | 0.85349 | 205.0553 | 52.3875 | 44148.8925 | 45053.7467 | 897.0749 | 797.0749 | 1015.1459
5389 | 89.4835 | 1.0261 | 1.0205 | 1.0631 | 191.5928 | 50.3516 | 59439.1107 | 58241.0891 | 1160.8218 | 1060.8218 | 1015.3172
5390 | 88.7402 | 0.97049 | 0.79033 | 0.99596 | 200.9015 | 52.4171 | 45974.8449 | 44951.0015 | 895.02 | 795.02 | 1015.3219
5391 | 130.8987 | 1.0739 | 0.90198 | 0.88357 | 192.1138 | 47.2863 | 44524.7921 | 43385.9192 | 863.7184 | 763.7184 | 1015.3347
5392 | 95.367 | 0.89436 | 0.98722 | 1.0171 | 180.3586 | 51.6033 | 56398.3429 | 57150.7954 | 1139.0159 | 1039.0159 | 1015.3952
5393 | 106.3619 | 0.9872 | 1.2268 | 0.9651 | 199.0794 | 50.8615 | 65759.9139 | 63481.1056 | 1265.6221 | 1165.6221 | 1015.4251
5394 | 108.6106 | 0.94507 | 0.95121 | 0.9447 | 212.3093 | 43.0567 | 43322.9504 | 42287.2344 | 841.7447 | 741.7447 | 1015.4636
5395 | 101.8955 | 1.0294 | 1.1237 | 0.9892 | 202.9511 | 50.7023 | 61880.7713 | 63242.7624 | 1260.8552 | 1160.8552 | 1015.48
5396 | 85.2534 | 0.97229 | 0.96401 | 1.0491 | 208.2183 | 56.2189 | 61726.4269 | 62292.7781 | 1241.8556 | 1141.8556 | 1015.4908
5397 | 83.1923 | 1.0292 | 1.0174 | 0.94199 | 202.609 | 47.5645 | 49860.1047 | 50821.9713 | 1012.4394 | 912.4394 | 1015.4934
5398 | 94.0469 | 1.0764 | 1.1757 | 0.92151 | 208.4717 | 52.8083 | 62767.7364 | 64617.4448 | 1288.3489 | 1188.3489 | 1015.5084
5399 | 90.7877 | 0.99352 | 0.96929 | 1.0346 | 211.101 | 49.6063 | 54432.2273 | 54882.1909 | 1093.6438 | 993.6438 | 1015.5192
5400 | 101.3342 | 0.99545 | 0.95031 | 1.0333 | 199.6287 | 45.9094 | 49912.9253 | 50435.2887 | 1004.7058 | 904.7058 | 1015.5975
5401 | 104.6939 | 1.0748 | 1.1337 | 0.93498 | 205.609 | 47.797 | 56247.7724 | 57302.8162 | 1142.0563 | 1042.0563 | 1015.7435
5402 | 100.4536 | 1.0076 | 0.91148 | 0.99108 | 187.5378 | 50.5913 | 51010.2631 | 51774.4282 | 1031.4886 | 931.4886 | 1015.7567
5403 | 105.2673 | 1.01 | 0.72558 | 1.0278 | 192.8743 | 49.7474 | 42580.0329 | 43023.5387 | 856.4708 | 756.4708 | 1015.893
5404 | 107.1469 | 0.98787 | 0.79038 | 1.0322 | 203.1956 | 54.1365 | 50097.9167 | 48795.4378 | 971.9088 | 871.9088 | 1015.9231
5405 | 85.9604 | 0.99852 | 1.0362 | 1.0151 | 186.9995 | 46.8393 | 53473.7862 | 52749.8581 | 1050.9972 | 950.9972 | 1015.9763
5406 | 89.9562 | 1.0107 | 1.0548 | 0.87919 | 196.6225 | 51.174 | 52308.3425 | 51615.5922 | 1028.3118 | 928.3118 | 1015.9894
5407 | 83.5598 | 0.9514 | 0.98447 | 1.0794 | 206.7442 | 52.0175 | 59620.467 | 59286.571 | 1181.7314 | 1081.7314 | 1016.055
5408 | 90.1613 | 0.99603 | 0.97394 | 0.9638 | 197.2464 | 49.1171 | 50713.6439 | 50142.7315 | 998.8546 | 898.8546 | 1016.077
5409 | 89.9406 | 1.0076 | 0.59379 | 1.0954 | 205.5405 | 48.8926 | 36437.0279 | 36021.8196 | 716.4364 | 616.4364 | 1016.2641
5410 | 112.1401 | 0.99475 | 1.0499 | 0.976 | 221.7942 | 50.9637 | 58128.2626 | 58578.091 | 1167.5618 | 1067.5618 | 1016.2737
5411 | 117.5941 | 1.1127 | 0.9397 | 1.0833 | 195.3812 | 54.015 | 62251.5857 | 64024.1771 | 1276.4835 | 1176.4835 | 1016.3001
5412 | 113.8942 | 0.99931 | 0.89743 | 1.0248 | 208.1523 | 52.5053 | 54470.4556 | 52836.303 | 1052.7261 | 952.7261 | 1016.3216
5413 | 101.6708 | 0.9426 | 1.0487 | 1.0083 | 197.3614 | 46.864 | 54244.9017 | 53668.5937 | 1069.3719 | 969.3719 | 1016.3798
5414 | 92.9714 | 0.98021 | 1.1571 | 1.0512 | 208.0182 | 45.0253 | 59075.2118 | 61703.2016 | 1230.064 | 1130.064 | 1016.4057
5415 | 95.3385 | 0.97912 | 1.1239 | 1.1317 | 198.3323 | 50.1874 | 68715.6215 | 66241.4297 | 1320.8286 | 1220.8286 | 1016.428
5416 | 93.4382 | 1.0211 | 0.87776 | 0.96432 | 210.3299 | 46.3702 | 43884.1419 | 42484.917 | 845.6983 | 745.6983 | 1016.447
5417 | 96.7953 | 0.96718 | 1.0076 | 0.96886 | 197.7575 | 47.8731 | 51412.4938 | 48715.814 | 970.3163 | 870.3163 | 1016.4579
5418 | 93.0784 | 0.91644 | 0.93951 | 1.0687 | 219.3057 | 50.5939 | 55332.2023 | 56112.5203 | 1118.2504 | 1018.2504 | 1016.5401
5419 | 100.4218 | 1.0401 | 0.78733 | 0.89862 | 179.0363 | 51.2717 | 41809.2781 | 41719.3253 | 830.3865 | 730.3865 | 1016.56
5420 | 96.6488 | 0.98211 | 0.82607 | 1.0177 | 190.6479 | 49.5859 | 46581.8043 | 45723.0066 | 910.4601 | 810.4601 | 1016.5826
5421 | 101.3869 | 0.95138 | 0.82278 | 1.0208 | 211.2493 | 50.5061 | 47504.7375 | 46862.0896 | 933.2418 | 833.2418 | 1016.5934
5422 | 104.9746 | 0.97924 | 0.99729 | 1.011 | 219.1281 | 51.4157 | 57343.1602 | 57080.2987 | 1137.606 | 1037.606 | 1016.6223
5423 | 114.7484 | 1.0266 | 0.8944 | 0.94346 | 186.7989 | 50.5828 | 48828.7903 | 47595.1425 | 947.9028 | 847.9028 | 1016.6364
5424 | 85.5636 | 1.0273 | 1.1343 | 0.85494 | 205.7952 | 50.6465 | 53773.3705 | 52756.9134 | 1051.1383 | 951.1383 | 1016.6765
5425 | 108.5263 | 0.96312 | 0.9956 | 0.97792 | 189.8923 | 46.6174 | 50449.8403 | 51384.0782 | 1023.6816 | 923.6816 | 1016.7126
5426 | 108.4403 | 1.0822 | 1.0793 | 1.0652 | 206.8205 | 49.4358 | 62841.8726 | 63242.6945 | 1260.8539 | 1160.8539 | 1016.7171
5427 | 96.4277 | 0.95949 | 0.85534 | 0.96407 | 203.0158 | 51.9192 | 47819.9101 | 47181.9515 | 939.639 | 839.639 | 1016.7541
5428 | 111.5465 | 0.98629 | 0.90529 | 0.93049 | 201.7382 | 51.1206 | 48887.8984 | 49528.1781 | 986.5636 | 886.5636 | 1016.8195
5429 | 108.1203 | 1.0376 | 0.90835 | 0.93836 | 204.8382 | 48.0242 | 46526.2409 | 45607.2249 | 908.1445 | 808.1445 | 1016.82
5430 | 106.1791 | 1.0108 | 0.93489 | 0.97658 | 193.0894 | 44.7093 | 45810.9852 | 47606.4947 | 948.1299 | 848.1299 | 1016.8546
5431 | 94.1952 | 0.97682 | 1.2002 | 0.99372 | 201.9392 | 52.9659 | 68247.5464 | 70639.791 | 1408.7958 | 1308.7958 | 1016.8752
5432 | 94.5235 | 0.92664 | 0.91414 | 1.0458 | 180.8587 | 54.8525 | 57425.1427 | 58826.2913 | 1172.5258 | 1072.5258 | 1016.9873
5433 | 88.3329 | 0.95235 | 1.1374 | 0.95807 | 209.0628 | 50.1532 | 59079.9081 | 60197.0435 | 1199.9409 | 1099.9409 | 1016.9899
5434 | 101.1726 | 1.0545 | 0.86124 | 0.8999 | 200.0864 | 52.1494 | 46180.9472 | 46530.7994 | 926.616 | 826.616 | 1017.0875
5435 | 89.7043 | 0.93443 | 0.99581 | 1.032 | 191.82 | 50.8034 | 56660.0694 | 57357.0709 | 1143.1414 | 1043.1414 | 1017.1648
5436 | 91.6194 | 1.0005 | 1.0942 | 0.99924 | 193.6315 | 48.7422 | 57955.2381 | 58571.5346 | 1167.4307 | 1067.4307 | 1017.2447
5437 | 103.3789 | 1.0164 | 0.8834 | 0.99026 | 205.01 | 50.0347 | 49232.1048 | 49010.6954 | 976.2139 | 876.2139 | 1017.2644
5438 | 92.7483 | 0.92587 | 1.0765 | 0.97228 | 194.955 | 48.4744 | 55093.7022 | 53435.9179 | 1064.7184 | 964.7184 | 1017.2702
5439 | 125.656 | 0.98443 | 1.1175 | 0.97834 | 193.1756 | 51.0128 | 62272.9319 | 63350.1469 | 1263.0029 | 1163.0029 | 1017.3299
5440 | 106.9578 | 0.98301 | 0.9676 | 0.95992 | 199.2246 | 46.5838 | 48364.9147 | 48495.0213 | 965.9004 | 865.9004 | 1017.3503
5441 | 101.7017 | 0.88829 | 0.98642 | 0.99851 | 229.3392 | 51.266 | 55355.2902 | 54977.6747 | 1095.5535 | 995.5535 | 1017.3805
5442 | 117.7167 | 1.0016 | 0.75862 | 0.99296 | 176.8961 | 53.9232 | 47153.5741 | 46633.8401 | 928.6768 | 828.6768 | 1017.4949
5443 | 118.8738 | 0.96289 | 0.95473 | 1.0125 | 196.1778 | 51.5872 | 55966.8725 | 56440.9268 | 1124.8185 | 1024.8185 | 1017.503
5444 | 95.7659 | 1.0655 | 1.0685 | 1.0221 | 197.402 | 43.9557 | 52687.0008 | 51930.8488 | 1034.617 | 934.617 | 1017.6129
5445 | 95.8304 | 1.0172 | 1.0648 | 0.97407 | 203.8368 | 54.2489 | 61758.7043 | 62016.9861 | 1236.3397 | 1136.3397 | 1017.639
5446 | 113.4229 | 0.97959 | 1.0336 | 1.0098 | 203.0087 | 47.6059 | 55179.8879 | 55641.0887 | 1108.8218 | 1008.8218 | 1017.6876
5447 | 100.1044 | 0.93364 | 0.88329 | 1.0854 | 201.244 | 52.1713 | 55092.9023 | 54604.7709 | 1088.0954 | 988.0954 | 1017.7821
5448 | 93.949 | 0.97069 | 0.96942 | 0.97781 | 229.5406 | 48.6532 | 50784.9686 | 49639.7224 | 988.7944 | 888.7944 | 1017.9452
5449 | 105.5992 | 1.0049 | 0.67597 | 1.0001 | 203.4533 | 48.1137 | 37835.0402 | 37391.5013 | 743.83 | 643.83 | 1017.9727
5450 | 104.4367 | 1.0047 | 1.0536 | 0.95879 | 206.616 | 53.1991 | 59528.7076 | 58222.6976 | 1160.454 | 1060.454 | 1018.0173
5451 | 94.906 | 0.99951 | 0.79072 | 1.0719 | 207.7547 | 44.1323 | 41799.2206 | 41763.1812 | 831.2636 | 731.2636 | 1018.0552
5452 | 112.3796 | 0.99046 | 1.0716 | 0.98523 | 201.325 | 52.0064 | 60895.2922 | 61010.9539 | 1216.2191 | 1116.2191 | 1018.0793
5453 | 98.7347 | 0.9658 | 0.90455 | 1.045 | 202.3657 | 47.69 | 49830.6785 | 47347.0548 | 942.9411 | 842.9411 | 1018.0856
5454 | 99.9407 | 0.96311 | 1.0144 | 1.0949 | 198.6732 | 53.7224 | 65036.59 | 65131.5426 | 1298.6309 | 1198.6309 | 1018.0881
5455 | 110.212 | 0.97694 | 0.71975 | 0.95867 | 200.1005 | 48.9983 | 39284.5489 | 39813.6528 | 792.2731 | 692.2731 | 1018.2218
5456 | 82.573 | 0.97357 | 1.0053 | 0.94002 | 202.9226 | 48.5141 | 49950.0506 | 50099.5808 | 997.9916 | 897.9916 | 1018.2401
5457 | 94.2375 | 0.98882 | 0.90858 | 1.0353 | 205.9959 | 53.7569 | 55782.3227 | 55948.0876 | 1114.9618 | 1014.9618 | 1018.2448
5458 | 111.9753 | 0.98346 | 1.0063 | 1.0259 | 193.08 | 49.0687 | 56251.7255 | 56987.3867 | 1135.7477 | 1035.7477 | 1018.2504
5459 | 105.9066 | 1.0382 | 1.0076 | 0.93446 | 186.1964 | 49.2331 | 51954.8786 | 53156.8145 | 1059.1363 | 959.1363 | 1018.2956
5460 | 98.4209 | 0.98575 | 1.1449 | 1.0623 | 201.6197 | 48.6257 | 64059.1428 | 63885.5571 | 1273.7111 | 1173.7111 | 1018.3001
5461 | 92.8026 | 0.94538 | 0.85781 | 1.01 | 188.5266 | 45.9108 | 43994.4739 | 44968.2429 | 895.3649 | 795.3649 | 1018.3126
5462 | 105.0886 | 0.98463 | 1.021 | 1.0077 | 175.218 | 49.6174 | 56354.5618 | 54431.5624 | 1084.6312 | 984.6312 | 1018.3202
5463 | 91.5684 | 0.96998 | 0.82808 | 1.0685 | 209.4353 | 51.5282 | 50377.2934 | 51768.7617 | 1031.3752 | 931.3752 | 1018.3568
5464 | 90.0423 | 0.91226 | 1.0614 | 1.077 | 188.6177 | 48.2194 | 59270.1146 | 58570.5828 | 1167.4117 | 1067.4117 | 1018.3983
5465 | 119.4581 | 1.0285 | 1.2555 | 0.97746 | 203.1433 | 50.1328 | 67885.3763 | 66080.838 | 1317.6168 | 1217.6168 | 1018.4184
5466 | 109.5891 | 0.95246 | 0.97988 | 1.0079 | 205.8588 | 53.2931 | 58403.3197 | 58823.5879 | 1172.4718 | 1072.4718 | 1018.5223
5467 | 100.9166 | 0.98833 | 1.1395 | 0.97502 | 199.8403 | 47.4308 | 57627.5638 | 58506.3365 | 1166.1267 | 1066.1267 | 1018.5537
5468 | 102.3005 | 0.93044 | 0.73814 | 0.90237 | 188.8814 | 53.4035 | 40843.0254 | 41470.3888 | 825.4078 | 725.4078 | 1018.5673
5469 | 105.0528 | 0.94667 | 0.96527 | 0.96443 | 190.8972 | 47.5445 | 49180.0267 | 49580.1384 | 987.6028 | 887.6028 | 1018.6145
5470 | 97.1059 | 0.93257 | 1.1159 | 1.1011 | 189.7476 | 48.6771 | 64410.162 | 63324.8542 | 1262.4971 | 1162.4971 | 1018.6617
5471 | 103.4032 | 0.92457 | 0.95476 | 0.9646 | 188.1347 | 51.1879 | 52224.069 | 53129.4359 | 1058.5887 | 958.5887 | 1018.706
5472 | 104.6268 | 0.9265 | 1.1479 | 0.95228 | 203.2873 | 49.1481 | 58690.4378 | 58302.6316 | 1162.0526 | 1062.0526 | 1018.7334
5473 | 118.918 | 0.99877 | 0.73721 | 1.0727 | 220.4351 | 53.4092 | 48801.8261 | 50644.7319 | 1008.8946 | 908.8946 | 1018.7734
5474 | 91.1403 | 1.0266 | 0.93019 | 1.024 | 202.4423 | 51.1795 | 53741.745 | 52618.2144 | 1048.3643 | 948.3643 | 1018.7833
5475 | 98.6449 | 0.93158 | 0.8631 | 0.98013 | 190.2521 | 51.5416 | 48528.7078 | 47160.4601 | 939.2092 | 839.2092 | 1018.8035
5476 | 97.606 | 0.95018 | 1.0201 | 0.89737 | 186.5705 | 53.0437 | 53662.7996 | 55064.1095 | 1097.2822 | 997.2822 | 1018.907
5477 | 111.8917 | 1.078 | 0.80952 | 1.0334 | 203.046 | 50.6132 | 48647.4831 | 49394.9286 | 983.8986 | 883.8986 | 1018.9071
5478 | 96.8031 | 1.0273 | 1.0687 | 1.0213 | 204.6551 | 47.0708 | 56260.2551 | 56198.8593 | 1119.9772 | 1019.9772 | 1018.9943
5479 | 90.8339 | 1.0131 | 0.8165 | 0.95698 | 201.1308 | 45.2381 | 39712.0364 | 39646.5168 | 788.9303 | 688.9303 | 1019.0274
5480 | 90.0183 | 0.96999 | 1.0058 | 1.075 | 190.9092 | 48.676 | 57073.0978 | 57333.76 | 1142.6752 | 1042.6752 | 1019.0288
5481 | 93.5343 | 1.1035 | 1.0127 | 0.9744 | 179.4557 | 47.9644 | 52461.7714 | 50335.9615 | 1002.7192 | 902.7192 | 1019.1242
5482 | 104.7275 | 0.95177 | 1.1553 | 0.98982 | 203.2601 | 50.1665 | 62569.7518 | 62374.9218 | 1243.4984 | 1143.4984 | 1019.1436
5483 | 93.4263 | 1.077 | 1.1053 | 1.0709 | 203.8569 | 47.7238 | 61499.1837 | 63712.3082 | 1270.2462 | 1170.2462 | 1019.1469
5484 | 95.5485 | 0.94345 | 1.1222 | 1.0348 | 201.7176 | 51.0037 | 64030.531 | 63236.5477 | 1260.731 | 1160.731 | 1019.156
5485 | 100.5217 | 1.0131 | 1.0024 | 0.9424 | 198.5116 | 49.9738 | 52498.0866 | 51615.6969 | 1028.3139 | 928.3139 | 1019.1815
5486 | 100.4433 | 0.96375 | 1.2689 | 0.93769 | 181.0983 | 46.4443 | 59938.2945 | 58918.6165 | 1174.3723 | 1074.3723 | 1019.1951
5487 | 95.8738 | 0.9785 | 1.2352 | 0.98034 | 196.7664 | 50.2797 | 65799.1208 | 65006.1162 | 1296.1223 | 1196.1223 | 1019.2726
5488 | 104.5175 | 1.0269 | 0.9144 | 0.94538 | 201.8987 | 51.5262 | 50273.9596 | 49448.1314 | 984.9626 | 884.9626 | 1019.2912
5489 | 116.7914 | 0.99309 | 1.0048 | 0.9508 | 217.41 | 48.5761 | 52259.6745 | 51299.9679 | 1021.9994 | 921.9994 | 1019.3492
5490 | 104.8479 | 0.94473 | 0.90362 | 1.0138 | 186.7922 | 51.9651 | 52938.1171 | 51545.1894 | 1026.9038 | 926.9038 | 1019.3791
5491 | 102.4469 | 1.0414 | 1.0699 | 0.92639 | 212.5513 | 47.3178 | 52161.3802 | 51624.4658 | 1028.4893 | 928.4893 | 1019.4261
5492 | 112.56 | 0.90784 | 0.99395 | 1.0447 | 183.7394 | 55.5377 | 63527.1569 | 63172.0106 | 1259.4402 | 1159.4402 | 1019.4627
5493 | 99.1352 | 1.0312 | 1.0673 | 1.0225 | 190.3504 | 48.2689 | 57797.9597 | 59196.1463 | 1179.9229 | 1079.9229 | 1019.5187
5494 | 107.5168 | 0.99973 | 0.78785 | 0.96731 | 196.9045 | 49.9048 | 43593.3357 | 41676.826 | 829.5365 | 729.5365 | 1019.547
5495 | 111.9254 | 1.0309 | 0.95293 | 1.0225 | 211.3296 | 45.9559 | 50289.8892 | 50576.2416 | 1007.5248 | 907.5248 | 1019.5496
5496 | 93.384 | 1.0273 | 1.1777 | 0.9966 | 197.8789 | 53.6993 | 68375.5995 | 68283.1467 | 1361.6629 | 1261.6629 | 1019.5609
5497 | 107.5415 | 1.043 | 0.943 | 0.9364 | 188.9708 | 47.5369 | 47497.0983 | 47063.4105 | 937.2682 | 837.2682 | 1019.5913
5498 | 89.1274 | 0.98449 | 1.0309 | 1.0529 | 187.2513 | 48.6933 | 57315.5716 | 56835.1713 | 1132.7034 | 1032.7034 | 1019.685
5499 | 114.2961 | 0.97238 | 1.0187 | 1.0208 | 205.4421 | 47.4445 | 54818.215 | 54005.3107 | 1076.1062 | 976.1062 | 1019.7912
5500 | 111.3562 | 0.97182 | 0.84865 | 0.94751 | 206.5522 | 50.9548 | 46693.8567 | 47361.8421 | 943.2368 | 843.2368 | 1019.8074
5501 | 89.3244 | 1.0086 | 0.81894 | 0.98673 | 199.9273 | 46.6382 | 42088.6052 | 42179.5116 | 839.5902 | 739.5902 | 1019.8362
5502 | 106.9018 | 1.0071 | 1.1061 | 0.99683 | 184.8382 | 47.0098 | 57079.3306 | 56577.814 | 1127.5563 | 1027.5563 | 1019.8469
5503 | 105.6736 | 0.97862 | 1.1131 | 0.95651 | 196.4341 | 48.7003 | 57084.1531 | 55733.3718 | 1110.6674 | 1010.6674 | 1019.903
5504 | 107.7659 | 0.95048 | 1.2434 | 0.99153 | 216.3807 | 50.3696 | 67473.113 | 66824.4236 | 1332.4885 | 1232.4885 | 1019.9772
5505 | 95.7738 | 1.0096 | 0.87434 | 0.93073 | 199.6786 | 53.6469 | 49043.538 | 48773.0055 | 971.4601 | 871.4601 | 1020.0026
5506 | 81.8222 | 0.99526 | 1.084 | 0.97569 | 213.5755 | 47.8251 | 54689.9142 | 53914.8935 | 1074.2979 | 974.2979 | 1020.0771
5507 | 90.6124 | 1.0521 | 0.91243 | 0.98648 | 191.9498 | 56.0968 | 56031.7332 | 57396.3818 | 1143.9276 | 1043.9276 | 1020.098
5508 | 115.7564 | 0.97354 | 0.80252 | 0.99786 | 209.3605 | 53.1738 | 48783.6885 | 49215.4535 | 980.3091 | 880.3091 | 1020.1205
5509 | 90.7779 | 1.0494 | 1.0579 | 0.95758 | 204.6639 | 49.9934 | 55610.4453 | 56211.9165 | 1120.2383 | 1020.2383 | 1020.1555
5510 | 90.0502 | 0.93251 | 0.85115 | 1.0807 | 192.9561 | 49.5528 | 49933.648 | 49977.2949 | 995.5459 | 895.5459 | 1020.1843
5511 | 86.4991 | 1.0137 | 0.84607 | 1.0412 | 198.5655 | 51.4499 | 50033.6117 | 50620.6555 | 1008.4131 | 908.4131 | 1020.1856
5512 | 91.9049 | 0.9516 | 0.78661 | 0.9526 | 199.2959 | 47.8793 | 40263.6753 | 39357.5441 | 783.1509 | 683.1509 | 1020.2358
5513 | 105.0768 | 0.91617 | 0.92634 | 1.016 | 208.7734 | 51.7014 | 53847.0241 | 54400.2588 | 1084.0052 | 984.0052 | 1020.2383
5514 | 118.3507 | 1.0533 | 0.86727 | 1.0056 | 207.4327 | 53.4578 | 53493.922 | 53114.8482 | 1058.297 | 958.297 | 1020.2834
5515 | 95.3561 | 0.94112 | 1.1718 | 0.93086 | 197.8188 | 48.361 | 57291.0889 | 58374.25 | 1163.485 | 1063.485 | 1020.2907
5516 | 93.7268 | 1.0095 | 1.0021 | 0.98509 | 200.3414 | 45.3171 | 49221.7646 | 49853.2981 | 993.066 | 893.066 | 1020.3316
5517 | 86.3206 | 1.0128 | 0.76417 | 0.94716 | 210.0441 | 48.9094 | 39886.0671 | 40506.7829 | 806.1357 | 706.1357 | 1020.3632
5518 | 89.2165 | 0.97959 | 1.1096 | 1.0089 | 182.2244 | 51.5183 | 62357.3908 | 64182.2377 | 1279.6448 | 1179.6448 | 1020.4421
5519 | 94.3127 | 0.90297 | 0.75725 | 1.0544 | 197.3441 | 47.8575 | 42483.3287 | 43468.0877 | 865.3618 | 765.3618 | 1020.5026
5520 | 111.2902 | 0.93694 | 1.0735 | 1.0704 | 192.7067 | 48.7468 | 61294.003 | 61218.0556 | 1220.3611 | 1120.3611 | 1020.5189
5521 | 87.5245 | 0.9389 | 0.99122 | 0.98705 | 191.1267 | 46.9778 | 50013.852 | 50377.9359 | 1003.5587 | 903.5587 | 1020.5585
5522 | 110.8538 | 1.0108 | 1.0859 | 1.0099 | 198.6603 | 47.8729 | 58064.6579 | 59672.1303 | 1189.4426 | 1089.4426 | 1020.8115
5523 | 100.3664 | 1.0188 | 0.99118 | 1.0595 | 203.8965 | 52.227 | 60392.8137 | 60395.7706 | 1203.9154 | 1103.9154 | 1020.8353
5524 | 96.2901 | 0.96349 | 0.7788 | 1.0555 | 207.4611 | 47.3289 | 43504.6308 | 43883.5332 | 873.6707 | 773.6707 | 1020.907
5525 | 104.6682 | 1.0003 | 0.9872 | 0.98148 | 178.7198 | 51.296 | 55251.2825 | 54372.7352 | 1083.4547 | 983.4547 | 1020.9082
5526 | 88.0461 | 1.0038 | 0.84706 | 0.98025 | 200.5514 | 50.6498 | 46733.2926 | 48752.9459 | 971.0589 | 871.0589 | 1021.0337
5527 | 94.7722 | 0.93602 | 0.72882 | 1.0333 | 193.182 | 50.3785 | 42601.8591 | 41777.9713 | 831.5594 | 731.5594 | 1021.0996
5528 | 124.432 | 1.0092 | 1.1637 | 1.0046 | 197.2427 | 47.2832 | 61410.803 | 61601.5193 | 1228.0304 | 1128.0304 | 1021.1421
5529 | 90.656 | 0.98694 | 1.0886 | 1.0373 | 199.0496 | 50.4461 | 61676.1162 | 60773.6896 | 1211.4738 | 1111.4738 | 1021.1434
5530 | 81.4778 | 1.0528 | 1.0901 | 1.0129 | 186.5678 | 49.6194 | 59226.8093 | 60242.0415 | 1200.8408 | 1100.8408 | 1021.154
5531 | 117.3776 | 0.99685 | 1.2786 | 0.98822 | 184.1815 | 52.9353 | 73261.7993 | 71855.4898 | 1433.1098 | 1333.1098 | 1021.1598
5532 | 88.9503 | 0.99063 | 1.0594 | 1.0662 | 204.3065 | 51.3176 | 62694.9453 | 63286.0213 | 1261.7204 | 1161.7204 | 1021.1665
5533 | 87.568 | 1.0214 | 0.87122 | 1.0159 | 192.1188 | 48.3353 | 47296.7934 | 48834.95 | 972.699 | 872.699 | 1021.2276
5534 | 102.713 | 1.091 | 1.0527 | 1.1233 | 205.2558 | 46.7295 | 60696.906 | 62702.6992 | 1250.054 | 1150.054 | 1021.297
5535 | 106.8133 | 1.0377 | 1.0476 | 0.94107 | 193.3149 | 52.2074 | 57449.2853 | 57133.7086 | 1138.6742 | 1038.6742 | 1021.4057
5536 | 147.1274 | 0.97628 | 0.96028 | 0.99177 | 195.8999 | 48.9883 | 53887.328 | 51925.4377 | 1034.5088 | 934.5088 | 1021.4134
5537 | 90.8748 | 0.98526 | 0.91423 | 1.0189 | 205.0529 | 52.8299 | 54146.7188 | 53916.483 | 1074.3297 | 974.3297 | 1021.4951
5538 | 98.4061 | 1.0614 | 1.0108 | 1.0663 | 206.5375 | 55.5121 | 65837.8719 | 64878.8404 | 1293.5768 | 1193.5768 | 1021.5522
5539 | 94.9074 | 0.97969 | 0.93238 | 0.9842 | 225.8072 | 48.6339 | 49376.3781 | 49834.9707 | 992.6994 | 892.6994 | 1021.5871
5540 | 100.3103 | 1.0564 | 1.0225 | 0.93451 | 205.7238 | 53.7109 | 57221.9924 | 58450.6933 | 1165.0139 | 1065.0139 | 1021.6802
5541 | 111.6739 | 0.94268 | 0.9515 | 1.0156 | 212.5544 | 46.3287 | 49860.0421 | 49307.125 | 982.1425 | 882.1425 | 1021.683
5542 | 101.074 | 1.0514 | 1.0271 | 0.87783 | 191.3623 | 47.8336 | 48404.5811 | 47403.2683 | 944.0654 | 844.0654 | 1021.7006
5543 | 106.3575 | 0.91041 | 1.0749 | 1.0343 | 194.3492 | 50.7787 | 61564.5752 | 59557.8512 | 1187.157 | 1087.157 | 1021.7188
5544 | 104.6606 | 1.0115 | 0.95862 | 0.94608 | 212.9204 | 52.3286 | 53210.8581 | 52437.9456 | 1044.7589 | 944.7589 | 1021.7393
5545 | 87.7441 | 1.0835 | 1.0263 | 0.96833 | 195.9629 | 51.9624 | 56773.6241 | 56328.2863 | 1122.5657 | 1022.5657 | 1021.7496
5546 | 99.1963 | 0.94402 | 1.1287 | 0.96708 | 211.9211 | 48.8349 | 58091.3109 | 58045.5927 | 1156.9119 | 1056.9119 | 1021.766
5547 | 97.9306 | 0.94744 | 1.1301 | 0.94363 | 203.237 | 48.5711 | 56507.3941 | 56556.7565 | 1127.1351 | 1027.1351 | 1021.7887
5548 | 92.2214 | 0.92237 | 0.75731 | 1.0024 | 192.024 | 46.6849 | 39603.5854 | 40695.4553 | 809.9091 | 709.9091 | 1021.7985
5549 | 118.1707 | 0.91595 | 0.68439 | 1.0009 | 197.6626 | 53.9625 | 43003.6187 | 42998.3469 | 855.9669 | 755.9669 | 1021.904
5550 | 86.0923 | 1.0975 | 1.0543 | 1.0056 | 197.8838 | 54.4958 | 63125.0707 | 61440.0551 | 1224.8011 | 1124.8011 | 1021.9213
5551 | 97.0278 | 0.99671 | 1.0879 | 0.98515 | 189.2385 | 57.1926 | 67015.2197 | 65077.0428 | 1297.5409 | 1197.5409 | 1021.955
5552 | 96.37 | 0.95751 | 0.98465 | 0.95526 | 203.3188 | 53.4577 | 55418.2294 | 52601.5274 | 1048.0305 | 948.0305 | 1022.019
5553 | 114.9995 | 0.97086 | 0.8861 | 1.0251 | 184.8659 | 50.7042 | 51900.8659 | 53100.2574 | 1058.0051 | 958.0051 | 1022.0253
5554 | 95.7539 | 0.96774 | 0.71816 | 0.94437 | 204.0228 | 49.6839 | 38504.1119 | 39168.213 | 779.3643 | 679.3643 | 1022.0423
5555 | 108.8495 | 1.0959 | 1.0307 | 0.97669 | 193.6974 | 48.2007 | 54464.2972 | 52995.8924 | 1055.9178 | 955.9178 | 1022.078
5556 | 113.2984 | 1.0557 | 1.2205 | 1.0053 | 213.014 | 47.6772 | 64410.1611 | 64366.3115 | 1283.3262 | 1183.3262 | 1022.2351
5557 | 110.0155 | 1.0646 | 1.0873 | 1.0392 | 208.3357 | 53.8192 | 67319.1711 | 68177.3306 | 1359.5466 | 1259.5466 | 1022.2729
5558 | 108.1673 | 0.97652 | 0.73988 | 1.0497 | 196.8872 | 48.5168 | 43003.5194 | 44202.1579 | 880.0432 | 780.0432 | 1022.4329
5559 | 123.579 | 1.0275 | 0.94839 | 0.96728 | 198.3519 | 47.3238 | 49620.8162 | 49383.091 | 983.6618 | 883.6618 | 1022.4452
5560 | 104.0918 | 1.1089 | 1.0695 | 0.88596 | 198.9231 | 48.9353 | 52214.3659 | 51098.1316 | 1017.9626 | 917.9626 | 1022.4799
5561 | 105.6738 | 1.0949 | 0.92208 | 1.0685 | 215.0698 | 47.5587 | 52575.2817 | 49344.9894 | 982.8998 | 882.8998 | 1022.5072
5562 | 92.5752 | 0.98635 | 0.83239 | 1.1222 | 209.2543 | 48.0998 | 49531.2391 | 51232.6162 | 1020.6523 | 920.6523 | 1022.5657
5563 | 107.5578 | 1.119 | 0.89803 | 0.9747 | 192.3748 | 52.4845 | 52448.9365 | 50964.2054 | 1015.2841 | 915.2841 | 1022.5671
5564 | 106.823 | 1.1021 | 0.90276 | 0.96011 | 194.3864 | 47.5908 | 47046.3523 | 47023.3556 | 936.4671 | 836.4671 | 1022.7263
5565 | 120.5001 | 1.0586 | 1.1147 | 0.91981 | 202.0383 | 49.4656 | 57230.8663 | 58296.6024 | 1161.932 | 1061.932 | 1022.8209
5566 | 89.444 | 0.96546 | 1.1149 | 1.0235 | 205.0153 | 55.2075 | 67971.275 | 67188.5706 | 1339.7714 | 1239.7714 | 1022.8241
5567 | 109.5007 | 0.98002 | 1.0306 | 1.007 | 184.3691 | 53.1368 | 61032.5469 | 60372.6204 | 1203.4524 | 1103.4524 | 1022.8288
5568 | 107.511 | 0.9349 | 0.97563 | 1.011 | 212.5235 | 48.343 | 52753.8563 | 51303.9465 | 1022.0789 | 922.0789 | 1022.8302
5569 | 88.1832 | 1.1566 | 1.081 | 1.0147 | 205.8257 | 49.5557 | 59616.2885 | 59402.5859 | 1184.0517 | 1084.0517 | 1022.8744
5570 | 103.1822 | 0.99973 | 1.0044 | 0.93074 | 207.3441 | 47.2889 | 49291.0034 | 48047.4006 | 956.948 | 856.948 | 1022.8751
5571 | 127.3174 | 0.93334 | 0.96468 | 1.0121 | 211.7543 | 50.8881 | 55944.7554 | 57230.3035 | 1140.6061 | 1040.6061 | 1022.8808
5572 | 99.2673 | 0.93099 | 0.99656 | 1.0494 | 183.192 | 51.4468 | 58738.3133 | 55937.0132 | 1114.7403 | 1014.7403 | 1022.8811
5573 | 77.2369 | 0.98765 | 0.86187 | 0.86582 | 220.3177 | 51.1205 | 42267.6105 | 41936.4068 | 834.7281 | 734.7281 | 1022.973
5574 | 104.7568 | 0.9765 | 0.87119 | 1.0171 | 195.7911 | 50.2728 | 49882.571 | 49750.3793 | 991.0076 | 891.0076 | 1023.0013
5575 | 104.509 | 1.1047 | 1.0718 | 1.0149 | 200.7993 | 49.9604 | 60313.7791 | 60402.8248 | 1204.0565 | 1104.0565 | 1023.03
5576 | 96.0262 | 1.0437 | 0.93217 | 1.0023 | 206.4892 | 51.8406 | 53839.4192 | 53750.2033 | 1071.0041 | 971.0041 | 1023.0538
5577 | 88.9239 | 0.98698 | 0.97263 | 1.0175 | 197.6322 | 42.3315 | 45807.6804 | 45726.7289 | 910.5346 | 810.5346 | 1023.0969
5578 | 100.2353 | 0.91273 | 1.1069 | 0.97254 | 187.7663 | 52.6293 | 61660.6871 | 61916.5163 | 1234.3303 | 1134.3303 | 1023.1238
5579 | 83.245 | 0.96565 | 1.1224 | 1.04 | 201.549 | 50.8111 | 63602.0628 | 63588.8264 | 1267.7765 | 1167.7765 | 1023.1604
5580 | 94.8525 | 1.0257 | 1.0885 | 0.92011 | 196.2974 | 45.4579 | 50148.6249 | 50510.567 | 1006.2113 | 906.2113 | 1023.3269
5581 | 111.7873 | 0.911 | 0.81426 | 0.98236 | 188.4912 | 49.0626 | 44430.1824 | 43198.7496 | 859.975 | 759.975 | 1023.3811
5582 | 109.1272 | 1.0981 | 0.7655 | 1.0111 | 197.906 | 48.1435 | 43230.6011 | 44123.397 | 878.4679 | 778.4679 | 1023.4195
5583 | 99.7141 | 1.0549 | 0.98382 | 1.0885 | 183.9128 | 50.9105 | 60057.3404 | 60562.1264 | 1207.2425 | 1107.2425 | 1023.439
5584 | 92.1093 | 1.1112 | 1.1854 | 1.0266 | 209.9928 | 48.8788 | 64693.6268 | 63809.0355 | 1272.1807 | 1172.1807 | 1023.5182
5585 | 100.6477 | 0.9918 | 1.3046 | 1.0326 | 213.7549 | 55.6232 | 80702.0023 | 80106.7255 | 1598.1345 | 1498.1345 | 1023.5545
5586 | 113.8282 | 0.94787 | 0.98462 | 1.0142 | 198.9668 | 48.3375 | 53684.9888 | 53433.9394 | 1064.6788 | 964.6788 | 1023.5677
5587 | 107.7926 | 1.0078 | 0.98545 | 1.0244 | 199.2951 | 47.4018 | 53200.0935 | 54757.2923 | 1091.1458 | 991.1458 | 1023.5732
5588 | 110.1333 | 1.0195 | 0.92919 | 1.0507 | 189.0726 | 47.0947 | 51454.1189 | 51838.956 | 1032.7791 | 932.7791 | 1023.6257
5589 | 99.0279 | 1.0501 | 0.99758 | 0.99028 | 217.3443 | 52.2578 | 57276.2605 | 58353.7789 | 1163.0756 | 1063.0756 | 1023.7608
5590 | 80.5512 | 0.9898 | 1.0281 | 1.0641 | 204.3906 | 47.4928 | 55951.5062 | 54715.7427 | 1090.3149 | 990.3149 | 1023.8796
5591 | 104.559 | 0.96097 | 1.1741 | 1.0317 | 202.8553 | 51.2181 | 67389.1118 | 67195.9196 | 1339.9184 | 1239.9184 | 1024.0337
5592 | 124.078 | 0.97993 | 0.94994 | 1.0113 | 187.7735 | 49.2727 | 53513.3257 | 51053.0197 | 1017.0604 | 917.0604 | 1024.0888
5593 | 129.7335 | 0.92447 | 1.2942 | 0.97465 | 207.815 | 49.273 | 68272.1862 | 68545.8082 | 1366.9162 | 1266.9162 | 1024.1101
5594 | 92.4706 | 1.1025 | 1.1776 | 0.98349 | 195.2898 | 44.0608 | 55718.043 | 58055.067 | 1157.1013 | 1057.1013 | 1024.1624
5595 | 106.3092 | 1.0965 | 1.0283 | 1.0855 | 198.0902 | 50.6702 | 62662.3011 | 61905.052 | 1234.101 | 1134.101 | 1024.2912
5596 | 97.1109 | 0.92359 | 0.81856 | 1.0237 | 203.9182 | 50.3758 | 46934.0825 | 46385.3727 | 923.7075 | 823.7075 | 1024.2914
5597 | 80.2295 | 0.87325 | 0.92839 | 1.0511 | 201.7639 | 49.6369 | 52117.5126 | 53485.9935 | 1065.7199 | 965.7199 | 1024.2976
5598 | 105.1833 | 1.0569 | 1.0204 | 1.041 | 199.477 | 53.7614 | 63282.8399 | 64023.6215 | 1276.4724 | 1176.4724 | 1024.3105
5599 | 94.7134 | 0.93855 | 0.94953 | 1.0387 | 195.0248 | 55.2758 | 59623.7427 | 57139.8936 | 1138.7979 | 1038.7979 | 1024.4255
5600 | 91.9638 | 1.0858 | 0.8951 | 1.0306 | 220.5171 | 47.2138 | 48489.588 | 48555.4934 | 967.1099 | 867.1099 | 1024.4378
5601 | 81.6942 | 1.0676 | 0.85852 | 0.98233 | 200.6391 | 47.4896 | 44392.8487 | 43523.5328 | 866.4707 | 766.4707 | 1024.4537
5602 | 88.8151 | 1.0046 | 0.93349 | 1.0161 | 186.5023 | 51.5374 | 53667.5557 | 55496.129 | 1105.9226 | 1005.9226 | 1024.4832
5603 | 108.7567 | 0.97223 | 1.2723 | 0.99649 | 168.3636 | 47.042 | 64783.5783 | 66315.7766 | 1322.3155 | 1222.3155 | 1024.6408
5604 | 86.1493 | 1.0334 | 1.0539 | 1.0917 | 194.3581 | 48.0295 | 59730.8648 | 62725.7638 | 1250.5153 | 1150.5153 | 1024.7077
5605 | 107.8537 | 1.0564 | 1.0406 | 1.0631 | 192.4199 | 48.0802 | 58861.5645 | 58846.0767 | 1172.9215 | 1072.9215 | 1024.7522
5606 | 105.3372 | 0.99635 | 0.85905 | 1.0063 | 204.9745 | 47.9195 | 46657.6278 | 48644.2184 | 968.8844 | 868.8844 | 1024.8185
5607 | 111.1218 | 0.98347 | 0.95466 | 0.98954 | 189.4749 | 51.8319 | 54818.3165 | 54027.9816 | 1076.5596 | 976.5596 | 1024.8223
5608 | 102.5428 | 0.94544 | 1.0471 | 0.9699 | 206.1811 | 49.0398 | 54764.2062 | 55144.4528 | 1098.8891 | 998.8891 | 1024.8566
5609 | 97.1403 | 1.0799 | 0.80773 | 1.0103 | 194.9642 | 45.2021 | 41822.539 | 41068.0049 | 817.3601 | 717.3601 | 1024.9231
5610 | 90.7977 | 0.96307 | 1.1694 | 0.9824 | 202.4974 | 51.9382 | 64409.7351 | 64676.1705 | 1289.5234 | 1189.5234 | 1024.943
5611 | 108.2032 | 0.97288 | 0.73417 | 1.0116 | 209.5895 | 53.6426 | 45696.9664 | 46204.1109 | 920.0822 | 820.0822 | 1024.9619
5612 | 92.915 | 0.95883 | 0.98802 | 0.98132 | 192.5665 | 52.3546 | 55617.5557 | 52774.6302 | 1051.4926 | 951.4926 | 1025.0336
5613 | 95.3328 | 1.0024 | 1.0136 | 0.96593 | 184.7745 | 51.6307 | 55669.5515 | 56103.963 | 1118.0793 | 1018.0793 | 1025.0505
5614 | 103.6768 | 0.98171 | 1.109 | 0.93399 | 198.932 | 47.2672 | 53969.1628 | 55500.6609 | 1106.0132 | 1006.0132 | 1025.0727
5615 | 109.8192 | 1.0625 | 1.0414 | 0.98445 | 194.0712 | 50.1909 | 57505.6853 | 57094.9679 | 1137.8994 | 1037.8994 | 1025.0832
5616 | 101.1319 | 0.97543 | 0.8165 | 1.0642 | 199.4829 | 49.5833 | 48174.56 | 48412.9266 | 964.2585 | 864.2585 | 1025.1529
5617 | 95.6643 | 0.96155 | 1.266 | 0.94489 | 196.5947 | 50.3992 | 65120.0032 | 64305.9142 | 1282.1183 | 1182.1183 | 1025.21
5618 | 109.2048 | 0.98433 | 1.1299 | 1.015 | 193.3763 | 49.8266 | 62695.3027 | 62636.7315 | 1248.7346 | 1148.7346 | 1025.2698
5619 | 105.0691 | 0.94776 | 0.86204 | 1.0427 | 194.6711 | 50.8742 | 50990.9769 | 52247.599 | 1040.952 | 940.952 | 1025.3097
5620 | 81.446 | 1.0685 | 1.0067 | 0.99503 | 206.0296 | 53.3246 | 58262.3751 | 59351.399 | 1183.028 | 1083.028 | 1025.3131
5621 | 112.9087 | 1.0025 | 1.0912 | 1.0445 | 191.061 | 48.9934 | 61575.3591 | 61711.248 | 1230.225 | 1130.225 | 1025.3849
5622 | 107.4658 | 0.98717 | 0.98994 | 0.91912 | 180.0165 | 46.1495 | 47066.0166 | 46221.4723 | 920.4294 | 820.4294 | 1025.3901
5623 | 96.9642 | 1.0771 | 1.1302 | 1.0311 | 214.4584 | 46.3891 | 59120.3867 | 58243.8524 | 1160.877 | 1060.877 | 1025.4126
5624 | 110.3973 | 1.0548 | 0.83522 | 1.0285 | 197.6299 | 56.4453 | 55260.0822 | 55620.4785 | 1108.4096 | 1008.4096 | 1025.4623
5625 | 93.7177 | 0.96559 | 1.09 | 0.98521 | 195.4187 | 47.4477 | 55439.9287 | 54896.6282 | 1093.9326 | 993.9326 | 1025.4625
5626 | 95.2702 | 0.9767 | 1.1785 | 1.0522 | 187.4458 | 51.5553 | 68909.6741 | 67774.2987 | 1351.486 | 1251.486 | 1025.4723
5627 | 85.7678 | 0.95995 | 0.90575 | 1.0441 | 204.8502 | 47.4961 | 49033.5648 | 48932.5643 | 974.6513 | 874.6513 | 1025.527
5628 | 106.0036 | 0.9117 | 1.0932 | 1.0425 | 207.1701 | 53.8961 | 66840.7015 | 67686.781 | 1349.7356 | 1249.7356 | 1025.5508
5629 | 105.8035 | 1.0075 | 1.0265 | 1.0656 | 194.0402 | 50.4306 | 60735.4254 | 61005.2437 | 1216.1049 | 1116.1049 | 1025.6226
5630 | 95.1663 | 0.99704 | 1.1808 | 0.96999 | 201.6518 | 52.4258 | 65224.2125 | 66555.6486 | 1327.113 | 1227.113 | 1025.6257
5631 | 95.4788 | 1.0451 | 0.96182 | 1.0595 | 196.2802 | 48.5556 | 54520.795 | 55241.5748 | 1100.8315 | 1000.8315 | 1025.6706
5632 | 96.1843 | 0.935 | 1.0488 | 0.97916 | 208.0002 | 53.7832 | 60276.8464 | 58376.6768 | 1163.5335 | 1063.5335 | 1025.68
5633 | 98.6878 | 1.0159 | 0.95243 | 1.0293 | 211.2467 | 45.9551 | 49870.0074 | 51794.2326 | 1031.8847 | 931.8847 | 1025.7027
5634 | 95.1263 | 1.0365 | 0.89238 | 1.0341 | 222.0127 | 48.2416 | 49495.58 | 49800.2137 | 992.0043 | 892.0043 | 1025.7974
5635 | 121.7907 | 1.0145 | 0.83743 | 0.97058 | 209.4994 | 50.3074 | 47314.7549 | 47750.9173 | 951.0183 | 851.0183 | 1025.8037
5636 | 105.0095 | 0.95743 | 0.88383 | 1.0207 | 204.257 | 52.8283 | 53172.5096 | 52813.5626 | 1052.2713 | 952.2713 | 1025.8112
5637 | 94.5677 | 1.0609 | 1.0447 | 0.92834 | 197.2736 | 50.7178 | 54473.3859 | 52672.5296 | 1049.4506 | 949.4506 | 1025.8472
5638 | 90.8157 | 1.0282 | 0.82436 | 0.99417 | 200.9556 | 49.7635 | 45631.6035 | 45715.9585 | 910.3192 | 810.3192 | 1025.9141
5639 | 111.2879 | 0.9889 | 0.90117 | 1.0084 | 204.078 | 53.3457 | 54554.3942 | 53914.3653 | 1074.2873 | 974.2873 | 1025.9413
5640 | 94.0682 | 1.0924 | 1.0883 | 1.0231 | 187.69 | 47.4982 | 57957.404 | 56898.3606 | 1133.9672 | 1033.9672 | 1025.9864
5641 | 111.2736 | 1.0752 | 0.86041 | 0.99542 | 225.5838 | 46.0604 | 45185.6171 | 46224.4746 | 920.4895 | 820.4895 | 1025.9865
5642 | 102.7349 | 1.008 | 0.67545 | 0.97491 | 188.6405 | 49.5449 | 37944.7551 | 38254.7445 | 761.0949 | 661.0949 | 1025.9946
5643 | 99.4876 | 1.0348 | 1.2141 | 1.0364 | 210.755 | 45.8097 | 62567.7375 | 63626.3811 | 1268.5276 | 1168.5276 | 1025.998
5644 | 106.3912 | 0.90618 | 0.92587 | 0.97229 | 207.8993 | 49.3698 | 49410.7336 | 48907.8127 | 974.1563 | 874.1563 | 1026.0486
5645 | 84.9227 | 0.9653 | 1.0054 | 0.94967 | 204.836 | 48.9974 | 51004.5156 | 52194.4556 | 1039.8891 | 939.8891 | 1026.133
5646 | 109.3015 | 0.96925 | 1.0232 | 1.0193 | 205.8308 | 50.1126 | 57777.2683 | 58690.9138 | 1169.8183 | 1069.8183 | 1026.1507
5647 | 106.431 | 1.09 | 0.86362 | 1.0037 | 222.9416 | 48.4044 | 47797.4092 | 48342.1236 | 962.8425 | 862.8425 | 1026.3287
5648 | 96.0521 | 0.93606 | 0.93851 | 1.0019 | 201.0406 | 50.4044 | 52127.7239 | 52202.7751 | 1040.0555 | 940.0555 | 1026.3983
5649 | 96.6646 | 0.98237 | 0.99355 | 0.98885 | 199.5227 | 46.4404 | 50235.6526 | 49486.5129 | 985.7303 | 885.7303 | 1026.4349
5650 | 94.5597 | 0.90202 | 1.1658 | 0.98324 | 204.0779 | 49.623 | 61318.6145 | 63004.551 | 1256.091 | 1156.091 | 1026.5468
5651 | 82.1634 | 1.0693 | 0.99554 | 0.99314 | 190.5237 | 47.3101 | 51123.1283 | 51375.9042 | 1023.5181 | 923.5181 | 1026.5937
5652 | 97.12 | 0.8667 | 1.048 | 1.048 | 183.0485 | 55.6321 | 65964.9984 | 66489.3407 | 1325.7868 | 1225.7868 | 1026.6309
5653 | 96.7534 | 1.042 | 1.267 | 0.92981 | 199.0742 | 53.131 | 68146.914 | 65762.444 | 1311.2489 | 1211.2489 | 1026.7097
5654 | 93.8489 | 0.97845 | 1.2421 | 1.0195 | 194.3451 | 49.8428 | 67885.8387 | 66276.1309 | 1321.5226 | 1221.5226 | 1026.8466
5655 | 87.0248 | 1.0158 | 0.94675 | 1.0428 | 210.3656 | 46.7332 | 50478.0337 | 52177.2154 | 1039.5443 | 939.5443 | 1026.8779
5656 | 93.5638 | 0.99378 | 1.0027 | 1.0271 | 200.125 | 51.4863 | 58011.078 | 59959.0141 | 1195.1803 | 1095.1803 | 1026.8941
5657 | 110.2703 | 0.9906 | 0.91631 | 0.99958 | 191.6031 | 45.4406 | 46775.2033 | 47879.2646 | 953.5853 | 853.5853 | 1026.939
5658 | 88.8829 | 0.94522 | 0.79172 | 0.99028 | 213.6281 | 47.4823 | 41430.0662 | 41737.2574 | 830.7451 | 730.7451 | 1026.9431
5659 | 85.6452 | 0.94951 | 0.81066 | 0.99731 | 192.975 | 48.5366 | 43380.5994 | 45025.7511 | 896.515 | 796.515 | 1026.9698
5660 | 102.8614 | 0.96522 | 1.1674 | 0.98251 | 200.9949 | 53.4058 | 66757.2973 | 66670.9228 | 1329.4185 | 1229.4185 | 1026.9891
5661 | 106.7964 | 0.99982 | 0.9371 | 1.1 | 212.2672 | 49.4817 | 56501.5104 | 57583.1112 | 1147.6622 | 1047.6622 | 1027.0077
5662 | 97.8908 | 1.0177 | 1.1063 | 1.0243 | 192.0693 | 47.7493 | 59059.3102 | 58689.1884 | 1169.7838 | 1069.7838 | 1027.0406
5663 | 120.3186 | 0.98554 | 1.0206 | 1.0664 | 204.9868 | 48.7565 | 59053.6319 | 60482.6447 | 1205.6529 | 1105.6529 | 1027.1229
5664 | 106.0696 | 0.90069 | 1.1531 | 1.019 | 188.3477 | 49.751 | 63403.4504 | 61020.5044 | 1216.4101 | 1116.4101 | 1027.1351
5665 | 79.2939 | 0.93729 | 1.1286 | 0.98331 | 201.4375 | 55.4402 | 65849.7094 | 64855.8899 | 1293.1178 | 1193.1178 | 1027.1351
5666 | 81.8866 | 0.98277 | 1.1467 | 0.96104 | 204.3941 | 50.3382 | 59730.1514 | 60742.9802 | 1210.8596 | 1110.8596 | 1027.1785
5667 | 101.8024 | 1.0623 | 1.0589 | 1.0371 | 192.1621 | 50.127 | 60659.5423 | 61251.7077 | 1221.0342 | 1121.0342 | 1027.1882
5668 | 111.6327 | 0.93432 | 0.83102 | 0.99908 | 216.7625 | 49.9141 | 46864.1246 | 48553.4214 | 967.0684 | 867.0684 | 1027.1896
5669 | 99.1145 | 0.99101 | 0.99294 | 0.94826 | 192.2991 | 49.1266 | 51273.5598 | 52994.6579 | 1055.8932 | 955.8932 | 1027.3549
5670 | 114.6688 | 0.97006 | 0.98992 | 0.95925 | 194.8212 | 45.5256 | 48488.7732 | 50002.0843 | 996.0417 | 896.0417 | 1027.3575
5671 | 90.5384 | 0.95667 | 1.089 | 1.0097 | 181.9842 | 43.9474 | 52312.6762 | 53622.6943 | 1068.4539 | 968.4539 | 1027.4616
5672 | 85.8843 | 1.0095 | 1.2538 | 1.0287 | 199.2879 | 52.564 | 72556.528 | 70962.1991 | 1415.244 | 1315.244 | 1027.4712
5673 | 80.2656 | 1.0101 | 0.93417 | 0.99503 | 197.9762 | 43.5609 | 44220.8332 | 45853.7035 | 913.0741 | 813.0741 | 1027.5144
5674 | 99.6784 | 1.0633 | 1.2485 | 1.0069 | 202.2212 | 52.9545 | 72389.6009 | 70890.6519 | 1413.813 | 1313.813 | 1027.5258
5675 | 93.943 | 0.99912 | 1.0387 | 0.99434 | 185.3162 | 54.5939 | 61697.3359 | 61184.3353 | 1219.6867 | 1119.6867 | 1027.5278
5676 | 99.5511 | 1.0228 | 0.76523 | 0.98426 | 224.5099 | 47.4756 | 40816.3599 | 41006.9952 | 816.1399 | 716.1399 | 1027.5563
5677 | 123.5137 | 1.0289 | 1.0672 | 1.0408 | 201.4423 | 43.6647 | 54252.6506 | 53189.8909 | 1059.7978 | 959.7978 | 1027.5873
5678 | 110.7284 | 0.994 | 1.1097 | 0.96376 | 196.6407 | 47.8483 | 56638.1403 | 57530.9543 | 1146.6191 | 1046.6191 | 1027.6209
5679 | 113.5878 | 0.95983 | 0.90542 | 0.98891 | 204.6004 | 48.8166 | 49236.1992 | 48962.2077 | 975.2442 | 875.2442 | 1027.6686
5680 | 106.9769 | 0.98482 | 1.1758 | 1.0658 | 216.0762 | 50.1712 | 68374.6126 | 69841.7437 | 1392.8349 | 1292.8349 | 1027.7962
5681 | 101.6611 | 1.0113 | 0.79418 | 0.9605 | 210.9189 | 52.8294 | 45941.0914 | 44819.3573 | 892.3871 | 792.3871 | 1027.8007
5682 | 108.9547 | 1.0709 | 0.93117 | 1.0693 | 190.2096 | 54.0732 | 60338.6029 | 61408.9472 | 1224.1789 | 1124.1789 | 1027.8092
5683 | 91.3134 | 1.0958 | 1.3965 | 1.1126 | 186.19 | 49.4228 | 81923.5081 | 80717.0057 | 1610.3401 | 1510.3401 | 1027.9435
5684 | 113.5249 | 1.0451 | 0.80662 | 0.90089 | 195.8377 | 45.5414 | 38693.3732 | 37838.4585 | 752.7692 | 652.7692 | 1028.0984
5685 | 117.5766 | 0.9672 | 1.1403 | 1.0271 | 181.2612 | 52.8039 | 68030.2509 | 66062.8952 | 1317.2579 | 1217.2579 | 1028.1712
5686 | 84.0743 | 0.9513 | 0.79718 | 0.99434 | 182.107 | 49.5579 | 43428.6327 | 42930.0642 | 854.6013 | 754.6013 | 1028.1724
5687 | 95.5988 | 0.91853 | 1.12 | 1.0347 | 202.6813 | 49.6897 | 62145.7967 | 63076.6397 | 1257.5328 | 1157.5328 | 1028.2523
5688 | 99.7848 | 0.94596 | 1.2599 | 0.97848 | 197.935 | 53.9144 | 71751.1961 | 68569.6802 | 1367.3936 | 1267.3936 | 1028.2555
5689 | 112.4892 | 1.1569 | 1.1982 | 0.91999 | 209.1835 | 47.0423 | 58189.5763 | 57255.8296 | 1141.1166 | 1041.1166 | 1028.414
5690 | 90.9673 | 1.01 | 0.84178 | 1.1147 | 207.7108 | 48.1705 | 49833.7252 | 48058.1969 | 957.1639 | 857.1639 | 1028.4517
5691 | 105.3619 | 0.87953 | 1.1235 | 1.016 | 201.9594 | 53.2117 | 65873.9172 | 67526.2501 | 1346.525 | 1246.525 | 1028.5201
5692 | 95.0966 | 0.97501 | 0.94524 | 0.96583 | 211.1613 | 51.0364 | 51536.8033 | 50463.8063 | 1005.2761 | 905.2761 | 1028.521
5693 | 99.1894 | 1.0859 | 1.0202 | 0.99406 | 200.5088 | 47.3116 | 53277.8731 | 54065.5245 | 1077.3105 | 977.3105 | 1028.5275
5694 | 106.0308 | 1.0031 | 1.1624 | 0.95875 | 213.2793 | 49.3128 | 60416.5028 | 60333.3267 | 1202.6665 | 1102.6665 | 1028.5573
5695 | 88.9774 | 0.97115 | 0.71752 | 1.0467 | 189.1645 | 47.7613 | 40186.0075 | 39074.0595 | 777.4812 | 677.4812 | 1028.632
5696 | 122.1306 | 0.96736 | 1.1589 | 1.0134 | 191.9667 | 50.6478 | 65660.2023 | 68241.759 | 1360.8352 | 1260.8352 | 1028.641
5697 | 89.1784 | 0.98983 | 1.0943 | 0.93898 | 199.875 | 54.2841 | 60769.2478 | 61110.8741 | 1218.2175 | 1118.2175 | 1028.6509
5698 | 96.5833 | 1.0855 | 0.98923 | 0.98488 | 209.7106 | 49.2179 | 53321.3401 | 53400.9663 | 1064.0193 | 964.0193 | 1028.7355
5699 | 92.1557 | 1.0446 | 1.122 | 0.99946 | 187.7875 | 53.4069 | 65219.2038 | 64798.7677 | 1291.9754 | 1191.9754 | 1028.7519
5700 | 81.6935 | 0.98987 | 1.0982 | 0.94947 | 211.115 | 52.7556 | 59487.3007 | 60018.4293 | 1196.3686 | 1096.3686 | 1028.767
5701 | 106.0167 | 1.0262 | 0.95476 | 1.0091 | 186.5169 | 50.9949 | 54864.3137 | 55310.7697 | 1102.2154 | 1002.2154 | 1028.7846
5702 | 112.5193 | 1.0552 | 1.0049 | 1.0146 | 210.8338 | 47.8668 | 54697.0841 | 54434.8027 | 1084.6961 | 984.6961 | 1028.7897
5703 | 84.6865 | 0.96664 | 1.0651 | 1.0196 | 189.8596 | 47.0995 | 55196.0981 | 55079.2573 | 1097.5851 | 997.5851 | 1028.8719
5704 | 101.6185 | 0.99969 | 1.0436 | 0.92857 | 214.3998 | 51.118 | 54942.3195 | 55763.6394 | 1111.2728 | 1011.2728 | 1028.8874
5705 | 93.7715 | 0.94814 | 0.58789 | 1.0662 | 206.3318 | 46.127 | 33221.2949 | 32812.5949 | 652.2519 | 552.2519 | 1028.9867
5706 | 98.3606 | 1.0317 | 1.215 | 0.98763 | 198.3479 | 53.3828 | 69671.02 | 69395.8503 | 1383.917 | 1283.917 | 1029.0174
5707 | 119.5438 | 0.98684 | 1.2152 | 0.96516 | 200.6807 | 46.306 | 59972.0806 | 61114.0574 | 1218.2811 | 1118.2811 | 1029.0191
5708 | 109.1963 | 1.0047 | 0.75883 | 0.97266 | 190.4993 | 51.3126 | 43693.4426 | 43196.081 | 859.9216 | 759.9216 | 1029.0192
5709 | 120.0408 | 0.99713 | 1.139 | 0.90782 | 192.5475 | 50.0551 | 57939.7991 | 60440.2968 | 1204.8059 | 1104.8059 | 1029.0773
5710 | 86.1249 | 1.0187 | 1.2444 | 0.96599 | 187.7533 | 52.5394 | 67952.916 | 68205.2229 | 1360.1045 | 1260.1045 | 1029.0795
5711 | 100.5801 | 1.0319 | 1.0165 | 1.0636 | 186.0203 | 51.5105 | 61221.5059 | 59936.1309 | 1194.7226 | 1094.7226 | 1029.0864
5712 | 97.6063 | 0.89827 | 1.1102 | 0.96734 | 183.0341 | 47.0629 | 54854.1127 | 55222.3687 | 1100.4474 | 1000.4474 | 1029.2093
5713 | 78.233 | 0.93087 | 1.0445 | 0.97693 | 212.1838 | 49.5509 | 54384.036 | 54776.4145 | 1091.5283 | 991.5283 | 1029.2855
5714 | 107.4804 | 0.97931 | 0.84212 | 0.94803 | 192.4645 | 49.6073 | 45018.1761 | 45245.5139 | 900.9103 | 800.9103 | 1029.3423
5715 | 94.668 | 1.0362 | 1.2706 | 0.98868 | 207.0592 | 54.0039 | 73343.9208 | 73194.9136 | 1459.8983 | 1359.8983 | 1029.3424
5716 | 114.4228 | 0.94254 | 0.922 | 0.96326 | 195.8838 | 54.614 | 54590.0503 | 53140.1153 | 1058.8023 | 958.8023 | 1029.3866
5717 | 94.3673 | 0.99496 | 1.0144 | 0.92952 | 212.8857 | 45.5149 | 47403.585 | 48483.9363 | 965.6787 | 865.6787 | 1029.4798
5718 | 102.7226 | 1.0121 | 0.98423 | 1.0342 | 210.4834 | 51.1054 | 57545.9505 | 57268.3913 | 1141.3678 | 1041.3678 | 1029.5497
5719 | 102.1236 | 0.99393 | 0.74754 | 0.90361 | 215.3339 | 50.6051 | 39534.938 | 39508.2609 | 786.1652 | 686.1652 | 1029.5537
5720 | 77.9456 | 0.97283 | 0.99353 | 1.0688 | 192.6613 | 52.2084 | 59590.7748 | 59613.298 | 1188.266 | 1088.266 | 1029.5796
5721 | 94.8746 | 1.0272 | 0.67631 | 0.99914 | 197.6712 | 54.211 | 42112.3829 | 42764.2952 | 851.2859 | 751.2859 | 1029.5899
5722 | 91.317 | 1.0252 | 1.0909 | 0.96161 | 211.2715 | 49.4656 | 56734.671 | 57824.1009 | 1152.482 | 1052.482 | 1029.6242
5723 | 104.826 | 0.98247 | 0.8368 | 0.95276 | 203.8461 | 49.1886 | 44486.6121 | 44355.1844 | 883.1037 | 783.1037 | 1029.6485
5724 | 104.3511 | 1.0383 | 1.0049 | 0.9356 | 196.8482 | 50.8509 | 53515.6048 | 52701.3027 | 1050.0261 | 950.0261 | 1029.6974
5725 | 108.3975 | 0.97344 | 0.90677 | 0.9806 | 197.6789 | 52.3895 | 52309.032 | 52149.902 | 1038.998 | 938.998 | 1029.7171
5726 | 113.4252 | 1.0459 | 1.2336 | 0.95496 | 203.9032 | 47.3098 | 61550.6469 | 61551.6362 | 1227.0327 | 1127.0327 | 1029.7728
5727 | 107.1795 | 1.0077 | 1.1639 | 0.97882 | 219.628 | 50.1435 | 62762.704 | 65230.5092 | 1300.6102 | 1200.6102 | 1029.8323
5728 | 104.0039 | 0.99588 | 1.0233 | 0.97269 | 202.9627 | 49.2661 | 54344.3683 | 54874.722 | 1093.4944 | 993.4944 | 1029.8663
5729 | 99.0522 | 1.0434 | 0.96558 | 1.0087 | 194.4331 | 49.3216 | 53332.1545 | 53387.3101 | 1063.7462 | 963.7462 | 1029.9056
5730 | 93.3733 | 0.93316 | 1.1387 | 1.0452 | 198.9471 | 50.8881 | 65199.2837 | 65197.5392 | 1299.9508 | 1199.9508 | 1029.9561
5731 | 81.7108 | 0.93436 | 0.90806 | 1.0117 | 189.1221 | 51.4966 | 51429.5147 | 52523.5243 | 1046.4705 | 946.4705 | 1029.9592
5732 | 112.5451 | 0.99493 | 1.3096 | 1.0462 | 193.3002 | 50.6983 | 75332.5377 | 74760.3782 | 1491.2076 | 1391.2076 | 1029.9621
5733 | 103.1331 | 1.0783 | 0.92451 | 0.98554 | 204.368 | 55.9396 | 57394.4101 | 58203.6094 | 1160.0722 | 1060.0722 | 1029.9635
5734 | 115.2934 | 1.0361 | 1.1812 | 1.0206 | 205.2219 | 50.3263 | 66885.1481 | 66177.5683 | 1319.5514 | 1219.5514 | 1030.0763
5735 | 94.431 | 0.95668 | 1.1193 | 1.0373 | 208.3243 | 49.2592 | 61850.9717 | 63029.8784 | 1256.5976 | 1156.5976 | 1030.115
5736 | 98.3706 | 0.92048 | 0.87145 | 0.91953 | 209.0058 | 53.2892 | 47736.3091 | 47880.7498 | 953.615 | 853.615 | 1030.1553
5737 | 102.8558 | 1.0187 | 0.8416 | 1.0424 | 196.2492 | 45.3788 | 44759.6058 | 44515.4153 | 886.3083 | 786.3083 | 1030.1707
5738 | 94.5576 | 1.0347 | 0.9362 | 0.99543 | 174.219 | 51.5806 | 53289.9812 | 52625.9575 | 1048.5191 | 948.5191 | 1030.2615
5739 | 94.2278 | 1.0098 | 1.0309 | 0.97156 | 199.5927 | 50.2335 | 55290.7525 | 54578.4282 | 1087.5686 | 987.5686 | 1030.2726
5740 | 88.031 | 1.0126 | 1.0167 | 0.99843 | 202.8748 | 50.501 | 55967.4597 | 55297.1809 | 1101.9436 | 1001.9436 | 1030.2879
5741 | 106.5738 | 0.98994 | 1.1271 | 0.995 | 210.7513 | 55.0059 | 67699.622 | 65951.7974 | 1315.0359 | 1215.0359 | 1030.3087
5742 | 86.5143 | 1.0996 | 0.85405 | 0.98303 | 208.0819 | 49.4228 | 46402.8676 | 46576.9469 | 927.5389 | 827.5389 | 1030.3468
5743 | 108.9034 | 0.96312 | 0.99795 | 0.9549 | 187.2687 | 58.4721 | 62040.5543 | 59600.9114 | 1188.0182 | 1088.0182 | 1030.3624
5744 | 102.8783 | 0.96699 | 1.0513 | 1.0408 | 208.46 | 51.121 | 61227.7994 | 62174.9477 | 1239.499 | 1139.499 | 1030.4795
5745 | 100.0661 | 1.0501 | 0.88628 | 0.92896 | 188.7244 | 47.5828 | 44364.3091 | 44631.9318 | 888.6386 | 788.6386 | 1030.504
5746 | 100.496 | 1.0113 | 1.0779 | 1.1328 | 189.943 | 48.7522 | 64670.0513 | 69133.8899 | 1378.6778 | 1278.6778 | 1030.5228
5747 | 100.6315 | 1.0349 | 1.0636 | 1.0501 | 202.9364 | 40.4787 | 49630.9142 | 49018.8631 | 976.3773 | 876.3773 | 1030.5499
5748 | 95.9146 | 0.93602 | 0.98354 | 0.94392 | 199.4937 | 49.533 | 50631.9897 | 49159.1993 | 979.184 | 879.184 | 1030.5562
5749 | 75.9415 | 0.99545 | 0.85785 | 0.95558 | 207.7598 | 49.1936 | 44252.6585 | 43333.6424 | 862.6728 | 762.6728 | 1030.5949
5750 | 93.6693 | 0.88385 | 1.2639 | 0.99214 | 193.3565 | 51.6982 | 69299.9174 | 69232.3259 | 1380.6465 | 1280.6465 | 1030.7275
5751 | 92.5253 | 0.97923 | 1.3394 | 0.96724 | 200.1039 | 48.2026 | 67016.4064 | 65547.079 | 1306.9416 | 1206.9416 | 1030.762
5752 | 96.6219 | 1.064 | 1.2781 | 0.95846 | 209.9296 | 53.1469 | 70779.9275 | 70695.3721 | 1409.9074 | 1309.9074 | 1030.7958
5753 | 102.5642 | 1.0263 | 0.75579 | 1.0013 | 192.6821 | 52.3325 | 45303.5642 | 45862.1331 | 913.2427 | 813.2427 | 1030.8042
5754 | 108.7016 | 0.9792 | 0.95486 | 1.0439 | 181.9158 | 46.2887 | 51247.9396 | 52074.2612 | 1037.4852 | 937.4852 | 1030.8145
5755 | 102.7771 | 1.0144 | 0.92844 | 0.99922 | 179.096 | 50.6012 | 52398.2525 | 51848.7319 | 1032.9746 | 932.9746 | 1030.8569
5756 | 93.7466 | 1.0873 | 0.63061 | 0.97002 | 208.145 | 50.6536 | 36356.6177 | 35418.8857 | 704.3777 | 604.3777 | 1030.8581
5757 | 110.2969 | 0.95654 | 1.2378 | 0.88814 | 196.5681 | 45.8027 | 55382.9302 | 54077.5235 | 1077.5505 | 977.5505 | 1030.9618
5758 | 94.0283 | 0.92382 | 0.90573 | 0.9395 | 195.5666 | 48.8428 | 46000.2138 | 45831.9207 | 912.6384 | 812.6384 | 1031.011
5759 | 81.9624 | 0.96121 | 0.94447 | 1.0515 | 201.2591 | 49.3957 | 53146.7806 | 53235.967 | 1060.7193 | 960.7193 | 1031.0134
5760 | 98.2956 | 1.0248 | 1.1236 | 0.96291 | 214.3595 | 48.8618 | 58000.8488 | 59163.9384 | 1179.2788 | 1079.2788 | 1031.0171
5761 | 102.3769 | 0.99838 | 1.1038 | 0.99895 | 194.3181 | 47.7488 | 57722.0194 | 56848.2649 | 1132.9653 | 1032.9653 | 1031.024
5762 | 111.4798 | 0.92061 | 1.2174 | 0.99838 | 207.5596 | 47.6508 | 63013.4918 | 62687.2802 | 1249.7456 | 1149.7456 | 1031.0279
5763 | 91.4144 | 1.12 | 1.1159 | 1.041 | 201.2316 | 47.7334 | 60539.4001 | 61231.8099 | 1220.6362 | 1120.6362 | 1031.0399
5764 | 124.8119 | 1.0281 | 0.99638 | 1.0113 | 200.6997 | 52.2327 | 59536.8995 | 60403.3261 | 1204.0665 | 1104.0665 | 1031.0532
5765 | 83.8379 | 0.99825 | 1.2074 | 1.0189 | 195.3453 | 56.2155 | 74055.2893 | 75230.8954 | 1500.6179 | 1400.6179 | 1031.1026
5766 | 85.983 | 1.0705 | 1.0349 | 0.96252 | 205.614 | 50.4083 | 55059.6375 | 56808.0868 | 1132.1617 | 1032.1617 | 1031.1318
5767 | 102.151 | 1.0682 | 0.97956 | 1.0131 | 200.4678 | 52.7189 | 58272.0397 | 57798.4332 | 1151.9687 | 1051.9687 | 1031.2232
5768 | 96.6149 | 1.0712 | 1.0353 | 1.0295 | 211.4755 | 53.3387 | 62579.8999 | 62972.9096 | 1255.4582 | 1155.4582 | 1031.2415
5769 | 96.3941 | 0.9026 | 1.3662 | 1.0107 | 187.2663 | 53.8209 | 79191.0257 | 80518.0124 | 1606.3602 | 1506.3602 | 1031.2646
5770 | 85.6269 | 1.0166 | 0.85139 | 0.95883 | 192.0561 | 46.2379 | 41962.484 | 42064.8517 | 837.297 | 737.297 | 1031.2756
5771 | 90.2951 | 0.94805 | 1.0389 | 1.0132 | 202.2111 | 51.6918 | 59036.3348 | 59689.4227 | 1189.7885 | 1089.7885 | 1031.2957
5772 | 113.1291 | 0.96665 | 0.94768 | 1.0269 | 213.42 | 49.4217 | 53711.4196 | 54467.0817 | 1085.3416 | 985.3416 | 1031.3147
5773 | 101.152 | 1.0112 | 1.0246 | 0.94274 | 208.1385 | 51.8148 | 55558.841 | 54334.779 | 1082.6956 | 982.6956 | 1031.3208
5774 | 92.6195 | 1.0178 | 0.79103 | 1.0518 | 210.329 | 56.5868 | 52625.2489 | 53030.8708 | 1056.6174 | 956.6174 | 1031.3568
5775 | 102.7113 | 0.90498 | 0.86062 | 0.90712 | 202.0884 | 49.3897 | 43350.7053 | 42846.6566 | 852.9331 | 752.9331 | 1031.3963
5776 | 101.1473 | 0.94562 | 1.0879 | 1.003 | 202.0209 | 49.9415 | 59470.7598 | 57942.0698 | 1154.8414 | 1054.8414 | 1031.4269
5777 | 100.0903 | 1.0548 | 1.0574 | 1.0245 | 192.21 | 52.6937 | 62836.9031 | 62110.5754 | 1238.2115 | 1138.2115 | 1031.5071
5778 | 111.5847 | 0.96204 | 0.8209 | 0.98076 | 201.8536 | 47.0996 | 43178.2773 | 43187.2835 | 859.7457 | 759.7457 | 1031.5408
5779 | 105.5535 | 0.99615 | 0.79227 | 0.95925 | 197.0231 | 49.8363 | 43311.7812 | 44144.468 | 878.8894 | 778.8894 | 1031.5452
5780 | 99.3292 | 1.0741 | 1.2934 | 0.88448 | 215.3714 | 55.1677 | 69213.0441 | 69635.6 | 1388.712 | 1288.712 | 1031.6198
5781 | 86.9909 | 0.98669 | 1.0122 | 0.97625 | 188.2884 | 53.474 | 57616.5552 | 56984.183 | 1135.6837 | 1035.6837 | 1031.6913
5782 | 84.5535 | 0.92526 | 0.7946 | 0.9694 | 196.8425 | 49.4697 | 42172.5438 | 40964.4738 | 815.2895 | 715.2895 | 1031.7001
5783 | 86.3852 | 0.97781 | 0.88457 | 1.0101 | 210.8615 | 49.0315 | 48161.3519 | 47031.8111 | 936.6362 | 836.6362 | 1031.7161
5784 | 98.0635 | 0.99159 | 0.71421 | 1.0025 | 202.3677 | 50.4484 | 41228.5941 | 40903.8183 | 814.0764 | 714.0764 | 1031.7558
5785 | 91.7208 | 1.0521 | 0.9771 | 0.93219 | 216.7688 | 55.6719 | 56297.5548 | 55560.7825 | 1107.2157 | 1007.2157 | 1031.8575
5786 | 117.9762 | 0.98491 | 1.2273 | 0.96881 | 186.752 | 50.0529 | 65514.2722 | 64316.6776 | 1282.3336 | 1182.3336 | 1031.9438
5787 | 103.5799 | 1.1149 | 1.1393 | 0.99374 | 208.6689 | 50.2043 | 62847.914 | 62278.6375 | 1241.5727 | 1141.5727 | 1031.9685
5788 | 83.0042 | 1.0373 | 1.3071 | 1.0192 | 205.234 | 44.0636 | 62701.417 | 63813.9494 | 1272.279 | 1172.279 | 1032.0053
5789 | 106.0122 | 0.98704 | 1.1037 | 0.95743 | 187.6144 | 47.6214 | 55490.9024 | 57464.2598 | 1145.2852 | 1045.2852 | 1032.0218
5790 | 102.6977 | 1.0741 | 1.1395 | 1.0499 | 188.4844 | 53.7072 | 70366.6316 | 71533.9808 | 1426.6796 | 1326.6796 | 1032.0996
5791 | 91.4092 | 1.0207 | 0.92697 | 1.0082 | 198.2147 | 54.4594 | 56174.2014 | 55996.1538 | 1115.9231 | 1015.9231 | 1032.1139
5792 | 110.6924 | 1.1112 | 0.83958 | 0.97022 | 205.3208 | 47.0682 | 44335.1668 | 45332.547 | 902.6509 | 802.6509 | 1032.1617
5793 | 91.0511 | 0.94495 | 0.9669 | 1.0315 | 190.4321 | 48.7163 | 52967.9992 | 52522.3973 | 1046.4479 | 946.4479 | 1032.1768
5794 | 95.5208 | 1.0377 | 1.1159 | 1.0847 | 207.2579 | 50.2584 | 66022.2978 | 64227.3652 | 1280.5473 | 1180.5473 | 1032.1911
5795 | 92.4004 | 0.92526 | 0.81203 | 0.99433 | 200.4445 | 48.6373 | 43629.5025 | 43502.1286 | 866.0426 | 766.0426 | 1032.3332
5796 | 95.754 | 0.96569 | 0.94553 | 0.89867 | 192.5186 | 48.4778 | 45867.7262 | 47685.7733 | 949.7155 | 849.7155 | 1032.4183
5797 | 116.5087 | 1.0512 | 1.0088 | 1.0317 | 190.9017 | 48.4341 | 56529.5262 | 57033.3119 | 1136.6662 | 1036.6662 | 1032.4845
5798 | 80.2784 | 0.97558 | 0.9779 | 0.92919 | 212.6389 | 51.6523 | 51191.9085 | 48779.8487 | 971.597 | 871.597 | 1032.4911
5799 | 105.3497 | 1.0205 | 0.98586 | 0.96549 | 197.2508 | 45.694 | 48603.1202 | 49172.1447 | 979.4429 | 879.4429 | 1032.5063
5800 | 112.8538 | 0.93846 | 1.1362 | 1.0281 | 180.4723 | 49.158 | 62811.693 | 62840.2496 | 1252.805 | 1152.805 | 1032.511
5801 | 98.5203 | 0.9443 | 0.90509 | 1.0101 | 188.274 | 51.5735 | 52137.1191 | 52673.603 | 1049.4721 | 949.4721 | 1032.5583
5802 | 99.7393 | 0.94836 | 1.0155 | 0.90802 | 210.3495 | 49.292 | 50322.4844 | 52080.9381 | 1037.6188 | 937.6188 | 1032.5648
5803 | 84.3589 | 0.99537 | 0.98822 | 0.92011 | 200.9243 | 44.6961 | 44594.5296 | 45112.3689 | 898.2474 | 798.2474 | 1032.5769
5804 | 96.3416 | 1.0122 | 1.0287 | 1.0392 | 195.2415 | 50.9961 | 59685.5798 | 58812.8312 | 1172.2566 | 1072.2566 | 1032.6057
5805 | 108.3692 | 0.97193 | 0.90091 | 0.94732 | 203.8547 | 52.9546 | 50975.4762 | 51228.2885 | 1020.5658 | 920.5658 | 1032.6441
5806 | 122.5671 | 0.91345 | 1.2143 | 0.90756 | 181.9898 | 53.5034 | 65135.8457 | 66890.3278 | 1333.8066 | 1233.8066 | 1032.6448
5807 | 88.8376 | 0.96137 | 0.83893 | 1.0826 | 198.314 | 44.7135 | 44627.5717 | 43719.022 | 870.3804 | 770.3804 | 1032.6571
5808 | 100.5708 | 1.0794 | 0.92838 | 1.0888 | 203.8286 | 54.664 | 61393.3387 | 60186.4562 | 1199.7291 | 1099.7291 | 1032.6903
5809 | 95.6061 | 0.92798 | 1.065 | 0.90943 | 195.6982 | 51.0866 | 54207.9298 | 55081.8402 | 1097.6368 | 997.6368 | 1032.7034
5810 | 103.3094 | 0.94783 | 1.026 | 1.0308 | 209.6312 | 55.2647 | 64072.9954 | 65576.0512 | 1307.521 | 1207.521 | 1032.7535
5811 | 97.5561 | 0.96803 | 0.93377 | 1.0512 | 197.4446 | 51.4532 | 55564.0967 | 55877.6773 | 1113.5535 | 1013.5535 | 1032.8036
5812 | 112.4392 | 0.95986 | 0.67676 | 1.0377 | 185.1255 | 48.9043 | 39806.8095 | 39054.5644 | 777.0913 | 677.0913 | 1032.8099
5813 | 91.4762 | 0.95513 | 1.1185 | 1.0318 | 195.7873 | 45.7966 | 57050.1167 | 55841.6008 | 1112.832 | 1012.832 | 1032.8459
5814 | 96.4803 | 0.97844 | 1.0835 | 1.0147 | 202.7291 | 50.1145 | 60031.2614 | 60049.0859 | 1196.9817 | 1096.9817 | 1032.873
5815 | 110.8573 | 1.0517 | 1.1769 | 1.0221 | 195.3852 | 47.7153 | 63156.8399 | 62623.9953 | 1248.4799 | 1148.4799 | 1032.8862
5816 | 96.3232 | 1.0479 | 0.9487 | 0.99452 | 191.7695 | 51.2127 | 53680.5744 | 55086.7658 | 1097.7353 | 997.7353 | 1032.929
5817 | 109.1902 | 1.015 | 1.0576 | 0.9493 | 212.9743 | 44.9458 | 50317.9683 | 50896.096 | 1013.9219 | 913.9219 | 1032.9343
5818 | 101.8335 | 1.0827 | 1.0119 | 1.018 | 206.3808 | 45.7689 | 52400.4978 | 53393.8101 | 1063.8762 | 963.8762 | 1032.9479
5819 | 123.0869 | 0.97025 | 1.167 | 0.93802 | 198.2353 | 50.261 | 61219.5396 | 60567.0233 | 1207.3405 | 1107.3405 | 1032.9653
5820 | 96.415 | 1.0995 | 1.0328 | 1.0205 | 200.8043 | 51.1595 | 59542.2449 | 60387.8656 | 1203.7573 | 1103.7573 | 1032.9654
5821 | 91.6904 | 0.97551 | 1.1263 | 0.93565 | 210.735 | 52.2903 | 59990.5093 | 58735.3832 | 1170.7077 | 1070.7077 | 1032.982
5822 | 93.1967 | 1.0601 | 0.85922 | 0.97701 | 206.9627 | 48.8573 | 46048.1953 | 46217.5963 | 920.3519 | 820.3519 | 1033.0312
5823 | 101.3956 | 0.98163 | 0.82903 | 0.91172 | 195.2911 | 51.1396 | 43938.8032 | 43640.1118 | 868.8022 | 768.8022 | 1033.0609
5824 | 110.0328 | 1.0005 | 0.59523 | 0.97094 | 187.1084 | 54.1725 | 37458.474 | 38376.7826 | 763.5357 | 663.5357 | 1033.1572
5825 | 105.2067 | 0.97527 | 1.0946 | 0.96008 | 192.672 | 49.0728 | 56796.461 | 56457.6426 | 1125.1529 | 1025.1529 | 1033.2003
5826 | 97.9761 | 0.98897 | 0.86296 | 0.97303 | 194.1476 | 54.7426 | 51464.6081 | 52648.4369 | 1048.9687 | 948.9687 | 1033.3011
5827 | 85.5939 | 0.9375 | 1.2222 | 1.0085 | 198.6041 | 51.2008 | 67416.7711 | 68070.7529 | 1357.4151 | 1257.4151 | 1033.3779
5828 | 110.1582 | 0.91102 | 1.1428 | 0.99828 | 199.9467 | 47.8674 | 59612.0194 | 58777.2344 | 1171.5447 | 1071.5447 | 1033.3866
5829 | 107.4096 | 0.92587 | 0.80721 | 0.97408 | 209.3557 | 49.0481 | 43653.0571 | 43157.4483 | 859.149 | 759.149 | 1033.3878
5830 | 87.9384 | 0.96077 | 1.0999 | 0.93678 | 191.9707 | 45.3247 | 50721.1341 | 49696.5219 | 989.9304 | 889.9304 | 1033.4287
5831 | 96.0164 | 0.91235 | 1.1799 | 1.0339 | 223.6255 | 49.376 | 64785.6089 | 64074.8289 | 1277.4966 | 1177.4966 | 1033.4376
5832 | 97.703 | 0.98342 | 1.0782 | 1.0519 | 205.8307 | 49.2533 | 60801.3363 | 61210.8229 | 1220.2165 | 1120.2165 | 1033.5076
5833 | 110.0213 | 0.92276 | 0.96551 | 1.044 | 197.0767 | 48.12 | 53587.1073 | 55952.2421 | 1115.0448 | 1015.0448 | 1033.5885
5834 | 120.0868 | 0.96941 | 1.1843 | 1.002 | 206.8243 | 50.5779 | 66111.0431 | 66328.5572 | 1322.5711 | 1222.5711 | 1033.6137
5835 | 94.2086 | 0.9488 | 0.96893 | 1.0148 | 213.723 | 48.611 | 52355.1539 | 53690.5061 | 1069.8101 | 969.8101 | 1033.6255
5836 | 101.1055 | 1.0033 | 0.87401 | 0.95574 | 202.2709 | 52.5984 | 49474.5767 | 49386.0428 | 983.7209 | 883.7209 | 1033.6391
5837 | 86.4429 | 1.017 | 0.85365 | 0.95839 | 203.933 | 49.2131 | 44792.8707 | 44066.6317 | 877.3326 | 777.3326 | 1033.6883
5838 | 86.659 | 0.99348 | 1.2977 | 1.1029 | 195.3294 | 50.9565 | 77512.0066 | 77261.3587 | 1541.2272 | 1441.2272 | 1033.692
5839 | 114.1316 | 0.93862 | 0.95493 | 1.0322 | 210.0098 | 54.9957 | 60308.1337 | 61185.1577 | 1219.7032 | 1119.7032 | 1033.7202
5840 | 97.7345 | 0.96128 | 1.0666 | 0.99997 | 191.2687 | 46.0629 | 53649.8985 | 53697.7402 | 1069.9548 | 969.9548 | 1033.7942
5841 | 94.7683 | 1.0162 | 1.0987 | 1.0017 | 209.7167 | 56.1966 | 67467.2074 | 67550.8888 | 1347.0178 | 1247.0178 | 1033.8104
5842 | 101.7309 | 1.0913 | 0.86194 | 0.97202 | 186.6835 | 45.5568 | 43413.2567 | 44979.0921 | 895.5818 | 795.5818 | 1033.8235
5843 | 72.8319 | 1.0959 | 1.1624 | 0.97228 | 202.1077 | 49.9896 | 60688.4443 | 61547.155 | 1226.9431 | 1126.9431 | 1033.884
5844 | 86.572 | 1.0382 | 0.93483 | 0.96162 | 209.4692 | 49.5788 | 49234.6639 | 49647.3451 | 988.9469 | 888.9469 | 1033.9672
5845 | 128.5285 | 0.97747 | 1.1895 | 1.1225 | 197.9345 | 51.1515 | 74922.8027 | 72894.9254 | 1453.8985 | 1353.8985 | 1033.9969
5846 | 99.7119 | 0.95234 | 1.0621 | 1.0352 | 181.8431 | 52.113 | 62427.6089 | 61332.9069 | 1222.6581 | 1122.6581 | 1034.0376
5847 | 91.1622 | 1.1323 | 1.0513 | 0.92248 | 223.4484 | 48.1308 | 51871.0756 | 53345.4419 | 1062.9088 | 962.9088 | 1034.0601
5848 | 88.9108 | 1.0066 | 0.93189 | 0.92852 | 204.9156 | 48.0398 | 46072.3732 | 44630.5975 | 888.612 | 788.612 | 1034.1269
5849 | 109.668 | 1.0264 | 1.1271 | 1.0433 | 210.4496 | 48.2066 | 62325.4483 | 62996.121 | 1255.9224 | 1155.9224 | 1034.1741
5850 | 87.7747 | 0.96158 | 0.99507 | 1.041 | 228.4761 | 48.3091 | 54348.3565 | 54031.4808 | 1076.6296 | 976.6296 | 1034.2046
5851 | 105.0593 | 1.0628 | 0.97816 | 0.99541 | 199.142 | 49.8967 | 54353.2425 | 55230.9298 | 1100.6186 | 1000.6186 | 1034.3201
5852 | 100.1825 | 1.0869 | 0.88923 | 1.0689 | 199.8668 | 48.6032 | 51689.2223 | 51397.2959 | 1023.9459 | 923.9459 | 1034.4688
5853 | 88.6343 | 1.0061 | 0.8464 | 0.98904 | 193.4191 | 52.9078 | 49201.8367 | 50685.4422 | 1009.7088 | 909.7088 | 1034.5136
5854 | 91.8699 | 1.0007 | 0.84461 | 1.0506 | 200.0187 | 51.7272 | 50855.2166 | 49698.5685 | 989.9714 | 889.9714 | 1034.5505
5855 | 114.3892 | 0.97809 | 1.0193 | 0.99285 | 209.093 | 48.4698 | 54683.4973 | 52896.3055 | 1053.9261 | 953.9261 | 1034.5687
5856 | 113.2571 | 0.9689 | 0.96065 | 0.9826 | 202.3931 | 45.4406 | 48081.7684 | 47466.8083 | 945.3362 | 845.3362 | 1034.5767
5857 | 90.9811 | 1.0066 | 1.0465 | 1.0426 | 191.8452 | 50.4624 | 59872.8443 | 59913.515 | 1194.2703 | 1094.2703 | 1034.6983
5858 | 87.6842 | 1.011 | 0.83992 | 1.0095 | 196.9797 | 47.1042 | 44311.3892 | 43675.7365 | 869.5147 | 769.5147 | 1034.7777
5859 | 106.9007 | 0.97196 | 1.0116 | 1.042 | 206.5438 | 53.3554 | 61994.5439 | 61456.3389 | 1225.1268 | 1125.1268 | 1034.8249
5860 | 95.2635 | 1.001 | 1.0143 | 0.92389 | 213.7339 | 53.2653 | 55209.6996 | 54992.3816 | 1095.8476 | 995.8476 | 1034.8645
5861 | 87.5726 | 1.0348 | 1.1264 | 1.0441 | 200.1738 | 49.7173 | 63176.5972 | 61462.5949 | 1225.2519 | 1125.2519 | 1034.887
5862 | 97.96 | 1.0149 | 1.0865 | 0.99963 | 204.4029 | 52.173 | 62057.0487 | 61004.9592 | 1216.0992 | 1116.0992 | 1034.9241
5863 | 105.4733 | 0.95247 | 1.2097 | 0.99831 | 199.9652 | 46.8879 | 61532.9028 | 60917.3702 | 1214.3474 | 1114.3474 | 1034.9477
5864 | 98.1814 | 0.98186 | 1.1284 | 0.94265 | 175.1327 | 49.8419 | 57993.7811 | 55582.6978 | 1107.654 | 1007.654 | 1034.9735
5865 | 84.4541 | 1.0318 | 1.0378 | 0.99605 | 185.8164 | 48.4674 | 54510.1244 | 52876.5169 | 1053.5303 | 953.5303 | 1035.0182
5866 | 105.2019 | 0.91756 | 1.0878 | 0.98965 | 205.5376 | 53.4297 | 62882.0105 | 62943.2164 | 1254.8643 | 1154.8643 | 1035.0445
5867 | 93.3016 | 0.9681 | 1.0128 | 0.98353 | 189.2127 | 46.409 | 50609.3392 | 51165.7057 | 1019.3141 | 919.3141 | 1035.0607
5868 | 104.0459 | 1.0561 | 0.74429 | 0.95581 | 185.767 | 50.0704 | 41307.985 | 41939.1428 | 834.7829 | 734.7829 | 1035.1791
5869 | 100.4933 | 0.97952 | 1.0134 | 0.95378 | 209.3327 | 52.1335 | 55733.0504 | 55823.8893 | 1112.4778 | 1012.4778 | 1035.2064
5870 | 96.6047 | 0.99938 | 1.0385 | 1.056 | 199.6473 | 49.4548 | 59206.0161 | 58262.2147 | 1161.2443 | 1061.2443 | 1035.3681
5871 | 95.5502 | 0.96362 | 1.1208 | 1.0084 | 200.2931 | 48.2475 | 59172.211 | 59008.7308 | 1176.1746 | 1076.1746 | 1035.3743
5872 | 100.9024 | 1.0777 | 1.1152 | 1.0743 | 182.3605 | 55.0807 | 72163.3215 | 74133.8034 | 1478.6761 | 1378.6761 | 1035.3842
5873 | 97.87 | 1.0509 | 0.99437 | 1.0087 | 196.668 | 48.0246 | 53303.4626 | 52350.4009 | 1043.008 | 943.008 | 1035.3863
5874 | 95.6255 | 0.97417 | 0.98735 | 1.0676 | 207.4593 | 49.2862 | 56750.2859 | 56848.27 | 1132.9654 | 1032.9654 | 1035.4109
5875 | 104.5881 | 1.0427 | 1.0325 | 0.99269 | 199.5098 | 49.3273 | 56134.5657 | 56740.7235 | 1130.8145 | 1030.8145 | 1035.4678
5876 | 103.3638 | 1.0231 | 0.97252 | 1.0303 | 187.9147 | 55.4577 | 61623.0886 | 62010.5494 | 1236.211 | 1136.211 | 1035.5339
5877 | 93.6595 | 1.0465 | 0.90682 | 1.0234 | 198.5704 | 50.0578 | 51559.5726 | 50430.492 | 1004.6098 | 904.6098 | 1035.5733
5878 | 96.1324 | 0.94904 | 0.93937 | 1.074 | 202.3692 | 47.7311 | 52712.1266 | 53418.9835 | 1064.3797 | 964.3797 | 1035.624
5879 | 101.1923 | 0.96375 | 1.123 | 1.0044 | 197.423 | 50.1912 | 61704.4328 | 59326.8835 | 1182.5377 | 1082.5377 | 1035.6837
5880 | 109.8695 | 0.96944 | 1.0106 | 1.0055 | 196.7361 | 51.913 | 58478.4646 | 56140.1753 | 1118.8035 | 1018.8035 | 1035.7336
5881 | 106.1764 | 0.95929 | 0.98241 | 1.0672 | 199.8644 | 45.5006 | 52540.2326 | 51511.7727 | 1026.2355 | 926.2355 | 1035.7477
5882 | 109.7886 | 0.98733 | 1.0242 | 1.0482 | 197.8581 | 47.819 | 56716.9927 | 56890.5207 | 1133.8104 | 1033.8104 | 1035.7644
5883 | 97.1016 | 0.97877 | 1.0274 | 0.97144 | 204.9424 | 50.5225 | 55431.2426 | 55560.9134 | 1107.2183 | 1007.2183 | 1035.7779
5884 | 102.8904 | 1.0019 | 0.92093 | 0.99117 | 187.8972 | 53.4549 | 54491.6196 | 53395.2076 | 1063.9042 | 963.9042 | 1035.7876
5885 | 106.2552 | 0.95116 | 1.1668 | 1.0246 | 211.9629 | 51.4999 | 66985.1776 | 64707.3181 | 1290.1464 | 1190.1464 | 1035.8103
5886 | 99.9406 | 1.0052 | 0.82358 | 0.89564 | 199.9908 | 43.3137 | 36500.6879 | 37371.2448 | 743.4249 | 643.4249 | 1035.8203
5887 | 94.6684 | 1.0455 | 1.0258 | 1.055 | 216.6371 | 48.4316 | 57420.8578 | 57260.9942 | 1141.2199 | 1041.2199 | 1035.9387
5888 | 92.0183 | 1.0197 | 0.93443 | 0.94718 | 196.0897 | 50.5657 | 49694.9866 | 49807.1691 | 992.1434 | 892.1434 | 1035.9603
5889 | 89.9252 | 1.006 | 0.90026 | 0.9987 | 207.1874 | 52.8827 | 52537.5746 | 53219.9289 | 1060.3986 | 960.3986 | 1036.0298
5890 | 98.9134 | 1.0552 | 1.312 | 1.0528 | 207.3378 | 48.7168 | 72584.5144 | 72713.4696 | 1450.2694 | 1350.2694 | 1036.0303
5891 | 126.9548 | 0.99684 | 0.64544 | 1.0353 | 211.7156 | 49.3818 | 39458.7944 | 39378.605 | 783.5721 | 683.5721 | 1036.0339
5892 | 106.0462 | 1.0572 | 0.9532 | 1.0094 | 200.8965 | 48.9276 | 52761.686 | 52955.7323 | 1055.1146 | 955.1146 | 1036.035
5893 | 108.1247 | 0.99808 | 1.0343 | 1.016 | 206.4495 | 47.9813 | 55802.6482 | 55765.4297 | 1111.3086 | 1011.3086 | 1036.058
5894 | 94.6267 | 1.0819 | 1.3087 | 0.96381 | 213.1464 | 51.1073 | 69907.0555 | 68869.5854 | 1373.3917 | 1273.3917 | 1036.1057
5895 | 84.4305 | 1.0156 | 0.80142 | 0.93596 | 194.4123 | 52.3957 | 43988.955 | 42695.473 | 849.9095 | 749.9095 | 1036.1203
5896 | 99.3848 | 1.0394 | 0.90702 | 0.98546 | 188.1788 | 51.6018 | 51641.9474 | 50094.2954 | 997.8859 | 897.8859 | 1036.1474
5897 | 90.4212 | 0.97419 | 0.83214 | 1.0076 | 193.5086 | 50.5124 | 46997.0904 | 46883.9521 | 933.679 | 833.679 | 1036.197
5898 | 114.4986 | 1.0506 | 1.1241 | 0.95631 | 203.1465 | 49.882 | 59827.4988 | 60732.7372 | 1210.6547 | 1110.6547 | 1036.2154
5899 | 106.9224 | 0.96632 | 1.1541 | 1.0192 | 216.0858 | 47.3117 | 60757.0684 | 63466.9819 | 1265.3396 | 1165.3396 | 1036.2322
5900 | 87.8323 | 0.96172 | 1.0459 | 1.0138 | 208.6478 | 53.8145 | 61811.8982 | 63044.1855 | 1256.8837 | 1156.8837 | 1036.2614
5901 | 107.8247 | 1.008 | 1.1342 | 1.0538 | 195.1793 | 52.9539 | 69243.1074 | 70792.3467 | 1411.8469 | 1311.8469 | 1036.2785
5902 | 97.7806 | 0.91056 | 0.69371 | 0.98858 | 199.2468 | 47.0322 | 36641.0779 | 37124.007 | 738.4801 | 638.4801 | 1036.2842
5903 | 106.5164 | 0.94163 | 0.79068 | 0.91878 | 188.4005 | 49.1804 | 40848.7813 | 41372.9377 | 823.4588 | 723.4588 | 1036.3552
5904 | 97.4941 | 1.0399 | 1.1834 | 1.04 | 201.6553 | 50.5285 | 67509.7152 | 66580.2338 | 1327.6047 | 1227.6047 | 1036.4012
5905 | 109.4951 | 1.0142 | 0.93025 | 1.0193 | 200.3077 | 51.8302 | 55100.6836 | 54450.805 | 1085.0161 | 985.0161 | 1036.4644
5906 | 103.5712 | 1.0024 | 1.004 | 1.0536 | 194.3809 | 51.2533 | 59730.3007 | 60708.9283 | 1210.1786 | 1110.1786 | 1036.4693
5907 | 104.9107 | 0.9731 | 0.91121 | 0.97521 | 207.913 | 49.1488 | 48900.1547 | 47989.4152 | 955.7883 | 855.7883 | 1036.4963
5908 | 120.4124 | 1.0665 | 1.1917 | 1.0866 | 202.2529 | 48.4362 | 69141.4858 | 68601.9146 | 1368.0383 | 1268.0383 | 1036.6515
5909 | 112.2534 | 1.0426 | 1.0956 | 1.0137 | 193.117 | 48.6184 | 59878.4262 | 60964.8989 | 1215.298 | 1115.298 | 1036.6581
5910 | 100.0249 | 0.97256 | 1.0413 | 0.97509 | 199.8914 | 50.7894 | 56708.5207 | 56350.0666 | 1123.0013 | 1023.0013 | 1036.6662
5911 | 100.5129 | 1.0107 | 1.1315 | 0.95988 | 195.2008 | 47.5716 | 56694.1052 | 56559.408 | 1127.1882 | 1027.1882 | 1036.6935
5912 | 103.1412 | 0.91634 | 0.96782 | 0.92726 | 193.5068 | 53.5254 | 53287.5603 | 54134.8186 | 1078.6964 | 978.6964 | 1036.7201
5913 | 106.3842 | 1.0586 | 0.88019 | 1.0693 | 195.8886 | 51.0106 | 53952.3687 | 53251.2242 | 1061.0245 | 961.0245 | 1036.7511
5914 | 114.9466 | 1.0793 | 1.1245 | 0.97408 | 190.4613 | 46.7745 | 57227.4851 | 56739.7895 | 1130.7958 | 1030.7958 | 1036.7749
5915 | 102.7748 | 0.97076 | 0.95997 | 0.94814 | 201.1316 | 53.0176 | 53746.3266 | 53041.9732 | 1056.8395 | 956.8395 | 1036.7763
5916 | 107.2949 | 1.0096 | 0.97341 | 1.0757 | 198.3084 | 49.6287 | 57539.8074 | 58684.4207 | 1169.6884 | 1069.6884 | 1036.8146
5917 | 114.6006 | 1.0502 | 0.98082 | 0.98203 | 182.6515 | 48.9832 | 53258.3729 | 52771.297 | 1051.4259 | 951.4259 | 1036.8479
5918 | 102.0147 | 1.0833 | 0.86491 | 1.0141 | 201.9379 | 50.0129 | 49594.4205 | 50470.9486 | 1005.419 | 905.419 | 1036.8624
5919 | 92.8077 | 0.98586 | 0.78906 | 0.93924 | 190.6175 | 52.9569 | 44283.3267 | 44087.1697 | 877.7434 | 777.7434 | 1036.8808
5920 | 99.3753 | 1.0394 | 1.1657 | 0.98616 | 183.5214 | 50.9253 | 63986.7606 | 61875.6082 | 1233.5122 | 1133.5122 | 1036.8965
5921 | 100.5958 | 0.98128 | 0.94834 | 1.0053 | 203.7097 | 49.9398 | 52742.1869 | 54772.5322 | 1091.4506 | 991.4506 | 1036.9046
5922 | 106.2001 | 1.0266 | 0.98025 | 1.0714 | 188.3163 | 50.1086 | 58278.9354 | 59024.7042 | 1176.4941 | 1076.4941 | 1037.0598
5923 | 98.6216 | 1.0044 | 1.0011 | 0.94235 | 188.1199 | 47.9408 | 50163.2652 | 51578.8267 | 1027.5765 | 927.5765 | 1037.1144
5924 | 107.4478 | 0.95528 | 1.05 | 0.99994 | 215.3927 | 50.9342 | 58921.9826 | 60629.7735 | 1208.5955 | 1108.5955 | 1037.129
5925 | 118.2907 | 1.0504 | 0.9645 | 0.9609 | 192.7757 | 51.2888 | 54099.5552 | 54898.3804 | 1093.9676 | 993.9676 | 1037.1639
5926 | 108.9265 | 0.9144 | 1.0108 | 1.0527 | 213.5979 | 49.6058 | 57937.6784 | 59143.1779 | 1178.8636 | 1078.8636 | 1037.2212
5927 | 103.7788 | 1.0703 | 0.61688 | 1.0311 | 198.8743 | 47.5212 | 35704.9752 | 35878.9282 | 713.5786 | 613.5786 | 1037.3109
5928 | 89.1046 | 1.0311 | 0.83465 | 0.96291 | 194.9913 | 51.0291 | 45895.3335 | 44898.8763 | 893.9775 | 793.9775 | 1037.3123
5929 | 110.114 | 0.99012 | 0.93844 | 1.0261 | 192.2052 | 48.9967 | 52716.7607 | 52790.1875 | 1051.8037 | 951.8037 | 1037.3376
5930 | 102.2149 | 1.0124 | 0.97704 | 1.0823 | 168.4339 | 47.0161 | 54752.2048 | 55040.2661 | 1096.8053 | 996.8053 | 1037.3448
5931 | 108.705 | 0.9784 | 0.89728 | 1.0402 | 181.0694 | 50.0201 | 52187.2792 | 53155.7596 | 1059.1152 | 959.1152 | 1037.47
5932 | 108.5428 | 1.039 | 0.86388 | 0.88953 | 177.2089 | 47.0336 | 41624.0692 | 42345.8121 | 842.9162 | 742.9162 | 1037.4758
5933 | 96.3231 | 0.96397 | 0.95353 | 1.0398 | 178.7585 | 48.637 | 52916.6344 | 50791.539 | 1011.8308 | 911.8308 | 1037.5623
5934 | 90.1077 | 1.0167 | 0.84943 | 0.97472 | 206.439 | 48.5161 | 44819.9311 | 46153.0385 | 919.0608 | 819.0608 | 1037.5781
5935 | 101.1661 | 1.0214 | 0.91462 | 0.88783 | 207.1514 | 49.4934 | 45511.6645 | 45626.9307 | 908.5386 | 808.5386 | 1037.606
5936 | 107.5282 | 0.94469 | 0.93976 | 1.0358 | 196.0647 | 51.2691 | 55308.3865 | 56442.8279 | 1124.8566 | 1024.8566 | 1037.6139
5937 | 107.5148 | 0.96963 | 1.1393 | 1.0631 | 189.563 | 48.6307 | 64160.2752 | 65328.517 | 1302.5703 | 1202.5703 | 1037.7106
5938 | 92.5683 | 1.0363 | 1.1129 | 1.0472 | 191.1415 | 51.6411 | 65329.1323 | 65581.0945 | 1307.6219 | 1207.6219 | 1037.7348
5939 | 92.9576 | 1.0188 | 1.0349 | 0.99772 | 205.9722 | 49.1619 | 55623.4284 | 55688.8467 | 1109.7769 | 1009.7769 | 1037.8983
5940 | 95.1011 | 1.0456 | 0.6356 | 1.0374 | 207.5531 | 47.1801 | 36007.136 | 34939.9557 | 694.7991 | 594.7991 | 1037.8994
5941 | 100.7774 | 0.95371 | 0.88594 | 1.0536 | 190.4793 | 49.7838 | 51446.4446 | 52434.8468 | 1044.6969 | 944.6969 | 1038.0489
5942 | 124.9122 | 0.91756 | 0.94131 | 1.1207 | 195.6748 | 50.5834 | 59354.0804 | 60738.7916 | 1210.7758 | 1110.7758 | 1038.1579
5943 | 105.5185 | 0.93261 | 0.99301 | 0.99323 | 193.5556 | 48.4267 | 52721.4289 | 53878.7377 | 1073.5748 | 973.5748 | 1038.2095
5944 | 109.4907 | 1.0528 | 1.0407 | 0.94974 | 196.4851 | 51.5808 | 57125.2534 | 58297.7432 | 1161.9549 | 1061.9549 | 1038.3085
5945 | 99.3891 | 1.0003 | 0.84973 | 1.0049 | 203.6461 | 49.23 | 47134.4902 | 46734.52 | 930.6904 | 830.6904 | 1038.3489
5946 | 92.9586 | 0.85425 | 0.79714 | 0.94756 | 186.5121 | 52.0886 | 43667.2127 | 43030.848 | 856.617 | 756.617 | 1038.4017
5947 | 103.0945 | 1.0503 | 0.99266 | 1.0396 | 205.8066 | 51.3454 | 58753.226 | 58804.7738 | 1172.0955 | 1072.0955 | 1038.4606
5948 | 105.4945 | 0.97038 | 1.0164 | 0.9407 | 206.6187 | 51.8314 | 55072.4266 | 53281.8971 | 1061.6379 | 961.6379 | 1038.4748
5949 | 117.3696 | 0.96762 | 0.93383 | 0.98183 | 187.5851 | 49.8411 | 51545.4311 | 52606.0524 | 1048.121 | 948.121 | 1038.5188
5950 | 78.7039 | 1.0241 | 1.0509 | 1.0156 | 196.1933 | 49.3752 | 56874.2146 | 55994.6516 | 1115.893 | 1015.893 | 1038.5867
5951 | 102.7258 | 1.0601 | 0.95982 | 0.92016 | 204.8378 | 52.351 | 52141.413 | 50531.1972 | 1006.6239 | 906.6239 | 1038.6201
5952 | 97.4479 | 0.98274 | 1.1179 | 1.0207 | 196.0339 | 47.5991 | 59069.7375 | 58155.1691 | 1159.1034 | 1059.1034 | 1038.6742
5953 | 92.7739 | 0.94503 | 0.92822 | 1.0563 | 213.9573 | 50.9924 | 54680.7072 | 55842.8338 | 1112.8567 | 1012.8567 | 1038.7276
5954 | 91.8609 | 0.9688 | 1.0847 | 1.024 | 200.0161 | 49.1878 | 59213.9552 | 60720.9294 | 1210.4186 | 1110.4186 | 1038.7979
5955 | 102.5365 | 1.0982 | 1.1011 | 0.91503 | 218.5261 | 48.0734 | 54066.9714 | 54208.2997 | 1080.166 | 980.166 | 1038.8125
5956 | 107.9604 | 0.92215 | 0.8049 | 0.90864 | 210.4201 | 48.7275 | 40698.9859 | 40433.7545 | 804.6751 | 704.6751 | 1038.8148
5957 | 70.6887 | 1.0039 | 1.007 | 1.0417 | 191.4369 | 51.5032 | 57871.8308 | 59802.3547 | 1192.0471 | 1092.0471 | 1038.8271
5958 | 104.3975 | 1.0267 | 1.2197 | 0.96481 | 198.2249 | 50.6201 | 65191.831 | 65542.3915 | 1306.8478 | 1206.8478 | 1038.8492
5959 | 77.5355 | 0.90276 | 0.84266 | 1.0115 | 194.2676 | 50.749 | 47000.2365 | 47688.5166 | 949.7703 | 849.7703 | 1038.8495
5960 | 91.374 | 1.0217 | 0.86193 | 1.0704 | 196.1207 | 53.0616 | 54107.3615 | 53957.1576 | 1075.1432 | 975.1432 | 1038.869
5961 | 105.1032 | 0.98462 | 1.0066 | 0.99525 | 198.6918 | 49.3698 | 54766.6544 | 54734.6915 | 1090.6938 | 990.6938 | 1038.9001
5962 | 116.375 | 1.0777 | 0.9471 | 0.97381 | 202.7889 | 45.4662 | 47838.8519 | 46913.8268 | 934.2765 | 834.2765 | 1038.9028
5963 | 93.9232 | 1.0006 | 0.96619 | 0.96961 | 204.7162 | 48.6965 | 50401.5731 | 49575.7637 | 987.5153 | 887.5153 | 1038.918
5964 | 111.9995 | 0.97293 | 0.69497 | 1.0379 | 198.3325 | 52.102 | 43456.2008 | 44325.9257 | 882.5185 | 782.5185 | 1038.9542
5965 | 97.0256 | 1.1375 | 0.99982 | 0.92407 | 212.6911 | 47.6803 | 49526.4995 | 48897.6638 | 973.9533 | 873.9533 | 1039.0028
5966 | 89.6957 | 0.97347 | 0.81301 | 1.0496 | 198.9932 | 53.571 | 50589.4963 | 50276.2347 | 1001.5247 | 901.5247 | 1039.0115
5967 | 99.6008 | 1.0334 | 1.0415 | 1.0287 | 211.8205 | 51.014 | 60119.5254 | 58507.2493 | 1166.145 | 1066.145 | 1039.0159
5968 | 106.1789 | 1.0424 | 1.1342 | 0.94925 | 216.8491 | 50.9029 | 60655.0775 | 59969.9022 | 1195.398 | 1095.398 | 1039.0866
5969 | 93.9615 | 0.92558 | 0.83814 | 1.0335 | 212.1106 | 50.3805 | 48232.819 | 48850.6505 | 973.013 | 873.013 | 1039.1547
5970 | 92.3487 | 0.97522 | 0.93773 | 1.045 | 194.2963 | 47.3944 | 50904.2807 | 49553.5327 | 987.0707 | 887.0707 | 1039.2505
5971 | 83.9113 | 0.99386 | 0.99443 | 1.0476 | 198.9541 | 49.0241 | 55358.1455 | 56627.8655 | 1128.5573 | 1028.5573 | 1039.2612
5972 | 101.0031 | 1.0013 | 0.98191 | 1.0586 | 193.539 | 53.0796 | 60736.4402 | 62653.6161 | 1249.0723 | 1149.0723 | 1039.2662
5973 | 109.0994 | 1.0728 | 0.8027 | 1.0987 | 202.9843 | 47.7147 | 47867.6035 | 47998.0418 | 955.9608 | 855.9608 | 1039.3173
5974 | 89.3682 | 1.0224 | 0.89286 | 1.0062 | 197.0253 | 51.6932 | 51363.5157 | 53392.2598 | 1063.8452 | 963.8452 | 1039.351
5975 | 115.8245 | 1.0375 | 0.96956 | 1.0061 | 194.489 | 51.3164 | 56421.2126 | 56473.1247 | 1125.4625 | 1025.4625 | 1039.4109
5976 | 88.1888 | 0.91769 | 0.91805 | 0.96251 | 208.4803 | 49.1654 | 47631.7144 | 48729.9999 | 970.6 | 870.6 | 1039.4473
5977 | 108.8221 | 0.93737 | 0.95502 | 1.0296 | 201.9854 | 52.1103 | 56759.6273 | 56673.9898 | 1129.4798 | 1029.4798 | 1039.4912
5978 | 106.723 | 1.0168 | 0.96323 | 0.98433 | 193.3415 | 46.3789 | 49200.1121 | 48527.8146 | 966.5563 | 866.5563 | 1039.5599
5979 | 94.0566 | 1.0094 | 1.023 | 1.0287 | 197.2316 | 46.966 | 54082.9112 | 54289.36 | 1081.7872 | 981.7872 | 1039.5985
5980 | 99.8385 | 0.97723 | 0.96262 | 1.0031 | 202.3826 | 53.2304 | 56796.9171 | 57023.4634 | 1136.4693 | 1036.4693 | 1039.7189
5981 | 115.6499 | 1.0613 | 1.0614 | 1.0155 | 198.3566 | 44.3012 | 53386.7796 | 53056.218 | 1057.1244 | 957.1244 | 1039.7233
5982 | 99.722 | 0.9326 | 1.1552 | 0.9899 | 215.8044 | 44.4864 | 55223.9466 | 53462.1999 | 1065.244 | 965.244 | 1039.7248
5983 | 98.2688 | 1.0491 | 1.1611 | 0.98821 | 197.6091 | 50.8472 | 63780.0957 | 61820.9281 | 1232.4186 | 1132.4186 | 1039.7525
5984 | 96.5274 | 0.93943 | 1.0522 | 1.0005 | 204.9782 | 49.0469 | 56282.957 | 56998.0174 | 1135.9603 | 1035.9603 | 1039.8261
5985 | 96.6128 | 0.91904 | 0.92326 | 0.96611 | 189.1829 | 52.0036 | 51192.426 | 51099.9049 | 1017.9981 | 917.9981 | 1039.8364
5986 | 110.7319 | 1.0281 | 0.88205 | 1.0525 | 182.8343 | 47.4765 | 49660.8276 | 48299.3765 | 961.9875 | 861.9875 | 1039.9407
5987 | 89.9913 | 0.99239 | 1.1108 | 1.066 | 215.8827 | 46.4829 | 59410.5728 | 61983.2773 | 1235.6655 | 1135.6655 | 1039.9445
5988 | 91.4581 | 1.0615 | 1.2307 | 1.0222 | 220.1599 | 51.6165 | 70161.3484 | 68701.0695 | 1370.0214 | 1270.0214 | 1039.9654
5989 | 102.723 | 0.96056 | 1.0707 | 0.93557 | 192.5418 | 48.8785 | 53978.6404 | 53976.164 | 1075.5233 | 975.5233 | 1039.9932
5990 | 99.3365 | 0.91826 | 0.78055 | 1.0601 | 203.2057 | 50.8685 | 46935.8341 | 47836.0019 | 952.72 | 852.72 | 1040.1503
5991 | 110.007 | 0.98736 | 1.1186 | 1.0158 | 201.8921 | 49.9708 | 62409.1555 | 62333.9063 | 1242.6781 | 1142.6781 | 1040.1891
5992 | 117.3678 | 0.96453 | 1.0335 | 0.93386 | 183.8761 | 50.8765 | 55045.8996 | 54308.7796 | 1082.1756 | 982.1756 | 1040.2049
5993 | 85.1385 | 0.97542 | 0.94115 | 0.97932 | 198.7368 | 45.4254 | 45839.2172 | 46087.5355 | 917.7507 | 817.7507 | 1040.2495
5994 | 100.1311 | 0.9604 | 1.1243 | 1.0096 | 195.4524 | 44.8627 | 55433.2426 | 55874.3492 | 1113.487 | 1013.487 | 1040.3044
5995 | 98.6025 | 0.89301 | 1.0686 | 0.97799 | 216.0011 | 44.6131 | 50767.4492 | 50808.2628 | 1012.1653 | 912.1653 | 1040.3341
5996 | 108.6723 | 0.97861 | 1.1307 | 1.0516 | 204.2269 | 51.5561 | 66991.5951 | 65701.4083 | 1310.0282 | 1210.0282 | 1040.3621
5997 | 93.5908 | 0.96779 | 1.3027 | 1.0002 | 188.3951 | 50.0248 | 69901.5781 | 68674.4481 | 1369.489 | 1269.489 | 1040.376
5998 | 96.2779 | 1.0012 | 0.71775 | 0.96864 | 203.9408 | 54.3451 | 43225.9824 | 43142.8161 | 858.8563 | 758.8563 | 1040.3837
5999 | 96.4651 | 0.97882 | 1.0752 | 1.0136 | 204.4648 | 49.4433 | 58758.2987 | 57730.3685 | 1150.6074 | 1050.6074 | 1040.4269
6000 | 90.1887 | 1.0682 | 1.2303 | 0.97049 | 197.5648 | 53.0128 | 68603.8324 | 67494.468 | 1345.8894 | 1245.8894 | 1040.5107
6001 | 104.6639 | 1.0831 | 1.0694 | 0.96637 | 201.1731 | 47.3972 | 54557.3823 | 54653.2082 | 1089.0642 | 989.0642 | 1040.515
6002 | 91.6008 | 1.0659 | 1.0147 | 1.0573 | 200.4024 | 48.3485 | 56792.2943 | 57259.7092 | 1141.1942 | 1041.1942 | 1040.5162
6003 | 100.4993 | 0.95743 | 1.0033 | 0.92305 | 221.0717 | 48.4137 | 49716.4934 | 49726.2186 | 990.5244 | 890.5244 | 1040.5678
6004 | 102.6481 | 1.0796 | 1.1804 | 0.99024 | 204.7783 | 49.5561 | 63619.951 | 64192.7746 | 1279.8555 | 1179.8555 | 1040.6061
6005 | 106.8741 | 0.96771 | 1.3473 | 0.95941 | 185.7233 | 50.1957 | 70258.8483 | 71708.3609 | 1430.1672 | 1330.1672 | 1040.6246
6006 | 87.3174 | 1.0378 | 1.0846 | 1.0573 | 205.1836 | 50.5679 | 62779.5558 | 63591.7985 | 1267.836 | 1167.836 | 1040.6349
6007 | 114.6534 | 1.0426 | 0.98306 | 0.97896 | 199.7089 | 50.0013 | 54296.7166 | 54505.1215 | 1086.1024 | 986.1024 | 1040.656
6008 | 85.0614 | 1.0223 | 1.0247 | 0.99326 | 195.0989 | 52.5538 | 58253.0097 | 56288.2998 | 1121.766 | 1021.766 | 1040.7116
6009 | 106.1124 | 0.93258 | 1.1042 | 0.96297 | 197.8245 | 45.7449 | 53364.2324 | 53662.3431 | 1069.2469 | 969.2469 | 1040.7456
6010 | 90.1024 | 0.99524 | 0.99209 | 0.99769 | 188.9327 | 49.7761 | 53920.8623 | 53378.7041 | 1063.5741 | 963.5741 | 1040.7649
6011 | 100.6986 | 0.95383 | 1.1383 | 1.1007 | 193.0063 | 52.1837 | 70590.1225 | 69593.4041 | 1387.8681 | 1287.8681 | 1040.8021
6012 | 109.8842 | 1.0172 | 0.7758 | 1.0403 | 206.13 | 46.6038 | 43026.12 | 43185.7104 | 859.7142 | 759.7142 | 1040.8077
6013 | 94.4424 | 0.95833 | 1.3492 | 1.011 | 201.7376 | 53.9497 | 78678.9176 | 79533.5605 | 1586.6712 | 1486.6712 | 1040.8078
6014 | 119.3662 | 0.98486 | 1.008 | 0.98411 | 218.8626 | 50.828 | 56616.4331 | 56681.2117 | 1129.6242 | 1029.6242 | 1040.9344
6015 | 109.9225 | 0.98002 | 1.1249 | 0.9864 | 223.3591 | 51.7711 | 63243.7554 | 61731.5041 | 1230.6301 | 1130.6301 | 1040.9489
6016 | 99.6943 | 0.89227 | 1.0261 | 0.95278 | 203.0043 | 50.548 | 54120.0619 | 53126.0422 | 1058.5208 | 958.5208 | 1040.9536
6017 | 101.8113 | 0.97217 | 1.0691 | 1.1041 | 196.7632 | 50.5832 | 64912.9747 | 65624.3549 | 1308.4871 | 1208.4871 | 1040.9873
6018 | 98.0922 | 1.0998 | 0.95644 | 0.88639 | 198.2586 | 48.8382 | 46870.5765 | 45619.3154 | 908.3863 | 808.3863 | 1041.0483
6019 | 96.2095 | 1.0827 | 1.2549 | 0.89617 | 219.4837 | 50.2697 | 61991.3333 | 60999.0468 | 1215.9809 | 1115.9809 | 1041.0665
6020 | 99.8555 | 1.0435 | 1.0392 | 0.98052 | 190.0953 | 51.7992 | 58371.2417 | 57423.5843 | 1144.4717 | 1044.4717 | 1041.0783
6021 | 91.4963 | 0.99333 | 1.0528 | 1.0035 | 201.2426 | 47.1103 | 54254.8833 | 56112.2413 | 1118.2448 | 1018.2448 | 1041.1166
6022 | 101.8917 | 0.93357 | 0.8996 | 1.0516 | 189.6328 | 49.1534 | 51365.4266 | 50941.4507 | 1014.829 | 914.829 | 1041.1714
6023 | 102.9003 | 1.0207 | 0.97633 | 0.97692 | 192.4527 | 46.0496 | 48950.7346 | 47286.3563 | 941.7271 | 841.7271 | 1041.1942
6024 | 91.7771 | 1.004 | 1.2313 | 0.98645 | 207.5661 | 44.3 | 58095.6059 | 56868.8929 | 1133.3779 | 1033.3779 | 1041.2199
6025 | 114.9542 | 0.9483 | 1.0733 | 0.99586 | 191.3678 | 54.3111 | 64164.449 | 65591.8459 | 1307.8369 | 1207.8369 | 1041.2338
6026 | 90.0979 | 0.88806 | 0.88923 | 0.9194 | 196.3664 | 48.2694 | 43521.6359 | 45216.5536 | 900.3311 | 800.3311 | 1041.2961
6027 | 87.1151 | 1.047 | 1.1836 | 0.91436 | 198.884 | 51.2558 | 60342.847 | 60312.5784 | 1202.2516 | 1102.2516 | 1041.331
6028 | 97.1037 | 0.97875 | 1.1434 | 1.0314 | 202.9385 | 52.1804 | 66699.7886 | 67591.55 | 1347.831 | 1247.831 | 1041.3678
6029 | 95.0548 | 1.0116 | 0.9422 | 0.92698 | 208.6834 | 47.2239 | 45995.147 | 46299.6518 | 921.993 | 821.993 | 1041.3886
6030 | 102.872 | 0.94598 | 1.0216 | 1.0017 | 185.99 | 49.7904 | 55984.637 | 55471.0026 | 1105.4201 | 1005.4201 | 1041.4161
6031 | 91.8282 | 0.95551 | 0.94777 | 0.93702 | 206.393 | 49.4403 | 48451.1666 | 48504.5107 | 966.0902 | 866.0902 | 1041.4982
6032 | 108.4276 | 0.9387 | 1.0555 | 0.98684 | 199.5795 | 48.2099 | 55321.9669 | 54721.6315 | 1090.4326 | 990.4326 | 1041.5056
6033 | 96.3749 | 0.94649 | 0.93228 | 1.0128 | 188.0853 | 49.1864 | 51117.1302 | 52446.3456 | 1044.9269 | 944.9269 | 1041.5166
6034 | 118.622 | 0.98266 | 1.0606 | 0.98456 | 192.0874 | 50.4122 | 58708.3718 | 57502.9125 | 1146.0583 | 1046.0583 | 1041.5245
6035 | 107.6761 | 1.0111 | 1.0243 | 1.047 | 216.7993 | 48.936 | 58024.2048 | 56151.37 | 1119.0274 | 1019.0274 | 1041.5251
6036 | 109.7738 | 1.0026 | 0.97143 | 0.92182 | 204.5583 | 52.5571 | 53052.9014 | 53628.503 | 1068.5701 | 968.5701 | 1041.7094
6037 | 110.8896 | 0.95391 | 1.1078 | 1.0915 | 188.4189 | 46.6486 | 61527.4902 | 60020.3095 | 1196.4062 | 1096.4062 | 1041.7126
6038 | 98.9933 | 0.97016 | 0.98872 | 1.0183 | 198.3352 | 53.8868 | 59629.7569 | 59692.8429 | 1189.8569 | 1089.8569 | 1041.7408
6039 | 79.1654 | 1.0034 | 1.2799 | 1.0906 | 196.0934 | 53.2508 | 78753.9177 | 81816.9673 | 1632.3393 | 1532.3393 | 1041.7718
6040 | 91.1514 | 0.88009 | 0.95549 | 1.0045 | 194.7036 | 52.5314 | 54827.6946 | 54227.1842 | 1080.5437 | 980.5437 | 1041.783
6041 | 108.0723 | 0.97596 | 1.1435 | 1.0597 | 201.4093 | 54.853 | 72453.4516 | 73882.7231 | 1473.6545 | 1373.6545 | 1041.7942
6042 | 115.9312 | 0.9913 | 1.176 | 1.0779 | 212.4146 | 51.5675 | 71504.8618 | 71184.5981 | 1419.692 | 1319.692 | 1041.8062
6043 | 93.1593 | 1.0212 | 0.96335 | 0.99376 | 185.6975 | 54.7522 | 57811.4424 | 57827.4786 | 1152.5496 | 1052.5496 | 1041.8089
6044 | 97.2909 | 1.0365 | 0.92429 | 1.0188 | 199.2784 | 44.3222 | 46406.5578 | 47180.1934 | 939.6039 | 839.6039 | 1041.8106
6045 | 101.3449 | 0.97735 | 1.0318 | 0.94685 | 208.0963 | 50.4191 | 54458.2523 | 54176.406 | 1079.5281 | 979.5281 | 1041.8462
6046 | 103.2477 | 1.0301 | 1.0067 | 1.0215 | 194.2299 | 48.4709 | 55193.3869 | 56527.3404 | 1126.5468 | 1026.5468 | 1041.847
6047 | 89.692 | 0.95095 | 0.72523 | 0.93982 | 199.4927 | 47.648 | 36739.4524 | 36895.9445 | 733.9189 | 633.9189 | 1041.918
6048 | 96.3431 | 0.98125 | 0.77831 | 0.93432 | 208.1347 | 51.5919 | 42602.4178 | 42849.2732 | 852.9855 | 752.9855 | 1041.9343
6049 | 95.66 | 1.037 | 1.0868 | 0.96346 | 192.6869 | 48.8383 | 56177.3167 | 56816.6584 | 1132.3332 | 1032.3332 | 1041.9887
6050 | 103.2398 | 0.96939 | 1.0467 | 0.99418 | 204.9234 | 51.2736 | 58693.1539 | 61395.3617 | 1223.9072 | 1123.9072 | 1041.991
6051 | 101.1732 | 0.9043 | 0.92391 | 1.0339 | 202.3563 | 52.088 | 54721.9389 | 55649.245 | 1108.9849 | 1008.9849 | 1041.9951
6052 | 101.0152 | 0.99076 | 0.69055 | 0.92016 | 196.4187 | 48.2012 | 35648.349 | 36636.1865 | 728.7237 | 628.7237 | 1042.0431
6053 | 104.348 | 1.0842 | 0.80751 | 0.96761 | 186.3362 | 48.6278 | 43683.7422 | 43706.9335 | 870.1387 | 770.1387 | 1042.0563
6054 | 93.9663 | 0.96033 | 0.99159 | 0.97299 | 226.5221 | 47.3921 | 50227.4031 | 52770.6456 | 1051.4129 | 951.4129 | 1042.1218
6055 | 103.3483 | 0.9922 | 1.0081 | 1.0864 | 182.1428 | 47.0267 | 56506.7345 | 55073.6302 | 1097.4726 | 997.4726 | 1042.1957
6056 | 96.9049 | 1.0199 | 1.1472 | 0.90435 | 219.488 | 47.4886 | 54179.7206 | 54105.1989 | 1078.104 | 978.104 | 1042.2613
6057 | 94.154 | 0.97428 | 0.83055 | 1.0663 | 186.8578 | 54.7418 | 53686.3628 | 52930.0166 | 1054.6003 | 954.6003 | 1042.3528
6058 | 87.2752 | 1.0331 | 1.2065 | 1.0145 | 202.9362 | 49.9915 | 65900.925 | 63809.4712 | 1272.1894 | 1172.1894 | 1042.3675
6059 | 105.2968 | 1.0524 | 1.1611 | 1.0374 | 221.5196 | 48.8206 | 64434.783 | 63131.7635 | 1258.6353 | 1158.6353 | 1042.4063
6060 | 104.8588 | 0.99606 | 1.2993 | 0.97195 | 216.9291 | 55.4985 | 76101.2271 | 76821.0977 | 1532.422 | 1432.422 | 1042.4325
6061 | 112.573 | 0.96535 | 1.1618 | 1.0371 | 208.7517 | 49.9265 | 65790.8099 | 64270.9007 | 1281.418 | 1181.418 | 1042.4979
6062 | 110.083 | 1.0134 | 1.1411 | 1.0111 | 192.3304 | 39.2714 | 49882.3291 | 50641.7123 | 1008.8342 | 908.8342 | 1042.5633
6063 | 99.5001 | 0.99853 | 0.81902 | 1.1093 | 181.2441 | 49.6133 | 50184.1487 | 51111.0159 | 1018.2203 | 918.2203 | 1042.6289
6064 | 99.1385 | 0.97149 | 1.024 | 0.97371 | 198.5516 | 52.1065 | 57170.7473 | 57290.3085 | 1141.8062 | 1041.8062 | 1042.6622
6065 | 106.6476 | 0.87637 | 0.87185 | 1.0079 | 207.4441 | 49.6898 | 48515.6203 | 50794.5361 | 1011.8907 | 911.8907 | 1042.6743
6066 | 110.5482 | 0.98478 | 0.93233 | 0.98433 | 190.7907 | 49.4667 | 50972.6366 | 50975.7723 | 1015.5154 | 915.5154 | 1042.6752
6067 | 112.91 | 1.0025 | 1.2383 | 0.94577 | 203.9507 | 46.2827 | 59644.8755 | 58315.7779 | 1162.3156 | 1062.3156 | 1042.7697
6068 | 88.4869 | 0.9746 | 1.1258 | 1.0339 | 206.5563 | 50.0088 | 62727.2351 | 61926.958 | 1234.5392 | 1134.5392 | 1042.7989
6069 | 114.0501 | 0.88943 | 0.93663 | 1.0315 | 200.8603 | 52.0259 | 55743.1058 | 56886.0124 | 1133.7202 | 1033.7202 | 1042.8007
6070 | 84.9743 | 0.90616 | 1.2555 | 1.0247 | 205.4596 | 50.6826 | 69317.0321 | 68727.8327 | 1370.5567 | 1270.5567 | 1042.8071
6071 | 97.7537 | 0.99026 | 0.95446 | 0.9577 | 198.2655 | 50.6065 | 51356.0841 | 52703.2761 | 1050.0655 | 950.0655 | 1042.8278
6072 | 103.8949 | 0.97076 | 0.73446 | 0.998 | 198.3045 | 50.3013 | 42141.6834 | 42752.3993 | 851.048 | 751.048 | 1042.9233
6073 | 107.1265 | 0.94035 | 0.9121 | 0.92138 | 184.106 | 48.5111 | 45839.0396 | 46206.5368 | 920.1307 | 820.1307 | 1042.9407
6074 | 102.8766 | 0.89669 | 0.87969 | 1.0337 | 212.3566 | 54.7663 | 55063.7864 | 52266.3621 | 1041.3272 | 941.3272 | 1042.9416
6075 | 97.9649 | 1.0986 | 1.0899 | 1.0059 | 196.4888 | 52.1435 | 62975.5826 | 62608.8022 | 1248.176 | 1148.176 | 1042.9495
6076 | 99.1175 | 1.0476 | 0.83346 | 1.0763 | 191.8752 | 47.9106 | 48145.4863 | 47120.9086 | 938.4182 | 838.4182 | 1042.9697
6077 | 111.1481 | 1.0154 | 1.2003 | 0.96295 | 180.868 | 50.9069 | 64763.7684 | 65523.4756 | 1306.4695 | 1206.4695 | 1043.0845
6078 | 112.9066 | 1.0081 | 1.1468 | 0.97395 | 228.3476 | 47.3017 | 58445.8737 | 58160.1873 | 1159.2037 | 1059.2037 | 1043.1414
6079 | 115.8345 | 0.98649 | 1.0269 | 1.0354 | 183.6795 | 53.0894 | 62699.9595 | 58929.7347 | 1174.5947 | 1074.5947 | 1043.1529
6080 | 103.6452 | 1.0098 | 1.3625 | 0.85443 | 205.6776 | 51.4706 | 65513.3333 | 66269.6388 | 1321.3928 | 1221.3928 | 1043.2568
6081 | 115.363 | 0.98653 | 1.1489 | 1.017 | 200.722 | 53.5806 | 68904.3296 | 68721.0603 | 1370.4212 | 1270.4212 | 1043.3136
6082 | 96.5184 | 1.0599 | 1.0936 | 0.92405 | 201.873 | 50.3129 | 56189.9053 | 57061.0597 | 1137.2212 | 1037.2212 | 1043.3458
6083 | 104.6663 | 1.0528 | 0.76728 | 1.0023 | 175.487 | 46.874 | 41388.552 | 41518.9559 | 826.3791 | 726.3791 | 1043.3561
6084 | 97.7545 | 1.0208 | 1.0871 | 0.95778 | 213.932 | 50.0712 | 57344.6764 | 57567.2048 | 1147.3441 | 1047.3441 | 1043.3798
6085 | 96.8598 | 1.0119 | 1.1291 | 1.0754 | 199.764 | 49.0356 | 64544.2792 | 63988.323 | 1275.7665 | 1175.7665 | 1043.5012
6086 | 94.5681 | 1.079 | 0.84917 | 0.99404 | 186.2318 | 49.5859 | 47101.5699 | 47229.918 | 940.5984 | 840.5984 | 1043.5104
6087 | 87.4948 | 1.0306 | 0.81248 | 0.97807 | 209.4561 | 49.6512 | 44142.562 | 44855.7437 | 893.1149 | 793.1149 | 1043.5913
6088 | 104.3766 | 1.055 | 1.1591 | 1.0462 | 195.731 | 51.3802 | 68156.9332 | 67302.45 | 1342.049 | 1242.049 | 1043.7063
6089 | 107.9658 | 1.0041 | 0.99197 | 1.0225 | 206.4346 | 50.1343 | 56492.4084 | 57170.5451 | 1139.4109 | 1039.4109 | 1043.7536
6090 | 86.224 | 0.97327 | 0.99248 | 1.0997 | 200.2218 | 50.5078 | 59563.6235 | 57755.7627 | 1151.1153 | 1051.1153 | 1043.7642
6091 | 87.1574 | 0.93893 | 1.142 | 0.99381 | 203.2946 | 51.1059 | 62386.7066 | 62474.2587 | 1245.4852 | 1145.4852 | 1043.7869
6092 | 117.9668 | 0.90426 | 0.90581 | 0.97349 | 199.5435 | 48.4483 | 48089.2034 | 47887.564 | 953.7513 | 853.7513 | 1043.8449
6093 | 97.4073 | 0.94131 | 0.95435 | 1.0462 | 196.0321 | 45.9369 | 50272.144 | 51336.715 | 1022.7343 | 922.7343 | 1043.8939
6094 | 99.5199 | 1.0045 | 0.91755 | 0.98264 | 200.7014 | 49.6567 | 49936.1161 | 50321.7863 | 1002.4357 | 902.4357 | 1043.9276
6095 | 107.9335 | 0.96644 | 0.93879 | 1.0197 | 193.8731 | 49.6056 | 52853.4826 | 53906.2975 | 1074.1259 | 974.1259 | 1043.9678
6096 | 102.0871 | 1.0202 | 1.0009 | 1.0966 | 198.2759 | 44.0926 | 53185.17 | 52590.843 | 1047.8169 | 947.8169 | 1044.001
6097 | 110.9519 | 1.021 | 1.0886 | 1.0512 | 189.2677 | 52.5189 | 66234.729 | 66984.0668 | 1335.6813 | 1235.6813 | 1044.0455
6098 | 107.5358 | 0.97027 | 0.75146 | 1.0854 | 177.7499 | 52.6303 | 48595.8079 | 48685.4036 | 969.7081 | 869.7081 | 1044.0647
6099 | 94.8684 | 1.0376 | 1.1655 | 1.1351 | 196.709 | 48.0953 | 68560.5144 | 68560.2286 | 1367.2046 | 1267.2046 | 1044.0697
6100 | 88.5604 | 1.0205 | 1.1895 | 0.95341 | 202.7664 | 53.017 | 65120.7766 | 64795.144 | 1291.9029 | 1191.9029 | 1044.0865
6101 | 99.7983 | 1.0219 | 0.91035 | 1.0407 | 182.764 | 47.7062 | 50243.3343 | 48965.4495 | 975.309 | 875.309 | 1044.1751
6102 | 87.6919 | 1.0933 | 1.1656 | 1.0479 | 213.7627 | 51.3251 | 67823.7457 | 69916.3398 | 1394.3268 | 1294.3268 | 1044.2068
6103 | 98.7025 | 1.0191 | 1.1074 | 1.0125 | 197.3619 | 46.8663 | 57459.6058 | 57410.341 | 1144.2068 | 1044.2068 | 1044.2701
6104 | 96.4477 | 0.94661 | 1.0809 | 1.0071 | 209.0907 | 53.6804 | 63545.0001 | 63902.9775 | 1274.0596 | 1174.0596 | 1044.3042
6105 | 90.5366 | 1.0644 | 1.0064 | 1.0219 | 188.8084 | 48.7883 | 55066.4334 | 55220.6801 | 1100.4136 | 1000.4136 | 1044.3949
6106 | 91.57 | 1.0032 | 0.99558 | 0.97286 | 211.3014 | 49.8783 | 53103.4783 | 54251.8099 | 1081.0362 | 981.0362 | 1044.4717
6107 | 106.4382 | 0.96687 | 1.1216 | 0.9687 | 204.249 | 44.9567 | 53677.592 | 53075.4079 | 1057.5082 | 957.5082 | 1044.5088
6108 | 111.468 | 0.95986 | 1.2033 | 1.0911 | 207.2099 | 47.9498 | 68294.3492 | 67568.3008 | 1347.366 | 1247.366 | 1044.5134
6109 | 104.132 | 1.0082 | 0.89656 | 0.90686 | 199.5037 | 53.1626 | 49005.2429 | 48324.1543 | 962.4831 | 862.4831 | 1044.5239
6110 | 85.4334 | 0.96717 | 1.0998 | 0.99333 | 201.8198 | 52.3346 | 61702.4734 | 59977.9682 | 1195.5594 | 1095.5594 | 1044.7033
6111 | 125.5352 | 0.94672 | 0.98305 | 1.1002 | 216.3444 | 47.3693 | 57079.5794 | 56688.6418 | 1129.7728 | 1029.7728 | 1044.7563
6112 | 90.3899 | 0.97253 | 1.0615 | 0.98968 | 194.6025 | 48.3977 | 55295.0246 | 55501.3226 | 1106.0265 | 1006.0265 | 1044.8145
6113 | 105.5884 | 0.96119 | 0.73156 | 0.98135 | 183.5532 | 52.185 | 42944.4437 | 44432.2278 | 884.6446 | 784.6446 | 1044.8854
6114 | 103.6286 | 1.0252 | 0.90945 | 1.027 | 216.4411 | 48.9825 | 51168.3411 | 51947.7647 | 1034.9553 | 934.9553 | 1044.9083
6115 | 112.3114 | 1.0018 | 1.4352 | 1.0221 | 210.3639 | 48.1912 | 76328.3231 | 75849.6317 | 1512.9926 | 1412.9926 | 1044.9108
6116 | 104.1118 | 1.0536 | 0.97936 | 0.96011 | 226.9662 | 47.2717 | 49861.6917 | 50257.2261 | 1001.1445 | 901.1445 | 1044.978
6117 | 94.9437 | 0.98184 | 0.96163 | 1.0578 | 196.3993 | 46.3167 | 51627.5825 | 52532.4116 | 1046.6482 | 946.6482 | 1045.0035
6118 | 97.8874 | 0.96821 | 1.0347 | 0.94338 | 201.8087 | 48.7197 | 52376.079 | 51440.232 | 1024.8046 | 924.8046 | 1045.0578
6119 | 96.3738 | 0.99126 | 1.0833 | 0.97111 | 197.4225 | 52.2268 | 60127.8697 | 60424.0938 | 1204.4819 | 1104.4819 | 1045.1657
6120 | 115.2819 | 1.0671 | 0.85293 | 1.0282 | 202.1631 | 53.3408 | 53541.1073 | 54162.0338 | 1079.2407 | 979.2407 | 1045.182
6121 | 104.2854 | 1.0404 | 0.87764 | 1.0116 | 190.1364 | 53.1819 | 53177.2732 | 52956.5133 | 1055.1303 | 955.1303 | 1045.2711
6122 | 77.0851 | 1.0378 | 0.93339 | 1.0512 | 195.8925 | 51.7715 | 55136.2496 | 56069.0237 | 1117.3805 | 1017.3805 | 1045.2852
6123 | 108.0409 | 1.0209 | 0.9645 | 1.0523 | 189.4123 | 50.4277 | 56931.0826 | 59025.4999 | 1176.51 | 1076.51 | 1045.3372
6124 | 123.5045 | 1.0166 | 1.1775 | 1.0438 | 179.5165 | 48.2883 | 65594.1203 | 65449.2071 | 1304.9841 | 1204.9841 | 1045.3514
6125 | 105.6977 | 1.0262 | 1.1414 | 1.0005 | 214.1857 | 52.4462 | 65799.1414 | 63997.4646 | 1275.9493 | 1175.9493 | 1045.3671
6126 | 98.3599 | 1.1122 | 1.0844 | 0.96632 | 183.8508 | 53.1256 | 61664.3893 | 60335.5023 | 1202.71 | 1102.71 | 1045.4069
6127 | 97.3192 | 0.99555 | 0.91914 | 0.87844 | 191.1723 | 52.1583 | 47357.364 | 47984.3593 | 955.6872 | 855.6872 | 1045.4232
6128 | 103.0623 | 1.0504 | 1.0986 | 0.91638 | 209.2452 | 48.7436 | 54559.6085 | 55138.7657 | 1098.7753 | 998.7753 | 1045.5136
6129 | 110.9096 | 1.0299 | 0.98854 | 0.98087 | 192.5297 | 50.9158 | 55378.3533 | 56115.6289 | 1118.3126 | 1018.3126 | 1045.5191
6130 | 83.9988 | 0.96568 | 0.80001 | 1.047 | 205.6449 | 48.852 | 45087.2786 | 45137.5279 | 898.7506 | 798.7506 | 1045.5273
6131 | 105.1367 | 1.0143 | 1.0682 | 0.99795 | 207.0359 | 52.5289 | 61802.4854 | 61540.2013 | 1226.804 | 1126.804 | 1045.5899
6132 | 87.4623 | 0.96477 | 0.93954 | 0.97814 | 196.5867 | 47.6464 | 48004.0642 | 46704.2297 | 930.0846 | 830.0846 | 1045.7027
6133 | 92.2531 | 0.98933 | 0.82172 | 1.0133 | 204.478 | 48.2171 | 44753.8738 | 44753.0473 | 891.0609 | 791.0609 | 1045.7212
6134 | 102.4172 | 0.93609 | 0.90186 | 1.0076 | 177.3191 | 51.2422 | 51656.8373 | 53557.7701 | 1067.1554 | 967.1554 | 1045.7577
6135 | 80.9493 | 1.0664 | 1.0452 | 0.99708 | 195.5476 | 48.2032 | 54591.6023 | 53717.0479 | 1070.341 | 970.341 | 1045.787
6136 | 95.6685 | 0.98764 | 1.0502 | 1.029 | 185.2055 | 48.5957 | 57293.8362 | 56206.0255 | 1120.1205 | 1020.1205 | 1045.8008
6137 | 118.7105 | 1.0624 | 1.1608 | 1.0108 | 196.7185 | 52.2124 | 68042.8595 | 70416.0211 | 1404.3204 | 1304.3204 | 1045.8345
6138 | 94.421 | 0.99939 | 1.146 | 0.9526 | 200.4995 | 45.4555 | 54111.4256 | 53709.8885 | 1070.1978 | 970.1978 | 1045.9162
6139 | 114.0956 | 1.0091 | 0.58342 | 0.92002 | 213.5511 | 46.0465 | 30231.0787 | 30415.8778 | 604.3176 | 504.3176 | 1046.011
6140 | 96.7292 | 0.97763 | 0.94764 | 1.0553 | 223.1978 | 52.2368 | 57403.8685 | 56542.3278 | 1126.8466 | 1026.8466 | 1046.0393
6141 | 100.3177 | 0.93487 | 1.101 | 0.99104 | 194.0162 | 52.3732 | 62249.5713 | 61338.4317 | 1222.7686 | 1122.7686 | 1046.0583
6142 | 104.5678 | 1.0549 | 1.1519 | 1.0558 | 206.9811 | 49.2942 | 65595.8279 | 65585.6021 | 1307.712 | 1207.712 | 1046.0927
6143 | 77.9907 | 1.0629 | 0.80143 | 1.0048 | 222.3384 | 52.2361 | 46615.7184 | 46610.723 | 928.2145 | 828.2145 | 1046.1006
6144 | 110.5639 | 0.98024 | 1.0275 | 1.0358 | 197.0287 | 53.9503 | 63465.3246 | 62710.7986 | 1250.216 | 1150.216 | 1046.2477
6145 | 106.0598 | 1.0817 | 1.0505 | 0.96858 | 203.0761 | 48.6791 | 55319.1327 | 55473.5908 | 1105.4718 | 1005.4718 | 1046.2537
6146 | 111.9288 | 1.0679 | 0.8409 | 0.95609 | 206.8466 | 50.28 | 46640.5026 | 47641.2723 | 948.8254 | 848.8254 | 1046.2592
6147 | 112.2495 | 1.0211 | 0.94927 | 0.97813 | 212.824 | 49.2687 | 51606.1048 | 50567.4844 | 1007.3497 | 907.3497 | 1046.3437
6148 | 95.8491 | 1.0616 | 1.0222 | 1.0058 | 198.478 | 53.4754 | 60620.9042 | 61259.4132 | 1221.1883 | 1121.1883 | 1046.345
6149 | 110.2154 | 1.009 | 1.0773 | 1.0708 | 194.6379 | 56.2689 | 71363.7505 | 69240.5449 | 1380.8109 | 1280.8109 | 1046.449
6150 | 79.7785 | 0.99446 | 0.88787 | 1.0544 | 176.0733 | 49.3159 | 50258.8308 | 52476.4325 | 1045.5286 | 945.5286 | 1046.4747
6151 | 76.9466 | 0.9593 | 1.0569 | 1.0308 | 202.4008 | 50.4066 | 58837.5109 | 57475.9565 | 1145.5191 | 1045.5191 | 1046.6191
6152 | 86.618 | 1.0225 | 1.0046 | 0.9851 | 202.6302 | 46.8771 | 50747.264 | 49646.2038 | 988.9241 | 888.9241 | 1046.634
6153 | 109.8036 | 1.0033 | 0.75765 | 1.1348 | 205.8828 | 49.5305 | 48246.5593 | 46789.2514 | 931.785 | 831.785 | 1046.6699
6154 | 96.5264 | 0.93261 | 0.99223 | 0.91516 | 194.6202 | 46.8267 | 46931.1612 | 46269.4014 | 921.388 | 821.388 | 1046.6991
6155 | 99.6137 | 1.0349 | 1.1869 | 1.0808 | 193.3945 | 48.131 | 66900.2402 | 67219.8982 | 1340.398 | 1240.398 | 1046.7103
6156 | 92.7197 | 1.0175 | 0.92852 | 1.0211 | 222.3204 | 50.7591 | 53137.7295 | 50909.0103 | 1014.1802 | 914.1802 | 1046.7222
6157 | 98.1568 | 0.95546 | 0.87577 | 1.0293 | 209.6528 | 50.1465 | 50117.6124 | 50500.9643 | 1006.0193 | 906.0193 | 1046.7611
6158 | 87.079 | 0.88415 | 1.0254 | 1.0638 | 214.3775 | 49.5259 | 58050.8166 | 58351.7775 | 1163.0355 | 1063.0355 | 1046.8247
6159 | 96.5242 | 1.0253 | 1.5015 | 0.96213 | 189.6803 | 48.3141 | 74768.2018 | 73459.7389 | 1465.1948 | 1365.1948 | 1046.8284
6160 | 115.1831 | 0.97825 | 1.2659 | 1.037 | 191.0181 | 49.0063 | 70048.075 | 70449.1543 | 1404.9831 | 1304.9831 | 1046.955
6161 | 121.3636 | 1.1098 | 1.0414 | 1.1009 | 223.9253 | 47.964 | 61675.1692 | 64318.6028 | 1282.3721 | 1182.3721 | 1046.9727
6162 | 103.4652 | 0.98521 | 0.86412 | 0.95961 | 198.7549 | 49.4298 | 46225.5077 | 45791.9503 | 911.839 | 811.839 | 1046.9855
6163 | 82.4026 | 0.85762 | 0.95138 | 1.0097 | 193.5008 | 53.6377 | 55507.5945 | 55915.6864 | 1114.3137 | 1014.3137 | 1047.0209
6164 | 98.5717 | 1.0259 | 0.80469 | 1.0136 | 205.8017 | 51.3427 | 47273.5301 | 47071.311 | 937.4262 | 837.4262 | 1047.115
6165 | 90.9042 | 0.96399 | 1.0117 | 1.0366 | 204.7466 | 46.9863 | 53600.7121 | 54389.1044 | 1083.7821 | 983.7821 | 1047.2573
6166 | 95.3665 | 0.96075 | 1.248 | 0.95069 | 206.491 | 50.4762 | 64717.9965 | 64181.303 | 1279.6261 | 1179.6261 | 1047.2969
6167 | 91.6752 | 0.95022 | 0.90958 | 0.96117 | 198.6521 | 51.4967 | 49706.1698 | 50337.5469 | 1002.7509 | 902.7509 | 1047.3441
6168 | 85.8075 | 0.97411 | 1.1079 | 0.9827 | 194.9889 | 51.7648 | 60880.0405 | 61095.0051 | 1217.9001 | 1117.9001 | 1047.3854
6169 | 107.9474 | 0.90094 | 1.0106 | 0.94093 | 214.5723 | 49.1909 | 51775.2572 | 50786.0388 | 1011.7208 | 911.7208 | 1047.3893
6170 | 106.7161 | 0.96859 | 0.98643 | 1.0595 | 196.2996 | 51.499 | 59341.0331 | 57547.7511 | 1146.955 | 1046.955 | 1047.425
6171 | 94.4091 | 1.0271 | 1.0195 | 0.93064 | 193.8411 | 50.0788 | 52564.8565 | 53776.0763 | 1071.5215 | 971.5215 | 1047.4463
6172 | 86.6792 | 0.91864 | 1.0222 | 1.0177 | 226.8045 | 49.3876 | 55534.0624 | 55372.1762 | 1103.4435 | 1003.4435 | 1047.4536
6173 | 112.8781 | 0.92161 | 0.78219 | 1.0758 | 176.5365 | 56.8101 | 53892.0704 | 54015.921 | 1076.3184 | 976.3184 | 1047.464
6174 | 91.104 | 1.0163 | 1.1823 | 1.0178 | 197.5942 | 51.4736 | 66899.403 | 64640.1754 | 1288.8035 | 1188.8035 | 1047.5006
6175 | 82.7492 | 0.97044 | 0.88562 | 1.0069 | 188.4852 | 46.1856 | 45080.977 | 46560.5542 | 927.2111 | 827.2111 | 1047.5038
6176 | 123.6829 | 0.98585 | 0.98459 | 1.0211 | 199.2846 | 49.8746 | 56423.2154 | 57754.0535 | 1151.0811 | 1051.0811 | 1047.5078
6177 | 107.206 | 1.066 | 1.1909 | 1.042 | 202.6648 | 46.1818 | 62790.7247 | 64988.7203 | 1295.7744 | 1195.7744 | 1047.545
6178 | 85.6588 | 0.95211 | 0.84338 | 1.0157 | 187.8941 | 47.6901 | 44927.6146 | 44024.2266 | 876.4845 | 776.4845 | 1047.6593
6179 | 98.7827 | 0.99305 | 0.8407 | 1.0075 | 194.3786 | 49.8167 | 47278.0107 | 46341.7566 | 922.8351 | 822.8351 | 1047.6622
6180 | 99.7802 | 1.0426 | 1.0675 | 1.0012 | 202.7606 | 47.9631 | 56457.2002 | 57751.4869 | 1151.0297 | 1051.0297 | 1047.6763
6181 | 108.9488 | 0.98367 | 1.1574 | 0.98772 | 192.022 | 50.5246 | 63364.7688 | 62334.8276 | 1242.6966 | 1142.6966 | 1047.7127
6182 | 101.4309 | 0.88328 | 1.0827 | 1.0119 | 217.5799 | 52.4911 | 62432.0064 | 63403.423 | 1264.0685 | 1164.0685 | 1047.7368
6183 | 112.8345 | 0.97513 | 1.0755 | 1.0839 | 209.4692 | 51.2158 | 65548.6903 | 64172.6673 | 1279.4533 | 1179.4533 | 1047.7574
6184 | 82.2263 | 0.95272 | 1.2371 | 1.0169 | 204.2353 | 50.2967 | 67413.7281 | 68956.9184 | 1375.1384 | 1275.1384 | 1047.7861
6185 | 98.5819 | 0.98307 | 1.0295 | 1.0958 | 200.1361 | 54.4394 | 66888.8245 | 64474.931 | 1285.4986 | 1185.4986 | 1047.8002
6186 | 100.086 | 1.0519 | 1.1392 | 0.95398 | 206.4099 | 50.5096 | 60417.7074 | 59303.9655 | 1182.0793 | 1082.0793 | 1047.8006
6187 | 104.2939 | 1.0648 | 1.227 | 0.95504 | 203.5827 | 52.1169 | 67065.8559 | 66624.1165 | 1328.4823 | 1228.4823 | 1047.8423
6188 | 110.5322 | 0.94676 | 0.83929 | 0.98481 | 187.9063 | 47.5595 | 44475.1558 | 43418.7829 | 864.3757 | 764.3757 | 1048.0277
6189 | 114.1432 | 1.0576 | 0.95845 | 0.93644 | 185.0477 | 50.7602 | 51871.7625 | 51532.2731 | 1026.6455 | 926.6455 | 1048.1325
6190 | 96.5827 | 0.96514 | 0.77808 | 1.0288 | 190.8913 | 51.8742 | 46550.291 | 48752.8609 | 971.0572 | 871.0572 | 1048.1713
6191 | 106.6333 | 0.97263 | 1.0743 | 1.0262 | 193.2342 | 51.4019 | 62190.1405 | 61230.2981 | 1220.606 | 1120.606 | 1048.2309
6192 | 110.8243 | 0.94959 | 0.96331 | 0.9716 | 196.2016 | 51.1689 | 53472.8527 | 54332.7739 | 1082.6555 | 982.6555 | 1048.2986
6193 | 112.1147 | 0.94505 | 1.0414 | 1.0574 | 217.7277 | 46.9578 | 56899.6152 | 56973.3884 | 1135.4678 | 1035.4678 | 1048.3271
6194 | 112.3238 | 1.048 | 1.1286 | 0.99375 | 198.2726 | 57.2627 | 71164.9656 | 68969.997 | 1375.3999 | 1275.3999 | 1048.3316
6195 | 99.4159 | 1.0108 | 0.72646 | 1.0239 | 189.3318 | 52.8024 | 44771.7478 | 43866.4742 | 873.3295 | 773.3295 | 1048.3867
6196 | 107.5065 | 0.98057 | 1.0059 | 0.96466 | 202.4536 | 49.9962 | 53986.8919 | 52441.2323 | 1044.8246 | 944.8246 | 1048.461
6197 | 99.3077 | 1.0215 | 1.0122 | 0.99244 | 199.842 | 50.8141 | 56399.0203 | 57034.6726 | 1136.6935 | 1036.6935 | 1048.5428
6198 | 109.5841 | 1.0089 | 1.393 | 1.0264 | 197.8656 | 52.2793 | 80729.0188 | 81910.2014 | 1634.204 | 1534.204 | 1048.6228
6199 | 102.3161 | 0.91255 | 0.97604 | 0.96431 | 193.4112 | 43.926 | 45638.2476 | 46453.5182 | 925.0704 | 825.0704 | 1048.6462
6200 | 114.0169 | 1.0001 | 1.3195 | 1.0293 | 201.0445 | 48.6702 | 71851.8181 | 72918.1089 | 1454.3622 | 1354.3622 | 1048.6881
6201 | 95.6218 | 1.0231 | 0.79848 | 0.99645 | 195.948 | 50.0303 | 44897.1311 | 44011.8416 | 876.2368 | 776.2368 | 1048.7365
6202 | 121.6565 | 1.0846 | 1.0941 | 1.003 | 201.4211 | 47.2096 | 58237.3853 | 58717.2356 | 1170.3447 | 1070.3447 | 1048.9169
6203 | 116.6671 | 1.0033 | 1.0489 | 0.97245 | 193.4686 | 48.975 | 55878.8016 | 54314.6509 | 1082.293 | 982.293 | 1048.9398
6204 | 99.1132 | 1.0898 | 1.0504 | 1.0096 | 203.6465 | 55.5245 | 65086.7801 | 65425.5183 | 1304.5104 | 1204.5104 | 1048.9717
6205 | 97.9876 | 1.0248 | 0.58578 | 0.97855 | 210.4476 | 49.0779 | 33271.0999 | 33657.7955 | 669.1559 | 569.1559 | 1048.9854
6206 | 85.1492 | 1.0089 | 1.0057 | 1.0156 | 196.5627 | 48.4103 | 53801.5585 | 53383.4884 | 1063.6698 | 963.6698 | 1049.035
6207 | 109.8413 | 1.0056 | 1.0835 | 0.99896 | 197.2808 | 45.1316 | 54030.8959 | 52400.5336 | 1044.0107 | 944.0107 | 1049.0405
6208 | 100.9058 | 1.013 | 1.1924 | 1.0078 | 199.3883 | 55.0064 | 71924.3722 | 71056.211 | 1417.1242 | 1317.1242 | 1049.0487
6209 | 89.2158 | 1.1324 | 0.97078 | 0.98216 | 191.2086 | 50.0016 | 52917.59 | 53747.5418 | 1070.9508 | 970.9508 | 1049.0869
6210 | 96.7365 | 1.053 | 1.0523 | 0.99849 | 194.3585 | 49.3603 | 57086.9517 | 55836.5234 | 1112.7305 | 1012.7305 | 1049.093
6211 | 83.7311 | 1.0586 | 1.0342 | 1.1104 | 215.97 | 48.7252 | 60493.9892 | 61074.8367 | 1217.4967 | 1117.4967 | 1049.1439
6212 | 109.8849 | 1.0125 | 0.90084 | 1.0534 | 190.2985 | 47.4704 | 50518.8331 | 50020.5314 | 996.4106 | 896.4106 | 1049.147
6213 | 108.1869 | 1.0623 | 0.90491 | 0.99221 | 198.5074 | 52.2608 | 53127.7392 | 53710.6429 | 1070.2129 | 970.2129 | 1049.1703
6214 | 88.6107 | 1.0258 | 0.97853 | 1.0166 | 204.5434 | 50.6759 | 55220.0026 | 54898.7272 | 1093.9745 | 993.9745 | 1049.1803
6215 | 99.3259 | 1.0276 | 1.0519 | 0.99882 | 200.9475 | 52.2941 | 60482.7745 | 59859.3886 | 1193.1878 | 1093.1878 | 1049.1968
6216 | 84.0486 | 0.97314 | 0.84914 | 0.98889 | 194.04 | 45.9472 | 42534.2899 | 45272.9785 | 901.4596 | 801.4596 | 1049.2156
6217 | 103.9413 | 1.0224 | 1.2062 | 0.96675 | 199.955 | 51.8335 | 66153.0253 | 65295.8214 | 1301.9164 | 1201.9164 | 1049.2158
6218 | 95.9496 | 1.0268 | 0.82494 | 1.0156 | 196.4615 | 49.5671 | 46605.7305 | 46245.4997 | 920.91 | 820.91 | 1049.2756
6219 | 89.8308 | 1.0116 | 0.90895 | 1.0056 | 208.7652 | 50.4809 | 50935.939 | 50449.4918 | 1004.9898 | 904.9898 | 1049.3024
6220 | 97.7745 | 0.97413 | 1.147 | 1.0425 | 188.8744 | 52.8997 | 68478.8145 | 67048.1409 | 1336.9628 | 1236.9628 | 1049.3394
6221 | 106.7509 | 1.0778 | 0.75964 | 1.0081 | 195.626 | 49.1572 | 43494.2953 | 43799.6075 | 871.9921 | 771.9921 | 1049.3657
6222 | 85.8296 | 1.0586 | 1.066 | 1.0726 | 197.6094 | 50.5298 | 62568.1022 | 60759.9604 | 1211.1992 | 1111.1992 | 1049.3774
6223 | 113.4296 | 0.96142 | 1.0282 | 1.0068 | 211.0956 | 48.4379 | 55639.0787 | 55134.3344 | 1098.6867 | 998.6867 | 1049.5498
6224 | 86.7315 | 1.0112 | 0.89908 | 0.99246 | 200.675 | 53.9984 | 53119.6889 | 54398.8924 | 1083.9778 | 983.9778 | 1049.5826
6225 | 97.7732 | 1.0662 | 0.91995 | 0.84948 | 219.9583 | 46.7103 | 41592.4302 | 41863.9313 | 833.2786 | 733.2786 | 1049.6353
6226 | 101.7229 | 0.97356 | 1.0893 | 1.0372 | 209.9028 | 52.519 | 64746.2362 | 63540.5091 | 1266.8102 | 1166.8102 | 1049.6524
6227 | 95.1531 | 1.0077 | 0.90997 | 0.96181 | 196.1526 | 50.8516 | 49578.0396 | 50585.3383 | 1007.7068 | 907.7068 | 1049.6592
6228 | 110.1143 | 0.96951 | 1.0553 | 0.92834 | 193.6648 | 53.9197 | 58773.5595 | 58798.6386 | 1171.9728 | 1071.9728 | 1049.8465
6229 | 116.5347 | 0.96283 | 0.79694 | 1.0417 | 212.0905 | 48.2095 | 45642.1969 | 44403.449 | 884.069 | 784.069 | 1049.8681
6230 | 106.3942 | 0.89824 | 0.93565 | 0.9058 | 192.5768 | 47.908 | 45373.9716 | 44817.41 | 892.3482 | 792.3482 | 1049.8908
6231 | 109.2163 | 1.0908 | 0.95772 | 0.94256 | 191.8802 | 49.2045 | 50470.804 | 50082.3569 | 997.6471 | 897.6471 | 1049.9014
6232 | 107.6594 | 0.99928 | 1.0875 | 1.0438 | 201.3794 | 51.7299 | 64484.1978 | 66791.944 | 1331.8389 | 1231.8389 | 1049.9502
6233 | 91.5368 | 0.9983 | 0.89309 | 0.98362 | 197.3348 | 50.0029 | 48691.9603 | 48212.7636 | 960.2553 | 860.2553 | 1050.0624
6234 | 103.8864 | 1.0348 | 1.1615 | 1.0375 | 191.7144 | 50.8758 | 66969.2985 | 66536.2836 | 1326.7257 | 1226.7257 | 1050.0889
6235 | 114.0481 | 1.0175 | 1.0049 | 0.97049 | 203.3413 | 52.2899 | 57267.1598 | 57385.3165 | 1143.7063 | 1043.7063 | 1050.1625
6236 | 115.0176 | 1.021 | 1.2434 | 1.0353 | 211.8694 | 50.6349 | 71335.4799 | 72747.1538 | 1450.9431 | 1350.9431 | 1050.1795
6237 | 95.1748 | 0.94213 | 1.2512 | 1.0394 | 213.8065 | 48.2092 | 67227.3745 | 69384.3118 | 1383.6862 | 1283.6862 | 1050.2098
6238 | 88.9288 | 0.95822 | 1.1698 | 0.97913 | 203.1826 | 46.3749 | 57271.646 | 57388.2102 | 1143.7642 | 1043.7642 | 1050.2137
6239 | 119.4969 | 0.96961 | 0.97101 | 0.95613 | 196.2355 | 47.8108 | 50124.165 | 51786.1063 | 1031.7221 | 931.7221 | 1050.3435
6240 | 87.3835 | 0.94554 | 1.136 | 0.95725 | 211.0439 | 50.5293 | 59331.5269 | 58398.3076 | 1163.9662 | 1063.9662 | 1050.3443
6241 | 114.6146 | 1.0454 | 1.0508 | 0.95829 | 204.7907 | 47.9028 | 54182.3807 | 56497.0642 | 1125.9413 | 1025.9413 | 1050.3602
6242 | 104.6155 | 0.92733 | 0.99614 | 0.99116 | 203.3625 | 50.327 | 54775.2706 | 54061.4656 | 1077.2293 | 977.2293 | 1050.39
6243 | 104.0246 | 1.0182 | 0.95895 | 1.0745 | 182.8923 | 49.7278 | 56688.205 | 57274.9107 | 1141.4982 | 1041.4982 | 1050.4127
6244 | 96.1451 | 1.0243 | 1.2265 | 0.94294 | 184.8635 | 50.4787 | 63534.3084 | 63113.723 | 1258.2745 | 1158.2745 | 1050.4895
6245 | 77.0987 | 0.94914 | 0.91174 | 0.98943 | 206.3684 | 48.9665 | 47962.3559 | 46488.9528 | 925.7791 | 825.7791 | 1050.5101
6246 | 93.1814 | 0.97629 | 0.85321 | 1.0012 | 182.2285 | 47.3555 | 44941.1251 | 45644.3427 | 908.8869 | 808.8869 | 1050.6074
6247 | 90.2876 | 0.914 | 0.76441 | 0.96878 | 195.0723 | 52.3731 | 43301.5843 | 43995.9119 | 875.9182 | 775.9182 | 1050.6184
6248 | 97.0384 | 0.984 | 1.215 | 0.98456 | 210.6829 | 48.4444 | 62787.7132 | 61646.2798 | 1228.9256 | 1128.9256 | 1050.661
6249 | 113.1826 | 1.0167 | 0.85605 | 0.96758 | 194.2724 | 48.5306 | 45976.981 | 46727.0629 | 930.5413 | 830.5413 | 1050.6663
6250 | 94.4585 | 1.0373 | 0.92937 | 0.99669 | 193.2174 | 55.4659 | 57005.403 | 57403.4842 | 1144.0697 | 1044.0697 | 1050.8159
6251 | 103.7862 | 1.1171 | 1.0884 | 1.0289 | 183.7727 | 44.6961 | 55415.7551 | 55921.841 | 1114.4368 | 1014.4368 | 1050.8188
6252 | 121.9677 | 1.0282 | 1.1667 | 0.99522 | 198.4218 | 50.8074 | 65561.6262 | 66326.474 | 1322.5295 | 1222.5295 | 1050.8464
6253 | 96.3945 | 1.0363 | 1.0972 | 1.0182 | 181.1483 | 52.8562 | 64510.5828 | 63335.7836 | 1262.7157 | 1162.7157 | 1050.8921
6254 | 83.9938 | 1.037 | 0.86736 | 0.96242 | 204.2949 | 49.9112 | 46215.475 | 46767.156 | 931.3431 | 831.3431 | 1050.8929
6255 | 97.4025 | 1.0139 | 0.86902 | 1.0086 | 189.6452 | 52.474 | 51366.0953 | 50992.9684 | 1015.8594 | 915.8594 | 1050.9483
6256 | 85.0075 | 1.0258 | 1.1068 | 1.0574 | 195.9119 | 54.1701 | 68315.7977 | 66552.1328 | 1327.0427 | 1227.0427 | 1050.9728
6257 | 84.4711 | 0.97875 | 1.1376 | 1.0828 | 191.8511 | 49.8825 | 65762.8596 | 66846.412 | 1332.9282 | 1232.9282 | 1051.0297
6258 | 95.8982 | 0.99279 | 0.66569 | 1.0591 | 165.1854 | 48.6965 | 39135.5306 | 39036.4153 | 776.7283 | 676.7283 | 1051.0397
6259 | 93.4321 | 1.0285 | 1.1469 | 1.0218 | 193.9209 | 51.447 | 65428.3958 | 65045.0257 | 1296.9005 | 1196.9005 | 1051.0461
6260 | 95.7187 | 1.0222 | 0.93127 | 0.9717 | 201.0188 | 47.6937 | 48026.345 | 47490.9918 | 945.8198 | 845.8198 | 1051.0615
6261 | 104.6992 | 1.1227 | 0.9385 | 1.0265 | 208.4837 | 53.2052 | 57720.2569 | 59476.6076 | 1185.5322 | 1085.5322 | 1051.0811
6262 | 103.5952 | 0.97597 | 0.90174 | 1.1229 | 191.1731 | 51.8102 | 57888.4157 | 58285.9488 | 1161.719 | 1061.719 | 1051.1153
6263 | 94.8099 | 0.93947 | 1.0958 | 1.0353 | 184.5064 | 46.9645 | 57648.7657 | 57123.031 | 1138.4606 | 1038.4606 | 1051.1255
6264 | 106.3292 | 1.0518 | 0.98544 | 1.0142 | 202.1237 | 48.7684 | 54397.0375 | 53310.1177 | 1062.2024 | 962.2024 | 1051.1408
6265 | 96.1845 | 0.98797 | 0.84598 | 0.91818 | 227.8179 | 49.0359 | 42976.8349 | 43655.555 | 869.1111 | 769.1111 | 1051.1446
6266 | 118.4028 | 0.92322 | 1.0488 | 0.9704 | 190.9669 | 46.2722 | 52342.3854 | 52461.799 | 1045.236 | 945.236 | 1051.1523
6267 | 101.5036 | 0.9446 | 1.0373 | 1.0177 | 190.1607 | 49.2722 | 56930.1646 | 57757.0415 | 1151.1408 | 1051.1408 | 1051.2131
6268 | 107.0034 | 0.98157 | 1.2679 | 1.0578 | 196.5092 | 54.0218 | 78329.0601 | 78338.9421 | 1562.7788 | 1462.7788 | 1051.2325
6269 | 106.4924 | 1.0218 | 1.0272 | 1.0132 | 207.6385 | 47.2502 | 54522.3039 | 53485.2065 | 1065.7041 | 965.7041 | 1051.2728
6270 | 89.2934 | 0.89708 | 0.94006 | 0.98364 | 207.0675 | 50.3336 | 50781.3676 | 50910.1716 | 1014.2034 | 914.2034 | 1051.2777
6271 | 98.535 | 0.93114 | 0.94172 | 1.0978 | 192.6273 | 49.8839 | 56340.7548 | 57824.5394 | 1152.4908 | 1052.4908 | 1051.312
6272 | 93.0347 | 1.0044 | 1.0631 | 1.0077 | 211.3372 | 47.7066 | 55773.5997 | 56558.9234 | 1127.1785 | 1027.1785 | 1051.3675
6273 | 92.8005 | 1.0805 | 0.6414 | 0.92009 | 205.3944 | 58.4647 | 40570.4495 | 38609.054 | 768.1811 | 668.1811 | 1051.385
6274 | 124.742 | 0.98738 | 0.8048 | 1.0051 | 191.4173 | 55.0728 | 51524.2996 | 51415.3952 | 1024.3079 | 924.3079 | 1051.3908
6275 | 87.646 | 0.97711 | 0.94758 | 0.94227 | 206.9568 | 54.8552 | 53883.9379 | 54454.9562 | 1085.0991 | 985.0991 | 1051.5322
6276 | 109.2853 | 0.99636 | 1.0644 | 1.0117 | 194.8861 | 51.6538 | 61442.2647 | 60391.1106 | 1203.8222 | 1103.8222 | 1051.5431
6277 | 102.9097 | 1.0269 | 1.1064 | 0.93498 | 214.3856 | 47.4567 | 54323.558 | 53337.3306 | 1062.7466 | 962.7466 | 1051.567
6278 | 108.1542 | 1.082 | 1.1026 | 0.99266 | 205.1234 | 49.621 | 60320.7116 | 62325.1894 | 1242.5038 | 1142.5038 | 1051.5978
6279 | 91.4638 | 0.94643 | 0.96626 | 1.0092 | 195.1669 | 45.024 | 47998.0838 | 47433.0683 | 944.6614 | 844.6614 | 1051.6227
6280 | 124.9889 | 0.99877 | 1.1337 | 0.89522 | 199.9924 | 47.3545 | 54173.3271 | 53492.2757 | 1065.8455 | 965.8455 | 1051.6728
6281 | 101.6939 | 1.0036 | 0.93414 | 1.1146 | 211.2811 | 49.316 | 56590.4143 | 57387.6806 | 1143.7536 | 1043.7536 | 1051.74
6282 | 112.6693 | 1.0258 | 1.0655 | 0.99459 | 197.2438 | 49.3103 | 58151.1216 | 59479.0097 | 1185.5802 | 1085.5802 | 1051.7459
6283 | 91.0832 | 0.93718 | 1.124 | 1.0163 | 199.616 | 50.4023 | 62078.3996 | 63188.3009 | 1259.766 | 1159.766 | 1051.8179
6284 | 108.9033 | 0.94975 | 1.0717 | 1.0488 | 202.8371 | 47.6143 | 58647.1505 | 59385.5243 | 1183.7105 | 1083.7105 | 1051.8271
6285 | 94.2739 | 0.93942 | 0.80978 | 1.0065 | 207.3243 | 46.412 | 42143.956 | 42215.8823 | 840.3176 | 740.3176 | 1051.8793
6286 | 108.4493 | 0.95786 | 1.0491 | 1.0755 | 195.6198 | 47.1911 | 58341.3427 | 57425.439 | 1144.5088 | 1044.5088 | 1051.9687
6287 | 97.2921 | 1.0581 | 0.93286 | 1.0814 | 207.3187 | 49.096 | 54791.1919 | 53754.9978 | 1071.1 | 971.1 | 1052.0409
6288 | 94.2631 | 0.93104 | 1.0934 | 0.98251 | 185.3438 | 49.1171 | 57261.1774 | 57564.8464 | 1147.2969 | 1047.2969 | 1052.0715
6289 | 99.2836 | 0.93249 | 0.76847 | 0.92084 | 193.125 | 48.9487 | 39363.0392 | 38528.1114 | 766.5622 | 666.5622 | 1052.1393
6290 | 84.3838 | 0.91923 | 0.94437 | 0.94514 | 207.8025 | 54.8564 | 53425.8788 | 52116.2071 | 1038.3241 | 938.3241 | 1052.2155
6291 | 120.0732 | 1.0916 | 0.99035 | 1.0141 | 191.5653 | 50.9561 | 58046.1119 | 58191.2138 | 1159.8243 | 1059.8243 | 1052.2326
6292 | 109.0229 | 0.9335 | 1.0011 | 0.93756 | 179.4982 | 50.7405 | 52966.0505 | 52749.4511 | 1050.989 | 950.989 | 1052.2396
6293 | 99.4524 | 0.93402 | 0.83761 | 0.97343 | 208.7731 | 46.3361 | 42293.4402 | 42963.2731 | 855.2655 | 755.2655 | 1052.2475
6294 | 96.17 | 1.0244 | 1.0437 | 1.1342 | 196.3513 | 45.9785 | 59153.1515 | 58103.1093 | 1158.0622 | 1058.0622 | 1052.2478
6295 | 103.9885 | 1.0162 | 1.0584 | 1.0084 | 208.0984 | 53.9201 | 63452.0722 | 63681.3283 | 1269.6266 | 1169.6266 | 1052.2619
6296 | 105.5657 | 1.0136 | 0.93103 | 1.0641 | 209.9201 | 48.6946 | 53661.9828 | 55281.4285 | 1101.6286 | 1001.6286 | 1052.2877
6297 | 102.8764 | 0.9872 | 0.81048 | 0.99234 | 213.6307 | 46.9311 | 42725.2159 | 42371.8841 | 843.4377 | 743.4377 | 1052.3483
6298 | 97.728 | 0.95686 | 0.88181 | 1.0727 | 202.1437 | 50.702 | 52905.0026 | 53072.55 | 1057.451 | 957.451 | 1052.3498
6299 | 104.2073 | 1.0517 | 1.0224 | 0.9379 | 207.8936 | 53.5733 | 57452.6227 | 57495.8088 | 1145.9162 | 1045.9162 | 1052.4163
6300 | 97.5863 | 1.003 | 0.90415 | 0.95305 | 208.9736 | 47.3911 | 45684.6459 | 46275.5504 | 921.511 | 821.511 | 1052.4416
6301 | 105.8501 | 1.0634 | 0.92009 | 1.0021 | 187.4572 | 55.5441 | 57650.5954 | 57608.5625 | 1148.1713 | 1048.1713 | 1052.482
6302 | 112.1135 | 0.99552 | 0.98261 | 1.0156 | 202.3505 | 53.6721 | 59755.3564 | 59694.6169 | 1189.8923 | 1089.8923 | 1052.4908
6303 | 103.4194 | 1.026 | 1.0588 | 1.0374 | 189.9988 | 47.1093 | 56935.4197 | 58530.6563 | 1166.6131 | 1066.6131 | 1052.497
6304 | 97.9695 | 0.92542 | 1.1752 | 0.95993 | 205.0325 | 51.5366 | 63017.3882 | 63851.9456 | 1273.0389 | 1173.0389 | 1052.5445
6305 | 92.8655 | 0.90627 | 0.92386 | 0.99501 | 208.6081 | 47.6869 | 48057.8437 | 47364.7878 | 943.2958 | 843.2958 | 1052.5496
6306 | 100.0199 | 1.0727 | 1.1957 | 1.0424 | 197.5398 | 47.598 | 64631.1435 | 64046.7558 | 1276.9351 | 1176.9351 | 1052.5704
6307 | 106.4606 | 1.0006 | 0.81587 | 0.95087 | 195.0295 | 46.9268 | 41598.7629 | 40673.718 | 809.4744 | 709.4744 | 1052.581
6308 | 89.6611 | 0.97907 | 0.7356 | 1.0311 | 208.4248 | 52.4953 | 44634.2524 | 43538.9065 | 866.7781 | 766.7781 | 1052.5829
6309 | 110.0466 | 0.99776 | 1.138 | 1.0083 | 225.4493 | 50.4066 | 63600.306 | 64097.694 | 1277.9539 | 1177.9539 | 1052.6226
6310 | 112.1141 | 0.97485 | 0.93324 | 0.90081 | 181.077 | 49.4385 | 47146.1392 | 46148.1858 | 918.9637 | 818.9637 | 1052.7121
6311 | 113.6874 | 0.97043 | 1.0075 | 0.99197 | 202.359 | 46.0491 | 51302.9454 | 51516.379 | 1026.3276 | 926.3276 | 1052.7334
6312 | 104.08 | 1.0377 | 1.0165 | 0.90269 | 198.0237 | 44.6349 | 45975.3807 | 45741.6262 | 910.8325 | 810.8325 | 1052.7519
6313 | 104.5528 | 1.0406 | 1.0272 | 0.9967 | 202.7533 | 49.0161 | 55719.0184 | 56301.2662 | 1122.0253 | 1022.0253 | 1052.7598
6314 | 78.2756 | 0.99452 | 0.83771 | 1.0313 | 205.3682 | 50.0338 | 47325.0171 | 46671.4574 | 929.4291 | 829.4291 | 1052.7627
6315 | 99.6346 | 0.92982 | 0.90172 | 1.0777 | 185.7712 | 51.8937 | 55421.0129 | 54510.6437 | 1086.2129 | 986.2129 | 1052.805
6316 | 102.878 | 0.98617 | 0.924 | 1.1043 | 198.5852 | 47.1412 | 53082.6851 | 50094.7576 | 997.8952 | 897.8952 | 1052.9659
6317 | 90.4814 | 0.98591 | 0.90825 | 0.96224 | 190.127 | 52.2338 | 50500.0079 | 50143.8498 | 998.877 | 898.877 | 1053.0585
6318 | 78.1688 | 1.0403 | 1.2356 | 1.0367 | 194.179 | 52.9613 | 72344.1753 | 71315.6491 | 1422.313 | 1322.313 | 1053.1321
6319 | 101.0817 | 1.0194 | 1.1583 | 1.05 | 197.4967 | 50.6848 | 67068.3441 | 65643.0689 | 1308.8614 | 1208.8614 | 1053.1421
6320 | 116.1083 | 0.92013 | 1.1026 | 0.98266 | 210.8632 | 55.0715 | 65761.0289 | 69930.2998 | 1394.606 | 1294.606 | 1053.175
6321 | 94.0859 | 1.0622 | 0.79994 | 1.0558 | 213.3479 | 50.0524 | 47489.4436 | 48095.4122 | 957.9082 | 857.9082 | 1053.1779
6322 | 117.2244 | 1.0064 | 1.0018 | 0.92608 | 199.1871 | 46.8834 | 49225.1431 | 49676.1925 | 989.5238 | 889.5238 | 1053.2127
6323 | 111.6745 | 1.0326 | 1.1413 | 0.97193 | 206.8508 | 53.7807 | 66068.2672 | 68105.9829 | 1358.1197 | 1258.1197 | 1053.2632
6324 | 105.527 | 1.0512 | 0.89114 | 0.98076 | 206.4874 | 48.1302 | 47611.2318 | 46247.7424 | 920.9548 | 820.9548 | 1053.2634
6325 | 109.6795 | 1.0266 | 0.9541 | 0.97907 | 202.7103 | 49.4955 | 52011.4236 | 52680.8495 | 1049.617 | 949.617 | 1053.2929
6326 | 102.097 | 1.1258 | 0.73055 | 1.0289 | 200.5734 | 51.1287 | 44508.7145 | 43803.5749 | 872.0715 | 772.0715 | 1053.3254
6327 | 108.6515 | 0.9725 | 1.1279 | 1.0093 | 193.3657 | 47.7401 | 59581.6142 | 58762.4078 | 1171.2482 | 1071.2482 | 1053.381
6328 | 97.5038 | 0.9632 | 0.98264 | 0.93636 | 181.8477 | 45.5978 | 46418.9357 | 46821.3281 | 932.4266 | 832.4266 | 1053.4316
6329 | 101.9593 | 1.0362 | 0.95628 | 1.0018 | 191.3597 | 45.727 | 48830.0779 | 49097.314 | 977.9463 | 877.9463 | 1053.5856
6330 | 113.0561 | 1.0039 | 0.8326 | 1.0689 | 180.0596 | 49.4943 | 49845.1081 | 50314.8602 | 1002.2972 | 902.2972 | 1053.6389
6331 | 111.0598 | 1.0173 | 1.0081 | 1.0339 | 188.5927 | 48.5981 | 56332.2101 | 57154.3321 | 1139.0866 | 1039.0866 | 1053.7105
6332 | 95.9824 | 1.0112 | 0.82711 | 1.0514 | 177.5285 | 49.0572 | 47600.0709 | 46183.0041 | 919.6601 | 819.6601 | 1053.8265
6333 | 108.5119 | 1.0098 | 1.2255 | 1.0144 | 194.6682 | 50.6537 | 68713.0939 | 68929.4197 | 1374.5884 | 1274.5884 | 1053.9632
6334 | 102.3766 | 1.0131 | 1.1868 | 1.0005 | 194.3582 | 45.7952 | 59323.1503 | 56408.1202 | 1124.1624 | 1024.1624 | 1053.9709
6335 | 102.6563 | 0.96109 | 1.0887 | 1.0708 | 198.5082 | 52.3945 | 66447.7272 | 65681.9935 | 1309.6399 | 1209.6399 | 1054.1637
6336 | 104.1979 | 0.99685 | 1.0735 | 0.93581 | 216.7082 | 49.5944 | 55190.2101 | 54054.5459 | 1077.0909 | 977.0909 | 1054.1932
6337 | 103.5347 | 1.0114 | 1.108 | 0.97884 | 213.763 | 52.7794 | 62984.6202 | 61992.9913 | 1235.8598 | 1135.8598 | 1054.3345
6338 | 97.7886 | 0.95927 | 0.90602 | 0.96384 | 192.8769 | 51.6108 | 50103.7639 | 48550.1823 | 967.0036 | 867.0036 | 1054.4994
6339 | 111.846 | 1.0214 | 1.0853 | 0.93977 | 203.5002 | 48.8728 | 55633.6292 | 54522.764 | 1086.4553 | 986.4553 | 1054.5409
6340 | 98.2737 | 1.0836 | 0.84096 | 0.92375 | 208.7011 | 44.5915 | 39597.7107 | 39243.9367 | 780.8787 | 680.8787 | 1054.5528
6341 | 118.5998 | 0.95827 | 1.0689 | 1.0029 | 210.0842 | 48.2865 | 57464.1335 | 56765.7339 | 1131.3147 | 1031.3147 | 1054.6511
6342 | 108.2693 | 0.9053 | 0.93592 | 1.0135 | 196.286 | 47.5242 | 49933.5786 | 49948.9523 | 994.979 | 894.979 | 1054.6523
6343 | 102.466 | 1.0394 | 0.89632 | 0.972 | 211.7628 | 48.2787 | 47415.1106 | 47438.2343 | 944.7647 | 844.7647 | 1054.6845
6344 | 85.7996 | 0.8877 | 0.98349 | 0.99382 | 212.9398 | 47.5024 | 50260.0048 | 49257.8662 | 981.1573 | 881.1573 | 1054.7715
6345 | 98.4395 | 1.0052 | 0.96908 | 0.97514 | 203.4294 | 48.1682 | 50488.1039 | 49981.0145 | 995.6203 | 895.6203 | 1054.8414
6346 | 102.8754 | 0.97002 | 0.82976 | 1.0444 | 197.6681 | 47.7913 | 46383.7075 | 45985.2471 | 915.7049 | 815.7049 | 1054.9096
6347 | 100.5158 | 1.0093 | 1.0581 | 1.0838 | 199.5706 | 54.2017 | 67853.8546 | 69025.4301 | 1376.5086 | 1276.5086 | 1054.9599
6348 | 97.3368 | 1.0483 | 1.052 | 0.94955 | 190.3579 | 48.2011 | 53259.697 | 53760.2634 | 1071.2053 | 971.2053 | 1055.0077
6349 | 94.3253 | 1.0596 | 1.1024 | 1.0074 | 201.5221 | 47.6766 | 57912.4423 | 57466.8608 | 1145.3372 | 1045.3372 | 1055.0135
6350 | 87.4608 | 0.95141 | 1.1424 | 1.0252 | 199.608 | 53.6729 | 67527.3743 | 67252.0634 | 1341.0413 | 1241.0413 | 1055.047
6351 | 103.2041 | 1.0658 | 1.0199 | 1.0377 | 203.3726 | 52.2549 | 61257.1908 | 61802.189 | 1232.0438 | 1132.0438 | 1055.0819
6352 | 77.6683 | 1.0348 | 1.0252 | 1.0484 | 199.3431 | 50.479 | 58511.6987 | 58204.9328 | 1160.0987 | 1060.0987 | 1055.0988
6353 | 110.3309 | 0.95009 | 1.0498 | 0.99107 | 184.1435 | 50.6956 | 58244.3704 | 58088.7812 | 1157.7756 | 1057.7756 | 1055.1247
6354 | 109.3362 | 0.94967 | 0.83591 | 1.0349 | 197.5696 | 50.0648 | 48708.3226 | 47647.2369 | 948.9447 | 848.9447 | 1055.1862
6355 | 103.2849 | 1.0286 | 0.79882 | 0.99167 | 208.4016 | 52.302 | 47197.0002 | 48903.0831 | 974.0617 | 874.0617 | 1055.2946
6356 | 105.189 | 0.99688 | 0.80452 | 1.0404 | 203.3459 | 47.4431 | 44888.072 | 43505.8653 | 866.1173 | 766.1173 | 1055.4221
6357 | 106.0575 | 0.97644 | 0.89734 | 1.0744 | 187.2576 | 49.0604 | 52565.2789 | 51992.7245 | 1035.8545 | 935.8545 | 1055.574
6358 | 98.0042 | 0.91047 | 1.0322 | 1.0176 | 180.6294 | 57.2383 | 65410.2656 | 65889.6936 | 1313.7939 | 1213.7939 | 1055.6014
6359 | 100.1527 | 0.92109 | 1.2132 | 0.96296 | 202.144 | 50.4137 | 63749.431 | 64250.3692 | 1281.0074 | 1181.0074 | 1055.6145
6360 | 92.2257 | 1.0631 | 0.80372 | 0.97005 | 229.5281 | 51.7402 | 45641.4566 | 46931.4918 | 934.6298 | 834.6298 | 1055.6588
6361 | 90.2536 | 0.99253 | 0.96343 | 0.99872 | 199.0762 | 49.8936 | 52675.9928 | 53416.4529 | 1064.3291 | 964.3291 | 1055.7066
6362 | 89.524 | 0.96453 | 0.88775 | 1.0897 | 205.5409 | 49.5588 | 52425.3973 | 51777.4208 | 1031.5484 | 931.5484 | 1055.8283
6363 | 101.386 | 0.94215 | 0.92758 | 1.0643 | 208.8854 | 51.9152 | 56419.5502 | 55877.2397 | 1113.5448 | 1013.5448 | 1055.8694
6364 | 94.0495 | 1.0414 | 1.0774 | 0.97969 | 201.4135 | 50.7133 | 58695.8632 | 56834.517 | 1132.6903 | 1032.6903 | 1055.8874
6365 | 98.5451 | 0.97745 | 1.0237 | 1.0486 | 192.5824 | 49.0223 | 57537.5635 | 57043.121 | 1136.8624 | 1036.8624 | 1055.8914
6366 | 82.7116 | 0.96923 | 1.002 | 0.94542 | 196.252 | 49.928 | 51494.6778 | 52937.5912 | 1054.7518 | 954.7518 | 1055.9988
6367 | 93.3217 | 0.98553 | 0.86552 | 0.95778 | 205.4002 | 53.2183 | 49217.0613 | 49376.9681 | 983.5394 | 883.5394 | 1056.04
6368 | 89.538 | 0.91483 | 1.1019 | 0.9027 | 205.3381 | 52.2615 | 56472.1573 | 57150.5736 | 1139.0115 | 1039.0115 | 1056.0615
6369 | 105.7422 | 0.87906 | 1.0863 | 0.95175 | 205.8454 | 49.6989 | 56208.6327 | 56846.451 | 1132.929 | 1032.929 | 1056.1005
6370 | 100.0748 | 0.92337 | 1.0279 | 1.0151 | 193.8666 | 52.8521 | 60228.9311 | 59670.0275 | 1189.4006 | 1089.4006 | 1056.23
6371 | 99.1347 | 1.0313 | 0.59292 | 1.0065 | 203.938 | 50.1441 | 35255.1864 | 35815.5403 | 712.3108 | 612.3108 | 1056.2639
6372 | 106.317 | 1.0122 | 0.90646 | 1.0334 | 210.2509 | 48.3297 | 50684.3129 | 50903.8461 | 1014.0769 | 914.0769 | 1056.2797
6373 | 110.005 | 0.9932 | 1.0112 | 1.0227 | 218.8345 | 47.6841 | 54742.8942 | 55116.505 | 1098.3301 | 998.3301 | 1056.312
6374 | 109.3536 | 0.97872 | 1.141 | 0.99955 | 201.208 | 45.5688 | 57049.7235 | 57437.8133 | 1144.7563 | 1044.7563 | 1056.4473
6375 | 94.8742 | 1.1056 | 0.94728 | 0.97263 | 213.8316 | 49.9998 | 51526.1522 | 50523.4709 | 1006.4694 | 906.4694 | 1056.4538
6376 | 116.7434 | 1.0511 | 1.201 | 0.99362 | 197.3483 | 50.1389 | 66180.4759 | 65835.9749 | 1312.7195 | 1212.7195 | 1056.4886
6377 | 107.6145 | 1.0954 | 0.95903 | 1.0309 | 192.96 | 48.2842 | 53622.0398 | 53567.921 | 1067.3584 | 967.3584 | 1056.5149
6378 | 96.9526 | 1.0506 | 0.91243 | 1.0957 | 188.9553 | 47.7017 | 52735.3188 | 53531.1299 | 1066.6226 | 966.6226 | 1056.5483
6379 | 88.8621 | 0.99898 | 1.0367 | 1.0105 | 192.4345 | 47.9697 | 54703.0226 | 54987.7209 | 1095.7544 | 995.7544 | 1056.5822
6380 | 106.0968 | 1.0755 | 1.0324 | 1.0098 | 182.7607 | 49.0715 | 56941.7733 | 57288.5888 | 1141.7718 | 1041.7718 | 1056.6307
6381 | 107.9648 | 1.0402 | 1.0462 | 1.0056 | 215.6395 | 48.2619 | 56410.3509 | 57708.1248 | 1150.1625 | 1050.1625 | 1056.6447
6382 | 98.3589 | 1.0345 | 1.0633 | 0.97481 | 193.3274 | 47.2695 | 53997.9726 | 52037.9021 | 1036.758 | 936.758 | 1056.6997
6383 | 115.005 | 0.97519 | 1.0055 | 1.0707 | 193.5308 | 51.7148 | 61669.6874 | 59855.1857 | 1193.1037 | 1093.1037 | 1056.7468
6384 | 97.2203 | 1.0217 | 0.75881 | 0.9972 | 197.5353 | 52.0997 | 44795.7197 | 44725.281 | 890.5056 | 790.5056 | 1056.7813
6385 | 90.8847 | 0.97217 | 1.1864 | 1.008 | 193.2373 | 46.846 | 60354.8716 | 62043.2489 | 1236.865 | 1136.865 | 1056.7995
6386 | 109.6977 | 1.0563 | 0.76431 | 1.0615 | 202.8898 | 52.4914 | 48872.9553 | 48788.9558 | 971.7791 | 871.7791 | 1056.9119
6387 | 88.7155 | 0.9677 | 1.0532 | 0.98365 | 217.0273 | 50.1258 | 56451.7713 | 56780.9911 | 1131.6198 | 1031.6198 | 1056.9882
6388 | 102.5554 | 1.0476 | 0.91699 | 1.0714 | 187.0717 | 52.6364 | 57555.4826 | 57398.3891 | 1143.9678 | 1043.9678 | 1057.0483
6389 | 100.2006 | 0.93173 | 1.1731 | 1.0672 | 208.9673 | 45.3264 | 61187.0626 | 61591.2819 | 1227.8256 | 1127.8256 | 1057.0636
6390 | 111.6525 | 1.0234 | 0.95516 | 0.95642 | 193.2522 | 46.4542 | 47939.0991 | 49095.8798 | 977.9176 | 877.9176 | 1057.1013
6391 | 82.9289 | 1.0114 | 1.3254 | 1.0587 | 227.1808 | 47.7574 | 71247.9696 | 72043.7537 | 1436.8751 | 1336.8751 | 1057.1182
6392 | 83.7961 | 1.0713 | 0.98666 | 0.98712 | 229.3986 | 49.2615 | 52629.811 | 52755.0895 | 1051.1018 | 951.1018 | 1057.17
6393 | 103.7324 | 1.0197 | 1.0612 | 1.0245 | 199.0105 | 52.1252 | 62382.9637 | 61158.7365 | 1219.1747 | 1119.1747 | 1057.216
6394 | 111.6538 | 1.0071 | 1.1768 | 1.0087 | 203.2647 | 53.0476 | 69135.607 | 69728.8546 | 1390.5771 | 1290.5771 | 1057.2722
6395 | 97.5347 | 0.93507 | 1.1976 | 1.0742 | 179.7063 | 50.9679 | 70397.2144 | 71903.9049 | 1434.0781 | 1334.0781 | 1057.3128
6396 | 97.2407 | 0.9978 | 1.0345 | 0.91771 | 193.2848 | 52.9739 | 55625.9076 | 55494.764 | 1105.8953 | 1005.8953 | 1057.3457
6397 | 107.2061 | 0.96342 | 1.1402 | 1.0911 | 201.5121 | 49.2346 | 66538.5137 | 67014.2444 | 1336.2849 | 1236.2849 | 1057.3614
6398 | 99.8975 | 0.94236 | 1.1324 | 1.0044 | 205.2464 | 48.5226 | 59965.0466 | 58478.2493 | 1165.565 | 1065.565 | 1057.3753
6399 | 104.8575 | 1.058 | 1.0951 | 1.0356 | 208.9177 | 47.5794 | 59447.3366 | 60793.6975 | 1211.8739 | 1111.8739 | 1057.5302
6400 | 108.0012 | 1.039 | 0.93454 | 1.0504 | 215.3575 | 48.7303 | 53517.9381 | 53559.2221 | 1067.1844 | 967.1844 | 1057.5617
6401 | 109.2594 | 0.95848 | 1.2825 | 1.0782 | 202.4363 | 49.8621 | 74373.7057 | 72951.3421 | 1455.0268 | 1355.0268 | 1057.5721
6402 | 115.4267 | 1.0064 | 1.078 | 1.0606 | 212.8313 | 54.3223 | 68629.5649 | 66987.3602 | 1335.7472 | 1235.7472 | 1057.6402
6403 | 93.584 | 0.96042 | 1.0204 | 1.1059 | 192.8066 | 49.9355 | 61031.2492 | 60689.6246 | 1209.7925 | 1109.7925 | 1057.6601
6404 | 98.8812 | 0.91241 | 1.0614 | 1.007 | 206.9396 | 51.1864 | 59534.4635 | 60322.6003 | 1202.452 | 1102.452 | 1057.6621
6405 | 100.0049 | 1.0664 | 0.77958 | 0.9773 | 208.9687 | 51.5698 | 44998.4022 | 45215.0296 | 900.3006 | 800.3006 | 1057.6906
6406 | 87.8082 | 0.99158 | 1.2463 | 0.96769 | 207.1007 | 48.7441 | 63239.3425 | 63600.7757 | 1268.0155 | 1168.0155 | 1057.6916
6407 | 87.8087 | 0.99289 | 0.86067 | 0.93307 | 212.3626 | 46.2392 | 41376.8555 | 41762.6223 | 831.2524 | 731.2524 | 1057.744
6408 | 106.2625 | 1.0329 | 1.175 | 1.029 | 192.5824 | 51.7021 | 68380.4268 | 68420.6886 | 1364.4138 | 1264.4138 | 1057.7587
6409 | 99.2483 | 1.0035 | 1.1674 | 1.1044 | 192.0877 | 54.3655 | 75701.1553 | 75575.3335 | 1507.5067 | 1407.5067 | 1057.7756
6410 | 99.0288 | 0.91441 | 1.2678 | 1.0552 | 189.5546 | 50.0525 | 71680.5737 | 71089.1836 | 1417.7837 | 1317.7837 | 1057.7826
6411 | 106.8116 | 1.07 | 1.132 | 1.0027 | 190.4694 | 50.0698 | 62744.377 | 61852.2421 | 1233.0448 | 1133.0448 | 1057.8284
6412 | 105.2815 | 1.0029 | 1.1149 | 1.015 | 214.4995 | 48.5412 | 60271.0593 | 60978.1563 | 1215.5631 | 1115.5631 | 1057.8335
6413 | 106.8596 | 0.96414 | 0.85869 | 1.0566 | 195.8829 | 50.6418 | 51361.7899 | 50046.857 | 996.9371 | 896.9371 | 1057.8532
6414 | 109.3373 | 1.0339 | 1.1892 | 0.98607 | 181.395 | 49.213 | 63455.4111 | 62668.9977 | 1249.38 | 1149.38 | 1057.8987
6415 | 96.5351 | 1.1239 | 1.2418 | 1.0477 | 208.6775 | 48.6289 | 68752.2863 | 67734.1964 | 1350.6839 | 1250.6839 | 1058.0208
6416 | 121.0472 | 0.97467 | 0.82247 | 1.0722 | 198.8017 | 51.5833 | 51772.1398 | 52463.2692 | 1045.2654 | 945.2654 | 1058.0622
6417 | 101.4559 | 1.0187 | 0.90011 | 0.98335 | 222.0713 | 48.656 | 48317.1496 | 47867.0903 | 953.3418 | 853.3418 | 1058.1614
6418 | 111.6708 | 0.89462 | 1.0064 | 1.0072 | 198.2156 | 46.9365 | 52465.1994 | 52350.1617 | 1043.0032 | 943.0032 | 1058.2301
6419 | 106.8339 | 0.94991 | 1.0897 | 0.98707 | 190.825 | 51.123 | 60364.5846 | 60254.5078 | 1201.0902 | 1101.0902 | 1058.2382
6420 | 102.1536 | 0.93608 | 0.93511 | 1.0861 | 214.4447 | 46.059 | 51398.5822 | 51809.4437 | 1032.1889 | 932.1889 | 1058.2451
6421 | 103.8529 | 1.0157 | 0.98036 | 0.93893 | 199.0835 | 48.4821 | 49940.5474 | 50801.2377 | 1012.0248 | 912.0248 | 1058.2504
6422 | 90.5585 | 0.95681 | 0.98036 | 1.0793 | 194.7207 | 52.3181 | 60086.986 | 61370.7886 | 1223.4158 | 1123.4158 | 1058.265
6423 | 98.7568 | 1.0168 | 1.2808 | 1.0427 | 195.8786 | 51.4134 | 74023.8768 | 72683.364 | 1449.6673 | 1349.6673 | 1058.2843
6424 | 94.5692 | 0.94901 | 1.0524 | 0.96908 | 212.7038 | 52.9066 | 58917.2765 | 59322.22 | 1182.4444 | 1082.4444 | 1058.2891
6425 | 108.6851 | 1.0389 | 1.1218 | 1.0007 | 205.4514 | 52.9662 | 65645.6085 | 62859.0777 | 1253.1816 | 1153.1816 | 1058.2962
6426 | 98.0115 | 0.97591 | 0.9673 | 0.99828 | 182.7059 | 47.7928 | 50904.4615 | 51706.7317 | 1030.1346 | 930.1346 | 1058.3151
6427 | 102.9264 | 0.97523 | 0.78358 | 1.0207 | 191.8553 | 49.6284 | 44866.7672 | 45098.9946 | 897.9799 | 797.9799 | 1058.3395
6428 | 113.3273 | 0.96402 | 0.91038 | 0.94923 | 204.1754 | 52.222 | 51037.3714 | 51794.6558 | 1031.8931 | 931.8931 | 1058.3873
6429 | 104.3122 | 1.013 | 0.83268 | 0.983 | 182.9982 | 49.5246 | 45953.1313 | 45193.0495 | 899.861 | 799.861 | 1058.3986
6430 | 94.3606 | 1.0925 | 1.1145 | 0.96397 | 205.0173 | 50.4568 | 59613.1598 | 60979.3397 | 1215.5868 | 1115.5868 | 1058.4096
6431 | 100.9342 | 1.0054 | 1.0951 | 1.0894 | 203.7713 | 54.396 | 70621.1112 | 72270.7736 | 1441.4155 | 1341.4155 | 1058.5897
6432 | 101.2594 | 1.0003 | 0.81687 | 0.93375 | 208.5815 | 51.1651 | 44416.9193 | 42940.5716 | 854.8114 | 754.8114 | 1058.6098
6433 | 105.8746 | 0.92387 | 1.0094 | 0.91434 | 205.2793 | 50.101 | 51345.8915 | 52215.3086 | 1040.3062 | 940.3062 | 1058.6351
6434 | 95.1096 | 0.95709 | 0.70393 | 0.95448 | 198.3965 | 47.8624 | 36713.2224 | 36456.9825 | 725.1397 | 625.1397 | 1058.7124
6435 | 109.3526 | 0.92378 | 0.82742 | 0.92865 | 208.69 | 51.4585 | 44946.6982 | 44766.8011 | 891.336 | 791.336 | 1058.7129
6436 | 98.8264 | 0.99204 | 0.87431 | 1.0273 | 196.5937 | 51.4526 | 51456.0157 | 51015.6596 | 1016.3132 | 916.3132 | 1058.7536
6437 | 103.6953 | 0.99732 | 0.89964 | 1.0804 | 209.6478 | 54.4884 | 58807.3318 | 59118.0473 | 1178.3609 | 1078.3609 | 1058.822
6438 | 88.2942 | 1.0988 | 0.80621 | 1.0129 | 201.5434 | 47.1435 | 43273.0299 | 42770.6503 | 851.413 | 751.413 | 1058.8868
6439 | 95.1984 | 0.9972 | 1.0554 | 0.91466 | 193.5109 | 50.2362 | 53457.6301 | 53592.4821 | 1067.8496 | 967.8496 | 1058.9255
6440 | 95.4275 | 0.97463 | 0.85272 | 1.0796 | 202.7878 | 51.6897 | 52597.939 | 52440.3983 | 1044.808 | 944.808 | 1058.9393
6441 | 100.0224 | 1.0502 | 0.85391 | 1.0279 | 199.1089 | 50.12 | 49454.5942 | 49142.176 | 978.8435 | 878.8435 | 1058.947
6442 | 78.1197 | 1.0263 | 1.0535 | 0.94573 | 193.6758 | 53.6591 | 57955.8702 | 58700.8408 | 1170.0168 | 1070.0168 | 1059.101
6443 | 95.123 | 1.014 | 1.0582 | 1.0429 | 181.594 | 50.647 | 60959.8831 | 62288.2822 | 1241.7656 | 1141.7656 | 1059.1034
6444 | 97.6829 | 0.99881 | 1.1274 | 0.96053 | 210.8183 | 50.1873 | 59455.5697 | 58959.7041 | 1175.1941 | 1075.1941 | 1059.1251
6445 | 89.6948 | 1.0215 | 0.9971 | 1.0911 | 198.5943 | 46.2352 | 54735.3433 | 54407.9509 | 1084.159 | 984.159 | 1059.133
6446 | 105.4269 | 1.0277 | 0.91737 | 1.009 | 191.3058 | 51.5353 | 53475.5349 | 52732.2983 | 1050.646 | 950.646 | 1059.2037
6447 | 72.6495 | 0.92875 | 0.95035 | 1.0254 | 189.2824 | 50.1183 | 52408.7791 | 52845.4791 | 1052.9096 | 952.9096 | 1059.2445
6448 | 77.4155 | 1.0486 | 1.1211 | 0.95493 | 208.6192 | 54.7272 | 63239.8736 | 61955.4152 | 1235.1083 | 1135.1083 | 1059.378
6449 | 92.1431 | 1.0204 | 0.79469 | 1.0227 | 198.6946 | 55.0425 | 50109.9705 | 48456.5199 | 965.1304 | 865.1304 | 1059.4452
6450 | 97.4775 | 1.0082 | 1.1398 | 0.95309 | 186.2652 | 48.8938 | 58106.7531 | 61400.7691 | 1224.0154 | 1124.0154 | 1059.4628
6451 | 82.0941 | 0.95402 | 1.0837 | 0.95409 | 203.8476 | 51.2069 | 57161.1306 | 56923.4422 | 1134.4688 | 1034.4688 | 1059.4648
6452 | 120.6085 | 1.0057 | 1.2079 | 0.99455 | 207.8653 | 52.4975 | 69643.218 | 70480.5351 | 1405.6107 | 1305.6107 | 1059.4822
6453 | 90.3252 | 0.99209 | 1.1711 | 0.99875 | 202.5349 | 48.7233 | 61557.1335 | 61767.9794 | 1231.3596 | 1131.3596 | 1059.5013
6454 | 92.7904 | 0.91261 | 0.92419 | 0.90638 | 194.8324 | 55.8169 | 51677.3339 | 53039.474 | 1056.7895 | 956.7895 | 1059.5218
6455 | 119.4809 | 0.96199 | 0.90336 | 0.99175 | 218.4986 | 50.7852 | 51554.6228 | 52095.5733 | 1037.9115 | 937.9115 | 1059.5285
6456 | 124.7004 | 1.009 | 1.0991 | 0.94053 | 203.5986 | 46.6916 | 54344.8417 | 54744.8717 | 1090.8974 | 990.8974 | 1059.6825
6457 | 94.3652 | 1.0406 | 0.88695 | 0.93361 | 195.4057 | 52.0635 | 48419.6241 | 50839.6414 | 1012.7928 | 912.7928 | 1059.7561
6458 | 101.5209 | 1.0327 | 1.2719 | 0.91684 | 185.5736 | 51.4055 | 65518.6727 | 65610.945 | 1308.2189 | 1208.2189 | 1059.757
6459 | 85.7785 | 1.0293 | 0.99977 | 0.95999 | 212.3018 | 52.7766 | 55525.3999 | 53414.8633 | 1064.2973 | 964.2973 | 1059.794
6460 | 108.4856 | 1.0163 | 0.79823 | 0.96968 | 194.4931 | 46.895 | 41662.6478 | 41225.2548 | 820.5051 | 720.5051 | 1059.8048
6461 | 118.0586 | 1.0724 | 1.0947 | 1.0768 | 208.0559 | 48.7063 | 63789.0376 | 63947.9927 | 1274.9599 | 1174.9599 | 1059.8146
6462 | 117.1039 | 1.0358 | 0.9561 | 0.95766 | 177.4837 | 50.5141 | 52556.4001 | 53730.6953 | 1070.6139 | 970.6139 | 1059.8196
6463 | 92.9993 | 1.0145 | 0.93496 | 0.96153 | 186.2263 | 51.2464 | 51091.3576 | 51292.2256 | 1021.8445 | 921.8445 | 1059.8243
6464 | 102.5672 | 0.99204 | 0.79647 | 1.0132 | 210.6234 | 50.6222 | 46211.9005 | 44781.4076 | 891.6282 | 791.6282 | 1059.8564
6465 | 107.1443 | 1.0758 | 1.1696 | 1.0344 | 203.9213 | 49.3723 | 65630.0358 | 64558.9521 | 1287.179 | 1187.179 | 1059.8651
6466 | 101.1187 | 1.0216 | 1.1693 | 1.0138 | 205.738 | 52.051 | 67284.7382 | 62515.3242 | 1246.3065 | 1146.3065 | 1059.9597
6467 | 94.8005 | 0.95812 | 1.0953 | 0.99114 | 200.9975 | 51.2656 | 60510.6345 | 56679.4944 | 1129.5899 | 1029.5899 | 1060.0315
6468 | 89.0964 | 0.98268 | 1.0742 | 0.97378 | 182.8703 | 51.5837 | 58658.0666 | 58389.2696 | 1163.7854 | 1063.7854 | 1060.0722
6469 | 87.7173 | 1.0675 | 0.93186 | 1.0116 | 197.4683 | 47.4331 | 49352.4502 | 49995.4758 | 995.9095 | 895.9095 | 1060.093
6470 | 94.2779 | 0.98702 | 0.7269 | 1.0106 | 191.6637 | 51.3828 | 42720.9259 | 42681.097 | 849.6219 | 749.6219 | 1060.0987
6471 | 103.2981 | 0.97698 | 1.0038 | 0.95254 | 206.687 | 52.8897 | 56113.8785 | 57177.994 | 1139.5599 | 1039.5599 | 1060.1651
6472 | 102.2177 | 0.97939 | 0.91585 | 0.92184 | 212.9706 | 48.7295 | 46232.2945 | 46002.6449 | 916.0529 | 816.0529 | 1060.2379
6473 | 92.5914 | 0.99505 | 0.90307 | 1.0006 | 196.0432 | 56.3499 | 56305.4031 | 57959.3091 | 1155.1862 | 1055.1862 | 1060.3389
6474 | 106.0921 | 0.95064 | 1.3302 | 0.98944 | 200.087 | 46.1489 | 65592.0408 | 65781.1754 | 1311.6235 | 1211.6235 | 1060.3758
6475 | 110.6665 | 1.0237 | 0.97392 | 1.0121 | 203.0719 | 51.634 | 56946.7918 | 56717.3381 | 1130.3468 | 1030.3468 | 1060.3979
6476 | 100.5196 | 1.0356 | 1.294 | 1.0264 | 194.8373 | 49.4421 | 71008.4301 | 69676.5838 | 1389.5317 | 1289.5317 | 1060.4262
6477 | 112.0876 | 0.97507 | 1.0185 | 0.92288 | 184.9906 | 47.29 | 49803.6119 | 49648.1062 | 988.9621 | 888.9621 | 1060.454
6478 | 94.63 | 0.95751 | 1.0062 | 1.0121 | 220.8695 | 51.3133 | 57124.475 | 55337.3757 | 1102.7475 | 1002.7475 | 1060.4619
6479 | 87.691 | 1.0582 | 0.91366 | 0.93317 | 195.4795 | 52.943 | 50247.698 | 49793.4196 | 991.8684 | 891.8684 | 1060.4695
6480 | 104.691 | 1.0483 | 1.0679 | 1.0388 | 197.4221 | 50.0691 | 61230.8575 | 61430.1901 | 1224.6038 | 1124.6038 | 1060.5623
6481 | 111.2177 | 1.1091 | 0.96445 | 1.0026 | 230.5356 | 47.4302 | 51942.6576 | 51604.893 | 1028.0979 | 928.0979 | 1060.5747
6482 | 92.8737 | 1.0397 | 1.1303 | 1.0218 | 219.7843 | 52.3764 | 65772.5625 | 64626.648 | 1288.533 | 1188.533 | 1060.5787
6483 | 105.115 | 0.95836 | 0.82684 | 1.0108 | 200.6106 | 54.2066 | 50967.4271 | 51969.9131 | 1035.3983 | 935.3983 | 1060.6026
6484 | 104.8449 | 1.0731 | 0.81446 | 1.0278 | 196.0567 | 54.5768 | 52021.7976 | 51109.5731 | 1018.1915 | 918.1915 | 1060.7833
6485 | 116.0498 | 1.0017 | 1.0857 | 0.96835 | 200.3327 | 50.7177 | 59416.6142 | 59408.0618 | 1184.1612 | 1084.1612 | 1060.8123
6486 | 98.1987 | 0.95471 | 1.093 | 0.97948 | 212.0792 | 52.5834 | 61436.4936 | 61068.7738 | 1217.3755 | 1117.3755 | 1060.8124
6487 | 95.4195 | 0.99511 | 1.1536 | 0.93139 | 208.1183 | 53.1182 | 62326.9704 | 61514.4517 | 1226.289 | 1126.289 | 1060.8133
6488 | 104.6119 | 1.0313 | 1.0752 | 1.0236 | 196.2528 | 50.772 | 61554.0116 | 60256.5319 | 1201.1306 | 1101.1306 | 1060.8214
6489 | 100.1443 | 1.0328 | 0.6962 | 0.98114 | 209.3798 | 51.546 | 40750.1278 | 40065.8569 | 797.3171 | 697.3171 | 1060.8218
6490 | 105.5529 | 1.0406 | 0.92155 | 1.017 | 207.9858 | 49.2065 | 51729.9416 | 50971.2381 | 1015.4248 | 915.4248 | 1060.8752
6491 | 90.6237 | 0.98504 | 0.84692 | 1.0221 | 192.3101 | 48.5291 | 46534.8716 | 47429.5497 | 944.591 | 844.591 | 1060.877
6492 | 83.7291 | 0.94947 | 1.3367 | 0.97341 | 200.7953 | 48.6014 | 67302.4986 | 65350.0296 | 1303.0006 | 1203.0006 | 1060.9226
6493 | 90.9866 | 0.98442 | 1.0689 | 1.0015 | 202.1128 | 51.5347 | 59987.3874 | 60318.3505 | 1202.367 | 1102.367 | 1060.9449
6494 | 95.0706 | 0.96431 | 0.94956 | 0.89918 | 221.2962 | 53.9683 | 51248.7226 | 52349.5562 | 1042.9911 | 942.9911 | 1060.9862
6495 | 90.8733 | 0.99593 | 1.1384 | 0.92839 | 201.3175 | 50.8107 | 58502.5763 | 59432.964 | 1184.6593 | 1084.6593 | 1061.1027
6496 | 94.9983 | 0.92911 | 0.99351 | 1.0533 | 186.9816 | 46.315 | 52740.2338 | 52848.4803 | 1052.9696 | 952.9696 | 1061.1263
6497 | 103.68 | 1.0652 | 1.0018 | 1.0919 | 190.3777 | 49.93 | 60321.7896 | 60394.0046 | 1203.8801 | 1103.8801 | 1061.1851
6498 | 110.7954 | 0.9448 | 0.94053 | 1.0176 | 207.2099 | 53.2288 | 56724.7647 | 57347.0369 | 1142.9407 | 1042.9407 | 1061.2045
6499 | 115.8845 | 0.99988 | 0.84501 | 0.97643 | 203.0745 | 50.3089 | 47541.7093 | 48218.677 | 960.3735 | 860.3735 | 1061.2363
6500 | 86.8284 | 0.9358 | 1.0489 | 1.0146 | 206.8118 | 53.2279 | 61182.0828 | 58787.6911 | 1171.7538 | 1071.7538 | 1061.2443
6501 | 97.6327 | 1.033 | 1.2278 | 1.0898 | 202.5247 | 54.7198 | 78944.3858 | 79081.8252 | 1577.6365 | 1477.6365 | 1061.277
6502 | 99.3172 | 1.0033 | 1.256 | 1.1007 | 201.4587 | 48.4477 | 72007.0452 | 71614.376 | 1428.2875 | 1328.2875 | 1061.2878
6503 | 97.3442 | 0.91122 | 0.95274 | 0.97325 | 204.974 | 48.6893 | 49671.3595 | 48963.5275 | 975.2705 | 875.2705 | 1061.2909
6504 | 96.6765 | 0.96772 | 0.8674 | 0.96085 | 192.9949 | 49.8113 | 46368.2474 | 45886.4456 | 913.7289 | 813.7289 | 1061.3249
6505 | 100.8911 | 1.0423 | 0.9166 | 1.0118 | 204.987 | 51.8284 | 53722.3896 | 53904.33 | 1074.0866 | 974.0866 | 1061.3358
6506 | 107.7585 | 0.99898 | 1.035 | 0.99724 | 195.9145 | 50.252 | 57471.0428 | 56989.3818 | 1135.7876 | 1035.7876 | 1061.394
6507 | 102.8352 | 1.0064 | 1.1365 | 0.94602 | 190.4828 | 50.0645 | 59200.3285 | 58924.497 | 1174.4899 | 1074.4899 | 1061.4411
6508 | 82.4004 | 1.0254 | 1.0429 | 1.0294 | 192.9385 | 50.8417 | 59071.3734 | 58875.3041 | 1173.5061 | 1073.5061 | 1061.6112
6509 | 96.3902 | 1.0353 | 1.1359 | 0.9923 | 187.8791 | 51.5117 | 63388.93 | 63077.5448 | 1257.5509 | 1157.5509 | 1061.6712
6510 | 81.5929 | 1.0147 | 1.1332 | 1.0257 | 214.6785 | 50.7011 | 63343.6385 | 63354.4891 | 1263.0898 | 1163.0898 | 1061.719
6511 | 112.7425 | 0.97999 | 0.85667 | 1.0117 | 188.9582 | 54.2147 | 53165.9609 | 55048.2065 | 1096.9641 | 996.9641 | 1061.7313
6512 | 81.5617 | 1.0127 | 0.93377 | 1.0468 | 197.387 | 50.2857 | 53502.7635 | 51198.5879 | 1019.9718 | 919.9718 | 1061.7685
6513 | 105.7552 | 0.88918 | 1.1832 | 1.0401 | 201.7341 | 45.6747 | 60706.6834 | 60048.7164 | 1196.9743 | 1096.9743 | 1061.8175
6514 | 103.3488 | 0.94649 | 1.1154 | 1.0028 | 202.5882 | 46.9431 | 57297.9857 | 56828.2378 | 1132.5648 | 1032.5648 | 1061.8258
6515 | 85.7085 | 1.0369 | 0.90999 | 0.99981 | 186.7998 | 46.9415 | 47066.5472 | 45103.6399 | 898.0728 | 798.0728 | 1061.8515
6516 | 106.5424 | 1.0528 | 1.0623 | 1.0585 | 202.7035 | 46.609 | 57840.9876 | 56769.8138 | 1131.3963 | 1031.3963 | 1061.932
6517 | 88.4972 | 0.9069 | 0.77725 | 1.0228 | 193.1142 | 49.2671 | 43314.7941 | 43431.3445 | 864.6269 | 764.6269 | 1061.9549
6518 | 95.1521 | 1.0951 | 1.1309 | 0.98222 | 193.6251 | 51.2981 | 62520.4291 | 62086.99 | 1237.7398 | 1137.7398 | 1061.9721
6519 | 102.1371 | 0.96883 | 0.96366 | 1.084 | 180.9887 | 53.7919 | 61694.7909 | 63993.9933 | 1275.8799 | 1175.8799 | 1062.0033
6520 | 97.0299 | 1.0355 | 1.1491 | 0.99728 | 188.026 | 48.6408 | 60814.3849 | 58755.1457 | 1171.1029 | 1071.1029 | 1062.0036
6521 | 83.8087 | 0.96259 | 1.108 | 0.94926 | 184.3019 | 50.8848 | 57807.1558 | 57238.2441 | 1140.7649 | 1040.7649 | 1062.0526
6522 | 122.8576 | 0.97677 | 0.98861 | 0.93507 | 192.1335 | 51.6384 | 54124.5898 | 54963.443 | 1095.2689 | 995.2689 | 1062.053
6523 | 98.08 | 1.084 | 1.3212 | 1.006 | 206.9431 | 53.0886 | 76415.3579 | 77190.8929 | 1539.8179 | 1439.8179 | 1062.0729
6524 | 91.2935 | 0.98952 | 1.0576 | 1.0244 | 227.1128 | 49.7965 | 58673.9334 | 58187.8058 | 1159.7561 | 1059.7561 | 1062.1012
6525 | 103.9712 | 1.0199 | 1.0184 | 0.96368 | 201.8777 | 51.2747 | 55959.7141 | 53819.2505 | 1072.385 | 972.385 | 1062.2523
6526 | 110.3478 | 1.0168 | 0.88162 | 1.0104 | 192.872 | 52.7062 | 53055.8625 | 52133.924 | 1038.6785 | 938.6785 | 1062.3156
6527 | 88.0706 | 1.0373 | 0.96469 | 0.94747 | 195.4636 | 54.161 | 54646.7857 | 56727.8114 | 1130.5562 | 1030.5562 | 1062.3187
6528 | 107.4109 | 0.92548 | 0.90162 | 1.0939 | 196.4921 | 49.6386 | 54087.0769 | 53135.7139 | 1058.7143 | 958.7143 | 1062.3835
6529 | 117.1231 | 0.98415 | 1.2679 | 1.0168 | 212.8387 | 50.6877 | 71401.6042 | 72095.3192 | 1437.9064 | 1337.9064 | 1062.4348
6530 | 95.1339 | 1.0519 | 1.0488 | 1.0045 | 199.3313 | 51.2954 | 59372.6128 | 59368.2447 | 1183.3649 | 1083.3649 | 1062.4651
6531 | 107.3982 | 0.95543 | 0.90045 | 0.95518 | 196.0402 | 45.906 | 44389.9154 | 44909.0022 | 894.18 | 794.18 | 1062.4958
6532 | 95.064 | 0.99447 | 0.83607 | 0.99694 | 190.8219 | 47.3436 | 44127.8715 | 43665.0197 | 869.3004 | 769.3004 | 1062.5124
6533 | 104.1777 | 1.0165 | 0.76719 | 0.93499 | 199.9939 | 50.6812 | 41921.3023 | 42303.6444 | 842.0729 | 742.0729 | 1062.5254
6534 | 103.4056 | 0.97663 | 0.76424 | 0.93234 | 193.1843 | 46.706 | 38189.5554 | 38954.7349 | 775.0947 | 675.0947 | 1062.5483
6535 | 113.8479 | 1.0066 | 1.0914 | 1.0206 | 198.495 | 50.1612 | 61824.5749 | 58713.3181 | 1170.2664 | 1070.2664 | 1062.7544
6536 | 96.4889 | 0.97952 | 1.1354 | 1.0692 | 197.6189 | 48.9868 | 64298.1435 | 63265.3509 | 1261.307 | 1161.307 | 1062.828
6537 | 89.7132 | 1.0125 | 1.1943 | 1.0012 | 202.51 | 45.7118 | 59009.7254 | 60986.0888 | 1215.7218 | 1115.7218 | 1062.8766
6538 | 94.8844 | 1.0206 | 1.1213 | 0.99284 | 209.5833 | 50.5587 | 61391.9484 | 63534.995 | 1266.6999 | 1166.6999 | 1062.8938
6539 | 117.8074 | 0.96495 | 1.102 | 0.98951 | 198.5939 | 49.6551 | 59990.9822 | 59898.7217 | 1193.9744 | 1093.9744 | 1063.0355
6540 | 101.6562 | 1.0632 | 1.0823 | 0.93964 | 191.6586 | 51.7743 | 58442.2974 | 56960.3195 | 1135.2064 | 1035.2064 | 1063.0756
6541 | 97.345 | 0.9379 | 1.0976 | 0.93987 | 198.0421 | 46.0673 | 51926.5114 | 51882.9369 | 1033.6587 | 933.6587 | 1063.079
6542 | 104.6534 | 0.97724 | 1.09 | 1.0552 | 194.2287 | 46.74 | 58733.7153 | 59086.3649 | 1177.7273 | 1077.7273 | 1063.139
6543 | 96.6424 | 0.9845 | 0.69895 | 0.94712 | 204.1412 | 50.3511 | 38326.5501 | 37874.3787 | 753.4876 | 653.4876 | 1063.158
6544 | 97.1819 | 0.97106 | 1.1408 | 1.1025 | 191.5857 | 53.3388 | 72310.4866 | 71178.6615 | 1419.5732 | 1319.5732 | 1063.284
6545 | 117.5679 | 1.0457 | 1.04 | 0.89811 | 192.332 | 47.4532 | 50351.0373 | 49581.9646 | 987.6393 | 887.6393 | 1063.2853
6546 | 109.7745 | 0.92161 | 0.97587 | 0.99294 | 197.3582 | 52.8579 | 56763.2055 | 56040.9741 | 1116.8195 | 1016.8195 | 1063.4013
6547 | 117.4987 | 0.92947 | 1.2289 | 1.015 | 189.1264 | 53.401 | 72626.8183 | 74562.0935 | 1487.2419 | 1387.2419 | 1063.4368
6548 | 103.4283 | 0.99986 | 0.68712 | 1.0209 | 204.2389 | 52.8852 | 42772.6675 | 41891.1724 | 833.8234 | 733.8234 | 1063.485
6549 | 89.6288 | 0.99825 | 0.86882 | 1.0091 | 190.4744 | 47.5932 | 46173.1355 | 46437.1049 | 924.7421 | 824.7421 | 1063.5335
6550 | 104.782 | 1.0172 | 0.65444 | 1.0083 | 201.4856 | 48.8282 | 37627.2748 | 38423.428 | 764.4686 | 664.4686 | 1063.5349
6551 | 93.8388 | 0.91267 | 1.056 | 1.0316 | 219.7469 | 46.4271 | 54772.9175 | 54997.7149 | 1095.9543 | 995.9543 | 1063.5482
6552 | 93.9532 | 0.92449 | 0.91968 | 0.95842 | 208.7704 | 45.9903 | 44740.9366 | 45218.1989 | 900.364 | 800.364 | 1063.5551
6553 | 98.3855 | 0.99648 | 1.1459 | 0.97697 | 201.1031 | 54.6366 | 66726.0798 | 66041.4377 | 1316.8288 | 1216.8288 | 1063.6357
6554 | 122.2517 | 1.0154 | 0.74343 | 0.99329 | 203.6092 | 51.3004 | 44454.0454 | 44546.7368 | 886.9347 | 786.9347 | 1063.6396
6555 | 98.9855 | 1.0423 | 1.0052 | 0.99841 | 185.4103 | 46.438 | 51579.7482 | 52175.4632 | 1039.5093 | 939.5093 | 1063.7474
6556 | 110.8817 | 1.0138 | 1.1536 | 0.90675 | 201.6579 | 53.4242 | 62091.0649 | 61883.1388 | 1233.6628 | 1133.6628 | 1063.7854
6557 | 104.5766 | 0.98236 | 0.74132 | 0.98432 | 205.8155 | 46.8271 | 39186.0856 | 39772.196 | 791.4439 | 691.4439 | 1063.82
6558 | 110.3863 | 1.0697 | 1.0329 | 1.142 | 208.5164 | 53.9834 | 70259.151 | 70775.3517 | 1411.507 | 1311.507 | 1063.8664
6559 | 95.7804 | 1.0551 | 1.0949 | 0.95704 | 185.454 | 52.7709 | 60816.4054 | 58211.8929 | 1160.2379 | 1060.2379 | 1063.9662
6560 | 96.8959 | 1.0659 | 1.2064 | 0.89265 | 193.3096 | 47.5727 | 56336.8429 | 56173.1325 | 1119.4627 | 1019.4627 | 1064.0059
6561 | 105.8839 | 0.9895 | 1.066 | 0.97562 | 206.9995 | 48.4348 | 55652.6619 | 55132.536 | 1098.6507 | 998.6507 | 1064.0509
6562 | 102.4613 | 1.0076 | 0.99544 | 1.0577 | 204.5903 | 53.2812 | 61801.5936 | 61894.3072 | 1233.8861 | 1133.8861 | 1064.1264
6563 | 90.2937 | 0.98617 | 0.91639 | 0.94591 | 222.9957 | 51.1807 | 49145.2784 | 48789.762 | 971.7952 | 871.7952 | 1064.1505
6564 | 92.7936 | 1.0659 | 0.73913 | 0.83446 | 190.5401 | 52.278 | 37604.6353 | 37404.5097 | 744.0902 | 644.0902 | 1064.1654
6565 | 94.8604 | 1.0467 | 1.1471 | 0.97871 | 192.9832 | 48.7098 | 59714.073 | 59321.7449 | 1182.4349 | 1082.4349 | 1064.1842
6566 | 101.9885 | 0.97134 | 1.0374 | 0.99288 | 199.5487 | 47.2254 | 53522.3985 | 52535.5512 | 1046.711 | 946.711 | 1064.1884
6567 | 77.7317 | 0.94581 | 1.0225 | 1.0871 | 200.1422 | 49.4892 | 58850.7006 | 57822.0805 | 1152.4416 | 1052.4416 | 1064.2838
6568 | 102.3153 | 0.94091 | 1.0001 | 1.0199 | 206.0121 | 49.5379 | 55504.8439 | 55282.9384 | 1101.6588 | 1001.6588 | 1064.376
6569 | 99.1762 | 1.0698 | 1.0827 | 1.0392 | 199.1911 | 53.4013 | 65951.6547 | 66074.6706 | 1317.4934 | 1217.4934 | 1064.4162
6570 | 84.643 | 1.0106 | 0.87905 | 0.92637 | 183.4835 | 50.148 | 45309.6508 | 45419.9618 | 904.3992 | 804.3992 | 1064.5366
6571 | 93.5559 | 0.99632 | 0.86464 | 1.0193 | 201.8345 | 47.8774 | 46859.7626 | 47756.3582 | 951.1272 | 851.1272 | 1064.5433
6572 | 93.9468 | 1.0302 | 1.0745 | 0.9767 | 206.0368 | 49.1606 | 56556.995 | 56203.8546 | 1120.0771 | 1020.0771 | 1064.5507
6573 | 103.0303 | 0.9889 | 1.0693 | 1.0505 | 194.9007 | 47.4005 | 58269.4459 | 55568.9296 | 1107.3786 | 1007.3786 | 1064.585
6574 | 80.2413 | 0.952 | 0.89461 | 1.1059 | 202.1646 | 52.8922 | 56573.7257 | 57040.7286 | 1136.8146 | 1036.8146 | 1064.6789
6575 | 91.4862 | 0.95112 | 1.1701 | 0.97965 | 198.7792 | 49.9039 | 61743.1904 | 62018.8073 | 1236.3761 | 1136.3761 | 1064.7332
6576 | 74.4724 | 1.0023 | 0.74256 | 0.95703 | 201.6609 | 49.2253 | 38858.259 | 38582.5814 | 767.6516 | 667.6516 | 1064.8221
6577 | 96.9991 | 0.98724 | 1.1742 | 1.0187 | 194.8482 | 53.6426 | 69495.912 | 68489.7938 | 1365.7959 | 1265.7959 | 1064.8402
6578 | 94.6462 | 1.11 | 0.83112 | 1.0544 | 204.4354 | 50.0684 | 49339.0892 | 49587.2764 | 987.7455 | 887.7455 | 1064.9574
6579 | 85.051 | 1.0332 | 1.0736 | 0.9932 | 187.6652 | 49.0867 | 56842.0513 | 54767.3433 | 1091.3469 | 991.3469 | 1065.0139
6580 | 107.6365 | 1.0176 | 1.1928 | 0.90685 | 198.7977 | 50.6852 | 60577.4431 | 60038.483 | 1196.7697 | 1096.7697 | 1065.0166
6581 | 101.0442 | 1.0538 | 0.95599 | 0.98036 | 198.2134 | 48.7412 | 51069.2467 | 50235.4154 | 1000.7083 | 900.7083 | 1065.0247
6582 | 91.0889 | 1.0575 | 1.0316 | 1.0105 | 186.1622 | 47.7523 | 54561.346 | 54752.0776 | 1091.0416 | 991.0416 | 1065.0862
6583 | 104.7372 | 1.0493 | 1.197 | 0.97654 | 212.9171 | 52.7226 | 67636.0667 | 65782.4351 | 1311.6487 | 1211.6487 | 1065.1184
6584 | 91.4838 | 1.0505 | 0.97601 | 0.92578 | 222.0046 | 47.9094 | 48116.1783 | 48452.6649 | 965.0533 | 865.0533 | 1065.236
6585 | 96.17 | 1.026 | 0.99535 | 1.0803 | 206.61 | 45.1299 | 53184.8892 | 52314.1738 | 1042.2835 | 942.2835 | 1065.2509
6586 | 89.6367 | 1.082 | 1.1142 | 0.94668 | 204.7039 | 54.769 | 63287.3399 | 61523.9969 | 1226.4799 | 1126.4799 | 1065.2841
6587 | 108.5432 | 1.0123 | 1.1867 | 1.0288 | 191.0423 | 53.3056 | 71128.0785 | 73305.2227 | 1462.1045 | 1362.1045 | 1065.299
6588 | 100.3712 | 0.9621 | 1.1193 | 1.0567 | 203.7552 | 51.6422 | 66273.524 | 65464.5741 | 1305.2915 | 1205.2915 | 1065.3268
6589 | 102.0081 | 0.97835 | 0.71232 | 1.0298 | 195.598 | 48.1685 | 40338.0493 | 40428.516 | 804.5703 | 704.5703 | 1065.4002
6590 | 102.7842 | 1.0534 | 0.96383 | 1.0664 | 172.6629 | 47.4203 | 54046.0617 | 53083.822 | 1057.6764 | 957.6764 | 1065.4076
6591 | 104.4385 | 0.97209 | 1.0255 | 0.99202 | 189.7314 | 44.1371 | 49572.0787 | 49336.4396 | 982.7288 | 882.7288 | 1065.4168
6592 | 90.5254 | 1.0686 | 1.0607 | 1.127 | 199.1919 | 51.7154 | 67024.7355 | 65661.0742 | 1309.2215 | 1209.2215 | 1065.4255
6593 | 111.9746 | 0.97096 | 1.0222 | 1.0158 | 196.3775 | 50.0074 | 57554.5907 | 54869.6352 | 1093.3927 | 993.3927 | 1065.565
6594 | 108.8616 | 1.0714 | 1.1136 | 0.97595 | 195.9459 | 45.2629 | 54666.445 | 52932.0347 | 1054.6407 | 954.6407 | 1065.6449
6595 | 104.4615 | 1.0344 | 1.1355 | 0.99292 | 209.2795 | 49.5339 | 61406.3982 | 61741.8924 | 1230.8378 | 1130.8378 | 1065.6922
6596 | 97.9368 | 0.95391 | 0.82076 | 0.94652 | 200.7724 | 48.3565 | 42284.5231 | 43036.3641 | 856.7273 | 756.7273 | 1065.7056
6597 | 102.995 | 0.99596 | 1.1369 | 1.0412 | 208.6403 | 49.1884 | 63480.9512 | 62367.5081 | 1243.3502 | 1143.3502 | 1065.7164
6598 | 87.5801 | 0.89036 | 0.96563 | 1.0087 | 212.9707 | 49.8879 | 52695.5942 | 52166.1533 | 1039.3231 | 939.3231 | 1065.7903
6599 | 96.2973 | 0.92094 | 0.9142 | 1.0018 | 200.4245 | 52.066 | 52504.0603 | 54614.3845 | 1088.2877 | 988.2877 | 1065.813
6600 | 91.097 | 1.0282 | 1.1308 | 0.9364 | 189.2942 | 51.6439 | 59711.0311 | 59169.3599 | 1179.3872 | 1079.3872 | 1065.8171
6601 | 113.3999 | 0.98983 | 0.92744 | 1.0323 | 198.7815 | 50.8398 | 54579.0399 | 53815.448 | 1072.309 | 972.309 | 1065.8387
6602 | 111.8124 | 0.9628 | 0.85196 | 1.0178 | 187.2526 | 54.2426 | 53060.0678 | 54079.5291 | 1077.5906 | 977.5906 | 1065.9061
6603 | 101.1188 | 0.96397 | 1.2515 | 1.004 | 195.0687 | 49.414 | 67104.7365 | 66090.0316 | 1317.8006 | 1217.8006 | 1065.9488
6604 | 93.2078 | 0.93155 | 1.0601 | 0.98692 | 206.6965 | 47.3143 | 53819.1796 | 51747.7287 | 1030.9546 | 930.9546 | 1066.0476
6605 | 110.7469 | 0.99734 | 0.77228 | 0.99706 | 193.5327 | 49.5446 | 43815.6248 | 42373.3757 | 843.4675 | 743.4675 | 1066.0495
6606 | 106.0098 | 1.0621 | 0.95105 | 1.0287 | 198.9638 | 49.8874 | 54622.7036 | 55039.4185 | 1096.7884 | 996.7884 | 1066.0598
6607 | 93.7724 | 1.0477 | 1.0317 | 0.98253 | 202.9173 | 52.5111 | 58589.3258 | 57448.9019 | 1144.978 | 1044.978 | 1066.1267
6608 | 99.1987 | 1.039 | 1.0941 | 0.95914 | 185.7971 | 50.3346 | 58196.3848 | 57582.9673 | 1147.6593 | 1047.6593 | 1066.145
6609 | 93.8975 | 1.0515 | 0.96575 | 0.97425 | 207.4285 | 51.2416 | 53479.02 | 53850.0159 | 1073.0003 | 973.0003 | 1066.1527
6610 | 106.1979 | 0.98767 | 0.90044 | 0.96957 | 184.4252 | 44.583 | 43783.5654 | 44531.5776 | 886.6316 | 786.6316 | 1066.2041
6611 | 88.3808 | 0.97331 | 0.9784 | 0.9974 | 211.1003 | 50.6249 | 53968.9059 | 55545.5429 | 1106.9109 | 1006.9109 | 1066.262
6612 | 80.2698 | 1.0179 | 1.0905 | 1.0374 | 204.9159 | 53.0256 | 64524.9759 | 64670.6562 | 1289.4131 | 1189.4131 | 1066.3169
6613 | 113.8729 | 1.038 | 1.1253 | 1.0761 | 198.7589 | 49.7141 | 66270.1263 | 68565.8828 | 1367.3177 | 1267.3177 | 1066.3415
6614 | 111.3451 | 1.0232 | 1.0812 | 1.0555 | 213.2034 | 52.339 | 65903.8857 | 68175.945 | 1359.5189 | 1259.5189 | 1066.3466
6615 | 99.6064 | 1.0374 | 1.2628 | 1.0221 | 202.2923 | 49.5791 | 69320.2036 | 67399.7013 | 1343.994 | 1243.994 | 1066.4209
6616 | 90.5697 | 0.96664 | 1.0821 | 0.90907 | 212.5521 | 49.7191 | 53475.6975 | 54080.2726 | 1077.6055 | 977.6055 | 1066.4344
6617 | 84.0332 | 0.99751 | 1.0175 | 0.99484 | 197.2257 | 53.4301 | 58758.8431 | 58490.8561 | 1165.8171 | 1065.8171 | 1066.5057
6618 | 89.1565 | 0.97764 | 1.2321 | 1.0944 | 201.0626 | 55.157 | 79380.0573 | 78835.1715 | 1572.7034 | 1472.7034 | 1066.5792
6619 | 114.8135 | 0.96939 | 0.9944 | 1.0146 | 196.0641 | 49.5406 | 55693.3889 | 56049.3632 | 1116.9873 | 1016.9873 | 1066.5817
6620 | 105.9402 | 1.0343 | 1.058 | 1.0267 | 217.6318 | 44.9997 | 54028.4968 | 53523.3242 | 1066.4665 | 966.4665 | 1066.6131
6621 | 119.2276 | 0.94648 | 1.0584 | 0.9537 | 196.3098 | 51.1466 | 57596.1478 | 57285.4707 | 1141.7094 | 1041.7094 | 1066.67
6622 | 107.073 | 1.011 | 1.2949 | 1.0316 | 189.7059 | 49.5232 | 71703.3827 | 69429.5298 | 1384.5906 | 1284.5906 | 1066.8364
6623 | 95.5399 | 1.01 | 1.0387 | 0.96411 | 216.2739 | 49.686 | 54769.0873 | 52673.3744 | 1049.4675 | 949.4675 | 1066.8713
6624 | 100.2574 | 0.97885 | 1.0535 | 1.0439 | 204.664 | 50.0808 | 60198.7507 | 57964.7282 | 1155.2946 | 1055.2946 | 1066.8922
6625 | 94.9128 | 1.0412 | 1.4232 | 1.0672 | 188.6059 | 48.549 | 78725.5463 | 79692.3667 | 1589.8473 | 1489.8473 | 1066.9007
6626 | 102.0625 | 1.0601 | 0.87293 | 0.9631 | 201.8796 | 49.661 | 47325.9239 | 45445.6953 | 904.9139 | 804.9139 | 1066.9334
6627 | 110.8067 | 0.96805 | 0.89137 | 0.89023 | 202.2762 | 50.6321 | 45811.362 | 45335.5475 | 902.711 | 802.711 | 1067.0027
6628 | 78.7901 | 0.97717 | 1.0877 | 1.0112 | 197.1252 | 53.5057 | 63165.2018 | 63115.3225 | 1258.3065 | 1158.3065 | 1067.0669
6629 | 98.9966 | 0.92796 | 0.97054 | 1.1355 | 202.3525 | 46.6115 | 55853.9149 | 56381.284 | 1123.6257 | 1023.6257 | 1067.1386
6630 | 94.757 | 1.0634 | 1.0029 | 0.99786 | 199.6429 | 50.9834 | 56361.2668 | 57024.815 | 1136.4963 | 1036.4963 | 1067.1769
6631 | 95.7728 | 1.0698 | 0.82906 | 1.0596 | 196.7029 | 48.9147 | 48179.2378 | 47529.321 | 946.5864 | 846.5864 | 1067.1825
6632 | 112.6652 | 1.0243 | 1.1987 | 0.91842 | 206.9994 | 49.1789 | 60023.9876 | 61695.0682 | 1229.9014 | 1129.9014 | 1067.2086
6633 | 104.5473 | 0.94809 | 0.76109 | 1.0734 | 175.4921 | 43.504 | 40029.8256 | 39593.5442 | 787.8709 | 687.8709 | 1067.2445
6634 | 110.0556 | 1.0172 | 1.0034 | 1.0427 | 203.8639 | 51.1721 | 59471.8842 | 59432.5853 | 1184.6517 | 1084.6517 | 1067.2793
6635 | 79.8969 | 1.0211 | 1.1745 | 1.0683 | 179.9203 | 50.3401 | 67446.1032 | 68736.2137 | 1370.7243 | 1270.7243 | 1067.2919
6636 | 106.7865 | 1.0186 | 0.85292 | 1.0428 | 196.9276 | 45.0524 | 45166.5366 | 44858.5639 | 893.1713 | 793.1713 | 1067.3067
6637 | 93.1564 | 0.99846 | 1.0931 | 0.88357 | 199.4979 | 55.2798 | 58733.6775 | 60053.2995 | 1197.066 | 1097.066 | 1067.3687
6638 | 85.6686 | 1.0111 | 0.83953 | 1.0107 | 203.4735 | 50.9661 | 47865.3348 | 46496.364 | 925.9273 | 825.9273 | 1067.4002
6639 | 94.5547 | 1.0243 | 1.1146 | 1.0224 | 186.3207 | 51.3924 | 63728.8616 | 67310.9711 | 1342.2194 | 1242.2194 | 1067.4117
6640 | 96.0051 | 1.0196 | 1.1341 | 0.96399 | 206.1321 | 47.8594 | 57214.4705 | 57475.6793 | 1145.5136 | 1045.5136 | 1067.4271
6641 | 89.1746 | 1.009 | 0.69346 | 1.0631 | 193.0218 | 46.6579 | 38787.973 | 39356.3748 | 783.1275 | 683.1275 | 1067.4307
6642 | 93.9587 | 0.9586 | 0.88031 | 0.96308 | 188.7156 | 44.0929 | 41542.4787 | 40887.2695 | 813.7454 | 713.7454 | 1067.4749
6643 | 105.6268 | 0.96146 | 1.2703 | 1.0064 | 207.7556 | 50.6912 | 70160.6439 | 72013.7764 | 1436.2755 | 1336.2755 | 1067.5618
6644 | 96.9763 | 0.94316 | 0.99312 | 0.96871 | 205.0614 | 49.5194 | 52374.2913 | 52270.453 | 1041.4091 | 941.4091 | 1067.6292
6645 | 102.2443 | 1.0374 | 0.84802 | 0.95512 | 192.7206 | 50.5312 | 46480.5341 | 45507.2732 | 906.1455 | 806.1455 | 1067.657
6646 | 101.1508 | 0.98988 | 1.1293 | 0.99379 | 185.4123 | 46.7839 | 57376.4864 | 57614.9287 | 1148.2986 | 1048.2986 | 1067.6989
6647 | 110.157 | 1.0242 | 0.87947 | 0.9582 | 204.597 | 47.3023 | 45403.7003 | 45076.1681 | 897.5234 | 797.5234 | 1067.9205
6648 | 112.9857 | 1.0275 | 0.94256 | 1.0144 | 210.4748 | 50.7432 | 54619.4487 | 53385.7727 | 1063.7155 | 963.7155 | 1068.0108
6649 | 105.9285 | 0.95661 | 1.2917 | 1.0406 | 203.9584 | 48.4427 | 70229.9842 | 69806.9041 | 1392.1381 | 1292.1381 | 1068.0561
6650 | 84.2119 | 0.94389 | 0.84557 | 0.95207 | 203.3267 | 50.6441 | 44999.7517 | 44495.4453 | 885.9089 | 785.9089 | 1068.108
6651 | 104.1787 | 0.87994 | 0.95753 | 1.0097 | 209.0941 | 48.8217 | 51885.98 | 51851.4721 | 1033.0294 | 933.0294 | 1068.1203
6652 | 114.2126 | 0.96708 | 0.96374 | 1.003 | 199.7722 | 52.3045 | 56537.7806 | 57187.6231 | 1139.7525 | 1039.7525 | 1068.2018
6653 | 103.0553 | 0.92239 | 1.2213 | 0.98916 | 194.4802 | 52.6186 | 68761.2216 | 69380.8407 | 1383.6168 | 1283.6168 | 1068.255
6654 | 87.5359 | 1.0434 | 0.81172 | 0.94366 | 200.9937 | 48.0684 | 41411.3459 | 41330.3991 | 822.608 | 722.608 | 1068.2722
6655 | 88.4275 | 0.97496 | 1.1374 | 0.9172 | 206.3123 | 49.8268 | 56484.1503 | 56531.5464 | 1126.6309 | 1026.6309 | 1068.4092
6656 | 101.6994 | 0.99551 | 0.9801 | 0.96919 | 193.9108 | 55.0884 | 58099.4535 | 57458.2853 | 1145.1657 | 1045.1657 | 1068.4517
6657 | 111.8992 | 0.9496 | 1.011 | 0.96805 | 210.8599 | 50.0595 | 54525.502 | 55027.2231 | 1096.5445 | 996.5445 | 1068.4647
6658 | 114.2258 | 0.95623 | 1.1618 | 1.0188 | 197.88 | 50.1578 | 65042.4723 | 65896.9765 | 1313.9395 | 1213.9395 | 1068.4693
6659 | 87.3682 | 1.0125 | 0.98111 | 1.0304 | 206.2594 | 41.8948 | 46263.9004 | 45114.7567 | 898.2951 | 798.2951 | 1068.5661
6660 | 103.1857 | 1.0025 | 1.2467 | 0.96572 | 209.161 | 52.4866 | 68830.7024 | 69975.1922 | 1395.5038 | 1295.5038 | 1068.5712
6661 | 104.0569 | 1.0342 | 1.182 | 0.98164 | 197.9402 | 50.764 | 64561.4205 | 64131.2411 | 1278.6248 | 1178.6248 | 1068.7955
6662 | 104.1471 | 0.99782 | 1.0323 | 1.017 | 209.9347 | 44.0833 | 51072.513 | 52591.4391 | 1047.8288 | 947.8288 | 1068.8056
6663 | 100.4262 | 1.0709 | 1.0873 | 1.0631 | 203.2432 | 46.8446 | 59389.2356 | 61105.6026 | 1218.1121 | 1118.1121 | 1068.8128
6664 | 94.8634 | 1.0356 | 1.2598 | 0.89189 | 191.9415 | 48.3269 | 59237.3754 | 58679.3732 | 1169.5875 | 1069.5875 | 1068.8147
6665 | 83.0881 | 0.98498 | 1.0766 | 1.0269 | 210.8 | 48.0452 | 57257.7576 | 57392.2451 | 1143.8449 | 1043.8449 | 1068.8384
6666 | 95.4701 | 1.0236 | 1.2825 | 1.0006 | 194.6534 | 50.3136 | 69675.5291 | 69216.045 | 1380.3209 | 1280.3209 | 1068.8649
6667 | 121.7492 | 1.0658 | 0.97392 | 1.0401 | 201.7957 | 48.2985 | 55392.5642 | 56742.8456 | 1130.8569 | 1030.8569 | 1068.8848
6668 | 91.0775 | 0.93644 | 0.99518 | 0.97782 | 221.9674 | 50.1128 | 53261.1383 | 54611.0656 | 1088.2213 | 988.2213 | 1068.9
6669 | 102.0044 | 0.96595 | 1.272 | 0.98153 | 199.022 | 50.5615 | 68308.9762 | 68397.7823 | 1363.9556 | 1263.9556 | 1069.0518
6670 | 106.7759 | 0.97667 | 0.88093 | 0.93212 | 200.4714 | 47.1111 | 43797.7576 | 42800.491 | 852.0098 | 752.0098 | 1069.0804
6671 | 106.1816 | 1.0342 | 1.1026 | 0.97626 | 189.9288 | 50.3409 | 59905.1275 | 61123.2681 | 1218.4654 | 1118.4654 | 1069.2239
6672 | 100.653 | 0.98276 | 1.0461 | 0.91041 | 193.3734 | 50.0389 | 52798.0478 | 54008.0993 | 1076.162 | 976.162 | 1069.3354
6673 | 100.4765 | 1.051 | 1.0651 | 1.001 | 199.9941 | 51.76 | 60854.2421 | 60960.777 | 1215.2155 | 1115.2155 | 1069.3474
6674 | 83.599 | 0.99263 | 0.89265 | 0.93826 | 215.6899 | 47.0977 | 43570.2944 | 43167.2115 | 859.3442 | 759.3442 | 1069.4274
6675 | 108.3674 | 1.0225 | 1.0838 | 1.0711 | 198.3719 | 52.0583 | 66398.7393 | 68668.0062 | 1369.3601 | 1269.3601 | 1069.4697
6676 | 101.0048 | 0.9232 | 0.865 | 1.0419 | 193.794 | 46.0081 | 45948.8173 | 45038.7492 | 896.775 | 796.775 | 1069.4984
6677 | 93.6559 | 0.9998 | 1.1526 | 1.0023 | 221.5094 | 52.7996 | 66162.7018 | 67343.4309 | 1342.8686 | 1242.8686 | 1069.501
6678 | 87.8419 | 0.98475 | 1.0761 | 1.0849 | 211.5683 | 50.3664 | 63373.3348 | 63618.3861 | 1268.3677 | 1168.3677 | 1069.544
6679 | 107.377 | 0.97503 | 1.1482 | 0.95932 | 197.9674 | 47.3669 | 57329.4565 | 58249.3098 | 1160.9862 | 1060.9862 | 1069.5875
6680 | 102.3977 | 1.0074 | 0.9698 | 1.0995 | 202.8107 | 53.995 | 63347.0484 | 63164.805 | 1259.2961 | 1159.2961 | 1069.6132
6681 | 101.999 | 1.0214 | 1.0334 | 1.0638 | 201.2842 | 52.7903 | 63732.2661 | 65512.8291 | 1306.2566 | 1206.2566 | 1069.6181
6682 | 108.7354 | 1.0241 | 0.76606 | 0.98978 | 193.04 | 46.5697 | 40689.232 | 40248.7264 | 800.9745 | 700.9745 | 1069.6284
6683 | 107.7097 | 0.93885 | 1.0983 | 0.87483 | 187.2947 | 52.6253 | 56071.9236 | 55544.0316 | 1106.8806 | 1006.8806 | 1069.662
6684 | 94.8027 | 1.0534 | 1.1201 | 0.97758 | 197.1862 | 53.2018 | 63763.6991 | 61829.4287 | 1232.5886 | 1132.5886 | 1069.6884
6685 | 100.6464 | 0.99987 | 0.94793 | 0.99452 | 213.0863 | 48.6923 | 51016.7489 | 51841.8623 | 1032.8372 | 932.8372 | 1069.7838
6686 | 93.5815 | 1 | 0.64756 | 0.95315 | 212.7667 | 49.1346 | 35137.7966 | 33955.5503 | 675.111 | 575.111 | 1069.7861
6687 | 104.9851 | 1.0124 | 1.0793 | 0.99654 | 189.7295 | 49.0154 | 58119.2125 | 57218.7995 | 1140.376 | 1040.376 | 1069.8183
6688 | 93.0739 | 1.0268 | 1.0481 | 0.97119 | 200.7577 | 51.3325 | 57358.701 | 59297.4774 | 1181.9495 | 1081.9495 | 1069.8361
6689 | 90.6912 | 0.94102 | 0.88122 | 0.98128 | 201.2233 | 51.4886 | 49118.8481 | 50323.588 | 1002.4718 | 902.4718 | 1069.911
6690 | 86.3731 | 0.96728 | 0.95616 | 1.0392 | 190.0782 | 49.7816 | 53813.488 | 52021.3815 | 1036.4276 | 936.4276 | 1070.0168
6691 | 79.0114 | 0.98994 | 0.761 | 0.97801 | 188.8732 | 50.2997 | 41559.5546 | 41249.1358 | 820.9827 | 720.9827 | 1070.0684
6692 | 97.405 | 0.99296 | 0.67605 | 0.97626 | 195.3133 | 54.4276 | 41381.892 | 41915.1888 | 834.3038 | 734.3038 | 1070.0806
6693 | 120.7728 | 0.96276 | 1.0722 | 1.0029 | 202.2264 | 43.7059 | 52281.3755 | 51192.3425 | 1019.8469 | 919.8469 | 1070.1798
6694 | 108.0955 | 0.89954 | 0.87557 | 0.93297 | 190.5722 | 49.6096 | 45539.7558 | 45704.6871 | 910.0937 | 810.0937 | 1070.2664
6695 | 95.6819 | 0.99185 | 1.0416 | 1.0118 | 190.2569 | 50.7711 | 58518.5599 | 58290.8766 | 1161.8175 | 1061.8175 | 1070.308
6696 | 92.3747 | 1.0686 | 0.96512 | 0.96219 | 213.1116 | 48.9627 | 50514.3882 | 50792.4049 | 1011.8481 | 911.8481 | 1070.3149
6697 | 100.0293 | 0.98752 | 0.88696 | 1.0048 | 193.0625 | 48.446 | 48154.4299 | 48643.9569 | 968.8791 | 868.8791 | 1070.318
6698 | 92.382 | 0.94798 | 0.90101 | 0.95292 | 207.8707 | 47.6307 | 45274.4633 | 45187.0027 | 899.7401 | 799.7401 | 1070.3291
6699 | 104.274 | 1.0626 | 1.0519 | 0.97069 | 202.7077 | 53.0142 | 60205.8162 | 57086.7378 | 1137.7348 | 1037.7348 | 1070.3314
6700 | 87.4489 | 1.0164 | 0.9892 | 1.0115 | 201.0361 | 55.8281 | 61020.9324 | 61579.3874 | 1227.5877 | 1127.5877 | 1070.3388
6701 | 100.6895 | 1.0647 | 1.1816 | 1.0121 | 200.5997 | 51.431 | 67221.663 | 68852.1279 | 1373.0426 | 1273.0426 | 1070.3389
6702 | 82.4221 | 1.0234 | 1.38 | 0.98213 | 173.3335 | 51.549 | 74389.8385 | 74642.3442 | 1488.8469 | 1388.8469 | 1070.3447
6703 | 96.5009 | 1.1141 | 0.98896 | 1.0479 | 208.0698 | 48.338 | 55500.9912 | 56016.078 | 1116.3216 | 1016.3216 | 1070.3628
6704 | 107.5061 | 0.92239 | 1.1161 | 0.93233 | 207.955 | 47.11 | 53899.053 | 53367.3632 | 1063.3473 | 963.3473 | 1070.4741
6705 | 102.7178 | 0.95683 | 1.0429 | 1.0909 | 188.4184 | 48.5881 | 60241.5145 | 60618.1595 | 1208.3632 | 1108.3632 | 1070.481
6706 | 90.0229 | 1.0177 | 0.98893 | 1.0195 | 202.3564 | 47.5041 | 52447.133 | 51064.3637 | 1017.2873 | 917.2873 | 1070.518
6707 | 112.0331 | 1.0009 | 1.0751 | 0.99741 | 188.7181 | 47.3954 | 56324.2348 | 57523.736 | 1146.4747 | 1046.4747 | 1070.5697
6708 | 93.9406 | 0.98967 | 1.0145 | 0.95753 | 191.1111 | 51.499 | 55006.1486 | 55800.3279 | 1112.0066 | 1012.0066 | 1070.5767
6709 | 105.5121 | 1.099 | 0.91903 | 0.96312 | 205.8361 | 48.815 | 49073.7066 | 48108.2568 | 958.1651 | 858.1651 | 1070.601
6710 | 90.3875 | 1.0689 | 1.0041 | 0.9887 | 173.9541 | 55.5285 | 60664.2601 | 60908.3766 | 1214.1675 | 1114.1675 | 1070.6997
6711 | 75.9849 | 1.045 | 0.93801 | 1.1285 | 194.6956 | 48.352 | 55217.2255 | 54165.2116 | 1079.3042 | 979.3042 | 1070.7077
6712 | 102.3626 | 0.9415 | 0.78902 | 0.93477 | 187.3264 | 48.973 | 41027.1458 | 40926.3169 | 814.5263 | 714.5263 | 1070.8537
6713 | 92.2783 | 1.0447 | 0.93281 | 0.92009 | 200.5056 | 53.0794 | 50874.055 | 51163.6665 | 1019.2733 | 919.2733 | 1070.8547
6714 | 101.3614 | 0.98812 | 0.85131 | 0.99713 | 198.934 | 51.0991 | 48693.3224 | 48069.7437 | 957.3949 | 857.3949 | 1070.9072
6715 | 93.3642 | 0.99053 | 0.8608 | 0.97832 | 220.3545 | 50.3464 | 47274.754 | 48935.7631 | 974.7153 | 874.7153 | 1071.0332
6716 | 104.8651 | 1.0143 | 0.86554 | 0.9503 | 210.5347 | 47.5559 | 44384.4838 | 44497.5114 | 885.9502 | 785.9502 | 1071.0535
6717 | 94.1214 | 1.0431 | 1.2796 | 0.99361 | 197.8177 | 49.1772 | 67550.6435 | 66546.353 | 1326.9271 | 1226.9271 | 1071.0964
6718 | 102.2361 | 0.97449 | 0.63585 | 0.91365 | 214.3722 | 50.1918 | 34373.7609 | 34567.7919 | 687.3558 | 587.3558 | 1071.1029
6719 | 88.9527 | 0.96131 | 0.95745 | 1.0478 | 191.6229 | 49.1399 | 53692.7084 | 54937.498 | 1094.75 | 994.75 | 1071.1165
6720 | 68.9887 | 0.99506 | 1.0736 | 0.99769 | 198.6245 | 49.8247 | 56988.6208 | 56736.3756 | 1130.7275 | 1030.7275 | 1071.203
6721 | 89.1446 | 1.0252 | 0.78009 | 0.9324 | 202.7216 | 55.6202 | 45741.5011 | 45604.0426 | 908.0809 | 808.0809 | 1071.2393
6722 | 78.0432 | 1.0138 | 1.0318 | 1.0456 | 205.6417 | 50.6091 | 58805.6831 | 58902.38 | 1174.0476 | 1074.0476 | 1071.2482
6723 | 114.4173 | 1.0184 | 1.1474 | 1.0566 | 199.4378 | 51.8133 | 69048.6063 | 69028.0204 | 1376.5604 | 1276.5604 | 1071.3229
6724 | 90.5874 | 0.97493 | 1.0101 | 1.0741 | 208.464 | 52.6629 | 61997.7538 | 63812.1488 | 1272.243 | 1172.243 | 1071.3763
6725 | 99.415 | 1.0563 | 1.08 | 1.1312 | 194.6732 | 50.8427 | 67648.5475 | 68587.952 | 1367.759 | 1267.759 | 1071.3957
6726 | 111.3548 | 0.95831 | 1.1274 | 0.99819 | 200.8426 | 49.6192 | 61334.0407 | 61573.5303 | 1227.4706 | 1127.4706 | 1071.4457
6727 | 97.5609 | 1.0168 | 1.0927 | 1.0062 | 208.3797 | 52.8222 | 63524.408 | 62891.3935 | 1253.8279 | 1153.8279 | 1071.4659
6728 | 104.2868 | 0.99671 | 0.95581 | 0.89854 | 177.9374 | 50.5959 | 48890.5612 | 49195.0892 | 979.9018 | 879.9018 | 1071.4714
6729 | 89.4145 | 1.0838 | 0.96163 | 1.0265 | 198.9805 | 50.8685 | 55343.3585 | 54423.1682 | 1084.4634 | 984.4634 | 1071.4988
6730 | 93.6621 | 1.0656 | 1.038 | 0.95361 | 179.0638 | 51.0913 | 55852.8398 | 57531.701 | 1146.634 | 1046.634 | 1071.5447
6731 | 99.7969 | 1.0594 | 0.89674 | 0.9705 | 199.5738 | 50.2373 | 49231.8242 | 50631.952 | 1008.639 | 908.639 | 1071.5637
6732 | 102.2722 | 1.0011 | 0.90435 | 1.0565 | 193.6051 | 52.1447 | 55352.0061 | 55512.7646 | 1106.2553 | 1006.2553 | 1071.5929
6733 | 102.1782 | 0.98547 | 1.0671 | 0.9988 | 209.1833 | 47.6312 | 55773.4488 | 55345.1786 | 1102.9036 | 1002.9036 | 1071.6291
6734 | 103.7719 | 1.0006 | 0.90204 | 0.82225 | 189.4234 | 48.7527 | 41411.1003 | 41605.3753 | 828.1075 | 728.1075 | 1071.6438
6735 | 102.4653 | 1.0038 | 1.2205 | 1.0693 | 217.9737 | 50.5622 | 71403.4538 | 71564.4372 | 1427.2887 | 1327.2887 | 1071.6788
6736 | 88.7389 | 0.912 | 0.96608 | 0.97879 | 213.2242 | 51.2383 | 52809.8708 | 52358.8952 | 1043.1779 | 943.1779 | 1071.7012
6737 | 100.5188 | 0.94294 | 0.97479 | 1.0008 | 206.3803 | 51.3565 | 55177.3929 | 54034.2456 | 1076.6849 | 976.6849 | 1071.7522
6738 | 114.6459 | 1.0671 | 0.98361 | 0.92431 | 201.4836 | 51.8645 | 53699.5647 | 55069.8856 | 1097.3977 | 997.3977 | 1071.7538
6739 | 114.5604 | 0.94731 | 0.90931 | 0.96401 | 197.2633 | 46.7612 | 46262.1491 | 46236.4608 | 920.7292 | 820.7292 | 1071.7588
6740 | 101.0032 | 0.99115 | 1.1479 | 0.96955 | 196.2098 | 51.2287 | 62339.8254 | 62430.8328 | 1244.6167 | 1144.6167 | 1071.7718
6741 | 95.5641 | 0.91873 | 0.78446 | 0.97147 | 192.8692 | 52.0595 | 44437.2803 | 44315.6098 | 882.3122 | 782.3122 | 1071.7799
6742 | 100.2376 | 1.0518 | 0.94111 | 1.0212 | 203.4715 | 53.6745 | 57446.4676 | 56950.9119 | 1135.0182 | 1035.0182 | 1071.7877
6743 | 93.1824 | 1.0273 | 1.2494 | 0.93032 | 201.2677 | 47.2491 | 59643.5942 | 59536.1438 | 1186.7229 | 1086.7229 | 1071.7947
6744 | 79.8583 | 0.98025 | 0.96783 | 1.0241 | 191.7595 | 52.5176 | 56353.3362 | 58300.1637 | 1162.0033 | 1062.0033 | 1071.8537
6745 | 118.0056 | 1.0115 | 1.1136 | 1.0264 | 205.6385 | 41.4287 | 52505.083 | 53292.6593 | 1061.8532 | 961.8532 | 1071.9728
6746 | 101.9436 | 0.92598 | 0.84841 | 1.0194 | 203.9973 | 50.9763 | 49103.5939 | 48897.6601 | 973.9532 | 873.9532 | 1072.0036
6747 | 104.1857 | 1.0756 | 1.0066 | 0.92517 | 183.7507 | 46.6193 | 48825.0412 | 47953.6405 | 955.0728 | 855.0728 | 1072.0233
6748 | 102.096 | 0.90953 | 0.92964 | 1.143 | 204.1608 | 51.4711 | 59676.6993 | 61240.5853 | 1220.8117 | 1120.8117 | 1072.0858
6749 | 102.255 | 0.99513 | 1.016 | 1.0307 | 192.6616 | 48.1307 | 55491.4553 | 54587.3861 | 1087.7477 | 987.7477 | 1072.0955
6750 | 97.8852 | 0.9567 | 0.66295 | 0.98258 | 193.9404 | 49.4209 | 37014.8026 | 36807.9351 | 732.1587 | 632.1587 | 1072.1333
6751 | 100.0506 | 1.038 | 0.83739 | 0.97491 | 204.3958 | 48.0088 | 44383.8871 | 44177.6664 | 879.5533 | 779.5533 | 1072.1454
6752 | 105.5079 | 0.97606 | 0.90662 | 0.93977 | 191.5948 | 49.0834 | 47065.9624 | 46340.5927 | 922.8119 | 822.8119 | 1072.2566
6753 | 104.484 | 1.1251 | 1.0715 | 0.90638 | 199.2705 | 54.1547 | 59161.0246 | 59711.6421 | 1190.2328 | 1090.2328 | 1072.3382
6754 | 108.1554 | 1.0044 | 0.98933 | 1.0146 | 197.5501 | 52.2372 | 58304.351 | 59241.1425 | 1180.8228 | 1080.8228 | 1072.3871
6755 | 100.7379 | 1.0776 | 1.0801 | 0.96118 | 200.7807 | 47.086 | 54195.2964 | 53983.9926 | 1075.6799 | 975.6799 | 1072.3969
6756 | 105.7278 | 1.0108 | 1.0832 | 1.0039 | 204.6173 | 51.7148 | 61969.5537 | 63693.398 | 1269.868 | 1169.868 | 1072.4063
6757 | 89.5372 | 0.95674 | 1.117 | 0.99133 | 183.7995 | 51.9474 | 62154.5463 | 62720.3431 | 1250.4069 | 1150.4069 | 1072.4337
6758 | 87.9095 | 1.0081 | 0.99718 | 0.99597 | 195.5128 | 47.8127 | 51918.7671 | 52036.4509 | 1036.729 | 936.729 | 1072.4358
6759 | 97.8007 | 1.0257 | 1.1986 | 1.0425 | 214.686 | 52.9507 | 71688.9601 | 71151.8345 | 1419.0367 | 1319.0367 | 1072.4718
6760 | 100.2051 | 1.0317 | 1.0034 | 0.88922 | 189.3312 | 45.622 | 45612.6983 | 46899.7602 | 933.9952 | 833.9952 | 1072.5258
6761 | 78.9405 | 0.98398 | 0.86266 | 0.9627 | 195.8323 | 46.663 | 42573.1766 | 41714.0626 | 830.2813 | 730.2813 | 1072.5869
6762 | 96.1971 | 0.94173 | 1.1052 | 1.0924 | 196.4936 | 53.1317 | 69156.0344 | 69087.3958 | 1377.7479 | 1277.7479 | 1072.5928
6763 | 94.5077 | 1.0145 | 0.98585 | 0.94493 | 198.7647 | 52.1748 | 53805.1119 | 52991.7713 | 1055.8354 | 955.8354 | 1072.6385
6764 | 104.6286 | 1.001 | 0.98196 | 0.99593 | 187.9916 | 47.6277 | 51754.4802 | 51305.0456 | 1022.1009 | 922.1009 | 1072.6831
6765 | 89.9185 | 0.89259 | 0.99373 | 1.0068 | 206.4518 | 52.0898 | 56502.3023 | 55656.9081 | 1109.1382 | 1009.1382 | 1072.743
6766 | 92.302 | 1.0199 | 0.96468 | 0.98212 | 185.3812 | 53.9338 | 56361.0748 | 55723.8346 | 1110.4767 | 1010.4767 | 1072.7551
6767 | 95.4873 | 1.0051 | 0.99548 | 1.0008 | 205.7343 | 55.3699 | 60684.7112 | 60040.1906 | 1196.8038 | 1096.8038 | 1072.8274
6768 | 89.3928 | 0.9646 | 1.2143 | 1.0069 | 198.6747 | 52.1998 | 68520.4093 | 69488.4508 | 1385.769 | 1285.769 | 1072.8647
6769 | 105.87 | 0.93714 | 0.96966 | 1.0168 | 201.7156 | 52.2091 | 56856.9877 | 56535.4851 | 1126.7097 | 1026.7097 | 1072.9215
6770 | 92.4965 | 1.0144 | 0.84395 | 0.92325 | 220.5976 | 47.8802 | 42020.4174 | 41529.497 | 826.5899 | 726.5899 | 1073.0224
6771 | 76.8572 | 1.0564 | 1.4559 | 1.0744 | 204.9601 | 49.4978 | 81645.2672 | 82512.7565 | 1646.2551 | 1546.2551 | 1073.0443
6772 | 84.1756 | 0.98727 | 1.2394 | 1.0381 | 194.9629 | 50.1648 | 68903.1178 | 69747.6283 | 1390.9526 | 1290.9526 | 1073.1444
6773 | 88.7942 | 1.0167 | 0.82716 | 0.9173 | 222.7875 | 51.8515 | 44246.3331 | 44646.0297 | 888.9206 | 788.9206 | 1073.1737
6774 | 103.9961 | 1.0109 | 1.0482 | 1.0499 | 198.5593 | 48.8326 | 59070.9615 | 60274.9485 | 1201.499 | 1101.499 | 1073.186
6775 | 102.5057 | 1.0544 | 1.0872 | 0.97321 | 207.8368 | 48.9742 | 57317.6833 | 56190.371 | 1119.8074 | 1019.8074 | 1073.3506
6776 | 109.7004 | 1.0123 | 0.69461 | 0.95354 | 202.535 | 51.3904 | 39947.2043 | 40424.9113 | 804.4982 | 704.4982 | 1073.3858
6777 | 100.1166 | 0.9965 | 1.0413 | 0.98212 | 201.2053 | 50.4685 | 56851.3422 | 56990.5135 | 1135.8103 | 1035.8103 | 1073.4168
6778 | 114.5094 | 1.0045 | 0.86413 | 1.0167 | 199.0423 | 45.9758 | 45880.2499 | 46550.2315 | 927.0046 | 827.0046 | 1073.4762
6779 | 112.6441 | 0.9878 | 0.89208 | 0.97387 | 205.8623 | 47.9798 | 47227.5781 | 47717.0906 | 950.3418 | 850.3418 | 1073.4777
6780 | 100.8018 | 0.97329 | 1.0631 | 1.0355 | 182.7319 | 50.9228 | 61232.4671 | 60147.0585 | 1198.9412 | 1098.9412 | 1073.5061
6781 | 102.8653 | 0.91592 | 1.0579 | 0.98801 | 198.4646 | 48.6673 | 55649.4498 | 55416.7662 | 1104.3353 | 1004.3353 | 1073.5233
6782 | 101.2389 | 0.99377 | 1.0062 | 1.0032 | 194.5124 | 53.4957 | 59578.261 | 60811.7342 | 1212.2347 | 1112.2347 | 1073.5269
6783 | 91.0404 | 1.0407 | 1.1467 | 0.95358 | 198.6657 | 50.8419 | 60609.9565 | 61848.0439 | 1232.9609 | 1132.9609 | 1073.6418
6784 | 94.7129 | 1.0319 | 1.0331 | 0.9811 | 202.4481 | 43.7324 | 48800.8231 | 49112.6597 | 978.2532 | 878.2532 | 1073.6456
6785 | 86.1945 | 0.99028 | 0.89819 | 0.98074 | 204.5983 | 49.311 | 47851.238 | 47748.9223 | 950.9784 | 850.9784 | 1073.6681
6786 | 120.0578 | 0.97177 | 1.0294 | 0.97149 | 200.7624 | 52.4213 | 58740.6343 | 59333.2671 | 1182.6653 | 1082.6653 | 1073.7279
6787 | 95.8009 | 0.97516 | 1.3278 | 1.0026 | 186.1763 | 46.7448 | 66782.8511 | 65148.5585 | 1298.9712 | 1198.9712 | 1073.7325
6788 | 98.898 | 1.0852 | 0.92325 | 0.9566 | 195.332 | 54.5061 | 54183.9049 | 55179.985 | 1099.5997 | 999.5997 | 1073.9208
6789 | 103.4792 | 1.0511 | 1.1629 | 0.98124 | 189.725 | 52.6088 | 65943.6182 | 64785.9566 | 1291.7191 | 1191.7191 | 1073.9296
6790 | 110.1605 | 0.9536 | 1.1512 | 1.102 | 172.2352 | 45.3105 | 62415.8009 | 63100.4937 | 1258.0099 | 1158.0099 | 1073.9878
6791 | 93.4331 | 0.96698 | 0.83324 | 0.99179 | 210.6566 | 47.7922 | 44023.7476 | 44831.9708 | 892.6394 | 792.6394 | 1073.9914
6792 | 87.3531 | 0.90601 | 1.0104 | 1.0145 | 192.3783 | 51.2729 | 56807.4779 | 55164.4422 | 1099.2888 | 999.2888 | 1074.0476
6793 | 100.0097 | 0.9911 | 0.98354 | 1.0243 | 201.9691 | 46.4798 | 51633.36 | 52215.2519 | 1040.305 | 940.305 | 1074.1106
6794 | 108.2022 | 1.1466 | 1.1607 | 1.056 | 208.5602 | 50.6793 | 68612.8774 | 70694.0752 | 1409.8815 | 1309.8815 | 1074.1343
6795 | 114.0054 | 1.0359 | 1.0403 | 1.0151 | 217.3461 | 48.6006 | 57275.779 | 55890.7781 | 1113.8156 | 1013.8156 | 1074.1436
6796 | 101.3923 | 1.0095 | 0.76714 | 0.9931 | 175.3158 | 47.926 | 41592.5706 | 41187.4333 | 819.7487 | 719.7487 | 1074.2867
6797 | 110.413 | 0.96607 | 0.98923 | 0.99179 | 192.2476 | 52.5191 | 57321.5631 | 56849.0992 | 1132.982 | 1032.982 | 1074.2969
6798 | 82.7734 | 1.0109 | 0.71863 | 0.97858 | 195.2682 | 48.6033 | 38441.4785 | 38029.0269 | 756.5805 | 656.5805 | 1074.3146
6799 | 114.4309 | 1.033 | 1.0978 | 1.0387 | 197.41 | 48.2778 | 60952.543 | 59663.447 | 1189.2689 | 1089.2689 | 1074.318
6800 | 120.9482 | 0.99407 | 0.91595 | 1.0112 | 225.7992 | 52.3689 | 55025.7032 | 55103.9002 | 1098.078 | 998.078 | 1074.3519
6801 | 102.4989 | 1.0414 | 0.8738 | 1.0506 | 200.3991 | 51.7135 | 53193.4083 | 51276.3334 | 1021.5267 | 921.5267 | 1074.3665
6802 | 107.1441 | 1.0376 | 0.93207 | 0.9246 | 187.024 | 47.6076 | 46507.4638 | 46514.9362 | 926.2987 | 826.2987 | 1074.3723
6803 | 99.7693 | 1.0246 | 0.98409 | 0.92968 | 205.0487 | 49.7383 | 50794.0304 | 50563.0974 | 1007.2619 | 907.2619 | 1074.3879
6804 | 112.6099 | 0.95813 | 1.096 | 1.0047 | 198.1676 | 49.664 | 60245.9067 | 59026.8023 | 1176.536 | 1076.536 | 1074.4478
6805 | 80.4189 | 0.9624 | 0.94983 | 1.0278 | 177.4773 | 52.0586 | 55027.5977 | 54436.0963 | 1084.7219 | 984.7219 | 1074.4899
6806 | 104.9074 | 0.99461 | 0.92111 | 0.91885 | 192.2059 | 50.3682 | 48077.3716 | 48522.8228 | 966.4565 | 866.4565 | 1074.5108
6807 | 98.3997 | 1.0284 | 0.93076 | 1.0862 | 198.0424 | 50.8418 | 56745.2729 | 58510.2066 | 1166.2041 | 1066.2041 | 1074.5597
6808 | 105.0531 | 0.94201 | 0.89034 | 1.017 | 209.1696 | 58.1485 | 58615.2612 | 59139.5829 | 1178.7917 | 1078.7917 | 1074.5947
6809 | 107.921 | 1.0375 | 1.1534 | 1.0194 | 207.1702 | 50.9878 | 65863.8219 | 65202.8571 | 1300.0571 | 1200.0571 | 1074.6207
6810 | 106.2954 | 0.97663 | 0.84104 | 0.99607 | 201.5764 | 50.3339 | 47592.9439 | 45973.8119 | 915.4762 | 815.4762 | 1074.7614
6811 | 91.6464 | 1.0883 | 1.0198 | 1.0562 | 190.7774 | 46.4973 | 54907.5704 | 56650.9529 | 1129.0191 | 1029.0191 | 1074.8452
6812 | 92.0702 | 0.94122 | 1.0699 | 1.0103 | 196.2489 | 47.2205 | 55328.3419 | 54479.4011 | 1085.588 | 985.588 | 1074.8497
6813 | 97.0278 | 0.99539 | 1.1876 | 1.0419 | 182.2278 | 48.4291 | 64785.779 | 64715.7714 | 1290.3154 | 1190.3154 | 1074.8661
6814 | 85.0318 | 0.91498 | 0.92024 | 0.99873 | 213.3262 | 48.6603 | 48721.2493 | 47143.8568 | 938.8771 | 838.8771 | 1074.9012
6815 | 119.1882 | 0.89893 | 1.0987 | 1.0614 | 198.0929 | 46.5289 | 59439.3172 | 60330.9873 | 1202.6197 | 1102.6197 | 1074.9929
6816 | 109.8745 | 1.0172 | 0.8844 | 1.0471 | 198.1802 | 51.8674 | 54025.2081 | 54647.2535 | 1088.9451 | 988.9451 | 1075.028
6817 | 94.8964 | 0.89822 | 0.88444 | 0.99057 | 209.1832 | 53.2547 | 51405.1025 | 51495.5118 | 1025.9102 | 925.9102 | 1075.1901
6818 | 96.8272 | 0.93622 | 1.1064 | 1.0149 | 203.427 | 52.7661 | 64233.2873 | 65574.3164 | 1307.4863 | 1207.4863 | 1075.1941
6819 | 93.7107 | 0.99357 | 1.1286 | 1.0741 | 206.9138 | 55.7744 | 73009.6991 | 74164.6797 | 1479.2936 | 1379.2936 | 1075.2325
6820 | 101.5861 | 1.1062 | 1.0861 | 1.0825 | 201.8315 | 50.4032 | 65123.8611 | 63757.9221 | 1271.1584 | 1171.1584 | 1075.2826
6821 | 107.2068 | 1.0878 | 1.0905 | 0.96114 | 195.2514 | 51.0482 | 59653.2051 | 58485.2784 | 1165.7056 | 1065.7056 | 1075.3058
6822 | 99.7515 | 0.95656 | 1.1387 | 0.96767 | 196.3792 | 52.0998 | 62577.1099 | 61347.1551 | 1222.9431 | 1122.9431 | 1075.3189
6823 | 74.3061 | 0.99388 | 0.88769 | 1.0171 | 199.6176 | 48.0692 | 47151.3066 | 46978.7919 | 935.5758 | 835.5758 | 1075.3395
6824 | 89.3072 | 0.91134 | 1.3049 | 1.0075 | 199.3378 | 52.2688 | 73174.0081 | 73431.2672 | 1464.6253 | 1364.6253 | 1075.3683
6825 | 103.6596 | 1.0555 | 1.1246 | 1.0111 | 190.8992 | 46.7394 | 58452.098 | 57009.851 | 1136.197 | 1036.197 | 1075.3793
6826 | 91.0151 | 1.0247 | 0.87377 | 1.0632 | 211.6 | 50.6032 | 51940.8678 | 50933.1145 | 1014.6623 | 914.6623 | 1075.3988
6827 | 110.8261 | 1.0957 | 0.99803 | 1.1384 | 192.929 | 50.0223 | 63100.5334 | 63931.4983 | 1274.63 | 1174.63 | 1075.4963
6828 | 88.1015 | 1.0186 | 1.0084 | 0.9567 | 191.5055 | 48.6696 | 51511.8288 | 52202.409 | 1040.0482 | 940.0482 | 1075.6441
6829 | 107.1925 | 0.97567 | 0.89977 | 0.94701 | 216.1962 | 49.2313 | 47314.3695 | 47886.0367 | 953.7207 | 853.7207 | 1075.6853
6830 | 113.1989 | 0.95603 | 0.92053 | 0.95885 | 191.6409 | 55.0094 | 54698.8668 | 53626.9246 | 1068.5385 | 968.5385 | 1075.7249
6831 | 96.994 | 0.91651 | 0.85985 | 1.0538 | 202.9648 | 49.5643 | 49520.1259 | 49541.7918 | 986.8358 | 886.8358 | 1075.7403
6832 | 98.394 | 0.92215 | 1.0687 | 0.97214 | 198.2842 | 52.1073 | 59059.771 | 59401.6806 | 1184.0336 | 1084.0336 | 1075.7657
6833 | 102.772 | 1.0186 | 1.2245 | 0.9891 | 212.706 | 51.3217 | 67745.8013 | 69502.84 | 1386.0568 | 1286.0568 | 1075.8472
6834 | 97.2864 | 0.90367 | 0.9741 | 1.0057 | 192.6965 | 48.488 | 51955.2732 | 50470.4593 | 1005.4092 | 905.4092 | 1075.9576
6835 | 92.1576 | 0.97719 | 0.84619 | 0.94135 | 199.1889 | 50.3788 | 44865.7814 | 44596.6071 | 887.9321 | 787.9321 | 1075.9919
6836 | 92.4063 | 0.97554 | 0.82957 | 1.0444 | 206.0361 | 51.9343 | 49883.3658 | 49626.0037 | 988.5201 | 888.5201 | 1076.0104
6837 | 107.0559 | 1.0855 | 1.2589 | 0.94189 | 202.6088 | 46.965 | 61351.0857 | 61646.1872 | 1228.9237 | 1128.9237 | 1076.1746
6838 | 101.4852 | 0.94911 | 0.75934 | 0.97301 | 209.6385 | 47.7952 | 40126.3001 | 39081.6527 | 777.6331 | 677.6331 | 1076.1875
6839 | 108.4757 | 1.0945 | 0.93522 | 0.931 | 193.2716 | 51.9936 | 51636.6777 | 51069.5601 | 1017.3912 | 917.3912 | 1076.3189
6840 | 104.1632 | 0.98875 | 0.96312 | 1.0553 | 202.4606 | 51.4616 | 57805.2354 | 57858.8953 | 1153.1779 | 1053.1779 | 1076.3309
6841 | 101.596 | 1.043 | 1.0191 | 0.91755 | 183.9111 | 49.0856 | 51283.0614 | 51334.3099 | 1022.6862 | 922.6862 | 1076.4941
6842 | 93.8083 | 0.98619 | 1.0032 | 1.0217 | 205.9244 | 49.0577 | 55024.1331 | 53138.099 | 1058.762 | 958.762 | 1076.4943
6843 | 94.6219 | 1.0069 | 0.83536 | 1.0186 | 193.0696 | 45.258 | 43013.2176 | 44718.8492 | 890.377 | 790.377 | 1076.51
6844 | 79.8753 | 1.0249 | 0.99451 | 0.97913 | 213.5759 | 48.6488 | 51568.0946 | 52610.5016 | 1048.21 | 948.21 | 1076.536
6845 | 93.2835 | 0.97763 | 0.98471 | 0.9987 | 195.0966 | 50.8523 | 54842.5986 | 54580.8477 | 1087.617 | 987.617 | 1076.589
6846 | 83.4785 | 1.07 | 1.0086 | 0.96344 | 199.1936 | 55.7938 | 59401.2842 | 59748.3406 | 1190.9668 | 1090.9668 | 1076.8238
6847 | 97.7169 | 0.91457 | 1.045 | 0.95533 | 206.5699 | 50.6897 | 55341.4457 | 56022.352 | 1116.447 | 1016.447 | 1076.8239
6848 | 92.5684 | 0.95623 | 0.93883 | 1.0068 | 197.0946 | 50.7994 | 52708.6502 | 54285.1915 | 1081.7038 | 981.7038 | 1076.8316
6849 | 99.0639 | 0.96581 | 1.0095 | 1.0481 | 202.8495 | 53.319 | 61716.2013 | 61174.932 | 1219.4986 | 1119.4986 | 1076.8358
6850 | 89.6165 | 0.95152 | 1.2038 | 1.0629 | 200.8823 | 50.0831 | 68556.6711 | 67257.6063 | 1341.1521 | 1241.1521 | 1076.9108
6851 | 82.5146 | 0.99235 | 1.1138 | 0.98242 | 202.6076 | 48.6981 | 57476.3817 | 57231.744 | 1140.6349 | 1040.6349 | 1076.943
6852 | 92.9455 | 0.92209 | 1.0834 | 1.0396 | 191.6302 | 53.2058 | 64675.9188 | 62779.7142 | 1251.5943 | 1151.5943 | 1076.9768
6853 | 96.9689 | 0.98218 | 1.1361 | 1.0165 | 202.1981 | 49.9694 | 62671.6271 | 62328.1388 | 1242.5628 | 1142.5628 | 1077.0409
6854 | 113.1503 | 1.0344 | 1.1632 | 1.099 | 199.6692 | 48.0782 | 67289.9296 | 65060.5013 | 1297.21 | 1197.21 | 1077.0943
6855 | 98.946 | 1.0246 | 0.90293 | 0.99785 | 204.9349 | 48.5095 | 48829.0811 | 49264.4587 | 981.2892 | 881.2892 | 1077.1115
6856 | 92.8931 | 0.96375 | 1.1634 | 1.0466 | 200.0006 | 50.7045 | 66478.7015 | 65457.7567 | 1305.1551 | 1205.1551 | 1077.134
6857 | 105.4382 | 1.03 | 0.96707 | 0.97829 | 194.5232 | 47.2806 | 50060.4396 | 49037.7746 | 976.7555 | 876.7555 | 1077.1775
6858 | 99.716 | 0.94214 | 1.2845 | 1.0367 | 193.4727 | 47.169 | 67438.3894 | 66339.2896 | 1322.7858 | 1222.7858 | 1077.2776
6859 | 109.4678 | 0.9642 | 0.67056 | 1.018 | 187.8527 | 53.7388 | 42543.7562 | 41918.6315 | 834.3726 | 734.3726 | 1077.33
6860 | 92.2426 | 0.9488 | 0.97508 | 1.0158 | 192.3007 | 52.2582 | 56528.2737 | 55869.0148 | 1113.3803 | 1013.3803 | 1077.4123
6861 | 99.8941 | 1.1003 | 1.1027 | 1.0128 | 205.444 | 50.3226 | 61934.7529 | 62912.0035 | 1254.2401 | 1154.2401 | 1077.4554
6862 | 94.5474 | 0.97946 | 1.0682 | 0.96617 | 202.293 | 46.9933 | 53055.3113 | 55377.9185 | 1103.5584 | 1003.5584 | 1077.4928
6863 | 119.3566 | 0.96605 | 1.3302 | 0.92958 | 201.8059 | 50.8167 | 68897.3274 | 68199.0121 | 1359.9802 | 1259.9802 | 1077.5013
6864 | 110.4531 | 0.96569 | 1.2021 | 1.069 | 199.1515 | 47.9384 | 66918.4324 | 66426.0531 | 1324.5211 | 1224.5211 | 1077.527
6865 | 103.4429 | 0.98608 | 1.1717 | 0.98156 | 194.7657 | 50.8416 | 63853.5839 | 62133.2059 | 1238.6641 | 1138.6641 | 1077.5468
6866 | 79.4641 | 0.96037 | 0.901 | 0.94758 | 207.8433 | 52.0757 | 48642.3318 | 48015.6527 | 956.3131 | 856.3131 | 1077.5589
6867 | 97.1944 | 0.99385 | 0.99163 | 0.99881 | 198.3787 | 50.8212 | 55443.8664 | 53893.447 | 1073.8689 | 973.8689 | 1077.6301
6868 | 113.104 | 1.0285 | 1.0762 | 0.98305 | 197.0631 | 54.2512 | 63901.2879 | 64548.4509 | 1286.969 | 1186.969 | 1077.6998
6869 | 98.229 | 1.0007 | 0.98184 | 0.97422 | 202.066 | 48.8464 | 51726.7243 | 51421.5237 | 1024.4305 | 924.4305 | 1077.7183
6870 | 99.2999 | 1.1087 | 0.84993 | 0.98952 | 199.0494 | 51.0733 | 48776.0334 | 47602.1237 | 948.0425 | 848.0425 | 1077.7273
6871 | 98.7812 | 1.0077 | 1.0086 | 1.0039 | 177.34 | 50.7668 | 56636.198 | 55104.7497 | 1098.095 | 998.095 | 1077.7842
6872 | 99.8956 | 0.96102 | 0.85102 | 1.0521 | 207.4571 | 55.2326 | 54960.6997 | 53943.4216 | 1074.8684 | 974.8684 | 1077.8035
6873 | 102.7131 | 1.0625 | 0.97103 | 0.90995 | 198.5062 | 54.5867 | 54388.2239 | 52784.3648 | 1051.6873 | 951.6873 | 1077.9199
6874 | 89.5503 | 1.0391 | 1.0913 | 0.92928 | 189.086 | 48.0656 | 53405.3634 | 52312.6041 | 1042.2521 | 942.2521 | 1077.9308
6875 | 95.0483 | 0.99997 | 0.97657 | 1.0533 | 213.1845 | 51.0171 | 57537.2683 | 55843.2001 | 1112.864 | 1012.864 | 1077.9431
6876 | 96.4303 | 1.1031 | 1.391 | 0.9975 | 198.9301 | 48.9023 | 73255.6017 | 71954.1644 | 1435.0833 | 1335.0833 | 1078.0079
6877 | 77.4185 | 0.98965 | 0.92061 | 0.97784 | 226.253 | 49.4792 | 48559.308 | 49130.0809 | 978.6016 | 878.6016 | 1078.1102
6878 | 117.6317 | 1.0188 | 0.99367 | 1.0268 | 177.513 | 47.7223 | 54586.3556 | 55797.4366 | 1111.9487 | 1011.9487 | 1078.3609
6879 | 96.4299 | 1.0257 | 1.0863 | 1.0223 | 182.3224 | 52.2587 | 63384.172 | 62193.9597 | 1239.8792 | 1139.8792 | 1078.485
6880 | 107.1997 | 0.98022 | 0.83812 | 0.96784 | 209.1224 | 49.1841 | 45273.4748 | 45706.1493 | 910.123 | 810.123 | 1078.499
6881 | 92.6288 | 1.024 | 1.169 | 0.99864 | 202.2581 | 50.1066 | 63448.5861 | 62730.2129 | 1250.6043 | 1150.6043 | 1078.5966
6882 | 100.5853 | 1.0784 | 0.96129 | 1.0797 | 207.6364 | 49.5788 | 57045.0853 | 57761.6249 | 1151.2325 | 1051.2325 | 1078.757
6883 | 102.6174 | 1.0254 | 1.0695 | 1.0402 | 219.7488 | 50.0751 | 61196.6083 | 61916.2229 | 1234.3245 | 1134.3245 | 1078.7917
6884 | 87.7221 | 1.0564 | 1.0582 | 0.96633 | 200.6472 | 53.7139 | 60106.49 | 60149.0274 | 1198.9805 | 1098.9805 | 1078.8392
6885 | 105.9955 | 0.92261 | 1.051 | 1.0182 | 200.5034 | 47.9475 | 56203.2882 | 57289.7087 | 1141.7942 | 1041.7942 | 1078.8636
6886 | 93.4271 | 1.0189 | 0.79819 | 0.93685 | 195.8524 | 49.0202 | 41518.6432 | 41874.8825 | 833.4976 | 733.4976 | 1078.9581
6887 | 93.1768 | 1.0096 | 0.96607 | 0.99022 | 185.6447 | 50.2076 | 52938.4836 | 51100.1003 | 1018.002 | 918.002 | 1079.0175
6888 | 98.1003 | 1.0004 | 1.2155 | 1.002 | 208.2303 | 51.6407 | 68170.3796 | 69542.6022 | 1386.852 | 1286.852 | 1079.031
6889 | 99.6011 | 0.99808 | 0.87602 | 0.94012 | 208.4194 | 46.0031 | 42667.715 | 43250.48 | 861.0096 | 761.0096 | 1079.0506
6890 | 103.4336 | 1.0487 | 0.98778 | 1.0043 | 189.9304 | 53.0984 | 58626.2677 | 56946.2048 | 1134.9241 | 1034.9241 | 1079.0883
6891 | 88.8021 | 1.0132 | 0.68975 | 1.0314 | 215.5463 | 51.0025 | 41089.3511 | 40558.2089 | 807.1642 | 707.1642 | 1079.1744
6892 | 95.2441 | 0.93141 | 0.82043 | 1.011 | 226.8665 | 45.6513 | 42141.2207 | 43035.138 | 856.7028 | 756.7028 | 1079.1849
6893 | 101.7007 | 0.96311 | 0.83162 | 1.0255 | 175.2877 | 49.8109 | 47535.6553 | 47902.8996 | 954.058 | 854.058 | 1079.196
6894 | 112.8855 | 0.93281 | 1.144 | 1.0242 | 208.7778 | 51.4878 | 65954.8093 | 67418.239 | 1344.3648 | 1244.3648 | 1079.2638
6895 | 83.8666 | 0.9663 | 0.91059 | 0.97676 | 183.6686 | 47.9882 | 46754.6664 | 47850.1806 | 953.0036 | 853.0036 | 1079.2681
6896 | 95.2657 | 0.98562 | 1.1984 | 0.93264 | 205.4198 | 52.9397 | 64343.4369 | 64780.5518 | 1291.611 | 1191.611 | 1079.2701
6897 | 86.9319 | 1.0133 | 0.74114 | 1.0768 | 195.4318 | 47.2453 | 42061.4456 | 42699.3586 | 849.9872 | 749.9872 | 1079.2788
6898 | 94.3763 | 1.0126 | 0.78034 | 1.0645 | 187.8289 | 51.5255 | 47913.7775 | 45609.6742 | 908.1935 | 808.1935 | 1079.3526
6899 | 116.7989 | 0.97901 | 1.3588 | 0.95823 | 193.657 | 47.1881 | 67030.9957 | 69355.5857 | 1383.1117 | 1283.1117 | 1079.3872
6900 | 97.8395 | 1.0879 | 1.0226 | 0.94404 | 203.4209 | 53.7345 | 57794.761 | 57753.0745 | 1151.0615 | 1051.0615 | 1079.3921
6901 | 109.2791 | 1.0618 | 1.1154 | 1.053 | 174.9503 | 45.5617 | 58971.5895 | 58761.9656 | 1171.2393 | 1071.2393 | 1079.4678
6902 | 86.1869 | 0.98276 | 1.267 | 0.9696 | 207.3587 | 48.82 | 64317.0492 | 63977.2718 | 1275.5454 | 1175.5454 | 1079.501
6903 | 111.3777 | 0.99585 | 0.78953 | 1.0033 | 198.1155 | 45.4198 | 41215.3965 | 43252.8656 | 861.0573 | 761.0573 | 1079.5289
6904 | 87.4796 | 0.99617 | 1.0219 | 1.0559 | 201.9415 | 45.8896 | 53714.431 | 54583.0176 | 1087.6604 | 987.6604 | 1079.5332
6905 | 97.9251 | 1.0283 | 0.89791 | 0.96064 | 206.8536 | 48.7377 | 47154.5056 | 47492.3238 | 945.8465 | 845.8465 | 1079.55
6906 | 122.694 | 1.0211 | 1.0727 | 1.0389 | 200.6122 | 49.9595 | 62139.8517 | 58986.245 | 1175.7249 | 1075.7249 | 1079.576
6907 | 92.8052 | 0.95449 | 0.98249 | 1.038 | 213.9308 | 52.7223 | 58649.7615 | 56254.9805 | 1121.0996 | 1021.0996 | 1079.7845
6908 | 90.7454 | 1.009 | 1.3038 | 0.92964 | 196.3233 | 51.1124 | 66826.455 | 65344.3613 | 1302.8872 | 1202.8872 | 1079.845
6909 | 94.2626 | 0.89777 | 0.9014 | 0.95553 | 191.005 | 47.3174 | 44950.4385 | 44200.1413 | 880.0028 | 780.0028 | 1079.8498
6910 | 119.2096 | 1.0273 | 0.93923 | 0.98097 | 202.0535 | 49.8151 | 52199.4729 | 52646.6258 | 1048.9325 | 948.9325 | 1079.8684
6911 | 110.4978 | 1.007 | 0.86789 | 0.96252 | 194.2091 | 51.5319 | 48976.3252 | 49039.3153 | 976.7863 | 876.7863 | 1079.9229
6912 | 99.7145 | 0.96611 | 1.1179 | 0.99174 | 206.139 | 51.7759 | 62598.4048 | 63504.2356 | 1266.0847 | 1166.0847 | 1079.9778
6913 | 106.9115 | 0.99237 | 0.92819 | 0.94039 | 209.9166 | 50.8444 | 49984.2178 | 49455.3641 | 985.1073 | 885.1073 | 1080.1629
6914 | 106.4016 | 1.0568 | 1.1545 | 1.0324 | 213.4652 | 48.6927 | 63727.106 | 64860.0782 | 1293.2016 | 1193.2016 | 1080.1992
6915 | 119.3214 | 0.84058 | 0.82047 | 0.97528 | 206.1404 | 51.6724 | 46736.6537 | 47236.9785 | 940.7396 | 840.7396 | 1080.2727
6916 | 88.6413 | 1.0084 | 0.74259 | 0.95885 | 205.8491 | 51.1656 | 41210.6645 | 42749.9604 | 850.9992 | 750.9992 | 1080.3015
6917 | 97.4779 | 0.98055 | 0.94413 | 1.0306 | 196.1324 | 50.6009 | 54266.1321 | 52391.4367 | 1043.8287 | 943.8287 | 1080.3033
6918 | 104.765 | 0.94162 | 1.0252 | 0.9499 | 186.9185 | 51.6864 | 55621.8264 | 55879.3125 | 1113.5863 | 1013.5863 | 1080.3264
6919 | 90.3236 | 1.0501 | 1.1201 | 0.97858 | 201.91 | 53.1269 | 63474.7708 | 64386.1943 | 1283.7239 | 1183.7239 | 1080.3523
6920 | 95.3653 | 0.96874 | 1.093 | 0.92144 | 217.443 | 47.4082 | 52343.8728 | 51873.9298 | 1033.4786 | 933.4786 | 1080.3568
6921 | 116.7076 | 1.0433 | 0.64819 | 1.0563 | 183.0842 | 48.5229 | 39315.25 | 38902.026 | 774.0405 | 674.0405 | 1080.4984
6922 | 125.3843 | 0.94613 | 0.78124 | 0.99684 | 186.6255 | 52.1273 | 46965.7304 | 47507.0318 | 946.1406 | 846.1406 | 1080.5495
6923 | 107.8086 | 1.047 | 1.1411 | 0.99282 | 198.3598 | 47.2791 | 59098.9698 | 58876.1666 | 1173.5233 | 1073.5233 | 1080.665
6924 | 115.602 | 1.0984 | 0.96285 | 0.91756 | 191.294 | 51.4326 | 52161.6241 | 53495.3924 | 1065.9078 | 965.9078 | 1080.7335
6925 | 117.0896 | 1.0196 | 1.2707 | 0.92648 | 188.4578 | 54.4055 | 70734.8716 | 68713.8887 | 1370.2778 | 1270.2778 | 1080.8228
6926 | 82.6841 | 1.0534 | 0.79795 | 0.97854 | 189.4251 | 50.9967 | 44450.64 | 44230.9057 | 880.6181 | 780.6181 | 1080.8911
6927 | 99.0285 | 1.0298 | 0.91648 | 0.98353 | 190.9834 | 47.8167 | 48168.3376 | 48262.1219 | 961.2424 | 861.2424 | 1080.9958
6928 | 86.0674 | 1.0321 | 1.0332 | 1.0093 | 199.1292 | 54.3111 | 61661.575 | 61309.4285 | 1222.1886 | 1122.1886 | 1081.0216
6929 | 85.6175 | 1.0028 | 1.0165 | 0.95606 | 197.3358 | 48.5445 | 51542.2341 | 53263.4685 | 1061.2694 | 961.2694 | 1081.0359
6930 | 108.6533 | 1.0187 | 0.84835 | 1.0655 | 183.039 | 51.1157 | 52044.1378 | 52344.9389 | 1042.8988 | 942.8988 | 1081.1041
6931 | 97.2041 | 0.9716 | 0.94363 | 0.98367 | 196.4203 | 48.7493 | 50050.525 | 49118.0613 | 978.3612 | 878.3612 | 1081.2173
6932 | 99.2733 | 1.0271 | 0.95596 | 1.0175 | 220.6518 | 50.2869 | 54263.4163 | 55588.4332 | 1107.7687 | 1007.7687 | 1081.3725
6933 | 84.9147 | 1.0545 | 1.0621 | 1.0173 | 183.9204 | 56.7376 | 66566.5403 | 66542.6967 | 1326.8539 | 1226.8539 | 1081.4805
6934 | 102.5497 | 1.0671 | 1.2996 | 1.0115 | 200.7579 | 47.8985 | 68410.0621 | 68115.337 | 1358.3067 | 1258.3067 | 1081.5103
6935 | 105.373 | 0.94151 | 1.1845 | 0.99758 | 217.3968 | 49.2847 | 63343.1228 | 61161.743 | 1219.2349 | 1119.2349 | 1081.5384
6936 | 96.3546 | 0.96586 | 0.99326 | 0.9232 | 206.9487 | 51.2856 | 52007.6745 | 51760.5197 | 1031.2104 | 931.2104 | 1081.5872
6937 | 101.2698 | 0.99078 | 0.93552 | 0.99069 | 199.5429 | 48.9392 | 50467.3961 | 51388.4773 | 1023.7695 | 923.7695 | 1081.6388
6938 | 81.7011 | 1.0549 | 1.2029 | 1.0259 | 194.4417 | 49.5122 | 65559.0897 | 65571.714 | 1307.4343 | 1207.4343 | 1081.692
6939 | 97.5796 | 0.94315 | 0.91315 | 0.95148 | 201.6761 | 52.382 | 50534.2517 | 49934.8034 | 994.6961 | 894.6961 | 1081.7314
6940 | 93.3761 | 1.0615 | 0.89452 | 1.0215 | 212.4943 | 50.2652 | 51124.1594 | 51142.8176 | 1018.8564 | 918.8564 | 1081.9495
6941 | 96.4306 | 0.95022 | 1.1784 | 0.99284 | 196.265 | 52.1149 | 65946.3168 | 65491.4516 | 1305.829 | 1205.829 | 1081.9688
6942 | 105.24 | 0.95927 | 0.96753 | 0.96553 | 196.6645 | 50.6894 | 52667.1115 | 52215.9618 | 1040.3192 | 940.3192 | 1081.9773
6943 | 96.7983 | 1.0193 | 1.0865 | 1.0431 | 195.304 | 54.4069 | 67224.6743 | 67326.4606 | 1342.5292 | 1242.5292 | 1081.9897
6944 | 91.6817 | 1.0371 | 1.4636 | 0.95641 | 209.4124 | 44.0079 | 65994.382 | 65843.0875 | 1312.8618 | 1212.8618 | 1082.0793
6945 | 103.7161 | 0.8925 | 1.0154 | 0.9562 | 186.1578 | 49.1568 | 52463.2679 | 53054.89 | 1057.0978 | 957.0978 | 1082.1202
6946 | 117.266 | 0.92205 | 0.98604 | 1.0286 | 205.3003 | 51.6675 | 58197.0943 | 58581.4611 | 1167.6292 | 1067.6292 | 1082.1463
6947 | 75.3394 | 1.0741 | 0.8631 | 1.0758 | 197.8974 | 50.3155 | 50988.9269 | 51985.1265 | 1035.7025 | 935.7025 | 1082.3951
6948 | 106.0998 | 1 | 1.1165 | 1.0563 | 196.3965 | 50.149 | 64666.2078 | 64675.3106 | 1289.5062 | 1189.5062 | 1082.4349
6949 | 94.4565 | 0.85657 | 1.1672 | 1.0294 | 205.4024 | 51.9986 | 66890.5334 | 68610.7949 | 1368.2159 | 1268.2159 | 1082.4382
6950 | 123.1904 | 0.95298 | 0.91864 | 0.96616 | 207.4305 | 49.1559 | 49606.7292 | 48670.274 | 969.4055 | 869.4055 | 1082.4444
6951 | 90.7161 | 0.9389 | 1.0809 | 1.0051 | 184.4914 | 53.9411 | 63380.2279 | 64274.9229 | 1281.4985 | 1181.4985 | 1082.4651
6952 | 94.0903 | 0.94771 | 1.0743 | 1.0421 | 196.1858 | 48.7155 | 59074.5919 | 60578.0977 | 1207.562 | 1107.562 | 1082.5377
6953 | 91.2579 | 0.95035 | 1.041 | 1.0046 | 214.9852 | 45.3719 | 51597.4256 | 51393.8914 | 1023.8778 | 923.8778 | 1082.655
6954 | 116.8491 | 0.9466 | 0.88081 | 1.0187 | 174.0415 | 53.8809 | 54479.684 | 56058.2411 | 1117.1648 | 1017.1648 | 1082.6653
6955 | 92.8169 | 1.0599 | 0.9554 | 0.95346 | 201.4013 | 50.1772 | 50845.9621 | 51452.9411 | 1025.0588 | 925.0588 | 1082.6671
6956 | 76.2314 | 1.0024 | 1.1582 | 0.95736 | 189.2954 | 50.0855 | 59549.7188 | 59457.1063 | 1185.1421 | 1085.1421 | 1082.7168
6957 | 96.8453 | 0.93179 | 1.3512 | 0.95684 | 196.1077 | 50.596 | 70177.0243 | 73399.8611 | 1463.9972 | 1363.9972 | 1082.7412
6958 | 117.1388 | 0.93189 | 0.88559 | 0.90896 | 199.1183 | 47.3553 | 43487.5471 | 42600.8214 | 848.0164 | 748.0164 | 1082.8075
6959 | 96.5228 | 1.0664 | 1.131 | 1.05 | 211.5618 | 51.7338 | 66972.8624 | 67375.8227 | 1343.5165 | 1243.5165 | 1082.8306
6960 | 88.5705 | 0.88329 | 0.98841 | 1.0084 | 191.8634 | 49.4173 | 53310.785 | 52405.119 | 1044.1024 | 944.1024 | 1082.8875
6961 | 80.7368 | 0.98212 | 1.1288 | 1.0606 | 183.1705 | 51.796 | 66299.3833 | 65075.4644 | 1297.5093 | 1197.5093 | 1082.9893
6962 | 94.3017 | 0.9937 | 1.0568 | 0.94726 | 193.2243 | 47.1791 | 51842.3527 | 50216.2771 | 1000.3255 | 900.3255 | 1083.028
6963 | 114.6321 | 1.0325 | 0.94223 | 0.9422 | 207.8594 | 47.7712 | 48271.5319 | 47073.8922 | 937.4778 | 837.4778 | 1083.1036
6964 | 96.7101 | 0.94461 | 1.1518 | 1.0187 | 198.3102 | 51.5359 | 65377.1953 | 64701.3335 | 1290.0267 | 1190.0267 | 1083.1176
6965 | 100.9938 | 0.98569 | 0.90898 | 0.96235 | 191.4175 | 45.3236 | 44350.5525 | 44795.896 | 891.9179 | 791.9179 | 1083.1454
6966 | 113.5738 | 1.0471 | 1.1467 | 0.99752 | 187.1434 | 50.9481 | 64520.767 | 65392.6599 | 1303.8532 | 1203.8532 | 1083.1707
6967 | 104.2396 | 0.93034 | 1.1241 | 1.0604 | 205.42 | 48.3491 | 62525.526 | 61729.7337 | 1230.5947 | 1130.5947 | 1083.2437
6968 | 77.76 | 1.0895 | 0.76179 | 1.0815 | 200.5728 | 51.1666 | 46691.5042 | 46766.7987 | 931.336 | 831.336 | 1083.2879
6969 | 95.8587 | 0.92916 | 1.0016 | 0.91949 | 195.33 | 52.8186 | 53543.3346 | 54716.8857 | 1090.3377 | 990.3377 | 1083.3074
6970 | 102.0793 | 0.96245 | 1.2897 | 0.98052 | 209.9663 | 50.282 | 68736.6864 | 68604.1383 | 1368.0828 | 1268.0828 | 1083.3395
6971 | 86.0468 | 1.0057 | 0.89801 | 0.93893 | 203.4759 | 48.8484 | 45618.4782 | 45217.8654 | 900.3573 | 800.3573 | 1083.3649
6972 | 106.3753 | 0.9164 | 1.2081 | 1.0212 | 198.5565 | 50.0033 | 66759.0757 | 64677.9729 | 1289.5595 | 1189.5595 | 1083.3763
6973 | 108.2989 | 0.98996 | 0.95982 | 1.0286 | 217.3453 | 49.6512 | 54558.3868 | 57339.9454 | 1142.7989 | 1042.7989 | 1083.3847
6974 | 110.4865 | 1.0237 | 1.0754 | 1.0922 | 210.4244 | 52.2213 | 67449.2755 | 67047.8388 | 1336.9568 | 1236.9568 | 1083.4137
6975 | 96.8435 | 1.0514 | 0.82004 | 1.0122 | 210.3004 | 47.9223 | 44866.4157 | 45078.7262 | 897.5745 | 797.5745 | 1083.471
6976 | 102.914 | 1.0105 | 1.2388 | 0.98165 | 211.9745 | 53.6555 | 71038.1324 | 70327.194 | 1402.5439 | 1302.5439 | 1083.483
6977 | 123.0704 | 1.0038 | 0.98663 | 0.98945 | 220.0478 | 49.4358 | 54588.0014 | 53267.6411 | 1061.3528 | 961.3528 | 1083.5468
6978 | 83.0817 | 0.97693 | 0.9516 | 1.0302 | 195.4166 | 47.6484 | 50773.3349 | 49991.2581 | 995.8252 | 895.8252 | 1083.5521
6979 | 107.2184 | 0.94701 | 1.0693 | 1.0452 | 193.4532 | 51.9633 | 63547.0073 | 62095.5831 | 1237.9117 | 1137.9117 | 1083.5732
6980 | 98.1002 | 1.1094 | 0.90315 | 0.95653 | 199.6175 | 51.7287 | 50516.9963 | 50229.0803 | 1000.5816 | 900.5816 | 1083.7021
6981 | 93.7837 | 0.93194 | 1.046 | 1.0367 | 192.27 | 43.2718 | 50901.3407 | 49419.483 | 984.3897 | 884.3897 | 1083.7105
6982 | 79.4354 | 1.0982 | 0.84596 | 1.0092 | 205.1381 | 53.4594 | 50510.6382 | 49916.6765 | 994.3335 | 894.3335 | 1083.7136
6983 | 94.8983 | 1.0674 | 1.0471 | 0.97425 | 204.0938 | 51.6155 | 58086.0578 | 57632.3088 | 1148.6462 | 1048.6462 | 1083.7749
6984 | 94.2513 | 0.97491 | 1.0252 | 1.0226 | 197.8708 | 47.4714 | 54326.1788 | 54582.5547 | 1087.6511 | 987.6511 | 1083.7803
6985 | 93.2316 | 1.0053 | 1.0688 | 0.92656 | 201.2743 | 50.1229 | 54533.8135 | 53844.6098 | 1072.8922 | 972.8922 | 1083.8928
6986 | 103.4363 | 0.95211 | 0.79901 | 0.98109 | 185.4388 | 49.6718 | 44014.9254 | 45793.2868 | 911.8657 | 811.8657 | 1083.9851
6987 | 99.3444 | 0.96067 | 1.341 | 1.0181 | 190.9359 | 54.2184 | 79389.0498 | 81679.0117 | 1629.5802 | 1529.5802 | 1083.9966
6988 | 88.2977 | 0.96168 | 0.6369 | 0.99393 | 182.691 | 48.9245 | 35308.0011 | 33824.9677 | 672.4994 | 572.4994 | 1084.0336
6989 | 113.3132 | 0.95592 | 1.3647 | 1.0171 | 197.1469 | 51.5859 | 77388.7645 | 80082.4986 | 1597.65 | 1497.65 | 1084.038
6990 | 107.0773 | 1.0284 | 1.1005 | 1.0644 | 215.2732 | 49.0508 | 63073.995 | 62419.6677 | 1244.3934 | 1144.3934 | 1084.0517
6991 | 95.7541 | 1.0195 | 0.94586 | 0.99942 | 188.1588 | 49.9493 | 52282.1894 | 51270.7153 | 1021.4143 | 921.4143 | 1084.0527
6992 | 110.6087 | 1.0826 | 0.77402 | 1.0643 | 191.7601 | 51.6746 | 48946.9546 | 48350.8319 | 963.0166 | 863.0166 | 1084.1286
6993 | 87.2591 | 0.99276 | 1.0182 | 1.014 | 197.3972 | 50.8756 | 57130.4499 | 57601.3833 | 1148.0277 | 1048.0277 | 1084.1518
6994 | 106.6732 | 1.0232 | 0.71768 | 0.97939 | 192.5003 | 50.4679 | 41174.2777 | 41575.0006 | 827.5 | 727.5 | 1084.1612
6995 | 102.4501 | 0.99752 | 0.90373 | 0.99012 | 201.8133 | 48.2019 | 48259.0367 | 48907.9302 | 974.1586 | 874.1586 | 1084.198
6996 | 89.3834 | 0.94637 | 1.0907 | 0.96186 | 198.1004 | 49.6166 | 56446.3555 | 57055.721 | 1137.1144 | 1037.1144 | 1084.2335
6997 | 122.91 | 1.0196 | 1.04 | 1.0511 | 205.2741 | 52.2656 | 63888.7708 | 62964.5543 | 1255.2911 | 1155.2911 | 1084.3278
6998 | 96.4051 | 1.0264 | 0.83869 | 0.99644 | 194.1162 | 50.018 | 46943.326 | 47340.6839 | 942.8137 | 842.8137 | 1084.3501
6999 | 111.1135 | 1.0303 | 1.1296 | 0.96914 | 198.0222 | 46.6489 | 56608.8438 | 56622.5833 | 1128.4517 | 1028.4517 | 1084.3524
7000 | 103.8608 | 0.94332 | 1.1927 | 0.96731 | 194.6435 | 49.6902 | 62392.7471 | 63006.5655 | 1256.1313 | 1156.1313 | 1084.3646
7001 | 101.5457 | 0.96162 | 0.84359 | 0.9841 | 218.9771 | 52.0397 | 48502.496 | 48640.2683 | 968.8054 | 868.8054 | 1084.4373
7002 | 101.8551 | 1.0462 | 0.99891 | 1.0283 | 194.5108 | 51.1073 | 58139.2795 | 59385.6804 | 1183.7136 | 1083.7136 | 1084.4476
7003 | 97.3021 | 1.0444 | 0.91555 | 0.95683 | 213.0619 | 48.9067 | 48026.3592 | 47509.1735 | 946.1835 | 846.1835 | 1084.474
7004 | 99.2893 | 0.9669 | 0.99429 | 1.1117 | 218.601 | 45.4835 | 54858.8676 | 55128.111 | 1098.5622 | 998.5622 | 1084.5935
7005 | 98.0516 | 1.0527 | 0.87381 | 1.0748 | 187.0794 | 47.3744 | 49568.6735 | 51208.437 | 1020.1687 | 920.1687 | 1084.6205
7006 | 111.7573 | 0.99218 | 1.1272 | 0.91028 | 187.3589 | 52.9197 | 60354.1242 | 59407.5904 | 1184.1518 | 1084.1518 | 1084.6517
7007 | 111.2498 | 0.96484 | 0.89354 | 0.98311 | 206.8367 | 51.6489 | 51121.6148 | 52208.889 | 1040.1778 | 940.1778 | 1084.6593
7008 | 90.5505 | 1.0037 | 1.1227 | 1.0006 | 182.5457 | 47.5325 | 57897.1577 | 57425.6721 | 1144.5134 | 1044.5134 | 1084.6967
7009 | 97.2205 | 1.1175 | 0.78642 | 1.042 | 216.1676 | 53.6607 | 50016.8227 | 49733.5254 | 990.6705 | 890.6705 | 1084.8504
7010 | 89.5908 | 1.0255 | 0.97751 | 0.92473 | 191.1037 | 48.7744 | 48761.1127 | 49299.6892 | 981.9938 | 881.9938 | 1084.861
7011 | 116.7799 | 1.0447 | 0.95289 | 1.0574 | 211.0945 | 49.5838 | 56220.183 | 56620.7007 | 1128.414 | 1028.414 | 1084.9247
7012 | 97.1401 | 0.9713 | 0.85352 | 0.99611 | 197.4038 | 52.0114 | 49324.6298 | 50540.0593 | 1006.8012 | 906.8012 | 1084.9633
7013 | 88.2096 | 1.0016 | 0.80957 | 1.0032 | 191.4313 | 53.1192 | 48024.1184 | 47681.2128 | 949.6243 | 849.6243 | 1085.1165
7014 | 106.5191 | 0.95047 | 0.88683 | 0.95747 | 195.8493 | 52.1308 | 49738.711 | 48629.0738 | 968.5815 | 868.5815 | 1085.1404
7015 | 93.9791 | 0.9804 | 1.1711 | 0.94753 | 182.0363 | 45.9554 | 55412.9906 | 53718.3171 | 1070.3663 | 970.3663 | 1085.1421
7016 | 90.0357 | 1.0382 | 1.1057 | 0.99797 | 205.4011 | 48.5738 | 58344.2782 | 57575.0275 | 1147.5006 | 1047.5006 | 1085.2186
7017 | 89.476 | 0.94596 | 1.0729 | 0.94775 | 183.7284 | 45.7019 | 50523.6987 | 50644.0448 | 1008.8809 | 908.8809 | 1085.232
7018 | 107.8176 | 1.0277 | 1.0402 | 0.93728 | 183.8445 | 49.2083 | 53613.5485 | 52670.5938 | 1049.4119 | 949.4119 | 1085.4358
7019 | 96.6328 | 0.95698 | 1.1781 | 1.1034 | 196.5193 | 48.4484 | 67654.1631 | 68655.6619 | 1369.1132 | 1269.1132 | 1085.4414
7020 | 110.6206 | 1.0472 | 1.0339 | 0.90363 | 219.0764 | 48.0868 | 50716.3562 | 50608.0354 | 1008.1607 | 908.1607 | 1085.4593
7021 | 96.9278 | 0.92818 | 0.93365 | 0.93654 | 202.8582 | 49.529 | 47967.2298 | 48139.7768 | 958.7955 | 858.7955 | 1085.5308
7022 | 96.1532 | 0.97354 | 0.96371 | 0.93784 | 193.5205 | 45.1142 | 45191.3582 | 44433.9556 | 884.6791 | 784.6791 | 1085.5322
7023 | 112.1644 | 1.005 | 1.0205 | 1.0295 | 197.1069 | 56.991 | 66496.2804 | 66946.5138 | 1334.9303 | 1234.9303 | 1085.5634
7024 | 89.6161 | 1.0178 | 1.0182 | 1.0584 | 211.5317 | 50.1177 | 58792.2028 | 60883.2931 | 1213.6659 | 1113.6659 | 1085.5737
7025 | 84.7208 | 0.99275 | 1.065 | 1.0226 | 180.4127 | 51.3194 | 60387.5799 | 60389.6692 | 1203.7934 | 1103.7934 | 1085.5802
7026 | 93.9152 | 0.96796 | 1.0543 | 0.99768 | 217.0208 | 48.253 | 55358.2024 | 53158.5027 | 1059.1701 | 959.1701 | 1085.6504
7027 | 96.4635 | 0.99355 | 1.0347 | 0.85051 | 208.2838 | 51.7143 | 50673.9241 | 50375.1969 | 1003.5039 | 903.5039 | 1085.6634
7028 | 106.7483 | 0.96088 | 0.99729 | 1.0823 | 183.733 | 51.631 | 61206.384 | 61411.4082 | 1224.2282 | 1124.2282 | 1085.6707
7029 | 96.5465 | 1.0599 | 0.95719 | 0.93092 | 210.5602 | 48.9907 | 48877.8907 | 48670.5302 | 969.4106 | 869.4106 | 1085.6891
7030 | 96.295 | 0.96197 | 0.97127 | 1.0049 | 186.4338 | 48.9398 | 52485.1602 | 52168.4219 | 1039.3684 | 939.3684 | 1085.7261
7031 | 100.665 | 0.96844 | 1.0334 | 0.94774 | 193.0096 | 48.2361 | 52138.0523 | 52563.0343 | 1047.2607 | 947.2607 | 1085.7925
7032 | 89.6274 | 0.99743 | 0.91574 | 1.0991 | 209.0552 | 51.557 | 56711.8899 | 58173.2417 | 1159.4648 | 1059.4648 | 1085.8849
7033 | 90.0348 | 1.0309 | 1.1819 | 1.0424 | 191.822 | 47.7029 | 63389.1774 | 62896.6665 | 1253.9333 | 1153.9333 | 1086.0162
7034 | 100.4044 | 0.95986 | 1.0528 | 1.0555 | 180.6848 | 52.9762 | 64153.5285 | 64645.8318 | 1288.9166 | 1188.9166 | 1086.0214
7035 | 108.0384 | 0.9874 | 0.89701 | 0.98879 | 201.4133 | 47.7757 | 47673.1014 | 48854.4255 | 973.0885 | 873.0885 | 1086.0342
7036 | 103.5515 | 1.0784 | 0.9304 | 1.0473 | 215.3382 | 49.0561 | 53494.3624 | 53634.8151 | 1068.6963 | 968.6963 | 1086.0363
7037 | 127.3345 | 1.0313 | 0.9723 | 0.9822 | 208.8349 | 51.8178 | 56499.1911 | 58076.5113 | 1157.5302 | 1057.5302 | 1086.0461
7038 | 92.5312 | 1.0222 | 1.1254 | 0.96059 | 200.6075 | 46.4034 | 54754.0353 | 57572.3161 | 1147.4463 | 1047.4463 | 1086.0744
7039 | 102.3469 | 0.98809 | 0.85699 | 1.0649 | 221.8115 | 46.7079 | 47573.4022 | 47104.2065 | 938.0841 | 838.0841 | 1086.1442
7040 | 91.9427 | 1.1663 | 0.87068 | 0.96514 | 194.7572 | 51.1264 | 48640.2008 | 49038.997 | 976.7799 | 876.7799 | 1086.1485
7041 | 108.2248 | 0.90329 | 0.71045 | 1.0371 | 180.1472 | 46.3855 | 38891.2095 | 39121.7259 | 778.4345 | 678.4345 | 1086.1571
7042 | 106.2012 | 1.0295 | 1.0556 | 0.98355 | 197.8366 | 47.0725 | 54217.4915 | 52517.1191 | 1046.3424 | 946.3424 | 1086.2143
7043 | 90.7977 | 1.0989 | 1.1345 | 1.0862 | 183.6432 | 52.4628 | 70067.239 | 71800.6084 | 1432.0122 | 1332.0122 | 1086.2185
7044 | 91.4487 | 1.0629 | 1.0958 | 0.96123 | 204.0402 | 49.8928 | 57605.6839 | 56031.1149 | 1116.6223 | 1016.6223 | 1086.287
7045 | 98.8438 | 1.0003 | 0.93698 | 0.97177 | 178.6262 | 48.6354 | 49271.5483 | 50061.8093 | 997.2362 | 897.2362 | 1086.3651
7046 | 92.5055 | 1.0027 | 1.0159 | 0.90011 | 207.7309 | 52.2569 | 52837.9643 | 53718.9072 | 1070.3781 | 970.3781 | 1086.409
7047 | 87.4527 | 0.99862 | 0.84229 | 1.0075 | 203.4882 | 50.3012 | 47283.6212 | 46155.2818 | 919.1056 | 819.1056 | 1086.4223
7048 | 103.4041 | 1.0771 | 0.99288 | 0.97677 | 197.1236 | 51.8781 | 56287.4816 | 58715.9004 | 1170.318 | 1070.318 | 1086.688
7049 | 108.0702 | 0.99764 | 1.1598 | 1.038 | 204.5058 | 54.0149 | 71056.4532 | 70744.2053 | 1410.8841 | 1310.8841 | 1086.7229
7050 | 98.4516 | 1.0198 | 1.1116 | 0.97335 | 197.4042 | 49.8262 | 59111.2068 | 58674.9211 | 1169.4984 | 1069.4984 | 1086.7535
7051 | 85.4268 | 0.99237 | 1.1382 | 1.0214 | 207.9386 | 51.1518 | 64012.7433 | 62044.8574 | 1236.8971 | 1136.8971 | 1086.8284
7052 | 91.6546 | 0.89386 | 1.1766 | 1.0023 | 198.3487 | 51.6365 | 65319.9356 | 65916.6983 | 1314.334 | 1214.334 | 1086.9458
7053 | 106.7323 | 1.0368 | 0.76345 | 0.95606 | 197.7832 | 46.1206 | 38965.1236 | 39336.5084 | 782.7302 | 682.7302 | 1086.9549
7054 | 101.1647 | 0.97438 | 1.1342 | 0.99603 | 193.753 | 49.4279 | 60903.6733 | 58326.268 | 1162.5254 | 1062.5254 | 1087.157
7055 | 115.8266 | 1.0105 | 0.94055 | 0.99566 | 191.8989 | 49.348 | 52180.4012 | 51829.4433 | 1032.5889 | 932.5889 | 1087.1689
7056 | 94.4972 | 1.0066 | 1.0691 | 1.0351 | 194.2666 | 51.1916 | 61712.7515 | 61167.6499 | 1219.353 | 1119.353 | 1087.3044
7057 | 94.6619 | 1.0416 | 0.99353 | 0.92512 | 196.5609 | 48.8822 | 49946.1196 | 50910.373 | 1014.2075 | 914.2075 | 1087.3066
7058 | 109.5602 | 0.97591 | 1.031 | 0.94948 | 192.7936 | 52.8938 | 57627.6284 | 55105.4876 | 1098.1098 | 998.1098 | 1087.3094
7059 | 100.9355 | 0.99487 | 1.0368 | 0.89226 | 197.7025 | 50.9668 | 52465.2254 | 52966.6061 | 1055.3321 | 955.3321 | 1087.3714
7060 | 90.8543 | 0.99738 | 0.91045 | 0.95514 | 181.3957 | 50.8406 | 48999.6866 | 48267.6926 | 961.3539 | 861.3539 | 1087.4315
7061 | 105.9145 | 1.0372 | 1.0881 | 1.0595 | 196.9925 | 53.2035 | 67379.1162 | 65754.7752 | 1311.0955 | 1211.0955 | 1087.4486
7062 | 110.5083 | 0.97643 | 1.0973 | 0.94201 | 200.2676 | 49.1549 | 56312.7131 | 55749.3071 | 1110.9861 | 1010.9861 | 1087.5549
7063 | 101.7798 | 1.0189 | 1.0636 | 1.09 | 207.1997 | 47.6777 | 60427.2349 | 59159.8017 | 1179.196 | 1079.196 | 1087.6446
7064 | 111.6696 | 0.96009 | 0.98306 | 1.0382 | 197.8902 | 46.6284 | 52787.1006 | 50963.5249 | 1015.2705 | 915.2705 | 1087.7106
7065 | 86.808 | 1.0234 | 1.1679 | 1.0493 | 203.9266 | 48.7797 | 64316.2998 | 65220.7025 | 1300.4141 | 1200.4141 | 1087.9494
7066 | 84.5817 | 1.1451 | 1.0576 | 1.0139 | 216.9025 | 48.3327 | 56723.745 | 55427.4244 | 1104.5485 | 1004.5485 | 1088.0017
7067 | 102.4661 | 0.91862 | 0.70571 | 0.93593 | 189.5698 | 51.7265 | 39223.4373 | 38663.7391 | 769.2748 | 669.2748 | 1088.0095
7068 | 92.4421 | 1.0156 | 0.97714 | 0.92923 | 206.4101 | 50.9853 | 51287.0971 | 50717.7818 | 1010.3556 | 910.3556 | 1088.0146
7069 | 94.0002 | 1.0125 | 1.0904 | 0.93916 | 201.1659 | 53.5113 | 60091.1656 | 62060.1001 | 1237.202 | 1137.202 | 1088.0182
7070 | 107.3211 | 1.0093 | 1.2788 | 1.0141 | 206.7596 | 49.6107 | 69918.7345 | 71001.2131 | 1416.0243 | 1316.0243 | 1088.0238
7071 | 97.5922 | 1.0009 | 1.0448 | 1.0048 | 210.617 | 53.3514 | 61431.9859 | 62245.0905 | 1240.9018 | 1140.9018 | 1088.0527
7072 | 109.0333 | 0.95565 | 1.0156 | 0.97879 | 208.0858 | 45.0836 | 49720.9593 | 51153.8848 | 1019.0777 | 919.0777 | 1088.0779
7073 | 114.6076 | 0.92188 | 1.0272 | 0.99731 | 194.2391 | 47.8777 | 54300.9666 | 55097.6816 | 1097.9536 | 997.9536 | 1088.0906
7074 | 78.9813 | 0.94131 | 1.0263 | 1.0184 | 189.4535 | 50.1562 | 56340.1994 | 57020.0582 | 1136.4012 | 1036.4012 | 1088.2126
7075 | 109.3007 | 1.077 | 1.0095 | 1.0453 | 196.0329 | 48.6961 | 57313.6953 | 55245.6464 | 1100.9129 | 1000.9129 | 1088.2326
7076 | 104.4619 | 1.0048 | 1.1512 | 0.96922 | 195.2842 | 49.7293 | 60902.8884 | 62426.889 | 1244.5378 | 1144.5378 | 1088.266
7077 | 101.3645 | 0.94225 | 1.1602 | 0.98388 | 202.4265 | 53.3258 | 66166.5411 | 66032.6968 | 1316.6539 | 1216.6539 | 1088.3389
7078 | 111.8437 | 1.0148 | 1.1066 | 0.99527 | 184.4145 | 50.1902 | 61158.5226 | 62218.8659 | 1240.3773 | 1140.3773 | 1088.4337
7079 | 102.716 | 1.0147 | 0.61663 | 1.0437 | 184.4866 | 51.3019 | 38546.8051 | 39150.4906 | 779.0098 | 679.0098 | 1088.5201
7080 | 98.6727 | 1.0223 | 1.1393 | 0.9965 | 185.6195 | 50.6322 | 62776.9895 | 61747.2755 | 1230.9455 | 1130.9455 | 1088.5333
7081 | 99.1353 | 1.0528 | 0.76439 | 1.0247 | 204.5031 | 48.4127 | 43175.7854 | 41526.8585 | 826.5372 | 726.5372 | 1088.6621
7082 | 112.7409 | 1.0585 | 1.2002 | 0.87215 | 197.7983 | 53.974 | 63136.5763 | 59276.9205 | 1181.5384 | 1081.5384 | 1088.6755
7083 | 105.1322 | 1.0131 | 0.82934 | 1.0195 | 199.6725 | 50.6034 | 48374.9724 | 48416.9734 | 964.3395 | 864.3395 | 1088.7477
7084 | 97.9311 | 0.98705 | 0.84887 | 1.0353 | 196.6487 | 47.8764 | 46898.3535 | 47712.2897 | 950.2458 | 850.2458 | 1088.7542
7085 | 103.2027 | 1.0412 | 1.1075 | 1.0384 | 196.9297 | 47.4659 | 59884.3899 | 59763.8411 | 1191.2768 | 1091.2768 | 1088.9289
7086 | 110.8579 | 1.0269 | 1.0388 | 0.98674 | 193.0644 | 50.3525 | 57540.5029 | 58882.088 | 1173.6418 | 1073.6418 | 1088.9885
7087 | 100.0457 | 1.0194 | 1.0783 | 1.0183 | 198.8068 | 45.5254 | 54829.2357 | 55552.8053 | 1107.0561 | 1007.0561 | 1089.041
7088 | 105.4141 | 0.96739 | 1.0652 | 1.0901 | 204.7585 | 54.6414 | 69223.0571 | 69325.1111 | 1382.5022 | 1282.5022 | 1089.259
7089 | 83.8395 | 1.043 | 1.006 | 1.0165 | 201.499 | 45.2413 | 50424.5153 | 50409.9815 | 1004.1996 | 904.1996 | 1089.2689
7090 | 95.1845 | 0.98865 | 1.3646 | 0.99078 | 187.7405 | 52.8849 | 76664.0153 | 75222.5889 | 1500.4518 | 1400.4518 | 1089.4006
7091 | 102.0245 | 0.94548 | 1.2062 | 1.0697 | 196.6949 | 48.4126 | 67335.1491 | 67254.5406 | 1341.0908 | 1241.0908 | 1089.4046
7092 | 104.2178 | 0.96218 | 1.0822 | 1.0183 | 206.3792 | 48.9798 | 59094.9023 | 60682.9288 | 1209.6586 | 1109.6586 | 1089.4162
7093 | 106.0282 | 0.91967 | 0.84953 | 1.0121 | 206.984 | 53.1209 | 51059.6882 | 51474.0426 | 1025.4809 | 925.4809 | 1089.4426
7094 | 91.3833 | 1.0417 | 1.16 | 1.0163 | 202.4118 | 53.8304 | 68790.2617 | 69941.6656 | 1394.8333 | 1294.8333 | 1089.5766
7095 | 115.2436 | 1.0542 | 0.93756 | 0.98276 | 200.2084 | 55.5949 | 58179.4975 | 57468.3574 | 1145.3671 | 1045.3671 | 1089.5814
7096 | 112.3418 | 0.95982 | 0.77684 | 1.034 | 203.6899 | 50.5976 | 46302.0073 | 46853.895 | 933.0779 | 833.0779 | 1089.7272
7097 | 85.1706 | 1.0274 | 0.84492 | 0.9725 | 202.1399 | 54.7623 | 49990.9871 | 50870.9324 | 1013.4186 | 913.4186 | 1089.7302
7098 | 79.6873 | 1.0329 | 0.88115 | 0.96817 | 176.2692 | 44.5749 | 41872.0811 | 41523.1976 | 826.464 | 726.464 | 1089.7885
7099 | 96.908 | 0.95438 | 0.96019 | 0.98468 | 205.8168 | 46.9775 | 48966.9456 | 49764.9079 | 991.2982 | 891.2982 | 1089.8477
7100 | 105.1123 | 1.0015 | 0.88944 | 0.97997 | 199.0069 | 49.9524 | 48997.2838 | 49322.8273 | 982.4565 | 882.4565 | 1089.8569
7101 | 93.0628 | 1.0224 | 1.0864 | 0.97156 | 190.3192 | 53.0376 | 61218.4721 | 61059.7658 | 1217.1953 | 1117.1953 | 1089.8923
7102 | 110.5199 | 0.95174 | 0.73067 | 0.97493 | 202.186 | 43.4667 | 35737.7714 | 36536.4545 | 726.7291 | 626.7291 | 1089.9879
7103 | 93.6474 | 0.94569 | 1.0471 | 1.0706 | 193.5824 | 46.6757 | 56653.108 | 55735.3277 | 1110.7066 | 1010.7066 | 1090.0529
7104 | 112.1717 | 0.98638 | 0.88362 | 1.0387 | 190.7219 | 44.2735 | 45726.337 | 45075.4624 | 897.5092 | 797.5092 | 1090.084
7105 | 97.0819 | 1.0134 | 0.87599 | 1.0246 | 207.361 | 53.0242 | 53016.2408 | 52761.8332 | 1051.2367 | 951.2367 | 1090.1219
7106 | 97.3185 | 0.99401 | 0.94591 | 1.052 | 186.6219 | 52.2499 | 57234.5966 | 58420.81 | 1164.4162 | 1064.4162 | 1090.1596
7107 | 91.3618 | 0.96344 | 1.1634 | 0.97651 | 174.9018 | 49.8733 | 61223.7512 | 62183.4643 | 1239.6693 | 1139.6693 | 1090.2328
7108 | 110.3057 | 1.0151 | 1.2078 | 0.94127 | 220.6296 | 46.2088 | 57927.217 | 58240.6157 | 1160.8123 | 1060.8123 | 1090.2442
7109 | 104.3671 | 0.91672 | 1.1147 | 1.0416 | 199.1747 | 48.0136 | 60538.831 | 61694.8966 | 1229.8979 | 1129.8979 | 1090.3013
7110 | 109.8999 | 1.0738 | 0.76267 | 1.0866 | 204.97 | 47.8277 | 45483.8802 | 44306.2168 | 882.1243 | 782.1243 | 1090.3204
7111 | 117.0858 | 0.93356 | 1.0919 | 0.96573 | 205.8356 | 51.4071 | 60033.6086 | 60221.0696 | 1200.4214 | 1100.4214 | 1090.4895
7112 | 112.9716 | 0.97798 | 1.2786 | 1.0002 | 201.9953 | 52.0444 | 72511.6284 | 73653.3717 | 1469.0674 | 1369.0674 | 1090.5706
7113 | 101.4898 | 1.0298 | 0.89792 | 1.0142 | 200.674 | 47.7842 | 48711.4688 | 49336.6863 | 982.7337 | 882.7337 | 1090.6118
7114 | 98.7329 | 0.92893 | 0.86658 | 1.0391 | 201.735 | 50.8129 | 50616.0471 | 51114.6119 | 1018.2922 | 918.2922 | 1090.629
7115 | 98.0678 | 0.97946 | 1.1613 | 1.0822 | 203.134 | 49.1835 | 66734.4104 | 65911.6562 | 1314.2331 | 1214.2331 | 1090.7527
7116 | 96.2928 | 0.98921 | 1.1 | 0.96357 | 187.9525 | 50.6218 | 58666.3076 | 57812.3881 | 1152.2478 | 1052.2478 | 1090.7817
7117 | 96.9714 | 1.0064 | 1.1328 | 0.91353 | 207.9618 | 46.581 | 52957.2468 | 53955.0527 | 1075.1011 | 975.1011 | 1090.8465
7118 | 93.6525 | 1.013 | 0.96014 | 1.0235 | 189.5286 | 50.6642 | 54786.2432 | 53750.4813 | 1071.0096 | 971.0096 | 1090.9044
7119 | 98.7203 | 1.0483 | 0.75385 | 0.99132 | 202.9945 | 54.9968 | 46993.9745 | 46975.6579 | 935.5132 | 835.5132 | 1090.9122
7120 | 99.5432 | 0.94912 | 0.97417 | 0.96887 | 181.6359 | 48.0721 | 50095.6422 | 49355.5011 | 983.11 | 883.11 | 1090.9611
7121 | 113.9478 | 1.0606 | 1.0983 | 1.0031 | 206.2894 | 50.8118 | 62328.1271 | 63605.2331 | 1268.1047 | 1168.1047 | 1090.9619
7122 | 104.6169 | 1.0447 | 0.90292 | 0.97777 | 204.0969 | 50.0916 | 49902.2659 | 51032.2377 | 1016.6448 | 916.6448 | 1090.9668
7123 | 108.5463 | 0.96843 | 1.1351 | 1.0812 | 204.0545 | 49.3351 | 65934.552 | 64420.8215 | 1284.4164 | 1184.4164 | 1091.0408
7124 | 102.1312 | 0.94898 | 0.75002 | 1.0169 | 208.3789 | 45.4927 | 39316.2824 | 39422.8086 | 784.4562 | 684.4562 | 1091.0854
7125 | 132.9207 | 1.0964 | 1.0325 | 0.97106 | 182.4265 | 54.2763 | 62512.0189 | 63449.8538 | 1264.9971 | 1164.9971 | 1091.107
7126 | 112.2185 | 0.96872 | 0.90876 | 0.99825 | 196.8611 | 49.1983 | 50176.4482 | 49756.6178 | 991.1324 | 891.1324 | 1091.1265
7127 | 103.8168 | 1.0161 | 1.0147 | 1.0639 | 180.0041 | 48.1734 | 57267.418 | 56156.2081 | 1119.1242 | 1019.1242 | 1091.1877
7128 | 94.1961 | 0.97743 | 1.0743 | 1.0374 | 207.9685 | 51.429 | 62259.516 | 60826.7078 | 1212.5342 | 1112.5342 | 1091.2355
7129 | 90.9843 | 1.0245 | 0.93984 | 1.0235 | 200.2469 | 48.9284 | 51827.1469 | 51910.1226 | 1034.2025 | 934.2025 | 1091.2768
7130 | 93.3442 | 1.0272 | 1.1139 | 1.0255 | 206.2408 | 48.8698 | 60715.2139 | 60873.4471 | 1213.4689 | 1113.4689 | 1091.3833
7131 | 87.2794 | 0.99103 | 0.89403 | 1.0497 | 196.7579 | 48.5763 | 49985.4303 | 48517.312 | 966.3462 | 866.3462 | 1091.3879
7132 | 94.7212 | 0.95625 | 0.98076 | 0.9558 | 192.4477 | 48.5298 | 50080.457 | 49770.7175 | 991.4144 | 891.4144 | 1091.4992
7133 | 103.7984 | 0.95842 | 0.85984 | 1.0624 | 216.305 | 55.5179 | 56455.1301 | 58461.8023 | 1165.236 | 1065.236 | 1091.5003
7134 | 109.494 | 1.0076 | 1.0245 | 1.0134 | 193.5417 | 50.853 | 58597.4722 | 60042.7283 | 1196.8546 | 1096.8546 | 1091.5571
7135 | 86.712 | 1.0738 | 0.97484 | 1.0489 | 198.9235 | 50.2664 | 56274.9386 | 53617.7941 | 1068.3559 | 968.3559 | 1091.6414
7136 | 90.9044 | 0.94577 | 1.2483 | 0.9887 | 200.548 | 50.2204 | 66500.2773 | 66450.1526 | 1325.0031 | 1225.0031 | 1091.7289
7137 | 104.6086 | 0.95179 | 1.0512 | 1.0666 | 195.9979 | 48.3661 | 59240.1537 | 58968.4125 | 1175.3683 | 1075.3683 | 1091.7722
7138 | 97.8005 | 1.0478 | 0.98464 | 1.1173 | 215.0927 | 51.6088 | 62279.8496 | 63477.0201 | 1265.5404 | 1165.5404 | 1091.7937
7139 | 95.3107 | 1.0661 | 1.0456 | 0.99772 | 199.6456 | 49.2966 | 56636.9439 | 53918.7319 | 1074.3746 | 974.3746 | 1091.8131
7140 | 110.2164 | 1.034 | 1.344 | 1.0154 | 199.02 | 50.1211 | 74307.2485 | 73816.7756 | 1472.3355 | 1372.3355 | 1091.82
7141 | 85.8262 | 1.0112 | 0.93842 | 0.98347 | 204.0798 | 49.9775 | 50666.2517 | 50069.2496 | 997.385 | 897.385 | 1091.8254
7142 | 94.106 | 1.0008 | 1.237 | 1.0384 | 202.6034 | 48.6829 | 67323.0415 | 66063.393 | 1317.2679 | 1217.2679 | 1091.8277
7143 | 102.576 | 1.122 | 1.0979 | 1.0858 | 184.437 | 46.2453 | 60638.2884 | 61043.1731 | 1216.8635 | 1116.8635 | 1091.9337
7144 | 123.4305 | 1.0141 | 0.87101 | 1.0124 | 217.2633 | 48.0319 | 48586.153 | 48733.6243 | 970.6725 | 870.6725 | 1092.0471
7145 | 109.5468 | 0.95026 | 1.1017 | 1.0283 | 208.7695 | 48.7797 | 60545.1127 | 60143.1861 | 1198.8637 | 1098.8637 | 1092.2519
7146 | 97.2643 | 1.0265 | 1.0876 | 1.0796 | 200.0303 | 48.4261 | 61896.2645 | 62189.6929 | 1239.7939 | 1139.7939 | 1092.2651
7147 | 103.4279 | 0.97849 | 1.0241 | 1.0828 | 208.5346 | 49.5108 | 60119.4295 | 59163.1904 | 1179.2638 | 1079.2638 | 1092.2778
7148 | 94.5714 | 0.9862 | 1.0853 | 0.99023 | 198.892 | 48.8387 | 57241.0574 | 55504.5632 | 1106.0913 | 1006.0913 | 1092.2843
7149 | 81.6964 | 1.026 | 1.0611 | 0.99024 | 201.1412 | 47.1142 | 53657.3945 | 53877.164 | 1073.5433 | 973.5433 | 1092.312
7150 | 93.7 | 0.97092 | 1.1739 | 1.1035 | 194.5922 | 46.3537 | 64459.1819 | 64974.2172 | 1295.4843 | 1195.4843 | 1092.3884
7151 | 94.3079 | 0.89193 | 0.72059 | 1.0769 | 191.5109 | 46.631 | 40299.815 | 39855.3567 | 793.1071 | 693.1071 | 1092.4831
7152 | 82.7484 | 1.0105 | 1.2465 | 1.0783 | 224.4316 | 46.9707 | 67284.0511 | 67207.6714 | 1340.1534 | 1240.1534 | 1092.4903
7153 | 120.0565 | 1.0166 | 0.9771 | 1.0031 | 195.401 | 45.2048 | 50017.4024 | 49592.7685 | 987.8554 | 887.8554 | 1092.5085
7154 | 109.6941 | 0.9891 | 1.1083 | 0.97999 | 198.0046 | 51.0071 | 61132.3181 | 61287.0184 | 1221.7404 | 1121.7404 | 1092.5691
7155 | 83.8396 | 0.85743 | 0.83057 | 1.011 | 199.1901 | 52.1287 | 47720.5438 | 47957.8611 | 955.1572 | 855.1572 | 1092.6015
7156 | 100.5965 | 0.98578 | 0.99228 | 1.0804 | 198.0499 | 51.4535 | 60463.3707 | 59985.0973 | 1195.7019 | 1095.7019 | 1092.6162
7157 | 105.6454 | 0.98705 | 1.0069 | 1.0348 | 193.3399 | 49.5347 | 56967.5572 | 56490.187 | 1125.8037 | 1025.8037 | 1092.7
7158 | 84.9726 | 1.0014 | 0.76324 | 0.99722 | 198.0609 | 50.5146 | 42944.2423 | 43864.7788 | 873.2956 | 773.2956 | 1092.799
7159 | 85.7014 | 0.97141 | 1.0199 | 0.97804 | 189.7728 | 51.2305 | 55558.8894 | 55384.7627 | 1103.6953 | 1003.6953 | 1092.8154
7160 | 96.195 | 0.99641 | 0.92572 | 0.96047 | 200.3925 | 50.7104 | 50148.7222 | 50956.5521 | 1015.131 | 915.131 | 1092.8528
7161 | 102.729 | 1.0262 | 0.82289 | 0.99339 | 207.2984 | 46.7015 | 43306.8115 | 43730.9143 | 870.6183 | 770.6183 | 1092.9735
7162 | 137.8132 | 0.8747 | 0.98255 | 0.9324 | 214.613 | 46.7923 | 48723.0263 | 48586.0169 | 967.7203 | 867.7203 | 1093.0474
7163 | 109.4819 | 0.97105 | 1.1772 | 0.99423 | 199.2985 | 55.6336 | 71225.8791 | 71142.682 | 1418.8536 | 1318.8536 | 1093.1037
7164 | 109.3062 | 1.1308 | 0.99538 | 1.0317 | 206.2012 | 48.4916 | 55999.282 | 54368.6994 | 1083.374 | 983.374 | 1093.1296
7165 | 96.3997 | 1.064 | 1.1407 | 1.0776 | 184.2948 | 47.5499 | 63511.9757 | 63478.0006 | 1265.56 | 1165.56 | 1093.1399
7166 | 109.4384 | 1.0454 | 1.315 | 0.99241 | 199.4123 | 50.107 | 71320.5986 | 70723.9216 | 1410.4784 | 1310.4784 | 1093.1691
7167 | 110.9455 | 0.98708 | 1.0442 | 0.93493 | 201.3174 | 53.0303 | 57780.2025 | 58886.625 | 1173.7325 | 1073.7325 | 1093.1781
7168 | 109.7075 | 0.96808 | 1.2074 | 0.90587 | 198.5491 | 49.9113 | 60088.3782 | 60041.0416 | 1196.8208 | 1096.8208 | 1093.1878
7169 | 89.5921 | 0.97875 | 0.81027 | 0.9939 | 190.4116 | 55.283 | 49559.0344 | 51792.6742 | 1031.8535 | 931.8535 | 1093.2261
7170 | 103.3978 | 1.0304 | 0.99044 | 1.0692 | 210.3201 | 51.2801 | 59977.6464 | 58025.7449 | 1156.5149 | 1056.5149 | 1093.3153
7171 | 80.6861 | 1.0182 | 0.93464 | 1.039 | 214.2927 | 48.24 | 51023.1237 | 50568.408 | 1007.3682 | 907.3682 | 1093.3515
7172 | 103.0716 | 0.97659 | 1.0788 | 0.97058 | 214.1523 | 47.9587 | 55257.1358 | 56289.9228 | 1121.7985 | 1021.7985 | 1093.4393
7173 | 111.4064 | 1.0441 | 1.0271 | 1.0554 | 219.4666 | 47.7229 | 57502.3868 | 56027.0031 | 1116.5401 | 1016.5401 | 1093.528
7174 | 98.8277 | 0.98357 | 0.94409 | 1.0554 | 216.3519 | 50.2808 | 55203.3075 | 54231.0184 | 1080.6204 | 980.6204 | 1093.6539
7175 | 91.9906 | 0.91708 | 0.94812 | 1.0144 | 191.3622 | 50.0559 | 52556.8454 | 51978.3689 | 1035.5674 | 935.5674 | 1093.6817
7176 | 91.5036 | 0.91388 | 1.2273 | 0.98977 | 195.356 | 48.033 | 62559.1629 | 62051.5443 | 1237.0309 | 1137.0309 | 1093.8662
7177 | 91.6441 | 1.0524 | 1.1328 | 0.97382 | 199.3403 | 49.0932 | 59089.4899 | 56529.6867 | 1126.5937 | 1026.5937 | 1093.9235
7178 | 88.4138 | 1.0088 | 1.0821 | 0.94516 | 206.2791 | 51.8265 | 57832.5995 | 58377.7571 | 1163.5551 | 1063.5551 | 1093.9744
7179 | 79.3742 | 1.0336 | 0.90643 | 0.95498 | 209.7161 | 50.5547 | 48118.257 | 49077.9922 | 977.5598 | 877.5598 | 1094.0637
7180 | 89.27 | 0.90388 | 0.75222 | 1.027 | 194.6782 | 51.0999 | 43795.8727 | 44534.835 | 886.6967 | 786.6967 | 1094.2046
7181 | 109.7893 | 1.0381 | 0.82017 | 0.99112 | 206.746 | 47.1682 | 43925.1677 | 45072.4419 | 897.4488 | 797.4488 | 1094.2703
7182 | 103.0545 | 1.0415 | 0.92173 | 1.0346 | 190.1735 | 47.839 | 50944.1896 | 50727.9421 | 1010.5588 | 910.5588 | 1094.3491
7183 | 91.6686 | 1.0278 | 0.92476 | 1.0426 | 197.1753 | 52.6077 | 55874.7619 | 55592.4137 | 1107.8483 | 1007.8483 | 1094.5427
7184 | 84.9913 | 0.99921 | 1.0031 | 1.0804 | 209.0407 | 50.7073 | 59466.3556 | 57763.8867 | 1151.2777 | 1051.2777 | 1094.553
7185 | 103.5435 | 1.0315 | 0.82593 | 0.98397 | 210.7534 | 50.487 | 46633.6822 | 45947.0362 | 914.9407 | 814.9407 | 1094.5822
7186 | 99.5705 | 0.9471 | 1.0327 | 1.0295 | 201.4628 | 53.2248 | 61803.0866 | 61543.806 | 1226.8761 | 1126.8761 | 1094.6078
7187 | 95.174 | 0.97429 | 0.81157 | 1.0599 | 198.7659 | 53.7472 | 51414.6597 | 52352.1878 | 1043.0438 | 943.0438 | 1094.7226
7188 | 92.8082 | 1.0069 | 0.9327 | 0.94051 | 196.6805 | 52.8882 | 51533.3792 | 52521.303 | 1046.4261 | 946.4261 | 1094.755
7189 | 97.3553 | 1.0149 | 0.99433 | 0.97918 | 196.8912 | 51.2197 | 55126.4399 | 54393.4281 | 1083.8686 | 983.8686 | 1094.8354
7190 | 99.8824 | 1.0655 | 0.99094 | 0.95239 | 193.2965 | 51.2517 | 54017.1567 | 52019.4881 | 1036.3898 | 936.3898 | 1094.8682
7191 | 101.0576 | 0.92843 | 0.94178 | 1.04 | 204.0301 | 53.8344 | 57983.2641 | 56608.5582 | 1128.1712 | 1028.1712 | 1094.9044
7192 | 80.9767 | 1.0595 | 1.1258 | 1.0373 | 207.9049 | 47.8804 | 60228.8409 | 57708.9775 | 1150.1795 | 1050.1795 | 1094.9116
7193 | 110.147 | 1.0672 | 0.99189 | 1.0469 | 184.8708 | 51.2891 | 59473.3467 | 57147.7099 | 1138.9542 | 1038.9542 | 1094.9523
7194 | 102.1296 | 0.96122 | 1.1609 | 1.0146 | 189.4514 | 50.838 | 65062.5976 | 65362.8212 | 1303.2564 | 1203.2564 | 1095.0375
7195 | 80.8346 | 1.0007 | 1.0437 | 1.035 | 197.0642 | 49.1915 | 57316.6866 | 58497.4405 | 1165.9488 | 1065.9488 | 1095.0383
7196 | 96.6388 | 0.97337 | 0.96341 | 0.98478 | 207.8171 | 48.2127 | 50484.7036 | 51140.4819 | 1018.8096 | 918.8096 | 1095.1076
7197 | 97.7686 | 1.0145 | 1.1184 | 0.9998 | 191.3623 | 46.3761 | 56648.2876 | 57299.4349 | 1141.9887 | 1041.9887 | 1095.1216
7198 | 99.693 | 1.0023 | 0.85692 | 0.89953 | 210.0698 | 49.7077 | 43492.5895 | 43231.1554 | 860.6231 | 760.6231 | 1095.1374
7199 | 87.7589 | 1.0162 | 1.0712 | 0.97976 | 203.1635 | 51.0762 | 58363.5024 | 58681.4204 | 1169.6284 | 1069.6284 | 1095.1429
7200 | 102.5758 | 1.0911 | 0.79691 | 0.99476 | 183.4064 | 50.8343 | 46171.1364 | 45244.0531 | 900.8811 | 800.8811 | 1095.1657
7201 | 96.5714 | 0.96421 | 1.0266 | 0.99184 | 177.5439 | 51.9609 | 57921.3846 | 56714.3954 | 1130.2879 | 1030.2879 | 1095.1727
7202 | 82.3935 | 1.0466 | 1.082 | 0.98183 | 183.4433 | 51.9065 | 59801.3611 | 61664.1115 | 1229.2822 | 1129.2822 | 1095.1803
7203 | 102.0499 | 0.93715 | 0.90652 | 0.96842 | 192.1903 | 52.769 | 51564.1398 | 52355.8215 | 1043.1164 | 943.1164 | 1095.1849
7204 | 115.304 | 0.95815 | 1.0032 | 0.99988 | 186.7665 | 49.6476 | 55471.3136 | 54533.6269 | 1086.6725 | 986.6725 | 1095.2288
7205 | 90.4935 | 1.0413 | 1.0574 | 1.0424 | 191.3838 | 49.1976 | 59054.1065 | 61112.832 | 1218.2566 | 1118.2566 | 1095.2483
7206 | 99.1554 | 0.9912 | 0.92933 | 1.0078 | 214.6317 | 47.6427 | 49519.8964 | 49577.5335 | 987.5507 | 887.5507 | 1095.2734
7207 | 93.6885 | 1.0095 | 0.85409 | 0.96265 | 200.6838 | 54.0234 | 49727.8018 | 48873.7655 | 973.4753 | 873.4753 | 1095.3135
7208 | 82.493 | 1.0296 | 0.67236 | 1.1142 | 213.4226 | 52.919 | 44350.1438 | 43467.3641 | 865.3473 | 765.3473 | 1095.3484
7209 | 91.2831 | 1.0111 | 1.1168 | 0.99794 | 203.263 | 46.6208 | 56466.397 | 57253.3255 | 1141.0665 | 1041.0665 | 1095.398
7210 | 91.3942 | 0.83655 | 1.0542 | 0.92033 | 208.2413 | 52.9523 | 55630.9503 | 56358.0189 | 1123.1604 | 1023.1604 | 1095.4911
7211 | 100.1467 | 0.98251 | 1.073 | 0.98621 | 183.6146 | 49.9176 | 57918.3818 | 57038.8148 | 1136.7763 | 1036.7763 | 1095.539
7212 | 105.9752 | 0.99923 | 1.1748 | 0.9363 | 202.2659 | 47.3612 | 57315.1243 | 56370.9755 | 1123.4195 | 1023.4195 | 1095.5471
7213 | 97.0851 | 1.1224 | 1.1302 | 1.0974 | 198.4203 | 54.2764 | 73432.9292 | 71957.4345 | 1435.1487 | 1335.1487 | 1095.5594
7214 | 109.6925 | 0.91348 | 1.2324 | 1.0011 | 199.163 | 48.1088 | 64375.8364 | 64813.4239 | 1292.2685 | 1192.2685 | 1095.5825
7215 | 110.2557 | 1.0107 | 0.82144 | 1.0443 | 194.5327 | 47.0419 | 45788.9678 | 45345.9204 | 902.9184 | 802.9184 | 1095.6095
7216 | 110.0541 | 1.0171 | 1.0348 | 1.0899 | 210.9913 | 51.9121 | 64571.1795 | 62968.3756 | 1255.3675 | 1155.3675 | 1095.6172
7217 | 105.7075 | 0.97173 | 0.94787 | 1.0134 | 194.7129 | 48.7178 | 51995.168 | 51969.6989 | 1035.394 | 935.394 | 1095.6741
7218 | 102.4927 | 0.99608 | 0.86549 | 0.9744 | 197.2112 | 53.2067 | 50500.2913 | 50816.0022 | 1012.32 | 912.32 | 1095.6898
7219 | 98.6373 | 1.0923 | 1.073 | 1.0177 | 203.1007 | 50.8131 | 61162.3434 | 63298.3931 | 1261.9679 | 1161.9679 | 1095.6909
7220 | 97.9643 | 0.95751 | 0.67486 | 1.0713 | 210.5933 | 46.1754 | 37927.2076 | 37267.397 | 741.3479 | 641.3479 | 1095.7019
7221 | 92.6853 | 1.0506 | 0.77636 | 1.0151 | 208.6983 | 53.0948 | 47222.5352 | 46465.4495 | 925.309 | 825.309 | 1095.7827
7222 | 90.7851 | 1.0723 | 0.9058 | 0.9967 | 229.2086 | 47.789 | 48025.8037 | 49046.7106 | 976.9342 | 876.9342 | 1095.8058
7223 | 92.0007 | 0.95167 | 1.175 | 1.1101 | 198.6358 | 47.4408 | 66234.0069 | 64826.6115 | 1292.5322 | 1192.5322 | 1095.9309
7224 | 111.4061 | 0.97089 | 1.1052 | 1.0305 | 228.2117 | 50.1488 | 62767.3316 | 62801.2485 | 1252.025 | 1152.025 | 1096.0091
7225 | 95.6382 | 0.95721 | 0.92317 | 1.0006 | 198.2028 | 51.4593 | 52442.2617 | 51501.3548 | 1026.0271 | 926.0271 | 1096.0618
7226 | 97.1637 | 1.0418 | 0.79369 | 1.05 | 183.0229 | 51.3886 | 48210.1553 | 48249.219 | 960.9844 | 860.9844 | 1096.1111
7227 | 106.0855 | 0.93592 | 0.9609 | 1.1447 | 197.6284 | 47.3181 | 56943.7597 | 58450.832 | 1165.0166 | 1065.0166 | 1096.1996
7228 | 102.231 | 0.98024 | 1.0139 | 1.0539 | 203.5048 | 50.966 | 59771.3563 | 60647.3407 | 1208.9468 | 1108.9468 | 1096.2709
7229 | 87.8243 | 1.0572 | 0.94484 | 1.1054 | 208.8864 | 55.0599 | 62826.4926 | 61707.0092 | 1230.1402 | 1130.1402 | 1096.3686
7230 | 101.224 | 1.0625 | 0.95255 | 0.97486 | 212.934 | 51.2403 | 53305.9409 | 51205.3858 | 1020.1077 | 920.1077 | 1096.4062
7231 | 114.2529 | 0.98135 | 1.0007 | 1.0596 | 198.67 | 51.9046 | 61053.4951 | 60865.0789 | 1213.3016 | 1113.3016 | 1096.5243
7232 | 105.0745 | 0.97981 | 0.8204 | 0.91266 | 183.0935 | 50.8642 | 43504.1148 | 43350.5912 | 863.0118 | 763.0118 | 1096.5294
7233 | 97.6377 | 1.0201 | 1.2363 | 0.9136 | 195.3761 | 51.5955 | 63612.0522 | 65154.3063 | 1299.0861 | 1199.0861 | 1096.5586
7234 | 107.9423 | 0.93637 | 1.228 | 1.0548 | 195.3609 | 49.1129 | 68774.9639 | 69892.2959 | 1393.8459 | 1293.8459 | 1096.5619
7235 | 90.5104 | 0.91251 | 1.0125 | 0.97895 | 193.9632 | 50.7127 | 54650.4076 | 54999.9995 | 1096 | 996 | 1096.5992
7236 | 101.9295 | 0.90561 | 0.98669 | 0.96809 | 198.8969 | 45.2367 | 47584.779 | 47149.6899 | 938.9938 | 838.9938 | 1096.6551
7237 | 106.3135 | 1.1064 | 0.85319 | 1.0139 | 192.0115 | 49.8883 | 49215.3633 | 48803.0816 | 972.0616 | 872.0616 | 1096.7352
7238 | 100.6048 | 1.0808 | 1.0894 | 1.0877 | 190.1575 | 49.3203 | 63995.173 | 63724.8718 | 1270.4974 | 1170.4974 | 1096.7557
7239 | 92.2006 | 0.98077 | 1.182 | 0.99463 | 196.9719 | 49.144 | 62418.9579 | 62526.6919 | 1246.5338 | 1146.5338 | 1096.7697
7240 | 84.5802 | 0.96801 | 0.62863 | 1.0233 | 196.4982 | 49.1462 | 35835.6999 | 37338.9844 | 742.7797 | 642.7797 | 1096.8038
7241 | 86.7754 | 1.0982 | 1.2033 | 0.93368 | 216.9103 | 51.0812 | 62472.4455 | 63210.7286 | 1260.2146 | 1160.2146 | 1096.8208
7242 | 122.3076 | 0.97042 | 0.89438 | 1.0548 | 208.2921 | 52.8944 | 56385.8153 | 56713.0742 | 1130.2615 | 1030.2615 | 1096.8393
7243 | 90.8637 | 0.99714 | 1.0132 | 0.94171 | 184.8298 | 50.9642 | 53426.8983 | 53774.8594 | 1071.4972 | 971.4972 | 1096.8405
7244 | 85.8941 | 0.99066 | 0.97468 | 1.0481 | 201.0244 | 49.4771 | 54954.2708 | 55469.3496 | 1105.387 | 1005.387 | 1096.8546
7245 | 105.9414 | 0.97456 | 0.96713 | 1.0039 | 205.9736 | 49.1195 | 52968.5222 | 53554.2584 | 1067.0852 | 967.0852 | 1096.8588
7246 | 81.2299 | 0.96251 | 0.97858 | 0.99407 | 189.5445 | 51.8458 | 54677.4242 | 53256.5326 | 1061.1307 | 961.1307 | 1096.9743
7247 | 120.4833 | 1.0757 | 0.99798 | 0.93296 | 198.1769 | 51.3099 | 54621.6705 | 54803.4005 | 1092.068 | 992.068 | 1096.9817
7248 | 93.4116 | 1.0069 | 1.1075 | 0.9931 | 191.4449 | 51.0793 | 61178.1944 | 61159.1736 | 1219.1835 | 1119.1835 | 1096.9863
7249 | 88.7166 | 0.92149 | 1.087 | 1.0223 | 203.8318 | 49.3341 | 59059.1323 | 59917.4569 | 1194.3491 | 1094.3491 | 1097.0091
7250 | 96.2644 | 0.98653 | 1.0733 | 0.90905 | 205.7063 | 52.1651 | 56055.6593 | 57435.1634 | 1144.7033 | 1044.7033 | 1097.0104
7251 | 94.8522 | 0.96056 | 0.80913 | 0.9841 | 212.7775 | 46.5597 | 41528.5948 | 40331.7049 | 802.6341 | 702.6341 | 1097.0156
7252 | 99.9445 | 0.88192 | 0.9823 | 1.014 | 187.5259 | 46.8087 | 50935.2769 | 50608.0286 | 1008.1606 | 908.1606 | 1097.066
7253 | 113.897 | 1.0384 | 1.014 | 1.0274 | 204.7224 | 48.0081 | 55896.2721 | 54748.9862 | 1090.9797 | 990.9797 | 1097.0974
7254 | 99.3574 | 1.0318 | 1.0208 | 1.0283 | 190.1624 | 47.5282 | 54947.8322 | 55969.759 | 1115.3952 | 1015.3952 | 1097.1053
7255 | 123.9969 | 0.92944 | 0.82075 | 1.077 | 184.7546 | 47.9691 | 48114.0223 | 46359.2574 | 923.1851 | 823.1851 | 1097.2063
7256 | 91.3598 | 0.96768 | 1.1299 | 1.0767 | 193.2874 | 50.8991 | 66617.0906 | 63934.8924 | 1274.6978 | 1174.6978 | 1097.2451
7257 | 113.214 | 0.96493 | 0.94259 | 0.97455 | 200.0321 | 52.1391 | 53790.7597 | 53890.5293 | 1073.8106 | 973.8106 | 1097.2569
7258 | 100.1421 | 0.99356 | 0.99064 | 0.96827 | 207.862 | 49.8707 | 53006.454 | 54117.351 | 1078.347 | 978.347 | 1097.3299
7259 | 103.9708 | 0.96745 | 0.97715 | 0.92235 | 182.8507 | 50.4751 | 50752.121 | 49947.7967 | 994.9559 | 894.9559 | 1097.353
7260 | 100.4618 | 0.99355 | 1.0009 | 1.0376 | 200.3928 | 47.9414 | 54773.9626 | 54345.1453 | 1082.9029 | 982.9029 | 1097.4138
7261 | 89.6177 | 1.0052 | 1.0005 | 1.0711 | 192.316 | 49.6156 | 57831.1641 | 56476.3512 | 1125.527 | 1025.527 | 1097.4449
7262 | 104.5823 | 1.0044 | 1.0654 | 0.95507 | 198.1505 | 50.8092 | 57234.0358 | 58623.463 | 1168.4693 | 1068.4693 | 1097.5222
7263 | 109.8448 | 0.95001 | 1.0581 | 1.0006 | 202.7949 | 52.821 | 61635.3447 | 62392.3692 | 1243.8474 | 1143.8474 | 1097.5508
7264 | 103.0999 | 1.0573 | 1.0803 | 1.0416 | 211.7436 | 53.2873 | 65978.601 | 66473.2424 | 1325.4648 | 1225.4648 | 1097.5518
7265 | 104.2894 | 0.97778 | 1.2477 | 1.0624 | 219.0426 | 50.9907 | 73011.1449 | 71763.0539 | 1431.2611 | 1331.2611 | 1097.5522
7266 | 91.2118 | 1.0641 | 1.0078 | 1.0742 | 204.9171 | 51.1616 | 60555.9925 | 62333.7215 | 1242.6744 | 1142.6744 | 1097.62
7267 | 90.6287 | 0.96572 | 0.67196 | 0.89514 | 200.4019 | 48.4803 | 33604.4611 | 33680.4164 | 669.6083 | 569.6083 | 1097.7315
7268 | 90.2228 | 1.0166 | 1.3032 | 0.99505 | 194.5501 | 51.9367 | 72306.3241 | 71695.2609 | 1429.9052 | 1329.9052 | 1097.9361
7269 | 104.9 | 1.0284 | 1.0697 | 1.0303 | 199.9464 | 54.4202 | 66048.55 | 64302.2983 | 1282.046 | 1182.046 | 1097.9521
7270 | 102.5073 | 1.008 | 0.83885 | 1.0123 | 219.1524 | 48.8234 | 46723.5547 | 45967.5637 | 915.3513 | 815.3513 | 1097.9627
7271 | 117.288 | 1.0036 | 1.0311 | 1.0451 | 190.7452 | 50.3406 | 60363.6621 | 58938.0688 | 1174.7614 | 1074.7614 | 1097.974
7272 | 98.1939 | 1.0181 | 0.95403 | 0.99408 | 193.9961 | 51.9575 | 54664.0361 | 55417.7637 | 1104.3553 | 1004.3553 | 1097.9941
7273 | 102.0115 | 0.99161 | 0.82942 | 1.0172 | 214.3021 | 52.4438 | 49767.1703 | 49109.9457 | 978.1989 | 878.1989 | 1098.0223
7274 | 95.4386 | 1.0541 | 0.83353 | 0.9853 | 199.3184 | 48.5529 | 44959.1494 | 44813.2906 | 892.2658 | 792.2658 | 1098.0609
7275 | 81.0094 | 0.97464 | 0.9519 | 0.96779 | 206.3098 | 58.1264 | 58343.9858 | 57814.3873 | 1152.2877 | 1052.2877 | 1098.127
7276 | 104.1897 | 1.0824 | 0.95757 | 0.97552 | 200.0247 | 48.6726 | 51155.7877 | 49555.6306 | 987.1126 | 887.1126 | 1098.3058
7277 | 109.5672 | 0.95612 | 1.1883 | 0.88127 | 205 | 46.4498 | 53712.964 | 54421.4527 | 1084.4291 | 984.4291 | 1098.5254
7278 | 95.2947 | 0.97805 | 1.1912 | 1.0309 | 199.9073 | 50.4462 | 66846.8681 | 67730.4235 | 1350.6085 | 1250.6085 | 1098.6005
7279 | 91.4717 | 1.0136 | 0.89952 | 0.95637 | 206.2506 | 50.0608 | 47913.839 | 50214.069 | 1000.2814 | 900.2814 | 1098.6707
7280 | 84.807 | 0.96524 | 1.0081 | 0.95108 | 212.7112 | 49.9758 | 52222.0203 | 53182.6474 | 1059.6529 | 959.6529 | 1098.6808
7281 | 109.8747 | 0.95784 | 0.84041 | 0.9946 | 189.7871 | 48.6528 | 45977.6255 | 46251.1126 | 921.0223 | 821.0223 | 1098.6821
7282 | 95.988 | 0.98633 | 0.8187 | 1.0482 | 200.3293 | 51.5618 | 49331.2869 | 49972.3529 | 995.4471 | 895.4471 | 1098.7001
7283 | 97.2731 | 0.94661 | 0.83788 | 1.0321 | 182.7512 | 53.6333 | 51504.1041 | 52290.9429 | 1041.8189 | 941.8189 | 1098.8637
7284 | 85.8399 | 1.0699 | 0.89663 | 1.0289 | 201.2236 | 50.3634 | 51290.7245 | 51114.4954 | 1018.2899 | 918.2899 | 1098.9412
7285 | 105.1736 | 1.0297 | 1.0606 | 0.99062 | 206.119 | 50.8557 | 59144.8031 | 60263.4238 | 1201.2685 | 1101.2685 | 1098.9414
7286 | 109.8158 | 0.88909 | 1.081 | 0.89786 | 187.8755 | 50.2352 | 53850.1848 | 53375.3299 | 1063.5066 | 963.5066 | 1098.9794
7287 | 107.1462 | 0.94353 | 1.2308 | 1.0691 | 208.6677 | 44.9305 | 63873.8763 | 63588.9027 | 1267.7781 | 1167.7781 | 1098.9805
7288 | 110.1622 | 1.018 | 1.1367 | 0.99634 | 205.2747 | 48.5002 | 60573.6041 | 60009.9801 | 1196.1996 | 1096.1996 | 1099.0165
7289 | 82.6231 | 0.99107 | 1.1832 | 1.0428 | 208.0933 | 50.8034 | 67053.3481 | 66962.4099 | 1335.2482 | 1235.2482 | 1099.3346
7290 | 106.5984 | 0.94383 | 1.1135 | 1.0591 | 224.8148 | 52.1933 | 67027.5606 | 65771.0107 | 1311.4202 | 1211.4202 | 1099.3477
7291 | 102.9254 | 1.0304 | 1.1119 | 0.93165 | 195.3407 | 47.1945 | 54090.2176 | 55448.2883 | 1104.9658 | 1004.9658 | 1099.3508
7292 | 88.1823 | 1.0019 | 0.92082 | 0.99555 | 195.3452 | 52.8093 | 53272.5061 | 52785.4349 | 1051.7087 | 951.7087 | 1099.3573
7293 | 85.0783 | 1.0049 | 1.1457 | 0.91767 | 204.0944 | 47.6479 | 54373.9895 | 53803.8008 | 1072.076 | 972.076 | 1099.3616
7294 | 108.6338 | 0.98223 | 0.83893 | 1.0021 | 177.8883 | 49.3263 | 46909.6906 | 47110.8261 | 938.2165 | 838.2165 | 1099.3719
7295 | 113.091 | 0.90487 | 0.9677 | 1.0069 | 204.0449 | 49.334 | 53322.5358 | 54040.2502 | 1076.805 | 976.805 | 1099.401
7296 | 98.2107 | 1.0179 | 1.2145 | 0.97434 | 200.4907 | 51.8264 | 66708.0958 | 65421.4084 | 1304.4282 | 1204.4282 | 1099.426
7297 | 97.5844 | 0.951 | 1.1583 | 1.0324 | 220.4338 | 48.612 | 62864.0013 | 61329.4783 | 1222.5896 | 1122.5896 | 1099.589
7298 | 96.3106 | 1.0421 | 0.81513 | 1.0282 | 219.3587 | 50.4976 | 47611.9735 | 46710.4766 | 930.2095 | 830.2095 | 1099.7291
7299 | 110.4398 | 0.97184 | 1.1255 | 1.0391 | 205.094 | 46.3151 | 59340.0982 | 58742.7346 | 1170.8547 | 1070.8547 | 1099.792
7300 | 97.3152 | 1.0129 | 0.80405 | 0.96736 | 197.3273 | 50.9694 | 44865.3209 | 44961.5967 | 895.2319 | 795.2319 | 1099.831
7301 | 101.4621 | 0.90582 | 1.1316 | 0.95858 | 201.0771 | 49.739 | 58727.2336 | 58292.5764 | 1161.8515 | 1061.8515 | 1099.8316
7302 | 111.0733 | 0.98592 | 1.1118 | 1.0146 | 202.6326 | 49.4004 | 61339.019 | 62767.6237 | 1251.3525 | 1151.3525 | 1099.8584
7303 | 97.624 | 0.91365 | 0.91447 | 0.93166 | 180.7253 | 50.7288 | 47925.2383 | 48370.1183 | 963.4024 | 863.4024 | 1099.9409
7304 | 88.7114 | 0.93478 | 0.81586 | 1.0201 | 206.4792 | 45.2344 | 41604.1933 | 40812.7498 | 812.255 | 712.255 | 1099.9585
7305 | 90.6847 | 0.99897 | 0.9937 | 0.88656 | 193.0112 | 44.8174 | 43736.1545 | 42559.8448 | 847.1969 | 747.1969 | 1100.0077
7306 | 95.7308 | 0.98068 | 0.93344 | 1.0644 | 190.839 | 51.4098 | 56094.8598 | 54684.2625 | 1089.6852 | 989.6852 | 1100.0392
7307 | 103.4029 | 1.0403 | 1.1511 | 1.0337 | 210.3992 | 48.7831 | 63506.068 | 64553.9876 | 1287.0798 | 1187.0798 | 1100.0677
7308 | 96.0611 | 1.0248 | 0.73883 | 1.0443 | 181.8073 | 45.692 | 39934.6054 | 39229.7125 | 780.5943 | 680.5943 | 1100.1118
7309 | 107.1347 | 0.92465 | 0.90609 | 1.053 | 198.1056 | 52.1649 | 55135.125 | 54648.5272 | 1088.9705 | 988.9705 | 1100.1151
7310 | 103.0358 | 1.0126 | 0.76052 | 1.0403 | 192.9231 | 45.7136 | 41129.6979 | 40460.143 | 805.2029 | 705.2029 | 1100.2436
7311 | 97.84 | 0.98374 | 0.86649 | 1.089 | 198.2568 | 50.8287 | 53051.8209 | 52525.184 | 1046.5037 | 946.5037 | 1100.3539
7312 | 101.2778 | 0.92881 | 0.83978 | 1 | 217.1556 | 48.6343 | 45634.361 | 45224.6442 | 900.4929 | 800.4929 | 1100.4127
7313 | 95.5812 | 0.94757 | 0.65653 | 0.97633 | 225.124 | 50.7524 | 37353.3028 | 36904.3286 | 734.0866 | 634.0866 | 1100.4214
7314 | 106.4064 | 1.0065 | 1.0425 | 1.0692 | 204.7897 | 52.6165 | 64489.8834 | 63713.8403 | 1270.2768 | 1170.2768 | 1100.5191
7315 | 94.1987 | 0.96995 | 0.83165 | 0.99441 | 192.3778 | 51.6532 | 47629.1825 | 47411.9546 | 944.2391 | 844.2391 | 1100.6349
7316 | 102.6786 | 0.98174 | 1.1358 | 1.0375 | 193.5228 | 52.4163 | 67246.0067 | 66154.0385 | 1319.0808 | 1219.0808 | 1100.6692
7317 | 97.77 | 1.0896 | 0.90779 | 0.99579 | 216.9407 | 53.4255 | 54203.455 | 52504.3196 | 1046.0864 | 946.0864 | 1100.7218
7318 | 121.5957 | 0.99629 | 0.9844 | 1.0431 | 192.8768 | 48.5987 | 55980.5024 | 55740.3237 | 1110.8065 | 1010.8065 | 1100.8066
7319 | 116.0809 | 0.99231 | 1.1239 | 1.1091 | 199.3821 | 46.9586 | 64144.9242 | 65233.1287 | 1300.6626 | 1200.6626 | 1100.824
7320 | 103.9622 | 1.0219 | 0.85784 | 1.0392 | 188.541 | 47.3952 | 47473.488 | 45978.5807 | 915.5716 | 815.5716 | 1100.8408
7321 | 112.7962 | 0.96803 | 1.2504 | 0.92884 | 218.1703 | 51.4148 | 65544.889 | 64919.9081 | 1294.3982 | 1194.3982 | 1100.9452
7322 | 97.1242 | 1.0513 | 1.0425 | 1.0217 | 195.6049 | 48.811 | 57168.7982 | 58370.064 | 1163.4013 | 1063.4013 | 1100.9467
7323 | 88.3513 | 1.1019 | 1.0589 | 1.0411 | 208.3937 | 51.9811 | 62574.6079 | 62306.3875 | 1242.1277 | 1142.1277 | 1100.9791
7324 | 99.0757 | 0.99156 | 1.1212 | 1.0009 | 194.2799 | 51.2668 | 62761.7229 | 60098.1349 | 1197.9627 | 1097.9627 | 1101.001
7325 | 98.6898 | 1.049 | 0.74287 | 1.0311 | 200.483 | 53.9039 | 47071.8777 | 47000.2817 | 936.0056 | 836.0056 | 1101.0879
7326 | 114.4477 | 0.96694 | 1.0489 | 0.90045 | 187.1575 | 49.3082 | 52215.6239 | 52426.6453 | 1044.5329 | 944.5329 | 1101.0902
7327 | 83.0624 | 0.96013 | 1.0671 | 1.0254 | 206.1146 | 54.3304 | 63987.2696 | 63869.4177 | 1273.3884 | 1173.3884 | 1101.1306
7328 | 94.6006 | 0.96957 | 0.82777 | 1.0395 | 210.0155 | 55.3618 | 52926.1111 | 53223.3399 | 1060.4668 | 960.4668 | 1101.2247
7329 | 124.1793 | 1.0143 | 1.1157 | 0.98703 | 205.4106 | 50.5338 | 62218.4724 | 62283.1225 | 1241.6625 | 1141.6625 | 1101.2333
7330 | 111.869 | 1.0242 | 0.83835 | 1.0227 | 201.7762 | 52.7385 | 51461.2175 | 51608.4184 | 1028.1684 | 928.1684 | 1101.2656
7331 | 100.2773 | 1.0393 | 1.1757 | 1.0117 | 189.8238 | 53.5436 | 69456.285 | 70674.1429 | 1409.4829 | 1309.4829 | 1101.2685
7332 | 102.4097 | 0.98428 | 1.0562 | 0.97227 | 204.3649 | 47.5822 | 53861.6051 | 56750.672 | 1131.0134 | 1031.0134 | 1101.2774
7333 | 88.5098 | 1.0907 | 0.79924 | 1.052 | 196.9475 | 54.7891 | 51554.3299 | 51355.3991 | 1023.108 | 923.108 | 1101.4071
7334 | 104.8429 | 1.0549 | 1.1071 | 0.93891 | 195.5636 | 49.2405 | 56826.8715 | 55660.2832 | 1109.2057 | 1009.2057 | 1101.4221
7335 | 94.7425 | 0.99111 | 1.1836 | 0.98236 | 208.6122 | 51.6656 | 65133.3847 | 63903.49 | 1274.0698 | 1174.0698 | 1101.4376
7336 | 91.2611 | 0.93906 | 0.80272 | 0.98902 | 202.5073 | 49.1614 | 43444.8519 | 43081.8881 | 857.6378 | 757.6378 | 1101.499
7337 | 105.7033 | 1.0529 | 1.0058 | 1.0147 | 186.0869 | 50.9305 | 57834.2532 | 60029.9579 | 1196.5992 | 1096.5992 | 1101.5884
7338 | 104.4044 | 0.9635 | 1.2254 | 0.89808 | 199.4297 | 49.0192 | 59077.6817 | 59409.8984 | 1184.198 | 1084.198 | 1101.6531
7339 | 86.1478 | 1.0773 | 1.1739 | 0.97574 | 200.2504 | 51.8784 | 64437.688 | 63301.4261 | 1262.0285 | 1162.0285 | 1101.7234
7340 | 98.55 | 1.0126 | 0.97768 | 1.067 | 193.54 | 55.7873 | 63955.2725 | 65393.6425 | 1303.8729 | 1203.8729 | 1101.8394
7341 | 98.875 | 0.97798 | 1.3935 | 1.0157 | 194.6328 | 49.0179 | 74309.779 | 74196.7616 | 1479.9352 | 1379.9352 | 1101.8949
7342 | 92.985 | 0.93168 | 1.1705 | 1.0075 | 182.609 | 48.0025 | 60946.2871 | 61804.3443 | 1232.0869 | 1132.0869 | 1101.9261
7343 | 103.4034 | 0.98135 | 1.0922 | 1.0276 | 197.9193 | 45.1633 | 55469.7845 | 55480.9617 | 1105.6192 | 1005.6192 | 1101.9755
7344 | 99.0477 | 0.99906 | 1.0542 | 0.95212 | 203.3517 | 50.796 | 56213.7941 | 55397.8101 | 1103.9562 | 1003.9562 | 1102.03
7345 | 114.6004 | 1.0014 | 1.0102 | 0.93226 | 200.3474 | 51.9843 | 55121.5124 | 56552.0305 | 1127.0406 | 1027.0406 | 1102.1361
7346 | 87.2955 | 0.99562 | 1.1972 | 1.0551 | 197.0467 | 53.7332 | 72740.8401 | 71541.3978 | 1426.828 | 1326.828 | 1102.2319
7347 | 97.6995 | 1.0049 | 1.1053 | 1.0473 | 197.4368 | 55.2131 | 69531.5912 | 68886.8927 | 1373.7379 | 1273.7379 | 1102.2516
7348 | 91.3466 | 0.99102 | 0.72449 | 1.0708 | 185.7272 | 52.635 | 45782.4161 | 46443.7119 | 924.8742 | 824.8742 | 1102.367
7349 | 95.8182 | 0.92995 | 1.1093 | 0.95824 | 207.7936 | 52.8498 | 61094.5144 | 61201.0864 | 1220.0217 | 1120.0217 | 1102.4018
7350 | 101.7456 | 0.9649 | 0.99143 | 1.036 | 195.659 | 50.1319 | 56606.702 | 56869.3291 | 1133.3866 | 1033.3866 | 1102.4347
7351 | 82.9249 | 0.86318 | 1.0267 | 0.94232 | 194.3215 | 49.4903 | 51617.353 | 51571.3981 | 1027.428 | 927.428 | 1102.4377
7352 | 93.6909 | 1.0494 | 1.1215 | 1.0229 | 211.6776 | 49.9651 | 62442.1498 | 62558.535 | 1247.1707 | 1147.1707 | 1102.452
7353 | 103.956 | 0.92577 | 1.2665 | 0.925 | 199.4366 | 51.3706 | 65322.8124 | 64362.2728 | 1283.2455 | 1183.2455 | 1102.5755
7354 | 116.8069 | 1.0591 | 0.96379 | 1.0815 | 179.3414 | 50.8196 | 59439.8621 | 57954.9383 | 1155.0988 | 1055.0988 | 1102.6197
7355 | 89.4525 | 0.86584 | 0.9769 | 0.9736 | 207.1471 | 46.2913 | 47820.8847 | 48177.068 | 959.5414 | 859.5414 | 1102.6665
7356 | 81.4948 | 1.1207 | 1.1338 | 0.96114 | 214.2815 | 49.4463 | 58615.3304 | 59443.0484 | 1184.861 | 1084.861 | 1102.71
7357 | 101.6246 | 0.97564 | 0.72011 | 1.0425 | 207.554 | 56.3443 | 48091.3185 | 46903.5442 | 934.0709 | 834.0709 | 1102.767
7358 | 99.7992 | 1.0347 | 0.93042 | 1.1079 | 194.7714 | 53.4943 | 60860.1076 | 59996.5438 | 1195.9309 | 1095.9309 | 1102.8286
7359 | 98.8163 | 1.0692 | 1.0775 | 1.0115 | 211.1303 | 52.8991 | 63453.6299 | 63923.5295 | 1274.4706 | 1174.4706 | 1102.8809
7360 | 99.5287 | 1.0045 | 1.1299 | 1.0164 | 200.0734 | 54.0855 | 67716.6086 | 66871.0767 | 1333.4215 | 1233.4215 | 1102.8869
7361 | 68.5629 | 1.0499 | 0.99679 | 0.99413 | 192.1682 | 53.5041 | 57062.6941 | 55966.0928 | 1115.3219 | 1015.3219 | 1102.9354
7362 | 92.1697 | 1.0386 | 1.0802 | 1.0008 | 211.3809 | 46.5116 | 54943.3483 | 53839.8927 | 1072.7979 | 972.7979 | 1102.9727
7363 | 104.7529 | 1.0501 | 1.0199 | 1.0268 | 201.0414 | 50.7896 | 58976.5082 | 60588.4019 | 1207.768 | 1107.768 | 1103.4091
7364 | 115.4669 | 1.0658 | 1.1362 | 1.0449 | 207.4397 | 52.1357 | 68516.4177 | 67514.4472 | 1346.2889 | 1246.2889 | 1103.4524
7365 | 87.3959 | 0.94441 | 0.7062 | 1.0169 | 199.9305 | 50.0673 | 40288.2932 | 40621.6256 | 808.4325 | 708.4325 | 1103.4703
7366 | 119.0855 | 1.0127 | 0.9966 | 1.0637 | 195.1723 | 49.2598 | 58357.3685 | 58162.2226 | 1159.2445 | 1059.2445 | 1103.4973
7367 | 93.4493 | 1.0223 | 1.1625 | 0.95147 | 189.9018 | 51.4998 | 62072.0408 | 60520.8944 | 1206.4179 | 1106.4179 | 1103.5605
7368 | 93.2665 | 1.1061 | 1.0222 | 0.96951 | 207.6676 | 53.8013 | 59078.56 | 59340.3771 | 1182.8075 | 1082.8075 | 1103.6502
7369 | 82.0397 | 1.0248 | 1.0747 | 1.0412 | 198.1665 | 52.5197 | 63380.6571 | 62713.0754 | 1250.2615 | 1150.2615 | 1103.7415
7370 | 107.5326 | 1.0198 | 0.93972 | 0.96332 | 218.0107 | 46.4252 | 47335.6238 | 46341.2517 | 922.825 | 822.825 | 1103.7573
7371 | 102.4479 | 0.93153 | 1.1208 | 0.9842 | 200.023 | 48.0046 | 57736.0639 | 58315.9361 | 1162.3187 | 1062.3187 | 1103.7666
7372 | 95.3722 | 1.0118 | 1.1552 | 1.0159 | 196.4668 | 49.0292 | 62464.532 | 64688.4517 | 1289.769 | 1189.769 | 1103.7934
7373 | 99.7261 | 0.92369 | 1.1244 | 0.91524 | 214.9573 | 48.2783 | 54347.3802 | 55050.385 | 1097.0077 | 997.0077 | 1103.8222
7374 | 107.6842 | 1.0718 | 0.93226 | 0.92551 | 195.7587 | 46.7298 | 45908.0577 | 45070.2776 | 897.4056 | 797.4056 | 1103.8801
7375 | 96.8069 | 0.95102 | 1.1091 | 0.99278 | 209.4664 | 47.8358 | 57284.1326 | 57681.7656 | 1149.6353 | 1049.6353 | 1103.9154
7376 | 98.7705 | 1.0744 | 1.1003 | 1.0116 | 208.9857 | 48.0814 | 58825.7417 | 58131.7572 | 1158.6351 | 1058.6351 | 1103.927
7377 | 95.3074 | 0.98362 | 1.0269 | 1.0234 | 193.1488 | 52.2376 | 59987.7047 | 58466.3391 | 1165.3268 | 1065.3268 | 1103.9604
7378 | 90.9497 | 1.0197 | 1.0112 | 1.0113 | 195.6852 | 49.7771 | 55712.2713 | 56241.7638 | 1120.8353 | 1020.8353 | 1104.0083
7379 | 107.0173 | 0.99278 | 0.96529 | 1.0101 | 191.1789 | 48.1984 | 52305.6482 | 52440.3662 | 1044.8073 | 944.8073 | 1104.0134
7380 | 110.9873 | 1.0728 | 1.0181 | 1.0676 | 194.4883 | 54.6251 | 66074.13 | 67225.6968 | 1340.5139 | 1240.5139 | 1104.0523
7381 | 93.5538 | 1.0594 | 1.1107 | 1.01 | 204.1707 | 51.4292 | 62999.6493 | 61404.5672 | 1224.0913 | 1124.0913 | 1104.0565
7382 | 96.895 | 1.0657 | 0.98143 | 1.0241 | 215.0835 | 48.9304 | 54446.1747 | 54846.6267 | 1092.9325 | 992.9325 | 1104.0665
7383 | 109.1909 | 0.9223 | 0.86266 | 0.9857 | 192.8286 | 52.9426 | 50542.9541 | 51296.7566 | 1021.9351 | 921.9351 | 1104.1037
7384 | 110.6198 | 0.959 | 0.88267 | 1.0187 | 187.2314 | 51.0143 | 51470.8974 | 51370.0662 | 1023.4013 | 923.4013 | 1104.1472
7385 | 99.478 | 0.94857 | 1.0751 | 0.9757 | 203.2724 | 52.8715 | 60654.9462 | 60271.1057 | 1201.4221 | 1101.4221 | 1104.2892
7386 | 91.971 | 0.96525 | 0.68219 | 1.0933 | 200.7333 | 49.9519 | 41890.0384 | 41199.5836 | 819.9917 | 719.9917 | 1104.4478
7387 | 105.6796 | 1.0971 | 0.71241 | 1.016 | 196.3135 | 49.2035 | 41514.0643 | 41118.4118 | 818.3682 | 718.3682 | 1104.4528
7388 | 109.247 | 1.0206 | 1.1365 | 1.0528 | 215.4043 | 48.8282 | 64081.0993 | 63314.2794 | 1262.2856 | 1162.2856 | 1104.4819
7389 | 95.7517 | 1.0376 | 1.2455 | 1.0215 | 212.365 | 50.312 | 69220.8086 | 69918.6354 | 1394.3727 | 1294.3727 | 1104.5797
7390 | 99.0022 | 0.9958 | 0.79203 | 1.0861 | 192.4605 | 49.3539 | 47514.8942 | 47750.2305 | 951.0046 | 851.0046 | 1104.592
7391 | 113.2002 | 1.0545 | 1.2442 | 1.0448 | 207.1361 | 51.7063 | 73593.2394 | 74523.8485 | 1486.477 | 1386.477 | 1104.7135
7392 | 95.6686 | 1.0129 | 0.87315 | 0.96047 | 198.7058 | 46.4485 | 43653.0545 | 43670.3916 | 869.4078 | 769.4078 | 1104.8059
7393 | 98.3092 | 0.99411 | 1.31 | 1.0019 | 205.7562 | 51.4812 | 72807.0363 | 71231.4481 | 1420.629 | 1320.629 | 1104.8091
7394 | 111.3269 | 0.99215 | 1.1167 | 1.0055 | 195.4145 | 52.0022 | 64328.3006 | 65459.3327 | 1305.1867 | 1205.1867 | 1105.0376
7395 | 101.739 | 0.99461 | 0.82696 | 1.0304 | 206.7537 | 49.7652 | 47646.7464 | 47265.9404 | 941.3188 | 841.3188 | 1105.0671
7396 | 104.9002 | 1.0384 | 1.1385 | 1.0522 | 206.9414 | 52.5988 | 68947.5774 | 69469.5454 | 1385.3909 | 1285.3909 | 1105.0763
7397 | 105.705 | 1.0272 | 0.97461 | 0.94577 | 210.5745 | 49.029 | 50727.2084 | 49548.2481 | 986.965 | 886.965 | 1105.0861
7398 | 92.2003 | 1.0448 | 0.74709 | 1.0015 | 180.5793 | 51.6445 | 43796.1087 | 43606.0073 | 868.1201 | 768.1201 | 1105.2143
7399 | 100.9389 | 1.0577 | 1.2862 | 0.96916 | 185.6924 | 49.1015 | 66636.42 | 68179.8625 | 1359.5972 | 1259.5972 | 1105.275
7400 | 102.8819 | 0.95264 | 0.999 | 0.93644 | 191.6643 | 51.1572 | 53063.1199 | 53629.0458 | 1068.5809 | 968.5809 | 1105.2937
7401 | 104.518 | 1.0219 | 0.89463 | 1.0144 | 203.4629 | 48.1595 | 49053.4736 | 48691.0841 | 969.8217 | 869.8217 | 1105.3812
7402 | 114.7442 | 0.93793 | 0.96599 | 1.0642 | 218.4574 | 50.2598 | 57294.195 | 57583.8171 | 1147.6763 | 1047.6763 | 1105.5825
7403 | 90.3567 | 0.87201 | 1.0137 | 1.0412 | 208.0575 | 52.7352 | 60019.2351 | 61174.7124 | 1219.4942 | 1119.4942 | 1105.5982
7404 | 106.99 | 1.0148 | 1.0502 | 0.99682 | 185.8018 | 46.7643 | 54218.164 | 54937.7429 | 1094.7549 | 994.7549 | 1105.6475
7405 | 97.7129 | 1.0577 | 0.97595 | 0.96712 | 191.4395 | 51.8166 | 54454.3664 | 55303.4596 | 1102.0692 | 1002.0692 | 1105.6529
7406 | 104.5422 | 1.023 | 0.93564 | 1.0236 | 201.7727 | 50.7631 | 54249.7845 | 53758.6908 | 1071.1738 | 971.1738 | 1105.7534
7407 | 107.7205 | 0.9602 | 1.0691 | 1.0417 | 207.0126 | 50.7677 | 61994.815 | 64527.8265 | 1286.5565 | 1186.5565 | 1105.9325
7408 | 121.2501 | 0.96419 | 0.90482 | 0.97726 | 209.6345 | 46.0638 | 46326.4366 | 46588.397 | 927.7679 | 827.7679 | 1106.1598
7409 | 84.7918 | 0.97565 | 0.95146 | 0.95344 | 202.4163 | 55.404 | 55046.2688 | 55254.8296 | 1101.0966 | 1001.0966 | 1106.2843
7410 | 106.6314 | 0.99962 | 0.84965 | 0.96796 | 195.1321 | 51.3442 | 47894.9105 | 47375.4097 | 943.5082 | 843.5082 | 1106.287
7411 | 88.1035 | 0.9428 | 0.93191 | 1.0596 | 199.4663 | 52.5377 | 56442.6264 | 56546.9476 | 1126.939 | 1026.939 | 1106.322
7412 | 103.7487 | 0.9114 | 0.82666 | 1.054 | 204.7931 | 52.2995 | 50716.6807 | 50229.5018 | 1000.59 | 900.59 | 1106.3232
7413 | 110.5583 | 1.0071 | 1.0065 | 1.0434 | 190.6274 | 52.8407 | 61566.4714 | 64001.1812 | 1276.0236 | 1176.0236 | 1106.3313
7414 | 100.6433 | 0.84126 | 1.0919 | 1.0057 | 205.7599 | 51.2003 | 60762.4012 | 58568.4358 | 1167.3687 | 1067.3687 | 1106.4028
7415 | 104.1962 | 0.98748 | 1.12 | 1.026 | 178.7914 | 50.4207 | 63307.3077 | 62839.5316 | 1252.7906 | 1152.7906 | 1106.4179
7416 | 87.0639 | 0.94967 | 1.0236 | 1.0133 | 203.4084 | 53.7098 | 60350.2879 | 63260.5426 | 1261.2109 | 1161.2109 | 1106.435
7417 | 97.3098 | 0.95867 | 1.0811 | 1.0446 | 201.9683 | 47.2921 | 58022.7004 | 59362.1847 | 1183.2437 | 1083.2437 | 1106.5329
7418 | 88.6218 | 1.0568 | 0.8972 | 1.0229 | 183.8198 | 46.4408 | 47153.7583 | 47911.7629 | 954.2353 | 854.2353 | 1106.7759
7419 | 98.8947 | 0.98809 | 1.044 | 1.0177 | 209.0101 | 52.9828 | 61676.9138 | 64094.3542 | 1277.8871 | 1177.8871 | 1106.7787
7420 | 95.4837 | 1.0391 | 1.1118 | 0.98659 | 216.3354 | 51.5038 | 61822.9907 | 60167.8657 | 1199.3573 | 1099.3573 | 1106.8223
7421 | 96.3431 | 1.037 | 1.0591 | 1.0417 | 203.9039 | 48.5429 | 58607.9251 | 58221.3075 | 1160.4262 | 1060.4262 | 1106.837
7422 | 130.6504 | 0.96428 | 0.97861 | 0.84922 | 196.6198 | 51.7558 | 49728.6903 | 49541.5304 | 986.8306 | 886.8306 | 1106.8463
7423 | 110.7513 | 1.0115 | 1.1757 | 0.96459 | 195.1351 | 46.6525 | 58330.7748 | 58673.4844 | 1169.4697 | 1069.4697 | 1106.8894
7424 | 103.4314 | 0.94193 | 1.1567 | 1.0272 | 186.0406 | 43.7382 | 56415.3763 | 56167.4581 | 1119.3492 | 1019.3492 | 1106.9732
7425 | 92.8009 | 1.0487 | 1.1221 | 0.97342 | 208.9374 | 57.4893 | 68597.2916 | 66906.5393 | 1334.1308 | 1234.1308 | 1107.0802
7426 | 104.947 | 1.0596 | 0.96856 | 0.97982 | 185.5641 | 46.4755 | 49459.6799 | 50583.2368 | 1007.6647 | 907.6647 | 1107.1167
7427 | 125.6158 | 0.93514 | 1.1329 | 0.9336 | 189.584 | 51.738 | 60990.7797 | 60541.8481 | 1206.837 | 1106.837 | 1107.1828
7428 | 84.4318 | 1.0403 | 1.1946 | 1.0749 | 184.4356 | 48.074 | 66138.4178 | 67228.9304 | 1340.5786 | 1240.5786 | 1107.1829
7429 | 86.0832 | 1.1071 | 1.1898 | 0.95095 | 224.6045 | 49.1074 | 60465.5341 | 59756.3268 | 1191.1265 | 1091.1265 | 1107.2396
7430 | 96.2721 | 0.9775 | 1.117 | 1.0054 | 193.8399 | 48.8425 | 59644.4823 | 59903.1837 | 1194.0637 | 1094.0637 | 1107.2425
7431 | 97.1704 | 1.045 | 0.92988 | 0.98673 | 188.6345 | 52.5881 | 53779.9712 | 54473.6919 | 1085.4738 | 985.4738 | 1107.2921
7432 | 117.7463 | 1.0953 | 0.88565 | 1.0478 | 217.6819 | 50.905 | 54020.9689 | 53380.0334 | 1063.6007 | 963.6007 | 1107.2948
7433 | 94.3994 | 1.06 | 0.97419 | 1.0368 | 201.7881 | 52.5541 | 58544.5303 | 58464.204 | 1165.2841 | 1065.2841 | 1107.3003
7434 | 85.7558 | 0.936 | 0.89843 | 0.98056 | 199.4142 | 54.0173 | 52122.3537 | 53018.1831 | 1056.3637 | 956.3637 | 1107.3405
7435 | 95.3034 | 1.0175 | 1.0647 | 1.0217 | 204.2783 | 45.8333 | 54504.8356 | 53417.8887 | 1064.3578 | 964.3578 | 1107.3567
7436 | 99.0241 | 1.0273 | 0.88372 | 0.95193 | 215.2083 | 47.173 | 44697.7848 | 43134.5062 | 858.6901 | 758.6901 | 1107.4496
7437 | 94.2592 | 0.95031 | 1.042 | 0.94916 | 203.3498 | 53.5096 | 57916.8833 | 57590.0091 | 1147.8002 | 1047.8002 | 1107.4688
7438 | 107.1144 | 1.0486 | 1.1405 | 0.98552 | 214.5047 | 51.4157 | 63780.1365 | 61942.218 | 1234.8444 | 1134.8444 | 1107.562
7439 | 104.9221 | 1.0925 | 0.85045 | 1.0194 | 203.6496 | 48.0623 | 47381.8059 | 46475.5998 | 925.512 | 825.512 | 1107.66
7440 | 101.8623 | 0.98705 | 1.222 | 1.0243 | 188.2332 | 47.3939 | 64275.0314 | 64502.3428 | 1286.0469 | 1186.0469 | 1107.7647
7441 | 92.7572 | 1.043 | 1.0942 | 0.98858 | 198.2091 | 51.19 | 60521.1314 | 60220.6349 | 1200.4127 | 1100.4127 | 1107.768
7442 | 90.9537 | 1.0293 | 0.89821 | 1.0219 | 209.5781 | 52.5889 | 53404.5956 | 52821.167 | 1052.4233 | 952.4233 | 1107.7991
7443 | 78.3174 | 0.99605 | 0.89987 | 1.0344 | 214.0297 | 51.0339 | 51696.5146 | 51983.1864 | 1035.6637 | 935.6637 | 1107.808
7444 | 81.1648 | 0.97389 | 1.0935 | 0.9649 | 174.9091 | 47.9379 | 54545.2431 | 54726.0214 | 1090.5204 | 990.5204 | 1107.8237
7445 | 93.2053 | 0.95966 | 1.1765 | 0.95291 | 199.1303 | 53.7666 | 65287.8019 | 65513.5484 | 1306.271 | 1206.271 | 1107.8879
7446 | 94.5699 | 1.0063 | 0.94033 | 0.96142 | 207.8718 | 52.4987 | 52665.5839 | 52208.9628 | 1040.1793 | 940.1793 | 1107.9407
7447 | 111.1239 | 0.93683 | 1.0647 | 0.97527 | 210.6162 | 50.3359 | 57715.6756 | 57408.7559 | 1144.1751 | 1044.1751 | 1108.0493
7448 | 90.7912 | 0.96273 | 1.0021 | 0.9465 | 189.6558 | 48.4514 | 50380.8226 | 50061.679 | 997.2336 | 897.2336 | 1108.2627
7449 | 92.8402 | 0.95371 | 1.0901 | 0.86555 | 201.7569 | 52.2076 | 54082.6567 | 52810.4846 | 1052.2097 | 952.2097 | 1108.2683
7450 | 92.7094 | 1.0295 | 1.1755 | 0.9942 | 205.488 | 47.8496 | 60692.9841 | 61243.7723 | 1220.8754 | 1120.8754 | 1108.3632
7451 | 108.6536 | 0.96846 | 0.72868 | 0.99849 | 194.9439 | 50.7657 | 42472.9017 | 43527.8192 | 866.5564 | 766.5564 | 1108.3711
7452 | 113.4412 | 0.95608 | 0.81372 | 1.0206 | 189.5563 | 48.9395 | 46142.7633 | 45552.581 | 907.0516 | 807.0516 | 1108.4784
7453 | 97.0836 | 0.94121 | 1.0463 | 1.0214 | 189.5746 | 48.9846 | 57011.6941 | 58319.1734 | 1162.3835 | 1062.3835 | 1108.5263
7454 | 103.8722 | 1.045 | 1.0949 | 0.96939 | 190.4115 | 52.1828 | 61241.5713 | 60967.853 | 1215.3571 | 1115.3571 | 1108.5955
7455 | 107.7302 | 0.99348 | 0.70933 | 1.0538 | 206.9305 | 49.2664 | 42307.5582 | 41692.1025 | 829.8421 | 729.8421 | 1108.6585
7456 | 94.3192 | 0.97187 | 1.0511 | 0.97625 | 191.1769 | 49.6952 | 55740.9335 | 58013.9836 | 1156.2797 | 1056.2797 | 1108.6892
7457 | 98.789 | 0.94554 | 1.1281 | 0.98867 | 186.0713 | 47.2154 | 57254.9594 | 57240.1028 | 1140.8021 | 1040.8021 | 1108.7329
7458 | 108.4983 | 1.002 | 0.85427 | 0.93039 | 202.9781 | 54.8005 | 49716.4792 | 49819.4552 | 992.3891 | 892.3891 | 1108.8019
7459 | 99.7554 | 0.95175 | 0.92654 | 0.95619 | 192.9399 | 50.1343 | 49369.4291 | 47676.329 | 949.5266 | 849.5266 | 1108.9093
7460 | 101.9216 | 1.0449 | 1.01 | 1.0009 | 210.0761 | 49.9903 | 56067.4119 | 56777.2597 | 1131.5452 | 1031.5452 | 1108.9389
7461 | 105.4861 | 1.0673 | 0.86753 | 0.92421 | 187.4018 | 55.3722 | 50818.2857 | 49119.9189 | 978.3984 | 878.3984 | 1108.9468
7462 | 89.7022 | 0.95603 | 1.0203 | 0.99108 | 195.9614 | 45.7838 | 50420.8062 | 50376.5738 | 1003.5315 | 903.5315 | 1109.0044
7463 | 83.4968 | 1.0174 | 1.0346 | 0.89911 | 199.8479 | 48.9383 | 49880.8377 | 49026.2402 | 976.5248 | 876.5248 | 1109.174
7464 | 103.6219 | 1.0586 | 0.93842 | 0.99293 | 196.8911 | 51.7816 | 54126.4935 | 55753.8499 | 1111.077 | 1011.077 | 1109.1778
7465 | 83.4743 | 0.97352 | 0.86825 | 0.88604 | 209.9735 | 51.1494 | 43715.6519 | 44477.7878 | 885.5558 | 785.5558 | 1109.329
7466 | 92.7363 | 1.0042 | 0.86596 | 0.98161 | 191.144 | 44.9668 | 42602.4078 | 42843.5575 | 852.8712 | 752.8712 | 1109.3428
7467 | 91.3743 | 1.0425 | 1.0138 | 1.0588 | 197.5243 | 50.2425 | 58914.1166 | 60070.6905 | 1197.4138 | 1097.4138 | 1109.3848
7468 | 86.5928 | 1.053 | 1.1565 | 0.94257 | 210.9311 | 49.4407 | 58611.522 | 58873.8846 | 1173.4777 | 1073.4777 | 1109.4308
7469 | 90.7835 | 0.90538 | 0.80336 | 1.1304 | 197.6733 | 50.6671 | 50372.8691 | 51192.4396 | 1019.8488 | 919.8488 | 1109.5111
7470 | 97.9836 | 1.079 | 0.83559 | 0.99451 | 193.0464 | 48.7559 | 45863.9996 | 45119.1737 | 898.3835 | 798.3835 | 1109.5468
7471 | 88.0157 | 0.99589 | 0.81308 | 0.99015 | 210.2365 | 47.806 | 42887.7241 | 42852.6029 | 853.0521 | 753.0521 | 1109.6104
7472 | 118.7882 | 1.0678 | 1.3127 | 0.99717 | 189.9673 | 48.7272 | 70154.8618 | 70471.5819 | 1405.4316 | 1305.4316 | 1109.6384
7473 | 95.4167 | 1 | 1.1362 | 0.97392 | 200.128 | 49.5853 | 59800.6557 | 59335.8397 | 1182.7168 | 1082.7168 | 1109.6402
7474 | 103.1583 | 0.99961 | 0.8806 | 0.98763 | 199.0573 | 51.0957 | 49906.444 | 50372.6788 | 1003.4536 | 903.4536 | 1109.6545
7475 | 83.6932 | 0.94871 | 1.2588 | 1.1018 | 198.3787 | 50.7921 | 74680.4301 | 73750.5408 | 1471.0108 | 1371.0108 | 1109.6586
7476 | 80.4323 | 0.90035 | 1.163 | 1.0102 | 189.0228 | 45.7377 | 57238.9462 | 56650.9591 | 1129.0192 | 1029.0192 | 1109.6788
7477 | 105.2558 | 1.0405 | 0.82377 | 1.0147 | 199.6144 | 46.5152 | 44176.433 | 45129.785 | 898.5957 | 798.5957 | 1109.6938
7478 | 101.345 | 0.94276 | 0.97282 | 1.0411 | 194.0194 | 49.2189 | 54747.4427 | 54626.0158 | 1088.5203 | 988.5203 | 1109.7203
7479 | 106.4139 | 1.0309 | 1.1194 | 0.97861 | 201.4733 | 51.6335 | 62426.8117 | 61018.0274 | 1216.3605 | 1116.3605 | 1109.7925
7480 | 95.7388 | 1.0665 | 0.7711 | 0.97257 | 201.6781 | 48.8073 | 41788.141 | 40727.8695 | 810.5574 | 710.5574 | 1109.8603
7481 | 97.0737 | 0.99844 | 1.0137 | 0.97243 | 185.3115 | 53.3102 | 57902.5752 | 58091.6747 | 1157.8335 | 1057.8335 | 1109.9192
7482 | 115.6879 | 0.89752 | 1.1968 | 1.0137 | 210.3934 | 50.6485 | 66915.9602 | 65609.1947 | 1308.1839 | 1208.1839 | 1110.0821
7483 | 79.814 | 1.0448 | 1.1839 | 1.0793 | 191.1523 | 50.8864 | 69456.0703 | 68762.7818 | 1371.2556 | 1271.2556 | 1110.1444
7484 | 100.006 | 0.94534 | 0.81668 | 1.0094 | 205.317 | 51.6871 | 47698.0781 | 48931.7406 | 974.6348 | 874.6348 | 1110.1786
7485 | 104.8433 | 0.98197 | 1.1609 | 1.0515 | 198.5016 | 54.6131 | 72485.0633 | 70697.2253 | 1409.9445 | 1309.9445 | 1110.2171
7486 | 100.308 | 1.1138 | 0.86422 | 0.97449 | 192.2715 | 44.6812 | 42813.661 | 43878.1437 | 873.5629 | 773.5629 | 1110.3044
7487 | 77.9411 | 1.0026 | 0.96518 | 1.0074 | 200.8303 | 49.0235 | 51698.1666 | 51704.3417 | 1030.0868 | 930.0868 | 1110.3301
7488 | 103.567 | 1.013 | 1.0876 | 0.92755 | 199.1602 | 50.6565 | 56614.5975 | 56323.9941 | 1122.4799 | 1022.4799 | 1110.34
7489 | 114.6602 | 1.0594 | 0.90805 | 0.98691 | 215.913 | 50.4492 | 51554.753 | 51403.1129 | 1024.0623 | 924.0623 | 1110.3433
7490 | 99.5678 | 0.95342 | 1.0495 | 0.86697 | 190.2855 | 50.2504 | 50684.0314 | 51984.6682 | 1035.6934 | 935.6934 | 1110.4186
7491 | 97.023 | 0.95773 | 1.1461 | 0.91845 | 221.4496 | 50.7409 | 58346.0717 | 59990.2919 | 1195.8058 | 1095.8058 | 1110.4932
7492 | 118.6741 | 1.0129 | 1.1315 | 1.0197 | 181.319 | 48.2849 | 61694.8948 | 61454.963 | 1225.0993 | 1125.0993 | 1110.5847
7493 | 106.5233 | 1.0203 | 0.77965 | 1.0166 | 197.5629 | 50.3053 | 45537.984 | 46118.1486 | 918.363 | 818.363 | 1110.6245
7494 | 91.6858 | 0.87966 | 1.0975 | 1.0092 | 200.3105 | 50.0281 | 59646.11 | 61060.1314 | 1217.2026 | 1117.2026 | 1110.6415
7495 | 120.132 | 0.92966 | 0.82537 | 0.9525 | 209.9143 | 53.7969 | 48511.5986 | 47983.61 | 955.6722 | 855.6722 | 1110.6495
7496 | 101.4006 | 0.99038 | 1.0229 | 0.94871 | 199.7388 | 44.5259 | 47880.8719 | 46507.191 | 926.1438 | 826.1438 | 1110.6547
7497 | 101.5629 | 0.94269 | 0.92339 | 0.98802 | 207.0194 | 47.8517 | 48445.165 | 47053.8765 | 937.0775 | 837.0775 | 1110.7011
7498 | 95.4805 | 0.91714 | 1.2209 | 1.0964 | 193.0457 | 52.6773 | 75322.5868 | 77714.3166 | 1550.2863 | 1450.2863 | 1110.7206
7499 | 110.0859 | 1.0348 | 1.2665 | 0.89619 | 188.0108 | 48.8886 | 61246.0647 | 60479.9088 | 1205.5982 | 1105.5982 | 1110.7758
7500 | 86.651 | 0.94419 | 1.098 | 1.0374 | 202.0702 | 50.9194 | 62368.102 | 59284.6008 | 1181.692 | 1081.692 | 1110.8177
7501 | 109.1893 | 1.0556 | 0.93602 | 1.0042 | 205.6858 | 55.0067 | 58250.206 | 57744.6457 | 1150.8929 | 1050.8929 | 1110.823
7502 | 88.6814 | 0.99491 | 0.84599 | 1.0115 | 200.7006 | 50.9097 | 48258.2955 | 47551.2465 | 947.0249 | 847.0249 | 1110.8275
7503 | 101.9339 | 0.9721 | 0.9303 | 1.1079 | 212.0726 | 52.3604 | 59367.0377 | 58695.549 | 1169.911 | 1069.911 | 1110.8596
7504 | 98.8294 | 1.0072 | 1.0364 | 0.95671 | 198.3429 | 50.2277 | 55001.9106 | 55703.1126 | 1110.0623 | 1010.0623 | 1111.0017
7505 | 110.8755 | 0.9768 | 0.88209 | 1.0204 | 211.8785 | 45.5881 | 46182.1328 | 44977.8614 | 895.5572 | 795.5572 | 1111.0386
7506 | 91.497 | 0.99111 | 1.0731 | 0.96087 | 164.7982 | 45.4313 | 51128.5171 | 49201.7718 | 980.0354 | 880.0354 | 1111.0975
7507 | 86.8645 | 1.0253 | 0.99919 | 1.0545 | 214.6626 | 49.4314 | 56701.1579 | 56207.7773 | 1120.1555 | 1020.1555 | 1111.1992
7508 | 75.3794 | 1.0466 | 1.0725 | 1.0246 | 213.0864 | 49.2848 | 58259.1781 | 57763.6424 | 1151.2728 | 1051.2728 | 1111.2587
7509 | 102.354 | 1.0151 | 1.0641 | 1.0016 | 195.2439 | 52.9095 | 62083.994 | 60841.7719 | 1212.8354 | 1112.8354 | 1111.305
7510 | 89.4338 | 1.0455 | 0.97706 | 1.0725 | 194.0756 | 50.0302 | 57298.3411 | 54394.136 | 1083.8827 | 983.8827 | 1111.3388
7511 | 108.1724 | 1.0426 | 1.1644 | 0.92276 | 199.4485 | 47.2817 | 56336.0376 | 55974.6678 | 1115.4934 | 1015.4934 | 1111.4738
7512 | 104.4835 | 0.97389 | 0.87696 | 1.0586 | 201.7569 | 53.997 | 55825.8091 | 53783.3957 | 1071.6679 | 971.6679 | 1111.4782
7513 | 102.7953 | 1.0583 | 0.87239 | 1.0125 | 196.7765 | 53.8449 | 53613.9347 | 51447.2965 | 1024.9459 | 924.9459 | 1111.6829
7514 | 86.2164 | 0.97698 | 0.99178 | 0.96433 | 192.2323 | 45.3999 | 47436.7359 | 48034.3648 | 956.6873 | 856.6873 | 1111.7019
7515 | 101.1817 | 1.0639 | 1.0655 | 1.031 | 204.1443 | 50.6398 | 61286.9948 | 60683.9393 | 1209.6788 | 1109.6788 | 1111.7713
7516 | 94.4395 | 1.0517 | 0.74494 | 0.92689 | 202.6939 | 46.7684 | 37140.4455 | 36525.4387 | 726.5088 | 626.5088 | 1111.7737
7517 | 89.1127 | 0.90888 | 1.0514 | 1.0947 | 203.3028 | 49.9815 | 61776.085 | 60913.1139 | 1214.2623 | 1114.2623 | 1111.8739
7518 | 103.5896 | 1.0173 | 1.3432 | 0.92415 | 195.8039 | 48.5554 | 65587.4738 | 65858.9004 | 1313.178 | 1213.178 | 1111.9147
7519 | 105.9741 | 0.99732 | 1.2984 | 1.011 | 192.8471 | 48.5783 | 69092.5493 | 68559.2278 | 1367.1846 | 1267.1846 | 1112.1171
7520 | 99.4555 | 1.0748 | 0.86717 | 0.99291 | 204.7884 | 47.6167 | 46293.8102 | 44338.8178 | 882.7764 | 782.7764 | 1112.1613
7521 | 102.8106 | 0.9329 | 0.76513 | 1.0614 | 192.6185 | 51.5223 | 46975.0269 | 47067.6121 | 937.3522 | 837.3522 | 1112.2347
7522 | 93.7972 | 1.0188 | 1.0625 | 1.0286 | 179.2869 | 53.1283 | 63319.4992 | 61661.2805 | 1229.2256 | 1129.2256 | 1112.2377
7523 | 103.0055 | 0.94447 | 0.97376 | 0.96374 | 181.558 | 49.5145 | 51465.6651 | 50826.4888 | 1012.5298 | 912.5298 | 1112.2621
7524 | 103.2328 | 0.98252 | 1.0205 | 1.0895 | 202.7334 | 45.0923 | 54912.6083 | 54980.1541 | 1095.6031 | 995.6031 | 1112.3088
7525 | 117.6728 | 0.9958 | 0.94881 | 1.0693 | 185.864 | 45.4776 | 51656.7578 | 51563.5373 | 1027.2707 | 927.2707 | 1112.3724
7526 | 80.6629 | 1.0015 | 0.97211 | 0.99928 | 200.2119 | 47.9856 | 50690.2168 | 51361.5551 | 1023.2311 | 923.2311 | 1112.385
7527 | 107.1003 | 0.99817 | 1.3174 | 1.0871 | 206.1723 | 50.8228 | 78423.7158 | 77261.5175 | 1541.2304 | 1441.2304 | 1112.4253
7528 | 107.0196 | 0.95542 | 0.92103 | 1.0871 | 180.6574 | 48.586 | 53797.1187 | 55190.6153 | 1099.8123 | 999.8123 | 1112.4817
7529 | 95.1291 | 1.0592 | 1.0101 | 0.96935 | 198.784 | 45.4952 | 49327.4592 | 49180.9097 | 979.6182 | 879.6182 | 1112.5069
7530 | 104.0379 | 1.0005 | 1.1327 | 1.0275 | 190.0997 | 49.2371 | 62617.7149 | 63167.7603 | 1259.3552 | 1159.3552 | 1112.5342
7531 | 96.8407 | 0.96011 | 1.2262 | 0.9709 | 191.3717 | 50.7679 | 65350.5613 | 67347.5206 | 1342.9504 | 1242.9504 | 1112.5569
7532 | 105.0758 | 1.0068 | 0.95038 | 1.0384 | 186.7723 | 47.6382 | 52240.6296 | 51070.6427 | 1017.4129 | 917.4129 | 1112.6779
7533 | 104.0667 | 0.98669 | 1.0726 | 0.93984 | 203.6126 | 48.6928 | 54288.3038 | 54364.6393 | 1083.2928 | 983.2928 | 1112.7324
7534 | 108.4113 | 1.0274 | 0.93752 | 0.97581 | 196.2442 | 54.7918 | 56424.5903 | 57285.6291 | 1141.7126 | 1041.7126 | 1112.8171
7535 | 93.9217 | 1.0904 | 0.97861 | 1.0822 | 223.5496 | 49.2279 | 57402.2831 | 56677.6839 | 1129.5537 | 1029.5537 | 1112.8354
7536 | 108.6568 | 1.0322 | 0.80158 | 1.075 | 196.0651 | 52.391 | 51218.9923 | 49866.0057 | 993.3201 | 893.3201 | 1112.9124
7537 | 102.4131 | 0.99618 | 0.7598 | 1.0431 | 198.9194 | 50.0151 | 44942.6922 | 43738.0053 | 870.7601 | 770.7601 | 1112.9743
7538 | 102.0025 | 1.0347 | 1.0362 | 1.0056 | 194.7301 | 44.6962 | 51485.2474 | 50500.6071 | 1006.0121 | 906.0121 | 1113.0422
7539 | 108.7213 | 1.0232 | 0.79011 | 0.93932 | 198.5883 | 49.87 | 42757.7367 | 42224.1292 | 840.4826 | 740.4826 | 1113.1364
7540 | 101.3824 | 0.95915 | 0.88591 | 0.99547 | 190.0748 | 48.3054 | 47487.8954 | 46820.6748 | 932.4135 | 832.4135 | 1113.3016
7541 | 98.0126 | 0.95694 | 0.69217 | 0.94016 | 187.2608 | 52.0574 | 38946.2479 | 38761.3647 | 771.2273 | 671.2273 | 1113.359
7542 | 98.7894 | 1.074 | 0.97476 | 0.92686 | 207.8841 | 49.7248 | 50408.3707 | 50866.3785 | 1013.3276 | 913.3276 | 1113.4689
7543 | 92.4979 | 0.9503 | 0.86097 | 1.0302 | 210.9541 | 52.8272 | 51710.3712 | 50410.1399 | 1004.2028 | 904.2028 | 1113.614
7544 | 93.4614 | 0.937 | 0.98048 | 1.0832 | 197.3445 | 48.0908 | 55481.8455 | 56225.9461 | 1120.5189 | 1020.5189 | 1113.6276
7545 | 90.0709 | 0.92532 | 0.58375 | 1.0821 | 179.7723 | 52.9816 | 38062.1202 | 38581.503 | 767.6301 | 667.6301 | 1113.6316
7546 | 98.7571 | 0.95845 | 1.1576 | 1.0564 | 221.2913 | 52.4835 | 69366.1284 | 68032.1867 | 1356.6437 | 1256.6437 | 1113.6659
7547 | 98.0723 | 1.0156 | 0.97021 | 1.0173 | 202.5665 | 52.0959 | 56807.7399 | 56401.6844 | 1124.0337 | 1024.0337 | 1113.8609
7548 | 101.5447 | 0.95232 | 0.93732 | 1.0022 | 207.1384 | 54.0177 | 56174.4673 | 54388.488 | 1083.7698 | 983.7698 | 1114.0845
7549 | 93.9231 | 0.96022 | 0.89656 | 0.94809 | 201.6462 | 49.8055 | 47028.8964 | 46259.4637 | 921.1893 | 821.1893 | 1114.0988
7550 | 116.805 | 1.0255 | 1.2748 | 0.88803 | 200.3212 | 50.0973 | 62914.7221 | 64059.1679 | 1277.1834 | 1177.1834 | 1114.1266
7551 | 101.4739 | 1.0131 | 0.99875 | 1.0259 | 212.2129 | 51.2305 | 57972.6579 | 57838.1373 | 1152.7627 | 1052.7627 | 1114.1675
7552 | 113.1468 | 0.92272 | 1.1913 | 0.99173 | 201.6305 | 48.5689 | 62653.5619 | 63524.157 | 1266.4831 | 1166.4831 | 1114.187
7553 | 92.1485 | 0.96652 | 0.98801 | 1.0478 | 196.1061 | 53.5514 | 60403.4034 | 59604.5319 | 1188.0906 | 1088.0906 | 1114.2199
7554 | 75.094 | 0.96943 | 1.3361 | 1.0042 | 213.3016 | 47.4207 | 67290.8048 | 67400.6569 | 1344.0131 | 1244.0131 | 1114.2623
7555 | 105.2943 | 0.95824 | 0.72085 | 1.0337 | 192.0885 | 50.0229 | 42515.062 | 43686.9023 | 869.738 | 769.738 | 1114.293
7556 | 112.8391 | 1.0133 | 0.74399 | 0.98493 | 195.3705 | 47.5516 | 40477.3767 | 40209.3919 | 800.1878 | 700.1878 | 1114.3474
7557 | 81.6042 | 1.0285 | 0.86104 | 1.0288 | 221.5188 | 47.0044 | 45802.9641 | 48201.191 | 960.0238 | 860.0238 | 1114.5099
7558 | 107.2823 | 0.95053 | 0.92215 | 0.99229 | 207.5668 | 51.7552 | 52843.5523 | 52444.0971 | 1044.8819 | 944.8819 | 1114.5513
7559 | 98.014 | 1.0306 | 1.1163 | 0.9296 | 193.5053 | 46.0166 | 52594.8338 | 50978.9564 | 1015.5791 | 915.5791 | 1114.5555
7560 | 100.4398 | 1.009 | 0.89647 | 0.9369 | 209.2956 | 47.2898 | 44720.9711 | 44967.7168 | 895.3543 | 795.3543 | 1114.5606
7561 | 98.5613 | 1.0272 | 1.0392 | 1.0337 | 205.1479 | 50.7669 | 59875.7994 | 59979.1239 | 1195.5825 | 1095.5825 | 1114.608
7562 | 99.325 | 0.94933 | 0.89028 | 1.0588 | 206.0104 | 46.9562 | 48896.131 | 51239.5464 | 1020.7909 | 920.7909 | 1114.6982
7563 | 107.7131 | 1.0002 | 0.86816 | 1.0094 | 189.5455 | 51.0121 | 50389.3605 | 49738.3149 | 990.7663 | 890.7663 | 1114.9157
7564 | 111.4389 | 1.0368 | 1.0163 | 1.0151 | 202.933 | 46.2687 | 53285.6426 | 52794.1835 | 1051.8837 | 951.8837 | 1114.9389
7565 | 117.6921 | 1.0997 | 0.763 | 1.0424 | 211.1782 | 47.7678 | 44384.6268 | 43539.5765 | 866.7915 | 766.7915 | 1114.9679
7566 | 90.5825 | 0.92504 | 0.78212 | 0.97352 | 195.2175 | 48.335 | 41047.7414 | 40829.8707 | 812.5974 | 712.5974 | 1114.9891
7567 | 97.0595 | 1.0676 | 0.88911 | 0.99454 | 187.2253 | 47.2669 | 46880.812 | 47471.7184 | 945.4344 | 845.4344 | 1115.0755
7568 | 116.4624 | 1.0248 | 0.91393 | 1.0073 | 185.2893 | 49.1739 | 51324.8846 | 50788.4795 | 1011.7696 | 911.7696 | 1115.0904
7569 | 87.2869 | 0.93998 | 1.1222 | 0.95546 | 203.6388 | 51.2198 | 59326.1732 | 56157.1821 | 1119.1436 | 1019.1436 | 1115.0945
7570 | 112.8252 | 1.0503 | 1.1541 | 1.0759 | 205.2258 | 48.9167 | 66740.882 | 64605.5535 | 1288.1111 | 1188.1111 | 1115.1065
7571 | 88.6301 | 1.0691 | 1.1067 | 1.0352 | 200.5524 | 50.6977 | 63083.1275 | 63741.867 | 1270.8373 | 1170.8373 | 1115.1316
7572 | 84.4216 | 1.1129 | 0.95227 | 1.0266 | 190.949 | 46.4474 | 49963.9383 | 49449.205 | 984.9841 | 884.9841 | 1115.1791
7573 | 94.2826 | 0.87116 | 0.94495 | 0.97072 | 221.7505 | 50.0035 | 50195.9588 | 50109.4954 | 998.1899 | 898.1899 | 1115.2155
7574 | 98.7765 | 1.0336 | 0.97087 | 1.0526 | 197.3258 | 50.8562 | 57359.0283 | 58032.2331 | 1156.6447 | 1056.6447 | 1115.2317
7575 | 93.5918 | 1.0628 | 0.71044 | 1.0925 | 181.559 | 49.543 | 43561.7615 | 43793.8243 | 871.8765 | 771.8765 | 1115.298
7576 | 104.2419 | 1.0797 | 1.0875 | 0.90943 | 209.5061 | 48.2087 | 53314.2527 | 52706.7749 | 1050.1355 | 950.1355 | 1115.3571
7577 | 92.5774 | 0.93662 | 0.98295 | 1.0289 | 203.8682 | 50.7397 | 55920.7025 | 57551.0473 | 1147.0209 | 1047.0209 | 1115.362
7578 | 109.4485 | 1.0435 | 0.93333 | 0.97718 | 195.9085 | 50.5038 | 52025.0462 | 52434.8207 | 1044.6964 | 944.6964 | 1115.4004
7579 | 97.1375 | 0.99256 | 1.0641 | 0.93378 | 195.7639 | 49.1729 | 53796.5557 | 54471.3767 | 1085.4275 | 985.4275 | 1115.4371
7580 | 93.5577 | 1.1061 | 0.81267 | 0.95485 | 198.7414 | 53.9883 | 47679.0238 | 48028.8467 | 956.5769 | 856.5769 | 1115.4539
7581 | 94.3462 | 0.99077 | 1.0714 | 1.0687 | 200.9301 | 47.4846 | 59012.5323 | 59984.4905 | 1195.6898 | 1095.6898 | 1115.5631
7582 | 94.2705 | 1.0618 | 0.93087 | 1.0452 | 200.6365 | 51.0978 | 55029.3143 | 54255.1498 | 1081.103 | 981.103 | 1115.5868
7583 | 106.2253 | 1.0855 | 0.80901 | 0.98705 | 203.551 | 47.4932 | 43604.7127 | 44850.9642 | 893.0193 | 793.0193 | 1115.6609
7584 | 105.0172 | 1.0531 | 0.89349 | 0.92957 | 196.4779 | 47.4596 | 44863.521 | 46745.6337 | 930.9127 | 830.9127 | 1115.7218
7585 | 91.8159 | 0.99494 | 1.1824 | 0.97607 | 192.9997 | 47.1959 | 58971.4757 | 60983.0441 | 1215.6609 | 1115.6609 | 1115.759
7586 | 103.6992 | 0.99016 | 1.1108 | 1.0731 | 214.9234 | 50.5317 | 65641.2985 | 65538.9186 | 1306.7784 | 1206.7784 | 1115.7796
7587 | 99.6891 | 1.0056 | 0.97045 | 1.0228 | 199.4412 | 52.6241 | 57708.2286 | 58092.6615 | 1157.8532 | 1057.8532 | 1115.818
7588 | 105.1432 | 1.0566 | 1.1233 | 1.1263 | 190.8133 | 47.0495 | 64945.4798 | 64831.3754 | 1292.6275 | 1192.6275 | 1115.8512
7589 | 108.1862 | 1.0345 | 0.99894 | 0.97746 | 180.8426 | 48.5712 | 53043.3908 | 52031.2797 | 1036.6256 | 936.6256 | 1115.8918
7590 | 117.6664 | 0.98168 | 0.91126 | 0.92843 | 216.5626 | 53.9287 | 52071.5981 | 51461.7585 | 1025.2352 | 925.2352 | 1115.9646
7591 | 97.8231 | 0.95443 | 0.97266 | 0.97079 | 205.2932 | 47.0298 | 49004.1802 | 50759.4071 | 1011.1881 | 911.1881 | 1115.9809
7592 | 127.3472 | 1.0013 | 1.0919 | 1.0599 | 207.662 | 51.1504 | 65928.1994 | 66322.3154 | 1322.4463 | 1222.4463 | 1116.0992
7593 | 103.4333 | 1.0131 | 1.0308 | 0.98309 | 214.9716 | 46.864 | 52617.3984 | 52356.8694 | 1043.1374 | 943.1374 | 1116.1049
7594 | 98.6836 | 0.99129 | 1.005 | 0.9802 | 210.841 | 54.7247 | 59471.4086 | 59951.8729 | 1195.0375 | 1095.0375 | 1116.1427
7595 | 100.554 | 0.95601 | 0.82591 | 1.0884 | 197.4 | 46.7942 | 46759.6582 | 48068.6356 | 957.3727 | 857.3727 | 1116.2191
7596 | 93.5776 | 0.98854 | 1.1744 | 1.0017 | 209.4342 | 52.2178 | 66466.8405 | 68513.419 | 1366.2684 | 1266.2684 | 1116.2857
7597 | 112.0022 | 1.0201 | 1.1591 | 0.99043 | 215.4542 | 51.0268 | 64626.2419 | 64978.6965 | 1295.5739 | 1195.5739 | 1116.3498
7598 | 92.3985 | 0.89875 | 1.2628 | 0.98968 | 204.07 | 47.2135 | 63129.7337 | 63698.2884 | 1269.9658 | 1169.9658 | 1116.354
7599 | 102.6266 | 1.0312 | 1.0278 | 1.0316 | 204.6981 | 52.7791 | 61750.9536 | 61467.3731 | 1225.3475 | 1125.3475 | 1116.3605
7600 | 109.493 | 1.0125 | 1.0446 | 1.0765 | 203.9664 | 50.8264 | 62992.2054 | 61777.1455 | 1231.5429 | 1131.5429 | 1116.3864
7601 | 87.4866 | 0.89364 | 1.2024 | 0.93955 | 218.8384 | 48.8196 | 59186.8853 | 57157.7327 | 1139.1547 | 1039.1547 | 1116.4101
7602 | 99.5517 | 0.93663 | 0.97641 | 0.93302 | 211.8971 | 50.3671 | 50792.8078 | 49836.7375 | 992.7347 | 892.7347 | 1116.4365
7603 | 92.7084 | 0.93248 | 0.98178 | 0.97969 | 203.3867 | 52.1383 | 54859.0491 | 53672.3117 | 1069.4462 | 969.4462 | 1116.4703
7604 | 98.0643 | 0.87551 | 0.72225 | 0.95282 | 210.3937 | 48.9473 | 38097.3397 | 37920.7503 | 754.415 | 654.415 | 1116.5685
7605 | 96.6673 | 1.0032 | 1.0254 | 1.0991 | 194.9414 | 49.4219 | 60685.0829 | 62191.5778 | 1239.8316 | 1139.8316 | 1116.5848
7606 | 95.9727 | 1.0002 | 0.98374 | 0.9707 | 191.5513 | 49.7285 | 52451.7994 | 53056.7918 | 1057.1358 | 957.1358 | 1116.7343
7607 | 98.7887 | 1.0934 | 0.71856 | 0.98065 | 198.7911 | 46.6065 | 38074.5806 | 37768.7409 | 751.3748 | 651.3748 | 1116.7457
7608 | 118.9844 | 1.0092 | 1.3014 | 0.9808 | 202.4067 | 48.2255 | 67550.0656 | 64925.7379 | 1294.5148 | 1194.5148 | 1116.7501
7609 | 103.686 | 0.9315 | 1.0411 | 0.91551 | 208.2719 | 51.9164 | 54704.8097 | 53296.3144 | 1061.9263 | 961.9263 | 1116.7676
7610 | 105.7074 | 0.95328 | 1.3416 | 0.975 | 185.7584 | 47.2336 | 66729.8825 | 66039.8058 | 1316.7961 | 1216.7961 | 1116.7784
7611 | 120.5432 | 1.0732 | 0.7046 | 0.93561 | 197.0863 | 51.2187 | 40588.0834 | 38948.9821 | 774.9796 | 674.9796 | 1116.7937
7612 | 106.1028 | 1.0132 | 1.1003 | 0.96868 | 204.679 | 53.7046 | 63218.6407 | 61374.7584 | 1223.4952 | 1123.4952 | 1116.8635
7613 | 108.8307 | 1.1136 | 0.79051 | 1.0206 | 218.774 | 49.19 | 45865.406 | 44741.8628 | 890.8373 | 790.8373 | 1116.9822
7614 | 106.1373 | 1.011 | 0.85421 | 1.0482 | 206.9737 | 48.3972 | 48733.6209 | 48694.2487 | 969.885 | 869.885 | 1117.0097
7615 | 113.1107 | 1.0327 | 0.80116 | 1.0157 | 203.8246 | 51.3075 | 47945.8483 | 48138.9948 | 958.7799 | 858.7799 | 1117.0149
7616 | 98.8506 | 0.99576 | 1.0871 | 0.99374 | 198.225 | 47.5366 | 56228.6387 | 55543.5025 | 1106.8701 | 1006.8701 | 1117.0278
7617 | 106.3957 | 0.96255 | 0.97672 | 1.0121 | 201.3991 | 49.0394 | 53702.1378 | 55449.708 | 1104.9942 | 1004.9942 | 1117.0475
7618 | 92.857 | 0.97815 | 0.85106 | 1.0173 | 204.2358 | 52.3317 | 50267.2146 | 50650.7444 | 1009.0149 | 909.0149 | 1117.0787
7619 | 97.8574 | 1.0015 | 1.0239 | 0.97081 | 207.0864 | 46.9893 | 51520.4434 | 52802.5686 | 1052.0514 | 952.0514 | 1117.1206
7620 | 91.6948 | 1.0404 | 0.90334 | 0.99989 | 193.4118 | 45.9226 | 46053.5188 | 46321.0351 | 922.4207 | 822.4207 | 1117.1334
7621 | 105.7404 | 0.92335 | 0.76056 | 1.0293 | 208.6682 | 52.4786 | 46416.6025 | 44853.314 | 893.0663 | 793.0663 | 1117.1953
7622 | 108.2634 | 1.0556 | 0.91636 | 0.99756 | 206.1276 | 54.2517 | 55999.1331 | 56296.0657 | 1121.9213 | 1021.9213 | 1117.2026
7623 | 127.3461 | 1.0258 | 1.0141 | 0.97364 | 202.1957 | 49.4678 | 55506.1355 | 54357.7346 | 1083.1547 | 983.1547 | 1117.3346
7624 | 113.7779 | 0.9959 | 0.78443 | 0.88908 | 195.5407 | 48.6263 | 39618.5313 | 39916.7826 | 794.3357 | 694.3357 | 1117.3755
7625 | 92.2974 | 1.0503 | 0.91609 | 1.0509 | 208.197 | 52.4008 | 55735.6634 | 56111.0904 | 1118.2218 | 1018.2218 | 1117.3943
7626 | 82.9491 | 0.93508 | 1.0396 | 0.98731 | 212.921 | 47.5365 | 52691.6755 | 51696.1646 | 1029.9233 | 929.9233 | 1117.4967
7627 | 84.1888 | 1.0062 | 1.3583 | 0.93225 | 184.3572 | 53.6984 | 72730.1273 | 70846.7677 | 1412.9354 | 1312.9354 | 1117.5613
7628 | 84.5727 | 1.0553 | 1.0722 | 1.0183 | 212.3014 | 51.9518 | 61572.4726 | 61505.9808 | 1226.1196 | 1126.1196 | 1117.5854
7629 | 120.8999 | 0.98824 | 1.0351 | 1.0045 | 182.4046 | 49.4914 | 57555.4237 | 56302.1127 | 1122.0423 | 1022.0423 | 1117.587
7630 | 100.1315 | 1.1235 | 1.0928 | 0.97402 | 188.9922 | 49.6175 | 58582.7334 | 58377.4086 | 1163.5482 | 1063.5482 | 1117.7495
7631 | 113.4052 | 0.90188 | 1.1636 | 1 | 201.5877 | 52.1769 | 66253.5237 | 67797.3368 | 1351.9467 | 1251.9467 | 1117.8369
7632 | 93.8601 | 1.0067 | 0.9882 | 1.0712 | 205.0162 | 48.7268 | 56388.6988 | 55670.7229 | 1109.4145 | 1009.4145 | 1117.8405
7633 | 94.4571 | 1.0354 | 1.0688 | 1.0641 | 190.0709 | 48.1642 | 59678.5982 | 57659.0154 | 1149.1803 | 1049.1803 | 1117.869
7634 | 103.5368 | 0.97445 | 0.7987 | 0.99522 | 193.9589 | 52.0533 | 46822.0961 | 47292.3197 | 941.8464 | 841.8464 | 1117.8869
7635 | 99.2508 | 1.0139 | 0.95508 | 1.0321 | 201.3891 | 48.4675 | 52856.4153 | 53012.727 | 1056.2545 | 956.2545 | 1117.9001
7636 | 106.0193 | 1.0176 | 1.0625 | 1.0548 | 203.9589 | 51.8724 | 63934.698 | 62959.4499 | 1255.189 | 1155.189 | 1117.9162
7637 | 95.3087 | 0.95555 | 1.0872 | 1.0768 | 220.8275 | 49.8205 | 63085.1608 | 63800.1551 | 1272.0031 | 1172.0031 | 1117.9886
7638 | 98.8949 | 1.013 | 0.88928 | 0.96973 | 189.5925 | 52.9372 | 51143.6113 | 51318.9785 | 1022.3796 | 922.3796 | 1117.9976
7639 | 108.7853 | 0.9893 | 1.159 | 0.95436 | 216.4521 | 47.4257 | 57780.3792 | 58082.0114 | 1157.6402 | 1057.6402 | 1118.0035
7640 | 88.3753 | 1.0246 | 1.0378 | 0.98948 | 200.0247 | 51.6071 | 57868.9236 | 57589.3043 | 1147.7861 | 1047.7861 | 1118.068
7641 | 83.7341 | 0.99454 | 1.064 | 0.98608 | 202.097 | 54.9293 | 62406.2128 | 60867.9503 | 1213.359 | 1113.359 | 1118.1121
7642 | 97.2056 | 1.0444 | 1.2364 | 0.98697 | 185.0085 | 53.2873 | 70622.2359 | 69557.4401 | 1387.1488 | 1287.1488 | 1118.1335
7643 | 115.2714 | 1.0219 | 0.89354 | 0.95577 | 210.3737 | 52.1285 | 50869.4675 | 50788.2475 | 1011.7649 | 911.7649 | 1118.1487
7644 | 105.3546 | 0.99904 | 0.95361 | 0.92795 | 195.0279 | 52.5219 | 52199.6137 | 51281.3821 | 1021.6276 | 921.6276 | 1118.1897
7645 | 95.2937 | 1.0295 | 1.0353 | 1.0183 | 189.44 | 50.0136 | 57818.0858 | 58562.2249 | 1167.2445 | 1067.2445 | 1118.2175
7646 | 98.3733 | 1.038 | 1.0324 | 1.0456 | 197.8358 | 51.0365 | 60504.0179 | 60994.5885 | 1215.8918 | 1115.8918 | 1118.2214
7647 | 98.6534 | 1.0052 | 1.08 | 1.1111 | 189.4771 | 51.8053 | 67493.7396 | 68632.1475 | 1368.6429 | 1268.6429 | 1118.2566
7648 | 92.3977 | 1.0696 | 0.66758 | 1.0707 | 193.9753 | 48.2644 | 39461.9052 | 40177.9036 | 799.5581 | 699.5581 | 1118.2811
7649 | 92.5945 | 1.0508 | 0.74018 | 0.94986 | 195.5608 | 52.4304 | 42158.8706 | 41713.4898 | 830.2698 | 730.2698 | 1118.3939
7650 | 91.9246 | 0.90067 | 1.0618 | 0.97486 | 186.7895 | 44.0427 | 49420.4889 | 49060.9418 | 977.2188 | 877.2188 | 1118.4412
7651 | 107.5081 | 0.89273 | 1.1992 | 1.0111 | 197.252 | 49.4159 | 64859.3751 | 65210.6886 | 1300.2138 | 1200.2138 | 1118.4654
7652 | 110.0968 | 1.0174 | 1.0027 | 0.9533 | 197.424 | 53.8046 | 57652.3628 | 60561.9821 | 1207.2396 | 1107.2396 | 1118.6824
7653 | 107.6245 | 0.93424 | 0.95306 | 1.1023 | 204.9016 | 53.0411 | 61263.043 | 61193.1937 | 1219.8639 | 1119.8639 | 1118.7333
7654 | 102.8839 | 1.0259 | 0.99442 | 1.0663 | 189.4525 | 52.401 | 61281.5756 | 60166.7301 | 1199.3346 | 1099.3346 | 1118.8236
7655 | 97.9068 | 0.95723 | 0.85681 | 1.0104 | 210.9382 | 51.4554 | 49581.0765 | 51451.2583 | 1025.0252 | 925.0252 | 1118.9308
7656 | 103.9115 | 0.97788 | 1.0121 | 1.0162 | 212.9926 | 48.9671 | 55555.2472 | 53886.9963 | 1073.7399 | 973.7399 | 1118.937
7657 | 97.1032 | 1.0721 | 0.88549 | 1.0247 | 198.7055 | 48.2932 | 49043.9833 | 48892.9371 | 973.8587 | 873.8587 | 1118.9868
7658 | 91.1614 | 1.0779 | 0.74814 | 0.92359 | 203.5474 | 50.0385 | 39695.8987 | 41634.6875 | 828.6937 | 728.6937 | 1119.0535
7659 | 96.287 | 1.037 | 0.93573 | 0.97992 | 198.5472 | 50.1923 | 51233.599 | 52218.4798 | 1040.3696 | 940.3696 | 1119.1008
7660 | 90.9476 | 1.0066 | 1.0528 | 0.99363 | 227.0852 | 49.2808 | 56288.5852 | 55093.1018 | 1097.862 | 997.862 | 1119.1155
7661 | 92.5441 | 1.0982 | 0.87122 | 0.95554 | 207.883 | 48.8546 | 45843.8515 | 44397.9895 | 883.9598 | 783.9598 | 1119.1747
7662 | 106.3689 | 0.95143 | 1.0526 | 0.9701 | 209.3855 | 50.8137 | 57239.527 | 57225.8116 | 1140.5162 | 1040.5162 | 1119.1835
7663 | 80.5932 | 1.0272 | 0.945 | 1.015 | 187.8816 | 53.5285 | 55963.0528 | 57354.2256 | 1143.0845 | 1043.0845 | 1119.2146
7664 | 83.0299 | 0.99146 | 1.0901 | 0.97797 | 204.5979 | 51.3278 | 59148.4484 | 59521.1146 | 1186.4223 | 1086.4223 | 1119.2349
7665 | 110.8603 | 0.95115 | 0.9262 | 1.021 | 211.3702 | 49.8362 | 52594.7935 | 52716.079 | 1050.3216 | 950.3216 | 1119.2656
7666 | 97.8896 | 1.0118 | 0.98255 | 1.0206 | 195.9722 | 50.8638 | 56240.9315 | 54423.9365 | 1084.4787 | 984.4787 | 1119.3075
7667 | 93.2912 | 1.0138 | 0.97629 | 1.061 | 185.5295 | 50.356 | 57111.0986 | 56632.0494 | 1128.641 | 1028.641 | 1119.353
7668 | 88.5963 | 1.0481 | 0.92443 | 0.92953 | 214.6607 | 50.7393 | 48526.078 | 48240.9628 | 960.8193 | 860.8193 | 1119.4942
7669 | 102.5797 | 1.0815 | 1.1235 | 0.96176 | 188.6761 | 50.4361 | 60283.8927 | 60819.2485 | 1212.385 | 1112.385 | 1119.4986
7670 | 104.9526 | 1.0406 | 0.92298 | 0.97389 | 199.8922 | 51.4943 | 52111.2596 | 52019.3333 | 1036.3867 | 936.3867 | 1119.6222
7671 | 106.5265 | 1.0222 | 0.99881 | 1.0428 | 205.9853 | 46.4082 | 53597.5001 | 53287.4059 | 1061.7481 | 961.7481 | 1119.6867
7672 | 106.5724 | 1.0134 | 0.82902 | 1.0465 | 194.9626 | 49.7312 | 48709.1494 | 47644.1606 | 948.8832 | 848.8832 | 1119.7032
7673 | 95.4058 | 0.98362 | 0.98504 | 1.0335 | 195.6831 | 51.7701 | 57758.2608 | 56127.687 | 1118.5537 | 1018.5537 | 1119.8639
7674 | 102.9175 | 1.035 | 0.81608 | 1.0503 | 185.7681 | 49.6714 | 48052.5825 | 46485.3956 | 925.7079 | 825.7079 | 1119.9515
7675 | 101.9674 | 0.94277 | 0.98008 | 1.024 | 191.3442 | 45.2421 | 49944.5255 | 48067.1605 | 957.3432 | 857.3432 | 1119.9627
7676 | 102.5456 | 1.0458 | 0.84794 | 0.96304 | 219.3584 | 48.8061 | 45308.8496 | 46099.0214 | 917.9804 | 817.9804 | 1119.9911
7677 | 87.9504 | 0.97522 | 1.0785 | 0.95045 | 187.8265 | 50.9006 | 56731.1799 | 59790.6536 | 1191.8131 | 1091.8131 | 1120.0217
7678 | 107.9963 | 0.96002 | 0.74715 | 1.029 | 182.4882 | 49.3058 | 43203.4017 | 43023.1945 | 856.4639 | 756.4639 | 1120.2146
7679 | 102.0606 | 1.0284 | 0.87677 | 0.99708 | 196.3146 | 46.6927 | 45916.5549 | 44549.9614 | 886.9992 | 786.9992 | 1120.2165
7680 | 91.1838 | 1.0858 | 1.2236 | 0.97987 | 177.8933 | 50.6923 | 65973.4558 | 66534.4846 | 1326.6897 | 1226.6897 | 1120.3098
7681 | 100.763 | 1.085 | 0.96263 | 1.011 | 214.7283 | 49.0101 | 53269.6872 | 54587.1415 | 1087.7428 | 987.7428 | 1120.3611
7682 | 90.1776 | 1.0201 | 1.1622 | 0.98073 | 211.8588 | 55.3441 | 68382.5616 | 67673.5081 | 1349.4702 | 1249.4702 | 1120.606
7683 | 98.1351 | 0.99298 | 1.063 | 0.99743 | 225.4595 | 43.5925 | 50694.1968 | 52071.9463 | 1037.4389 | 937.4389 | 1120.6123
7684 | 98.0218 | 0.92176 | 1.1496 | 0.96605 | 211.8478 | 55.0589 | 66334.2547 | 66774.8374 | 1331.4967 | 1231.4967 | 1120.6362
7685 | 94.2045 | 0.98615 | 0.98865 | 0.89915 | 186.0688 | 47.549 | 46871.6285 | 48062.5604 | 957.2512 | 857.2512 | 1120.662
7686 | 107.8847 | 0.93317 | 1.1601 | 1.0122 | 195.7656 | 52.9921 | 67758.2375 | 66114.0822 | 1318.2816 | 1218.2816 | 1120.7439
7687 | 95.6687 | 1.0185 | 0.88928 | 1.0116 | 193.7157 | 50.1883 | 50233.6377 | 51163.0243 | 1019.2605 | 919.2605 | 1120.8117
7688 | 101.2758 | 0.92488 | 0.89359 | 0.98461 | 198.2107 | 44.6864 | 43700.6045 | 43104.9092 | 858.0982 | 758.0982 | 1120.8252
7689 | 106.0709 | 0.97312 | 1.222 | 0.96479 | 195.2118 | 53.2598 | 68486.5746 | 68299.2159 | 1361.9843 | 1261.9843 | 1120.8754
7690 | 113.8597 | 1.0531 | 0.92571 | 1.0521 | 219.6458 | 47.53 | 52210.1792 | 51708.9099 | 1030.1782 | 930.1782 | 1120.8771
7691 | 112.9178 | 1.0328 | 1.0527 | 0.98633 | 176.0818 | 50.212 | 58169.4788 | 58111.908 | 1158.2382 | 1058.2382 | 1120.9577
7692 | 91.577 | 0.98348 | 0.83507 | 0.95483 | 191.989 | 47.8884 | 42688.9514 | 42638.1414 | 848.7628 | 748.7628 | 1121.0342
7693 | 108.4499 | 0.98316 | 1.0332 | 1.0523 | 196.5724 | 50.991 | 61075.6211 | 61646.5229 | 1228.9305 | 1128.9305 | 1121.0375
7694 | 91.3973 | 1.0219 | 1.3919 | 1.0328 | 186.2517 | 51.2923 | 78710.2391 | 78944.1666 | 1574.8833 | 1474.8833 | 1121.0714
7695 | 106.896 | 0.98277 | 0.67573 | 1.0045 | 208.2849 | 50.778 | 40009.9334 | 39020.4809 | 776.4096 | 676.4096 | 1121.1329
7696 | 114.198 | 1.0502 | 1.0676 | 1.0145 | 202.0676 | 50.9519 | 61495.4566 | 60588.2369 | 1207.7647 | 1107.7647 | 1121.1883
7697 | 93.8038 | 1.0844 | 1.3429 | 0.94356 | 188.3038 | 51.1581 | 70216.8656 | 68559.8023 | 1367.196 | 1267.196 | 1121.2433
7698 | 113.8563 | 0.99743 | 1.0758 | 1.0861 | 194.8146 | 48.4915 | 62361.5778 | 63212.7318 | 1260.2546 | 1160.2546 | 1121.3069
7699 | 79.0233 | 0.94223 | 0.95314 | 1.0096 | 199.6092 | 53.5138 | 55680.1709 | 55217.845 | 1100.3569 | 1000.3569 | 1121.3999
7700 | 111.1051 | 0.98034 | 0.96998 | 1.0485 | 217.7944 | 50.5165 | 57098.188 | 56808.8378 | 1132.1768 | 1032.1768 | 1121.4036
7701 | 94.9945 | 0.94482 | 1.1439 | 0.91645 | 201.9595 | 52.6137 | 60079.7008 | 61760.3136 | 1231.2063 | 1131.2063 | 1121.4537
7702 | 102.3505 | 0.99995 | 1.0441 | 1.0084 | 192.2979 | 48.3188 | 56010.6085 | 53937.6175 | 1074.7524 | 974.7524 | 1121.5755
7703 | 95.5823 | 1.0309 | 1.1754 | 0.97864 | 185.9724 | 48.7808 | 61104.3371 | 62003.2838 | 1236.0657 | 1136.0657 | 1121.6427
7704 | 86.6612 | 1.0827 | 0.88714 | 0.982 | 190.0164 | 49.5152 | 47972.1979 | 47878.7861 | 953.5757 | 853.5757 | 1121.7091
7705 | 97.5548 | 0.98041 | 1.0324 | 1.0147 | 210.9571 | 47.9593 | 55038.8144 | 54265.4625 | 1081.3092 | 981.3092 | 1121.7193
7706 | 116.3093 | 0.98618 | 0.93618 | 0.97223 | 213.6419 | 44.2283 | 45542.5595 | 44335.2807 | 882.7056 | 782.7056 | 1121.7404
7707 | 99.86 | 0.95445 | 1.1104 | 0.91978 | 179.9039 | 49.9365 | 55943.1648 | 55659.3836 | 1109.1877 | 1009.1877 | 1121.8237
7708 | 74.7296 | 1.0067 | 1.0342 | 0.98493 | 174.6763 | 48.9793 | 53749.879 | 53763.7897 | 1071.2758 | 971.2758 | 1121.8337
7709 | 104.1024 | 1.0142 | 1.2576 | 1.0504 | 215.2829 | 48.4485 | 69329.5547 | 69009.7636 | 1376.1953 | 1276.1953 | 1121.8579
7710 | 100.5892 | 1.0443 | 1.0877 | 1.0408 | 221.0514 | 52.5778 | 65268.566 | 66080.1714 | 1317.6034 | 1217.6034 | 1121.87
7711 | 107.9929 | 0.94429 | 0.92037 | 0.90355 | 206.6083 | 43.9637 | 41250.0055 | 40623.9657 | 808.4793 | 708.4793 | 1121.995
7712 | 99.4533 | 1.062 | 0.94208 | 1.0266 | 193.5664 | 47.8081 | 51482.2219 | 52213.7943 | 1040.2759 | 940.2759 | 1122.0491
7713 | 97.5156 | 0.92901 | 1.2039 | 1.0024 | 179.073 | 46.8261 | 60930.181 | 61291.1829 | 1221.8237 | 1121.8237 | 1122.1402
7714 | 89.0413 | 1.0804 | 1.1139 | 1.0049 | 203.5382 | 50.9502 | 62139.0131 | 62282.0839 | 1241.6417 | 1141.6417 | 1122.177
7715 | 107.0891 | 1.0297 | 0.93919 | 1.063 | 191.7984 | 51.2916 | 57056.6284 | 56969.3136 | 1135.3863 | 1035.3863 | 1122.1886
7716 | 101.4555 | 0.9727 | 1.1847 | 0.94451 | 191.3199 | 47.6538 | 58216.3498 | 57916.7226 | 1154.3345 | 1054.3345 | 1122.2739
7717 | 94.8512 | 0.95108 | 1.0318 | 1.0486 | 199.8634 | 45.2153 | 53200.5241 | 53450.5048 | 1065.0101 | 965.0101 | 1122.5012
7718 | 84.7232 | 1.0416 | 0.80988 | 1.0541 | 207.669 | 47.3481 | 44806.8322 | 44300.2894 | 882.0058 | 782.0058 | 1122.5467
7719 | 100.6494 | 1.0232 | 1.0605 | 1.0292 | 203.1281 | 49.3449 | 59144.8133 | 57837.9916 | 1152.7598 | 1052.7598 | 1122.5896
7720 | 111.1641 | 1.1302 | 0.78554 | 1.0179 | 200.514 | 44.7328 | 41590.0082 | 42528.3355 | 846.5667 | 746.5667 | 1122.5921
7721 | 108.0445 | 0.87754 | 0.93198 | 1.0921 | 199.2927 | 49.1775 | 54913.7501 | 56875.3805 | 1133.5076 | 1033.5076 | 1122.6581
7722 | 92.8955 | 1.0415 | 0.98905 | 1.0452 | 198.0797 | 51.1308 | 58002.3263 | 57185.9475 | 1139.7189 | 1039.7189 | 1122.6754
7723 | 96.5044 | 1.0431 | 1.0288 | 0.98296 | 216.3371 | 53.2332 | 59408.6112 | 60695.9594 | 1209.9192 | 1109.9192 | 1122.7686
7724 | 85.5773 | 0.94195 | 0.84082 | 0.91913 | 193.846 | 50.1942 | 43031.142 | 44352.5711 | 883.0514 | 783.0514 | 1122.9431
7725 | 84.5718 | 0.99349 | 1.2882 | 0.92767 | 201.6664 | 50.1778 | 64383.3895 | 63625.1174 | 1268.5023 | 1168.5023 | 1123.0586
7726 | 88.3222 | 1.0274 | 0.85733 | 1.0202 | 193.1287 | 46.9698 | 45538.1598 | 44206.9116 | 880.1382 | 780.1382 | 1123.059
7727 | 87.7569 | 1.0175 | 1.2167 | 1.0485 | 193.5201 | 44.873 | 61447.5136 | 61831.1307 | 1232.6226 | 1132.6226 | 1123.0625
7728 | 113.72 | 1.0128 | 0.9184 | 1.0323 | 200.4111 | 49.7334 | 53079.4013 | 53778.8647 | 1071.5773 | 971.5773 | 1123.0694
7729 | 91.4368 | 0.92897 | 1.1254 | 0.98623 | 192.1278 | 49.6549 | 59521.7734 | 58606.0133 | 1168.1203 | 1068.1203 | 1123.0736
7730 | 93.566 | 0.9635 | 0.84343 | 1.0143 | 180.7105 | 52.168 | 49514.3709 | 49554.948 | 987.099 | 887.099 | 1123.0772
7731 | 107.3395 | 0.96979 | 0.9027 | 0.99407 | 189.799 | 54.3756 | 54643.6546 | 53926.1609 | 1074.5232 | 974.5232 | 1123.1852
7732 | 89.3863 | 0.93148 | 0.99041 | 1.011 | 204.9222 | 48.9396 | 53283.4974 | 53757.1734 | 1071.1435 | 971.1435 | 1123.2026
7733 | 92.3882 | 0.99174 | 0.80526 | 0.98854 | 208.8152 | 55.8699 | 49802.3681 | 48713.5183 | 970.2704 | 870.2704 | 1123.2064
7734 | 109.9967 | 1.0245 | 0.98831 | 0.9929 | 212.5189 | 49.7152 | 54599.8366 | 55630.0062 | 1108.6001 | 1008.6001 | 1123.2372
7735 | 98.8023 | 0.9565 | 0.82266 | 0.98546 | 181.7381 | 47.8109 | 43460.2579 | 42887.0094 | 853.7402 | 753.7402 | 1123.3253
7736 | 111.3995 | 1.0571 | 1.2945 | 1.0347 | 199.5137 | 50.2408 | 73407.8852 | 73358.2287 | 1463.1646 | 1363.1646 | 1123.4158
7737 | 99.5451 | 0.90545 | 1.1054 | 0.97033 | 196.1968 | 51.5017 | 60076.756 | 57299.7535 | 1141.9951 | 1041.9951 | 1123.4952
7738 | 109.2104 | 1.0628 | 1.0235 | 1.0344 | 197.622 | 49.0909 | 57870.6065 | 59176.6604 | 1179.5332 | 1079.5332 | 1123.5954
7739 | 102.8314 | 0.9154 | 1.1553 | 0.97263 | 199.3964 | 49.9649 | 61047.1467 | 61832.1783 | 1232.6436 | 1132.6436 | 1123.7732
7740 | 106.8426 | 1.1336 | 0.95899 | 1.0368 | 189.0684 | 51.0877 | 57171.4852 | 58541.8213 | 1166.8364 | 1066.8364 | 1123.9006
7741 | 98.5435 | 0.91904 | 1.1109 | 1.0403 | 205.3623 | 46.5995 | 58278.0434 | 58997.8811 | 1175.9576 | 1075.9576 | 1123.9072
7742 | 84.9472 | 0.97136 | 0.94041 | 0.99016 | 195.4578 | 54.3786 | 55317.2922 | 55269.1683 | 1101.3834 | 1001.3834 | 1123.9102
7743 | 102.1344 | 1.0252 | 1.0381 | 0.9819 | 196.4477 | 51.8346 | 58458.6492 | 56825.3158 | 1132.5063 | 1032.5063 | 1123.9121
7744 | 93.2243 | 0.93304 | 0.89273 | 0.97767 | 215.6453 | 46.8292 | 45161.242 | 43972.2592 | 875.4452 | 775.4452 | 1124.0152
7745 | 91.8531 | 1.0346 | 1.1497 | 0.91444 | 209.1124 | 47.1919 | 54309.4402 | 53887.7449 | 1073.7549 | 973.7549 | 1124.0154
7746 | 74.1111 | 0.95248 | 0.89305 | 0.93305 | 189.2516 | 53.0546 | 48142.2282 | 49369.5287 | 983.3906 | 883.3906 | 1124.0879
7747 | 107.8507 | 1.0108 | 0.9358 | 0.99304 | 199.8186 | 50.0236 | 52139.5859 | 52953.2897 | 1055.0658 | 955.0658 | 1124.0913
7748 | 106.599 | 1.0146 | 1.2187 | 1.1005 | 198.1513 | 45.5794 | 66255.4594 | 67176.8373 | 1339.5367 | 1239.5367 | 1124.1789
7749 | 89.2997 | 0.93702 | 1.0591 | 1.0215 | 192.5063 | 48.5586 | 56789.4712 | 56927.5249 | 1134.5505 | 1034.5505 | 1124.2282
7750 | 97.8377 | 1.0032 | 0.92254 | 0.9694 | 198.2894 | 46.6303 | 46476.888 | 47687.2125 | 949.7443 | 849.7443 | 1124.2594
7751 | 107.112 | 0.88844 | 1.0291 | 0.94026 | 198.8923 | 54.1969 | 57798.5081 | 58113.2506 | 1158.265 | 1058.265 | 1124.315
7752 | 110.8343 | 0.99544 | 0.93223 | 1.0451 | 178.8837 | 50.2827 | 54714.9586 | 54502.7862 | 1086.0557 | 986.0557 | 1124.35
7753 | 100.1564 | 0.9354 | 0.9892 | 0.97404 | 201.4485 | 51.3293 | 54467.2104 | 54431.9453 | 1084.6389 | 984.6389 | 1124.4535
7754 | 81.9365 | 0.91173 | 0.80988 | 1.0403 | 178.9914 | 51.2839 | 47217.0756 | 47013.4051 | 936.2681 | 836.2681 | 1124.5256
7755 | 105.3939 | 0.96055 | 0.99979 | 1.0278 | 200.7428 | 47.1766 | 53454.5203 | 54041.037 | 1076.8207 | 976.8207 | 1124.6038
7756 | 104.2345 | 1.0014 | 0.96575 | 1.0531 | 196.9247 | 54.6614 | 61493.0415 | 61056.0321 | 1217.1206 | 1117.1206 | 1124.7017
7757 | 99.8292 | 0.97797 | 0.77088 | 1.0568 | 203.7458 | 51.4266 | 47120.2606 | 46258.3531 | 921.1671 | 821.1671 | 1124.8011
7758 | 97.1307 | 1.015 | 0.91665 | 0.91984 | 197.778 | 51.1645 | 48382.2492 | 48367.7164 | 963.3543 | 863.3543 | 1124.8629
7759 | 85.6269 | 1.004 | 1.1086 | 1.0128 | 199.7683 | 49.7585 | 60345.2321 | 61415.7495 | 1224.315 | 1124.315 | 1125.0055
7760 | 109.7978 | 0.89552 | 0.82779 | 0.90741 | 206.4824 | 53.3583 | 45532.818 | 45736.4346 | 910.7287 | 810.7287 | 1125.0167
7761 | 106.8407 | 0.88753 | 1.2599 | 1.0257 | 175.1927 | 50.2112 | 69818.2103 | 69645.8008 | 1388.916 | 1288.916 | 1125.0993
7762 | 91.8204 | 1.0295 | 0.76146 | 0.9894 | 189.7807 | 51.7244 | 44047.7945 | 44614.7083 | 888.2942 | 788.2942 | 1125.1268
7763 | 103.9466 | 1.0101 | 1.2445 | 0.95409 | 194.3518 | 53.5789 | 69437.3349 | 69326.2833 | 1382.5257 | 1282.5257 | 1125.129
7764 | 109.0147 | 1.1009 | 0.90163 | 0.98403 | 205.7471 | 47.2779 | 47826.1363 | 46154.337 | 919.0867 | 819.0867 | 1125.1597
7765 | 85.1269 | 1.0439 | 0.73758 | 0.97157 | 203.2104 | 51.4298 | 41628.6473 | 41924.593 | 834.4919 | 734.4919 | 1125.1911
7766 | 101.1394 | 0.98105 | 0.73206 | 1.0065 | 208.9288 | 53.6331 | 45046.9869 | 45077.0004 | 897.54 | 797.54 | 1125.1965
7767 | 89.3562 | 1.008 | 0.93051 | 0.96438 | 180.922 | 49.5941 | 49152.3984 | 50417.818 | 1004.3564 | 904.3564 | 1125.2185
7768 | 102.031 | 0.9721 | 0.78082 | 1.0203 | 205.7982 | 48.6809 | 43817.8961 | 41722.3973 | 830.4479 | 730.4479 | 1125.2519
7769 | 108.6059 | 0.94176 | 1.2116 | 0.95021 | 201.0721 | 51.3767 | 64603.0023 | 64723.0002 | 1290.46 | 1190.46 | 1125.3475
7770 | 102.4444 | 1.0444 | 0.93748 | 0.96451 | 200.515 | 51.4644 | 52241.0904 | 51637.0418 | 1028.7408 | 928.7408 | 1125.5611
7771 | 85.4232 | 1.0634 | 0.74483 | 0.86542 | 202.708 | 51.5637 | 38123.6751 | 37479.7995 | 745.596 | 645.596 | 1125.6997
7772 | 84.1488 | 0.93541 | 1.1024 | 1.0005 | 214.4427 | 48.336 | 57332.7491 | 59177.4986 | 1179.55 | 1079.55 | 1125.7338
7773 | 99.1019 | 1.02 | 0.94044 | 1.0262 | 188.2655 | 48.7915 | 52208.78 | 51621.499 | 1028.43 | 928.43 | 1125.7421
7774 | 105.88 | 0.95138 | 1.0025 | 1.0178 | 212.3707 | 50.3833 | 56695.8755 | 56928.4358 | 1134.5687 | 1034.5687 | 1125.7423
7775 | 91.6721 | 0.93612 | 0.85717 | 1.0443 | 183.882 | 47.1719 | 46458.0229 | 47227.7273 | 940.5545 | 840.5545 | 1125.8089
7776 | 87.6309 | 0.94387 | 0.92362 | 1.0458 | 203.5707 | 54.3918 | 57239.9706 | 55681.7692 | 1109.6354 | 1009.6354 | 1125.924
7777 | 110.1514 | 1.0022 | 0.89449 | 1.0753 | 199.3741 | 46.3422 | 49888.5884 | 51469.3465 | 1025.3869 | 925.3869 | 1125.9431
7778 | 94.3547 | 0.95028 | 0.96444 | 1.0638 | 215.2823 | 49.6268 | 55582.683 | 53915.1322 | 1074.3026 | 974.3026 | 1126.0013
7779 | 95.1682 | 1.0451 | 0.89449 | 0.991 | 197.6022 | 50.7371 | 50219.5155 | 50807.9027 | 1012.1581 | 912.1581 | 1126.1196
7780 | 92.7239 | 1.1078 | 1.1384 | 0.91947 | 202.5982 | 54.3799 | 62708.2786 | 62265.2049 | 1241.3041 | 1141.3041 | 1126.2111
7781 | 103.5872 | 1.061 | 0.90859 | 0.97579 | 203.9406 | 47.2341 | 47273.0524 | 46772.4818 | 931.4496 | 831.4496 | 1126.2146
7782 | 92.1816 | 0.97909 | 0.98503 | 0.97388 | 194.195 | 46.0491 | 48525.1908 | 49444.1804 | 984.8836 | 884.8836 | 1126.2872
7783 | 86.4181 | 1.0067 | 1.0345 | 1.0242 | 224.0049 | 51.6641 | 59458.3522 | 61353.4721 | 1223.0694 | 1123.0694 | 1126.289
7784 | 100.8953 | 0.99544 | 0.96223 | 0.98809 | 201.8127 | 49.7459 | 52494.9718 | 49730.4392 | 990.6088 | 890.6088 | 1126.4499
7785 | 108.1885 | 0.98902 | 0.90305 | 0.88504 | 194.5945 | 49.5254 | 45076.2091 | 45093.8762 | 897.8775 | 797.8775 | 1126.4799
7786 | 100.9669 | 0.97346 | 0.93998 | 1.1049 | 195.5244 | 51.936 | 59241.4246 | 58628.3069 | 1168.5661 | 1068.5661 | 1126.5356
7787 | 91.4716 | 0.95956 | 1.0456 | 1.1239 | 193.4289 | 49.579 | 62805.006 | 62539.4681 | 1246.7894 | 1146.7894 | 1126.545
7788 | 87.8446 | 0.99608 | 0.97358 | 0.92028 | 189.195 | 53.7786 | 53078.6006 | 53221.3645 | 1060.4273 | 960.4273 | 1126.6064
7789 | 97.1157 | 1.0297 | 0.96865 | 0.95569 | 208.3456 | 44.7506 | 46110.5119 | 46944.255 | 934.8851 | 834.8851 | 1126.7391
7790 | 97.1997 | 1.0357 | 0.78984 | 0.89576 | 191.8592 | 52.1893 | 42370.3608 | 43684.8758 | 869.6975 | 769.6975 | 1126.804
7791 | 90.7802 | 0.91881 | 1.2197 | 1.0275 | 187.6873 | 51.2516 | 68691.4831 | 70436.5648 | 1404.7313 | 1304.7313 | 1126.8122
7792 | 101.1254 | 0.94984 | 0.85093 | 1.0435 | 201.7844 | 45.6868 | 45156.7947 | 42557.8876 | 847.1578 | 747.1578 | 1126.8761
7793 | 93.5733 | 1.0409 | 1.1309 | 1.0349 | 196.4989 | 46.4047 | 59028.358 | 59761.775 | 1191.2355 | 1091.2355 | 1126.9301
7794 | 104.7612 | 1.1144 | 0.98599 | 0.99606 | 205.7481 | 47.4562 | 52352.9259 | 50788.4484 | 1011.769 | 911.769 | 1126.9431
7795 | 100.0404 | 1.0105 | 1.1058 | 0.96377 | 195.3828 | 51.6987 | 60518.8733 | 61395.606 | 1223.9121 | 1123.9121 | 1127.0206
7796 | 92.4309 | 0.92626 | 1.0435 | 0.98674 | 207.1011 | 48.5434 | 54347.2607 | 54679.0528 | 1089.5811 | 989.5811 | 1127.0265
7797 | 81.8551 | 0.94745 | 1.0312 | 1.0584 | 202.6915 | 49.3222 | 57856.4173 | 59565.4689 | 1187.3094 | 1087.3094 | 1127.0327
7798 | 89.9503 | 0.92532 | 0.85427 | 1.0503 | 197.0762 | 53.884 | 53027.36 | 53493.0854 | 1065.8617 | 965.8617 | 1127.0891
7799 | 95.2208 | 0.92618 | 1.068 | 1.0012 | 197.4561 | 45.4282 | 52781.7031 | 53472.2951 | 1065.4459 | 965.4459 | 1127.197
7800 | 103.6078 | 0.98913 | 0.73608 | 0.9738 | 200.6298 | 49.9131 | 41093.0134 | 42729.6552 | 850.5931 | 750.5931 | 1127.221
7801 | 90.9731 | 1.02 | 0.75395 | 1.0079 | 192.4732 | 51.9565 | 44495.8126 | 44320.7507 | 882.415 | 782.415 | 1127.2426
7802 | 104.3441 | 0.96727 | 1.2426 | 1.0763 | 187.921 | 52.7254 | 76026.2762 | 78158.6036 | 1559.1721 | 1459.1721 | 1127.2724
7803 | 115.6727 | 0.89301 | 1.1595 | 1.0524 | 207.3279 | 52.7853 | 70069.3294 | 69061.8235 | 1377.2365 | 1277.2365 | 1127.3509
7804 | 115.7667 | 1.0194 | 0.77899 | 0.96504 | 211.8337 | 58.0869 | 50734.1776 | 51719.9885 | 1030.3998 | 930.3998 | 1127.4706
7805 | 106.2755 | 0.97746 | 1.0516 | 1.0273 | 214.0559 | 46.2827 | 55019.2 | 55484.9378 | 1105.6988 | 1005.6988 | 1127.5592
7806 | 91.091 | 0.99058 | 0.77754 | 1.007 | 200.7122 | 47.6975 | 41851.699 | 41559.9952 | 827.1999 | 727.1999 | 1127.5877
7807 | 118.5478 | 0.98164 | 0.89104 | 1.0529 | 213.9742 | 51.1092 | 54109.207 | 54032.4891 | 1076.6498 | 976.6498 | 1127.8256
7808 | 83.9947 | 0.98865 | 0.92718 | 0.9989 | 202.3439 | 49.2046 | 49859.8042 | 47849.5044 | 952.9901 | 852.9901 | 1127.893
7809 | 100.3541 | 1.0119 | 0.64217 | 1.0615 | 222.7881 | 51.4393 | 40510.0454 | 41128.4759 | 818.5695 | 718.5695 | 1127.9182
7810 | 96.2514 | 1.0631 | 1.0766 | 1.0294 | 201.3431 | 50.281 | 61071.6642 | 57898.5467 | 1153.9709 | 1053.9709 | 1128.0304
7811 | 98.3505 | 0.98259 | 0.82972 | 0.99838 | 210.8777 | 50.9077 | 47301.2573 | 46388.1132 | 923.7623 | 823.7623 | 1128.0848
7812 | 92.1251 | 0.95901 | 0.84991 | 0.93666 | 179.7232 | 49.3539 | 43829.7061 | 41758.8501 | 831.177 | 731.177 | 1128.1368
7813 | 94.1555 | 1.0187 | 1.0126 | 1.0365 | 196.4573 | 55.2507 | 63480.5926 | 64055.3598 | 1277.1072 | 1177.1072 | 1128.1902
7814 | 109.29 | 1.0526 | 1.3053 | 1.0083 | 187.352 | 48.9133 | 70190.6714 | 70428.9596 | 1404.5792 | 1304.5792 | 1128.2505
7815 | 102.225 | 0.95126 | 0.99819 | 1.0676 | 191.5826 | 46.6226 | 54409.6765 | 54235.8364 | 1080.7167 | 980.7167 | 1128.377
7816 | 93.6286 | 0.98504 | 1.015 | 0.92754 | 204.3051 | 52.9249 | 54910.3921 | 55784.6887 | 1111.6938 | 1011.6938 | 1128.4356
7817 | 99.4757 | 0.97583 | 0.967 | 1.0322 | 195.9408 | 47.803 | 52550.1314 | 53606.1631 | 1068.1233 | 968.1233 | 1128.4439
7818 | 114.5952 | 0.96271 | 0.95527 | 0.95426 | 181.2443 | 46.752 | 47957.168 | 47480.7098 | 945.6142 | 845.6142 | 1128.4555
7819 | 96.1812 | 1.0784 | 0.97919 | 0.8567 | 200.3379 | 48.7882 | 46187.5136 | 46598.8693 | 927.9774 | 827.9774 | 1128.4581
7820 | 82.5226 | 0.94951 | 0.94674 | 0.90038 | 194.8981 | 53.5162 | 50006.9007 | 49752.0666 | 991.0413 | 891.0413 | 1128.4772
7821 | 86.7475 | 0.9444 | 1.0295 | 0.97043 | 217.2071 | 52.2863 | 56735.8399 | 56828.8446 | 1132.5769 | 1032.5769 | 1128.4953
7822 | 113.7693 | 0.9895 | 1.0578 | 1.0582 | 213.0652 | 47.0386 | 58163.9995 | 57740.7957 | 1150.8159 | 1050.8159 | 1128.5106
7823 | 112.5457 | 0.95736 | 1.0237 | 1.018 | 195.6875 | 46.9545 | 54186.3623 | 56685.8538 | 1129.7171 | 1029.7171 | 1128.5748
7824 | 107.6883 | 0.95903 | 0.99944 | 1.0204 | 208.2444 | 44.7575 | 50473.9702 | 49899.6428 | 993.9929 | 893.9929 | 1128.5902
7825 | 108.9038 | 1.0311 | 0.93921 | 1.0026 | 190.8432 | 50.3837 | 53293.9686 | 53223.3483 | 1060.467 | 960.467 | 1128.6306
7826 | 92.1852 | 0.96386 | 1.1021 | 0.99503 | 194.6318 | 51.6708 | 61447.8152 | 60191.552 | 1199.831 | 1099.831 | 1128.6365
7827 | 104.7074 | 1.0312 | 1.185 | 1.0207 | 195.5587 | 54.4665 | 71951.854 | 74365.943 | 1483.3189 | 1383.3189 | 1128.6452
7828 | 99.1153 | 0.96637 | 1.197 | 0.9455 | 201.5307 | 50.1547 | 61767.1188 | 59777.8535 | 1191.5571 | 1091.5571 | 1128.7655
7829 | 96.7466 | 0.99513 | 1.0406 | 1.0503 | 195.6857 | 48.6362 | 58035.8188 | 57577.2522 | 1147.545 | 1047.545 | 1128.7803
7830 | 93.4297 | 0.9415 | 1.2605 | 1.0276 | 205.2334 | 55.6146 | 77132.9623 | 80428.0543 | 1604.5611 | 1504.5611 | 1128.7906
7831 | 100.7983 | 1.0299 | 0.89882 | 1.003 | 202.9582 | 52.5571 | 53038.2961 | 50791.1611 | 1011.8232 | 911.8232 | 1128.8367
7832 | 94.7533 | 0.97168 | 0.95611 | 1.0323 | 194.6143 | 48.6097 | 52648.9291 | 52670.9534 | 1049.4191 | 949.4191 | 1128.8588
7833 | 99.8838 | 0.98165 | 1.2424 | 0.98607 | 207.1645 | 51.1715 | 67915.0868 | 67549.0176 | 1346.9804 | 1246.9804 | 1128.8891
7834 | 77.4776 | 1.0258 | 1.3545 | 0.97964 | 203.7424 | 53.7845 | 75844.9062 | 74911.6631 | 1494.2333 | 1394.2333 | 1128.9004
7835 | 95.0328 | 1.065 | 1.0494 | 1.0487 | 186.7969 | 52.5755 | 63367.7794 | 64989.4738 | 1295.7895 | 1195.7895 | 1128.9006
7836 | 100.0229 | 1.0164 | 1.0032 | 1.0195 | 195.4788 | 46.7406 | 52749.002 | 51750.3875 | 1031.0078 | 931.0078 | 1128.9237
7837 | 101.7689 | 0.99409 | 1.0963 | 1.0844 | 205.4567 | 48.7074 | 63038.544 | 62103.955 | 1238.0791 | 1138.0791 | 1128.9256
7838 | 91.1349 | 1.053 | 1.1841 | 0.99879 | 205.8219 | 51.4533 | 65994.5166 | 64926.2317 | 1294.5246 | 1194.5246 | 1128.9305
7839 | 103.2114 | 1.0211 | 1.1541 | 1.0539 | 204.0177 | 48.8273 | 64740.8098 | 63865.6146 | 1273.3123 | 1173.3123 | 1128.9672
7840 | 89.1555 | 1.0827 | 1.1119 | 1.0468 | 203.5648 | 47.65 | 60263.5446 | 59955.3791 | 1195.1076 | 1095.1076 | 1129.0089
7841 | 113.5113 | 0.99766 | 0.87809 | 1.025 | 193.8309 | 50.5538 | 51420.4198 | 50107.6244 | 998.1525 | 898.1525 | 1129.1191
7842 | 98.1578 | 0.9991 | 1.0195 | 0.95755 | 205.7843 | 51.108 | 55111.8326 | 53614.2034 | 1068.2841 | 968.2841 | 1129.2256
7843 | 112.5119 | 1.0046 | 1.0261 | 1.0534 | 191.0621 | 50.7715 | 60807.3991 | 62243.9109 | 1240.8782 | 1140.8782 | 1129.2296
7844 | 105.8703 | 1.0743 | 1.192 | 0.91958 | 206.8358 | 48.3565 | 58712.4544 | 56678.9792 | 1129.5796 | 1029.5796 | 1129.2822
7845 | 93.7476 | 0.93502 | 1.0942 | 1.0268 | 205.7796 | 47.9081 | 58230.2864 | 59824.5162 | 1192.4903 | 1092.4903 | 1129.3359
7846 | 95.8548 | 0.99083 | 0.96646 | 1.0366 | 207.7857 | 47.6709 | 52491.6964 | 51872.6961 | 1033.4539 | 933.4539 | 1129.4267
7847 | 100.977 | 0.93808 | 1.0724 | 0.96398 | 216.2061 | 52.752 | 59748.6409 | 59957.1439 | 1195.1429 | 1095.1429 | 1129.4617
7848 | 84.8232 | 1.0461 | 0.99494 | 1.0503 | 203.6748 | 46.4034 | 52810.7117 | 53152.7847 | 1059.0557 | 959.0557 | 1129.5166
7849 | 100.5087 | 1.0323 | 0.7778 | 1.1433 | 205.0268 | 48.5055 | 48371.6343 | 47735.5036 | 950.7101 | 850.7101 | 1129.544
7850 | 81.8121 | 0.98638 | 1.0372 | 1.0418 | 187.9445 | 47.6517 | 55528.1875 | 56682.427 | 1129.6485 | 1029.6485 | 1129.7452
7851 | 98.5867 | 1.0299 | 0.98603 | 1.0936 | 197.0211 | 52.7396 | 62424.3367 | 62180.305 | 1239.6061 | 1139.6061 | 1129.8375
7852 | 108.8473 | 0.94169 | 0.85858 | 1.0036 | 196.4889 | 44.6374 | 43234.0429 | 45220.0887 | 900.4018 | 800.4018 | 1129.8979
7853 | 101.5041 | 0.91774 | 1.0248 | 1.0328 | 180.0254 | 51.1029 | 59032.5547 | 58393.322 | 1163.8664 | 1063.8664 | 1129.9014
7854 | 104.7757 | 0.99972 | 1.076 | 0.919 | 202.2231 | 51.4826 | 56500.6815 | 54632.7796 | 1088.6556 | 988.6556 | 1130.064
7855 | 98.4623 | 0.9983 | 0.8107 | 1.0061 | 184.2555 | 50.2621 | 46121.4547 | 45482.7611 | 905.6552 | 805.6552 | 1130.1402
7856 | 101.5692 | 1.0362 | 0.97845 | 0.97696 | 214.5907 | 53.3299 | 56805.9853 | 57275.8307 | 1141.5166 | 1041.5166 | 1130.1997
7857 | 103.2669 | 0.96466 | 0.96234 | 1.0653 | 186.3535 | 54.5692 | 61567.1941 | 60541.1157 | 1206.8223 | 1106.8223 | 1130.225
7858 | 96.8365 | 0.96975 | 0.90084 | 0.99226 | 205.1194 | 50.6106 | 50196.9329 | 51945.0528 | 1034.9011 | 934.9011 | 1130.2794
7859 | 98.837 | 0.97538 | 1.058 | 1.0404 | 197.4759 | 49.3268 | 59246.1796 | 58772.2862 | 1171.4457 | 1071.4457 | 1130.3619
7860 | 106.4964 | 0.95655 | 1.0968 | 0.93359 | 199.5068 | 49.3453 | 55754.4655 | 55890.0274 | 1113.8005 | 1013.8005 | 1130.3878
7861 | 104.899 | 0.98284 | 1.0877 | 0.99023 | 223.9708 | 48.4277 | 57376.1873 | 56171.3038 | 1119.4261 | 1019.4261 | 1130.4673
7862 | 90.9417 | 0.97865 | 0.94435 | 1.0716 | 185.3216 | 48.7243 | 53828.4545 | 55210.6979 | 1100.214 | 1000.214 | 1130.5947
7863 | 101.7427 | 1.1143 | 0.93343 | 1.0453 | 200.7054 | 50.5982 | 55306.7619 | 56837.6773 | 1132.7535 | 1032.7535 | 1130.6137
7864 | 87.9214 | 0.96621 | 0.6919 | 1.0459 | 188.7597 | 48.4292 | 39347.282 | 39823.7059 | 792.4741 | 692.4741 | 1130.6301
7865 | 113.1072 | 0.96735 | 0.8956 | 0.9824 | 200.6034 | 49.2808 | 48951.5577 | 49278.9304 | 981.5786 | 881.5786 | 1130.6835
7866 | 113.4704 | 0.99485 | 0.84532 | 1.0628 | 193.2041 | 50.743 | 51510.3978 | 53430.282 | 1064.6056 | 964.6056 | 1130.8378
7867 | 104.5821 | 1.0593 | 0.96312 | 1.0034 | 190.2485 | 49.1949 | 53179.4182 | 52106.1708 | 1038.1234 | 938.1234 | 1130.908
7868 | 112.0148 | 1.0463 | 1.0627 | 0.97694 | 216.3564 | 47.364 | 54939.1873 | 54544.5859 | 1086.8917 | 986.8917 | 1130.9455
7869 | 108.2847 | 1.0146 | 0.96514 | 0.99514 | 201.4211 | 48.1266 | 51712.1872 | 50416.9978 | 1004.34 | 904.34 | 1130.9506
7870 | 106.2894 | 1.024 | 1.1508 | 1.1273 | 200.414 | 49.1236 | 69275.1296 | 66161.106 | 1319.2221 | 1219.2221 | 1130.9514
7871 | 120.9807 | 1.0601 | 0.9965 | 0.99468 | 207.1943 | 53.353 | 59933.0197 | 58781.4544 | 1171.6291 | 1071.6291 | 1131.0041
7872 | 88.6536 | 0.95358 | 0.85569 | 1.0812 | 213.2771 | 47.5707 | 48247.2887 | 49258.8482 | 981.177 | 881.177 | 1131.0271
7873 | 93.6181 | 1.0105 | 0.89757 | 0.8987 | 195.3939 | 50.7537 | 45937.0832 | 46055.6581 | 917.1132 | 817.1132 | 1131.0889
7874 | 91.4065 | 0.95117 | 1.1491 | 0.99223 | 184.2427 | 50.3597 | 61983.7954 | 61910.7341 | 1234.2147 | 1134.2147 | 1131.1851
7875 | 85.2662 | 1.0305 | 0.98429 | 0.94495 | 198.1047 | 48.1557 | 49219.1008 | 49770.6341 | 991.4127 | 891.4127 | 1131.2063
7876 | 93.8373 | 0.99922 | 1.1737 | 0.97582 | 209.102 | 47.8044 | 59443.1322 | 58029.1097 | 1156.5822 | 1056.5822 | 1131.2523
7877 | 77.4345 | 1.025 | 0.97673 | 1.0282 | 195.3617 | 47.4371 | 51599.1161 | 51877.8426 | 1033.5569 | 933.5569 | 1131.3532
7878 | 101.7702 | 1.0006 | 0.98503 | 1.0088 | 196.413 | 50.8325 | 55885.1532 | 55648.7929 | 1108.9759 | 1008.9759 | 1131.3596
7879 | 103.5501 | 0.91017 | 1.206 | 1.0183 | 182.2326 | 51.2001 | 67885.7495 | 67348.7281 | 1342.9746 | 1242.9746 | 1131.375
7880 | 120.3446 | 1.1393 | 1.2981 | 1.0702 | 197.1032 | 50.8287 | 77777.0978 | 76872.238 | 1533.4448 | 1433.4448 | 1131.3776
7881 | 98.3166 | 1.0404 | 0.91638 | 0.97831 | 214.3898 | 49.7034 | 49857.764 | 49549.9326 | 986.9987 | 886.9987 | 1131.5429
7882 | 104.4162 | 0.97222 | 0.88201 | 1.0935 | 210.7983 | 49.7487 | 53244.4632 | 53278.0439 | 1061.5609 | 961.5609 | 1131.6049
7883 | 86.3774 | 1.0336 | 1.093 | 1.0035 | 212.5566 | 53.8158 | 64044.2388 | 64827.5491 | 1292.551 | 1192.551 | 1131.6864
7884 | 90.8972 | 1.0804 | 0.8781 | 0.95931 | 199.5872 | 51.8151 | 48935.2451 | 47716.7007 | 950.334 | 850.334 | 1131.7609
7885 | 108.5034 | 1.0121 | 0.94224 | 0.96273 | 213.1546 | 50.0409 | 51101.6463 | 51042.0857 | 1016.8417 | 916.8417 | 1132.0198
7886 | 111.7333 | 0.92135 | 1.2929 | 0.97222 | 195.2211 | 50.0551 | 68265.295 | 66728.1117 | 1330.5622 | 1230.5622 | 1132.0438
7887 | 85.499 | 1.0232 | 0.99419 | 1.0438 | 204.6565 | 49.0741 | 55422.899 | 54542.4349 | 1086.8487 | 986.8487 | 1132.0438
7888 | 95.855 | 1.0784 | 1.0441 | 1.0944 | 211.063 | 48.5542 | 60710.1879 | 61642.9401 | 1228.8588 | 1128.8588 | 1132.0869
7889 | 98.2727 | 1.0699 | 0.81944 | 0.97539 | 188.2222 | 49.8953 | 45314.1027 | 44095.4306 | 877.9086 | 777.9086 | 1132.1671
7890 | 107.3058 | 0.88671 | 0.91805 | 1.0223 | 201.1221 | 48.6456 | 50483.9446 | 49371.86 | 983.4372 | 883.4372 | 1132.2044
7891 | 88.6842 | 0.97145 | 0.7914 | 1.0127 | 237.5527 | 48.91 | 43650.6688 | 44026.7989 | 876.536 | 776.536 | 1132.3256
7892 | 92.7802 | 1.0218 | 0.81709 | 0.97332 | 191.5011 | 51.768 | 46269.9788 | 45897.884 | 913.9577 | 813.9577 | 1132.3805
7893 | 110.1914 | 0.98699 | 1.0911 | 1.0189 | 204.9124 | 52.1113 | 63805.8811 | 66453.959 | 1325.0792 | 1225.0792 | 1132.4186
7894 | 99.6886 | 0.97565 | 0.89558 | 1.0305 | 169.7939 | 50.7442 | 51934.9925 | 51576.9269 | 1027.5385 | 927.5385 | 1132.4509
7895 | 101.4024 | 1.0591 | 0.97557 | 0.95759 | 216.1245 | 48.2131 | 50434.1832 | 49970.9007 | 995.418 | 895.418 | 1132.4947
7896 | 110.4441 | 1.0135 | 0.96586 | 0.98267 | 176.3232 | 50.5942 | 53859.302 | 56211.7894 | 1120.2358 | 1020.2358 | 1132.5886
7897 | 100.1674 | 1.1209 | 0.8821 | 1.0153 | 182.8264 | 51.6043 | 52192.3116 | 52614.1793 | 1048.2836 | 948.2836 | 1132.6115
7898 | 103.2663 | 1.0941 | 0.94485 | 0.94914 | 203.0434 | 49.9229 | 50613.8527 | 51642.9512 | 1028.859 | 928.859 | 1132.6226
7899 | 94.5312 | 1.0512 | 1.0932 | 1.002 | 192.3007 | 48.8476 | 58557.8414 | 59637.708 | 1188.7542 | 1088.7542 | 1132.6436
7900 | 105.1607 | 0.92824 | 1.111 | 0.94597 | 192.1259 | 47.1047 | 54295.5533 | 53361.2127 | 1063.2243 | 963.2243 | 1132.6534
7901 | 94.5944 | 1.0299 | 0.91318 | 1.0094 | 181.3654 | 48.3546 | 49465.4646 | 49806.7148 | 992.1343 | 892.1343 | 1132.7191
7902 | 121.0226 | 1.0075 | 0.75594 | 0.98782 | 205.6732 | 55.7105 | 48599.3837 | 47961.6177 | 955.2324 | 855.2324 | 1132.7863
7903 | 92.0208 | 0.97671 | 1.0001 | 0.95927 | 189.0513 | 50.5817 | 53262.9733 | 52858.1845 | 1053.1637 | 953.1637 | 1132.7904
7904 | 109.5 | 1.0599 | 1.0842 | 0.97158 | 203.7296 | 51.0013 | 59845.3645 | 59565.3287 | 1187.3066 | 1087.3066 | 1132.9609
7905 | 118.731 | 0.93035 | 0.93477 | 1.0341 | 205.989 | 50.303 | 54385.7409 | 56519.9132 | 1126.3983 | 1026.3983 | 1133.0017
7906 | 97.1663 | 0.91989 | 1.0303 | 0.94351 | 200.3337 | 46.7908 | 49865.7537 | 49769.9101 | 991.3982 | 891.3982 | 1133.0263
7907 | 94.4697 | 0.9949 | 1.053 | 0.96044 | 200.9341 | 49.0645 | 54433.3855 | 54898.7466 | 1093.9749 | 993.9749 | 1133.0448
7908 | 100.0716 | 0.95199 | 0.95461 | 0.942 | 184.0722 | 51.4077 | 51309.6454 | 52431.726 | 1044.6345 | 944.6345 | 1133.1683
7909 | 81.6614 | 0.98282 | 1.091 | 1.0618 | 196.4714 | 47.6968 | 59274.5394 | 59167.6285 | 1179.3526 | 1079.3526 | 1133.2526
7910 | 93.197 | 1.0487 | 1.0322 | 1.1052 | 183.4169 | 48.0868 | 59739.0647 | 59411.674 | 1184.2335 | 1084.2335 | 1133.2671
7911 | 103.2303 | 0.9505 | 1.0442 | 1.0334 | 218.5169 | 52.3099 | 61794.0607 | 63900.0906 | 1274.0018 | 1174.0018 | 1133.3739
7912 | 85.5129 | 1.015 | 0.96848 | 1.0851 | 189.2389 | 54.0023 | 61624.7099 | 63063.9307 | 1257.2786 | 1157.2786 | 1133.4065
7913 | 106.4518 | 0.97408 | 0.92406 | 1.0519 | 205.8123 | 44.0041 | 47543.0527 | 48263.5224 | 961.2704 | 861.2704 | 1133.5122
7914 | 108.8515 | 0.94424 | 1.1323 | 0.97113 | 191.9599 | 49.3029 | 59474.4662 | 58174.1078 | 1159.4822 | 1059.4822 | 1133.5393
7915 | 92.6071 | 1.0091 | 0.73745 | 1.0094 | 203.6103 | 46.769 | 39389.1694 | 39481.1203 | 785.6224 | 685.6224 | 1133.6628
7916 | 112.4908 | 0.96026 | 1.1158 | 1.0082 | 204.0657 | 54.2501 | 67088.7225 | 66172.5552 | 1319.4511 | 1219.4511 | 1133.7597
7917 | 97.9789 | 1.0873 | 1.2746 | 1.017 | 175.4389 | 48.057 | 67592.0787 | 67569.1006 | 1347.382 | 1247.382 | 1133.8353
7918 | 86.0332 | 1.0112 | 1.1245 | 0.96194 | 213.128 | 46.0991 | 54089.9808 | 52008.4697 | 1036.1694 | 936.1694 | 1133.8798
7919 | 88.9625 | 1.0411 | 1.036 | 1.0626 | 204.5406 | 49.4183 | 59180.3508 | 60233.4615 | 1200.6692 | 1100.6692 | 1133.8861
7920 | 85.6837 | 0.93929 | 1.1573 | 1.0123 | 189.9865 | 51.785 | 65028.3987 | 65605.7916 | 1308.1158 | 1208.1158 | 1134.0162
7921 | 110.4523 | 0.93176 | 0.93533 | 0.99397 | 200.3402 | 50.1618 | 51997.4357 | 53617.0101 | 1068.3402 | 968.3402 | 1134.0489
7922 | 105.4989 | 0.9708 | 0.97959 | 0.95239 | 193.1953 | 49.4136 | 51354.4482 | 51807.6185 | 1032.1524 | 932.1524 | 1134.0695
7923 | 104.3618 | 1.1132 | 1.177 | 0.90791 | 207.7732 | 48.4204 | 57576.2387 | 55258.7248 | 1101.1745 | 1001.1745 | 1134.0808
7924 | 99.3784 | 0.91556 | 1.0568 | 1.0172 | 205.2973 | 53.0782 | 62097.025 | 63945.018 | 1274.9004 | 1174.9004 | 1134.101
7925 | 120.1717 | 0.93097 | 0.92018 | 0.92692 | 197.1888 | 46.954 | 45499.2563 | 44965.6374 | 895.3127 | 795.3127 | 1134.1245
7926 | 85.3757 | 1.0129 | 0.76727 | 1.0915 | 211.06 | 49.2561 | 45722.5972 | 45572.8673 | 907.4573 | 807.4573 | 1134.1603
7927 | 92.1668 | 0.98709 | 0.90669 | 1.0121 | 193.0766 | 49.6786 | 50298.9647 | 51108.2483 | 1018.165 | 918.165 | 1134.1649
7928 | 88.3179 | 1.0197 | 0.96468 | 0.97783 | 201.8131 | 48.4205 | 50237.4677 | 49545.9506 | 986.919 | 886.919 | 1134.1936
7929 | 93.2014 | 0.98387 | 0.98477 | 0.99139 | 210.0483 | 51.5732 | 55289.3731 | 53628.3806 | 1068.5676 | 968.5676 | 1134.2147
7930 | 99.0741 | 1.0016 | 1.0712 | 1.0398 | 185.9396 | 53.1754 | 64692.5273 | 63687.8765 | 1269.7575 | 1169.7575 | 1134.258
7931 | 93.2738 | 1.0305 | 1.0952 | 1.0125 | 191.4732 | 53.2312 | 64334.9199 | 67658.8639 | 1349.1773 | 1249.1773 | 1134.3245
7932 | 112.7226 | 1.0135 | 1.0048 | 0.99959 | 189.1361 | 49.5745 | 55643.1398 | 55324.9119 | 1102.4982 | 1002.4982 | 1134.3303
7933 | 113.0885 | 1.0487 | 0.95512 | 0.99434 | 204.4813 | 49.2101 | 52776.4403 | 53402.0003 | 1064.04 | 964.04 | 1134.3544
7934 | 104.249 | 1.0085 | 0.89572 | 0.88679 | 208.4471 | 49.1062 | 44377.0309 | 46770.8963 | 931.4179 | 831.4179 | 1134.536
7935 | 78.0909 | 0.99121 | 0.91026 | 0.91105 | 187.095 | 51.7232 | 47084.1376 | 47718.8553 | 950.3771 | 850.3771 | 1134.5392
7936 | 101.8851 | 0.99918 | 0.91361 | 0.97475 | 188.9639 | 48.8023 | 48617.9463 | 48349.01 | 962.9802 | 862.9802 | 1134.6867
7937 | 123.9395 | 1.0707 | 0.76402 | 1.0595 | 203.4835 | 51.6666 | 48882.8864 | 47747.0573 | 950.9411 | 850.9411 | 1134.8444
7938 | 91.5217 | 1.0727 | 0.91805 | 1.0722 | 202.6871 | 49.3697 | 53644.6703 | 51708.0695 | 1030.1614 | 930.1614 | 1134.8902
7939 | 106.1977 | 1.0534 | 1.0792 | 1.0294 | 176.2345 | 48.5709 | 59565.9653 | 59281.9405 | 1181.6388 | 1081.6388 | 1134.9332
7940 | 104.3759 | 1.0344 | 1.0017 | 1.0018 | 192.782 | 54.928 | 61239.6217 | 61641.8333 | 1228.8367 | 1128.8367 | 1134.9976
7941 | 106.4763 | 0.99962 | 1.0124 | 0.90495 | 195.5145 | 51.0479 | 52399.1999 | 50988.4852 | 1015.7697 | 915.7697 | 1135.0267
7942 | 121.1289 | 1.0857 | 0.80503 | 1.0479 | 197.0206 | 47.3996 | 46417.4795 | 46184.062 | 919.6812 | 819.6812 | 1135.1083
7943 | 94.7224 | 0.97297 | 1.2067 | 0.9295 | 200.6139 | 47.6936 | 58092.9496 | 58517.328 | 1166.3466 | 1066.3466 | 1135.198
7944 | 99.6859 | 0.97468 | 0.99737 | 1.1349 | 215.0497 | 47.9916 | 59200.04 | 58792.6845 | 1171.8537 | 1071.8537 | 1135.5014
7945 | 98.8413 | 1.051 | 1.0619 | 1.0109 | 196.5725 | 52.1173 | 61554.8698 | 59842.6424 | 1192.8528 | 1092.8528 | 1135.5306
7946 | 98.6366 | 0.94058 | 1.0235 | 0.95674 | 189.2666 | 51.1846 | 55057.246 | 57375.0587 | 1143.5012 | 1043.5012 | 1135.563
7947 | 100.0549 | 1.0398 | 0.85539 | 1.0516 | 208.7568 | 51.1704 | 51561.9281 | 51991.2307 | 1035.8246 | 935.8246 | 1135.5669
7948 | 103.7642 | 1.0536 | 0.98373 | 1.109 | 206.6076 | 53.4008 | 64302.1305 | 63713.4787 | 1270.2696 | 1170.2696 | 1135.6655
7949 | 98.652 | 0.96334 | 0.84605 | 0.99085 | 203.39 | 49.5639 | 46463.4316 | 46300.1491 | 922.003 | 822.003 | 1135.7019
7950 | 115.5743 | 1.0928 | 0.91744 | 1.0922 | 214.0654 | 47.9659 | 54333.7887 | 52486.916 | 1045.7383 | 945.7383 | 1135.76
7951 | 113.1524 | 1.003 | 0.79167 | 1.0123 | 191.7344 | 52.702 | 48407.7724 | 48300.776 | 962.0155 | 862.0155 | 1135.764
7952 | 84.1134 | 0.94615 | 0.75985 | 1.0933 | 213.5656 | 44.974 | 41153.0555 | 41309.4575 | 822.1891 | 722.1891 | 1135.8141
7953 | 81.9844 | 0.94898 | 0.90335 | 0.97228 | 192.0249 | 49.9649 | 47963.9592 | 48032.1858 | 956.6437 | 856.6437 | 1135.8429
7954 | 105.7147 | 0.97283 | 1.1556 | 0.94697 | 186.1738 | 46.3397 | 55663.3416 | 54991.0674 | 1095.8213 | 995.8213 | 1135.8598
7955 | 90.0817 | 0.96969 | 0.85011 | 0.99168 | 200.4233 | 50.0538 | 46769.6131 | 47491.9103 | 945.8382 | 845.8382 | 1135.9768
7956 | 114.0608 | 1.0427 | 0.92477 | 0.95455 | 191.2841 | 51.3378 | 51614.5932 | 51087.7206 | 1017.7544 | 917.7544 | 1136.0112
7957 | 100.119 | 1.0123 | 0.75108 | 1.012 | 206.9186 | 52.1143 | 45101.0836 | 46197.9586 | 919.9592 | 819.9592 | 1136.0418
7958 | 97.777 | 0.95499 | 0.98754 | 1.1131 | 203.2323 | 50.9674 | 60985.1321 | 61094.3447 | 1217.8869 | 1117.8869 | 1136.0657
7959 | 101.5956 | 1.0165 | 1.0378 | 1.0674 | 207.1069 | 51.0344 | 62010.6314 | 62568.1524 | 1247.363 | 1147.363 | 1136.1575
7960 | 96.3323 | 1.0618 | 0.96066 | 1.0075 | 209.9911 | 51.3664 | 55177.7723 | 54743.637 | 1090.8727 | 990.8727 | 1136.211
7961 | 106.6578 | 1.0372 | 0.91964 | 0.97929 | 200.0479 | 51.2383 | 52013.1231 | 53145.3877 | 1058.9078 | 958.9078 | 1136.2589
7962 | 91.1412 | 0.96082 | 1.2968 | 1.0619 | 197.103 | 46.7613 | 68683.8665 | 68308.5716 | 1362.1714 | 1262.1714 | 1136.3167
7963 | 112.4813 | 0.98309 | 1.0467 | 1.0958 | 216.1593 | 52.3085 | 65997.6414 | 65468.4235 | 1305.3685 | 1205.3685 | 1136.3208
7964 | 91.3295 | 0.97659 | 0.92795 | 1.0181 | 199.4674 | 47.244 | 49046.5783 | 48374.2929 | 963.4859 | 863.4859 | 1136.3397
7965 | 99.4379 | 1.0299 | 1.0231 | 1.0277 | 199.534 | 51.9544 | 60146.9756 | 59319.7566 | 1182.3951 | 1082.3951 | 1136.3761
7966 | 115.1004 | 1.0193 | 0.98843 | 0.99299 | 200.5839 | 49.8368 | 54962.9289 | 55705.3713 | 1110.1074 | 1010.1074 | 1136.3777
7967 | 101.3225 | 0.92961 | 0.92128 | 1.0203 | 192.8414 | 47.7243 | 49548.5001 | 49747.4116 | 990.9482 | 890.9482 | 1136.441
7968 | 103.1016 | 1.0253 | 1.0997 | 0.97186 | 217.7782 | 49.7026 | 58592.9687 | 59417.6211 | 1184.3524 | 1084.3524 | 1136.4566
7969 | 88.9192 | 0.9441 | 1.0912 | 0.88583 | 201.2477 | 45.9072 | 48428.0333 | 48328.8277 | 962.5766 | 862.5766 | 1136.4893
7970 | 100.8035 | 1.0234 | 1.0039 | 0.98102 | 179.6734 | 47.686 | 52061.7374 | 53004.8338 | 1056.0967 | 956.0967 | 1136.5311
7971 | 91.1804 | 1.0377 | 0.96755 | 1.1148 | 196.3747 | 49.6567 | 58453.2837 | 59394.64 | 1183.8928 | 1083.8928 | 1136.7938
7972 | 115.3633 | 0.97886 | 0.95039 | 0.98062 | 206.4388 | 48.6337 | 51023.5149 | 50438.2907 | 1004.7658 | 904.7658 | 1136.865
7973 | 100.4316 | 1.0173 | 1.3614 | 1.1266 | 193.4617 | 54.5304 | 89400.9146 | 88665.5286 | 1769.3106 | 1669.3106 | 1136.8971
7974 | 96.626 | 1.0267 | 1.2434 | 0.96767 | 202.4904 | 49.2638 | 64365.1288 | 64120.0174 | 1278.4003 | 1178.4003 | 1136.9761
7975 | 100.8266 | 0.99513 | 0.98461 | 1.0128 | 185.5772 | 50.0571 | 55126.8839 | 55987.8357 | 1115.7567 | 1015.7567 | 1137.0309
7976 | 102.1979 | 0.97861 | 1.138 | 0.94411 | 200.5428 | 50.0229 | 58949.6074 | 57575.3923 | 1147.5078 | 1047.5078 | 1137.0606
7977 | 74.2218 | 1.0247 | 0.7559 | 1.0542 | 200.9112 | 48.9613 | 42940.9302 | 44937.1642 | 894.7433 | 794.7433 | 1137.0781
7978 | 93.5958 | 0.99315 | 1.1929 | 0.95285 | 195.6047 | 47.9853 | 59198.6927 | 58819.844 | 1172.3969 | 1072.3969 | 1137.1865
7979 | 104.3904 | 1.0093 | 1.123 | 1.0095 | 198.0688 | 51.0754 | 63483.2433 | 63862.9877 | 1273.2598 | 1173.2598 | 1137.202
7980 | 101.9149 | 0.96438 | 1.2026 | 0.90373 | 202.4075 | 49.8168 | 59239.7937 | 59100.3968 | 1178.0079 | 1078.0079 | 1137.2523
7981 | 116.627 | 0.95639 | 0.81294 | 0.92498 | 204.3325 | 49.0754 | 42580.5491 | 42594.0113 | 847.8802 | 747.8802 | 1137.3528
7982 | 115.7123 | 0.97102 | 1.0074 | 1.0042 | 201.8004 | 51.4164 | 57993.2054 | 54969.2716 | 1095.3854 | 995.3854 | 1137.4168
7983 | 117.8528 | 0.96547 | 0.9461 | 0.93153 | 199.7757 | 56.6695 | 56591.6977 | 56567.7446 | 1127.3549 | 1027.3549 | 1137.4447
7984 | 99.5749 | 1.0228 | 0.71387 | 0.99703 | 192.2188 | 45.1493 | 36925.5286 | 36728.0837 | 730.5617 | 630.5617 | 1137.4738
7985 | 98.5643 | 0.93457 | 0.97226 | 0.9758 | 181.282 | 49.9348 | 52155.9131 | 53587.5561 | 1067.7511 | 967.7511 | 1137.4812
7986 | 93.2527 | 1.0576 | 0.81073 | 1.006 | 200.9514 | 52.0236 | 47763.1492 | 47626.463 | 948.5293 | 848.5293 | 1137.6666
7987 | 108.9951 | 0.99337 | 0.99125 | 1.023 | 183.1886 | 53.1864 | 59873.2781 | 58546.6713 | 1166.9334 | 1066.9334 | 1137.6688
7988 | 104.1374 | 0.97059 | 0.84536 | 1.0177 | 196.0136 | 49.6921 | 47969.9702 | 47694.3606 | 949.8872 | 849.8872 | 1137.7245
7989 | 96.5132 | 1.0016 | 1.0508 | 0.94587 | 196.3887 | 49.0992 | 53743.0787 | 52958.8022 | 1055.176 | 955.176 | 1137.7398
7990 | 103.8877 | 1.0145 | 1.1835 | 0.95264 | 202.4668 | 52.1281 | 64467.76 | 64229.1128 | 1280.5823 | 1180.5823 | 1137.781
7991 | 106.1186 | 1.0061 | 1.0618 | 1.0638 | 206.9443 | 48.3362 | 59966.0541 | 60785.0942 | 1211.7019 | 1111.7019 | 1137.8309
7992 | 90.8249 | 0.93458 | 1.1832 | 0.91494 | 198.4648 | 47.3177 | 55439.2621 | 55788.4607 | 1111.7692 | 1011.7692 | 1137.9117
7993 | 100.0689 | 1.0206 | 0.78603 | 0.96284 | 214.7669 | 51.0862 | 44095.1296 | 44103.6367 | 878.0727 | 778.0727 | 1137.9698
7994 | 82.8921 | 1.0104 | 1.1255 | 1.0474 | 207.8211 | 49.6729 | 62923.2161 | 61160.7305 | 1219.2146 | 1119.2146 | 1138.024
7995 | 82.0256 | 1.0304 | 0.95361 | 1.0193 | 186.8807 | 49.1304 | 52096.4431 | 52374.0502 | 1043.481 | 943.481 | 1138.0383
7996 | 97.6877 | 0.97184 | 1.1549 | 0.96922 | 181.7449 | 48.7245 | 59345.9111 | 60848.716 | 1212.9743 | 1112.9743 | 1138.0791
7997 | 116.1401 | 1.0953 | 0.9861 | 1.0201 | 202.1498 | 49.2358 | 55994.6347 | 55561.0631 | 1107.2213 | 1007.2213 | 1138.1897
7998 | 98.7849 | 0.99016 | 0.86364 | 1.0077 | 198.4002 | 50.4098 | 48999.7525 | 47148.6385 | 938.9728 | 838.9728 | 1138.1925
7999 | 99.5763 | 0.99635 | 0.76121 | 1.0483 | 188.0068 | 47.4891 | 42794.9255 | 41321.4572 | 822.4291 | 722.4291 | 1138.2115
8000 | 102.9522 | 1.0912 | 1.376 | 1.0032 | 191.3658 | 49.0195 | 73364.256 | 74955.5752 | 1495.1115 | 1395.1115 | 1138.2131
8001 | 91.7011 | 1.0995 | 1.1961 | 1.0137 | 184.4283 | 48.5918 | 64001.4805 | 64794.7634 | 1291.8953 | 1191.8953 | 1138.2159
8002 | 92.4982 | 0.93715 | 0.84492 | 0.92956 | 205.5261 | 48.1949 | 42236.0217 | 44455.3474 | 885.1069 | 785.1069 | 1138.2886
8003 | 81.1637 | 1.0166 | 0.84125 | 1.0794 | 203.6818 | 46.822 | 46585.2978 | 48929.7366 | 974.5947 | 874.5947 | 1138.4741
8004 | 99.2996 | 1.0701 | 0.78226 | 1.0782 | 216.1161 | 51.0693 | 48717.4858 | 50447.022 | 1004.9404 | 904.9404 | 1138.4846
8005 | 89.76 | 1.096 | 1.0092 | 0.94494 | 221.0143 | 52.5445 | 55497.5971 | 56336.3135 | 1122.7263 | 1022.7263 | 1138.5369
8006 | 84.0321 | 1.0919 | 0.91473 | 0.93955 | 197.7656 | 50.5236 | 48255.1193 | 49296.7449 | 981.9349 | 881.9349 | 1138.5509
8007 | 115.6768 | 1.0041 | 0.98289 | 1.0218 | 184.3006 | 45.2347 | 50866.391 | 50714.4151 | 1010.2883 | 910.2883 | 1138.5916
8008 | 111.0936 | 1.0343 | 1.0786 | 0.90821 | 181.0435 | 47.6391 | 52322.3856 | 53141.4159 | 1058.8283 | 958.8283 | 1138.597
8009 | 123.0509 | 1.001 | 1.0104 | 1.0855 | 189.8885 | 53.813 | 65839.8643 | 66582.4562 | 1327.6491 | 1227.6491 | 1138.6641
8010 | 78.1446 | 0.92337 | 1.2353 | 0.92786 | 198.2525 | 49.6084 | 60636.9288 | 62791.9632 | 1251.8393 | 1151.8393 | 1138.6786
8011 | 96.3091 | 1.0147 | 0.83571 | 0.94941 | 201.2816 | 49.4019 | 44226.3791 | 42606.9028 | 848.1381 | 748.1381 | 1138.6857
8012 | 99.9362 | 1.0153 | 1.3878 | 1.1365 | 201.2981 | 49.4895 | 83277.5496 | 83065.4737 | 1657.3095 | 1557.3095 | 1138.707
8013 | 94.3398 | 0.99044 | 1.1328 | 0.95532 | 198.1975 | 55.1731 | 65059.9483 | 64812.7162 | 1292.2543 | 1192.2543 | 1138.8024
8014 | 108.9974 | 0.95594 | 1.0189 | 1.0558 | 193.8104 | 47.5612 | 56311.0014 | 56200.1288 | 1120.0026 | 1020.0026 | 1138.8915
8015 | 101.2612 | 0.98292 | 0.76199 | 0.97968 | 174.911 | 50.9038 | 43241.3925 | 42456.3143 | 845.1263 | 745.1263 | 1139.0172
8016 | 92.7837 | 1.0969 | 0.94616 | 0.93693 | 199.3497 | 52.4166 | 52000.8629 | 52192.3121 | 1039.8462 | 939.8462 | 1139.0643
8017 | 109.0768 | 0.95905 | 0.98733 | 1.0517 | 186.5319 | 46.8929 | 53782.5171 | 54367.4029 | 1083.3481 | 983.3481 | 1139.1378
8018 | 103.6201 | 1.064 | 0.91077 | 1.0196 | 206.1459 | 50.5809 | 52751.0142 | 52527.578 | 1046.5516 | 946.5516 | 1139.1516
8019 | 103.9152 | 0.93494 | 0.98991 | 1.0254 | 190.6874 | 51.1665 | 57097.061 | 58013.194 | 1156.2639 | 1056.2639 | 1139.1694
8020 | 124.4746 | 0.93687 | 1.2715 | 1.0181 | 189.7389 | 50.5384 | 71506.1539 | 72265.8018 | 1441.316 | 1341.316 | 1139.1845
8021 | 84.0026 | 0.9874 | 1.2641 | 1.0725 | 210.52 | 47.959 | 69204.3938 | 67518.0451 | 1346.3609 | 1246.3609 | 1139.2442
8022 | 106.7954 | 0.98692 | 0.77317 | 0.92794 | 191.3966 | 47.6425 | 39394.0882 | 39030.3554 | 776.6071 | 676.6071 | 1139.2738
8023 | 111.2438 | 0.95592 | 0.98635 | 1.046 | 216.2986 | 50.1739 | 57314.8976 | 56988.8943 | 1135.7779 | 1035.7779 | 1139.499
8024 | 90.5924 | 1.0238 | 1.048 | 0.83004 | 214.6289 | 49.4613 | 47828.9407 | 48923.4143 | 974.4683 | 874.4683 | 1139.6061
8025 | 97.3463 | 0.93158 | 1.0938 | 0.92321 | 205.3134 | 56.8732 | 62793.6011 | 62483.8805 | 1245.6776 | 1145.6776 | 1139.6311
8026 | 86.3999 | 0.97518 | 1.0476 | 0.91753 | 200.211 | 51.5329 | 54077.3526 | 54296.1287 | 1081.9226 | 981.9226 | 1139.6675
8027 | 96.2225 | 1.0429 | 0.9768 | 1.028 | 201.6248 | 52.6427 | 58344.5768 | 57817.4889 | 1152.3498 | 1052.3498 | 1139.6693
8028 | 119.0844 | 0.91769 | 1.0121 | 0.99476 | 185.0786 | 50.5612 | 56615.1077 | 56847.3935 | 1132.9479 | 1032.9479 | 1139.7509
8029 | 93.8206 | 1.046 | 1.0832 | 0.95861 | 195.0005 | 46.8328 | 53422.5785 | 54832.5037 | 1092.6501 | 992.6501 | 1139.7939
8030 | 111.0119 | 0.94009 | 0.97046 | 1.0374 | 205.5899 | 52.8191 | 58895.5322 | 60072.2442 | 1197.4449 | 1097.4449 | 1139.795
8031 | 104.7966 | 1.0897 | 0.91978 | 0.99785 | 216.8241 | 49.3802 | 51176.8079 | 50096.6561 | 997.9331 | 897.9331 | 1139.8316
8032 | 99.3273 | 1.0906 | 1.0765 | 0.99837 | 195.3911 | 49.4787 | 58729.8322 | 58175.0649 | 1159.5013 | 1059.5013 | 1139.8792
8033 | 98.0778 | 0.98805 | 1.1745 | 1.0412 | 184.9067 | 46.7799 | 61920.732 | 63105.607 | 1258.1121 | 1158.1121 | 1139.9472
8034 | 98.4971 | 0.8593 | 0.80928 | 0.94349 | 196.9519 | 47.9143 | 40837.3786 | 41479.682 | 825.5936 | 725.5936 | 1139.9579
8035 | 106.2784 | 0.98981 | 0.8882 | 0.95808 | 202.8202 | 49.4833 | 47516.6372 | 46068.3088 | 917.3662 | 817.3662 | 1140.0427
8036 | 100.453 | 1.0179 | 0.93342 | 1.0349 | 206.6476 | 46.9629 | 50374.3355 | 50191.1293 | 999.8226 | 899.8226 | 1140.0823
8037 | 107.7322 | 1.0129 | 0.90492 | 1.0509 | 198.4189 | 48.652 | 51775.7291 | 52732.9367 | 1050.6587 | 950.6587 | 1140.294
8038 | 99.3638 | 0.9248 | 0.95568 | 0.9713 | 194.3005 | 49.5957 | 50788.6223 | 51878.1263 | 1033.5625 | 933.5625 | 1140.3396
8039 | 89.0702 | 0.97324 | 0.9162 | 0.96224 | 207.6701 | 53.3941 | 51908.2354 | 51827.4328 | 1032.5487 | 932.5487 | 1140.3659
8040 | 97.4707 | 1.0145 | 1.2376 | 1.1373 | 181.2118 | 49.5728 | 74859.8448 | 73776.4429 | 1471.5289 | 1371.5289 | 1140.3773
8041 | 92.7484 | 0.99334 | 0.86638 | 0.987 | 209.1639 | 44.8417 | 42685.4425 | 42445.0849 | 844.9017 | 744.9017 | 1140.4241
8042 | 100.2518 | 0.9434 | 1.2254 | 1.0267 | 213.7583 | 50.4176 | 68414.1397 | 70666.6831 | 1409.3337 | 1309.3337 | 1140.6989
8043 | 98.2942 | 1.0022 | 1.048 | 0.99894 | 214.4766 | 52.9776 | 60896.3789 | 59055.5755 | 1177.1115 | 1077.1115 | 1140.7678
8044 | 106.0435 | 0.92449 | 1.0067 | 1.0226 | 199.3726 | 51.2734 | 58010.6418 | 58230.131 | 1160.6026 | 1060.6026 | 1140.8399
8045 | 94.1302 | 0.96739 | 1.1442 | 1.0703 | 193.4833 | 48.7047 | 64275.4242 | 64035.0352 | 1276.7007 | 1176.7007 | 1140.8782
8046 | 97.303 | 1.1227 | 0.769 | 1.0538 | 196.1052 | 54.2738 | 50106.2206 | 50615.4725 | 1008.3094 | 908.3094 | 1140.8929
8047 | 90.4684 | 1.046 | 1.2525 | 0.94045 | 202.0791 | 50.8524 | 64913.1818 | 64465.4138 | 1285.3083 | 1185.3083 | 1140.9018
8048 | 87.868 | 0.91011 | 1.0857 | 1.0722 | 200.5641 | 50.2273 | 62686.7119 | 61902.4438 | 1234.0489 | 1134.0489 | 1141.0373
8049 | 79.1919 | 0.91851 | 1.0724 | 0.97663 | 194.6029 | 50.2743 | 56507.5181 | 56697.9601 | 1129.9592 | 1029.9592 | 1141.0528
8050 | 95.7859 | 0.92726 | 0.96539 | 0.98862 | 195.3319 | 52.0216 | 54465.3462 | 55079.4411 | 1097.5888 | 997.5888 | 1141.0677
8051 | 98.4843 | 0.98542 | 0.95588 | 1.116 | 226.616 | 48.6776 | 56879.3256 | 55579.4413 | 1107.5888 | 1007.5888 | 1141.0855
8052 | 100.052 | 1.0918 | 0.76864 | 1.0526 | 199.7509 | 48.822 | 45032.8787 | 45395.5405 | 903.9108 | 803.9108 | 1141.2338
8053 | 89.8506 | 1.0308 | 0.99784 | 1.027 | 184.3287 | 49.4291 | 55415.016 | 56543.8937 | 1126.8779 | 1026.8779 | 1141.3041
8054 | 96.756 | 0.96456 | 1.0014 | 1.0625 | 197.9092 | 50.0405 | 58110.7064 | 58807.2714 | 1172.1454 | 1072.1454 | 1141.3325
8055 | 85.7062 | 1.0346 | 1.0023 | 1.0194 | 179.0782 | 50.43 | 56175.222 | 56832.8554 | 1132.6571 | 1032.6571 | 1141.3486
8056 | 91.2877 | 1.0463 | 0.93017 | 1.0279 | 204.9944 | 49.5955 | 52362.4459 | 52240.5698 | 1040.8114 | 940.8114 | 1141.3603
8057 | 87.4818 | 0.96369 | 1.0414 | 1.0568 | 194.4063 | 50.4832 | 60007.84 | 59755.3508 | 1191.107 | 1091.107 | 1141.4392
8058 | 104.6034 | 0.9499 | 1.0206 | 1.0089 | 190.1153 | 46.9048 | 53147.3063 | 53281.7435 | 1061.6349 | 961.6349 | 1141.5453
8059 | 92.075 | 0.98594 | 0.95964 | 0.97206 | 193.207 | 53.3758 | 54829.0431 | 55518.1213 | 1106.3624 | 1006.3624 | 1141.5727
8060 | 98.5954 | 1.0234 | 1.0722 | 1.0066 | 219.1493 | 50.9561 | 60353.5459 | 60731.2254 | 1210.6245 | 1110.6245 | 1141.5736
8061 | 101.27 | 1.0246 | 0.82535 | 0.98571 | 211.8819 | 51.0969 | 47084.055 | 44737.3935 | 890.7479 | 790.7479 | 1141.5743
8062 | 100.6828 | 0.9367 | 1.2732 | 0.89122 | 187.2979 | 50.7365 | 62544.7903 | 62391.5622 | 1243.8312 | 1143.8312 | 1141.6417
8063 | 103.6037 | 1.0355 | 0.9582 | 1.0171 | 209.046 | 53.2496 | 57816.8896 | 58083.0042 | 1157.6601 | 1057.6601 | 1141.6625
8064 | 111.0052 | 1.0136 | 1.237 | 1.0362 | 208.2434 | 54.9346 | 76801.6791 | 76775.2312 | 1531.5046 | 1431.5046 | 1141.7271
8065 | 117.0473 | 1.007 | 0.82092 | 1.0136 | 192.6934 | 50.226 | 47904.352 | 47882.0802 | 953.6416 | 853.6416 | 1141.7656
8066 | 92.5245 | 1.0162 | 0.78463 | 0.99175 | 212.094 | 47.808 | 41909.0407 | 44184.1422 | 879.6828 | 779.6828 | 1141.8556
8067 | 124.5426 | 0.94667 | 1.2835 | 1.0063 | 187.733 | 49.2138 | 69554.0654 | 71117.3073 | 1418.3461 | 1318.3461 | 1141.8622
8068 | 91.8347 | 0.99989 | 1.1161 | 1.0567 | 204.3605 | 51.4746 | 65639.104 | 65561.0982 | 1307.222 | 1207.222 | 1141.8974
8069 | 106.5553 | 1.0036 | 0.86469 | 0.99136 | 210.0111 | 52.4687 | 50797.8376 | 51344.9306 | 1022.8986 | 922.8986 | 1142.1277
8070 | 92.1905 | 0.98821 | 1.2339 | 1.0915 | 183.9529 | 50.2454 | 72428.3297 | 72937.1672 | 1454.7433 | 1354.7433 | 1142.4225
8071 | 92.3862 | 0.8255 | 1.04 | 1.0395 | 188.8988 | 50.8555 | 59047.9368 | 61366.2625 | 1223.3253 | 1123.3253 | 1142.5038
8072 | 98.4348 | 0.94673 | 1.1386 | 0.98681 | 193.6834 | 54.8928 | 66985.179 | 66901.193 | 1334.0239 | 1234.0239 | 1142.5628
8073 | 102.7084 | 0.93215 | 0.87404 | 1.0208 | 206.7302 | 49.6647 | 49273.478 | 50175.0499 | 999.501 | 899.501 | 1142.6744
8074 | 103.9237 | 1.0507 | 1.2444 | 1.0219 | 195.4574 | 50.0216 | 69268.1091 | 69341.0028 | 1382.8201 | 1282.8201 | 1142.6781
8075 | 88.8024 | 0.97023 | 1.239 | 1.032 | 200.004 | 45.1689 | 61846.5251 | 62593.3223 | 1247.8664 | 1147.8664 | 1142.6966
8076 | 102.2613 | 1.0336 | 1.1224 | 0.99423 | 207.3334 | 46.9011 | 57504.1075 | 57402.2749 | 1144.0455 | 1044.0455 | 1142.6974
8077 | 99.6506 | 0.98197 | 0.84535 | 0.9422 | 211.5173 | 46.9321 | 42184.9751 | 41663.7071 | 829.2741 | 729.2741 | 1142.7762
8078 | 103.161 | 1.0154 | 0.86745 | 0.97223 | 204.5162 | 50.9727 | 48532.0892 | 49458.7401 | 985.1748 | 885.1748 | 1143.0614
8079 | 118.6541 | 0.96063 | 0.87956 | 1.0222 | 201.7911 | 49.7665 | 50617.7331 | 52050.5434 | 1037.0109 | 937.0109 | 1143.3502
8080 | 97.137 | 0.97736 | 1.0651 | 0.9748 | 181.4926 | 51.2735 | 58285.981 | 58325.6223 | 1162.5124 | 1062.5124 | 1143.4984
8081 | 113.5599 | 1.0336 | 0.64288 | 1.098 | 209.0973 | 47.4556 | 39276.275 | 39073.7216 | 777.4744 | 677.4744 | 1143.6008
8082 | 91.7806 | 0.97966 | 0.77351 | 0.89395 | 199.2192 | 49.9938 | 39264.291 | 39106.4716 | 778.1294 | 678.1294 | 1143.766
8083 | 114.5285 | 0.99929 | 0.92557 | 0.94188 | 203.7034 | 48.6778 | 48211.0129 | 49480.8113 | 985.6162 | 885.6162 | 1143.8312
8084 | 111.4609 | 1.1255 | 1.2044 | 0.99645 | 199.3476 | 49.8668 | 66301.4505 | 63110.5639 | 1258.2113 | 1158.2113 | 1143.8456
8085 | 91.0541 | 0.96375 | 0.95324 | 0.9797 | 215.8375 | 49.8369 | 51131.7951 | 51155.0905 | 1019.1018 | 919.1018 | 1143.8474
8086 | 99.5126 | 1.0071 | 0.93206 | 1.0884 | 183.7256 | 49.1557 | 54978.4524 | 56343.7212 | 1122.8744 | 1022.8744 | 1143.9579
8087 | 97.1939 | 0.92116 | 0.8979 | 0.98348 | 196.3354 | 53.4501 | 52181.7295 | 52458.8715 | 1045.1774 | 945.1774 | 1144.0235
8088 | 104.3106 | 1.0014 | 0.72667 | 0.99568 | 176.232 | 49.2254 | 40934.1725 | 40547.9039 | 806.9581 | 706.9581 | 1144.1354
8089 | 126.9372 | 1.0343 | 1.0474 | 0.99467 | 206.4193 | 50.0608 | 58931.4339 | 58716.943 | 1170.3389 | 1070.3389 | 1144.2797
8090 | 85.4385 | 1.0924 | 1.1351 | 1.0184 | 194.1769 | 51.861 | 64982.045 | 65609.673 | 1308.1935 | 1208.1935 | 1144.3934
8091 | 103.0468 | 0.90786 | 1.1804 | 1.0216 | 195.4087 | 48.7218 | 63504.658 | 66058.5967 | 1317.1719 | 1217.1719 | 1144.5378
8092 | 93.7834 | 1.0673 | 0.74588 | 1.0359 | 212.8878 | 51.9338 | 45537.6777 | 44509.2634 | 886.1853 | 786.1853 | 1144.6019
8093 | 119.2732 | 0.91839 | 0.74171 | 1.1022 | 187.8597 | 50.2112 | 46738.1207 | 46328.4865 | 922.5697 | 822.5697 | 1144.6109
8094 | 104.6724 | 0.93325 | 0.90219 | 1.0296 | 220.8432 | 47.2775 | 48754.0881 | 50186.6223 | 999.7324 | 899.7324 | 1144.6167
8095 | 82.7489 | 1.0633 | 0.93857 | 1.0361 | 194.4638 | 50.6626 | 53918.3275 | 52287.4203 | 1041.7484 | 941.7484 | 1144.9353
8096 | 89.1701 | 1.0304 | 0.86026 | 1.0474 | 216.1097 | 51.6959 | 51547.4958 | 53288.3735 | 1061.7675 | 961.7675 | 1145.1988
8097 | 114.9522 | 0.97358 | 0.69061 | 0.93809 | 198.446 | 53.9143 | 41160.887 | 40191.2042 | 799.8241 | 699.8241 | 1145.2119
8098 | 94.4418 | 0.91905 | 0.98829 | 1.0833 | 176.8509 | 46.9699 | 54538.195 | 54905.4218 | 1094.1084 | 994.1084 | 1145.479
8099 | 112.0546 | 0.97391 | 0.67247 | 0.94214 | 204.4124 | 51.3215 | 38320.2451 | 37713.2444 | 750.2649 | 650.2649 | 1145.4852
8100 | 98.6575 | 0.99148 | 0.78095 | 1.0071 | 195.3879 | 49.458 | 43931.5564 | 45011.0098 | 896.2202 | 796.2202 | 1145.5748
8101 | 99.5847 | 1.0871 | 1.1709 | 0.97101 | 201.8619 | 54.3385 | 67865.3132 | 67208.9033 | 1340.1781 | 1240.1781 | 1145.5917
8102 | 105.8941 | 0.99139 | 1.2141 | 1.0158 | 194.9353 | 51.2437 | 68772.3789 | 69746.7475 | 1390.9349 | 1290.9349 | 1145.6776
8103 | 117.4326 | 0.96129 | 0.90779 | 1.0188 | 219.1418 | 51.3426 | 53500.065 | 52353.9956 | 1043.0799 | 943.0799 | 1145.716
8104 | 98.237 | 1.0108 | 1.0954 | 0.90417 | 197.1447 | 51.9632 | 56824.9447 | 57538.0565 | 1146.7611 | 1046.7611 | 1145.8025
8105 | 109.1614 | 1.0185 | 1.0222 | 0.97772 | 202.1414 | 49.6354 | 55326.7909 | 55911.0396 | 1114.2208 | 1014.2208 | 1146.099
8106 | 116.2673 | 0.97656 | 0.89313 | 1.059 | 189.4324 | 52.2509 | 55540.429 | 55442.1171 | 1104.8423 | 1004.8423 | 1146.2812
8107 | 92.9533 | 1.0291 | 0.89771 | 0.9733 | 175.7122 | 50.3975 | 49030.9776 | 50807.0678 | 1012.1414 | 912.1414 | 1146.2866
8108 | 93.1325 | 1.003 | 1.0209 | 1.0462 | 178.2746 | 46.6383 | 54345.4547 | 55445.151 | 1104.903 | 1004.903 | 1146.3065
8109 | 94.4206 | 1.0888 | 1.1428 | 1.1351 | 205.4995 | 52.0133 | 73025.1353 | 74005.3391 | 1476.1068 | 1376.1068 | 1146.4158
8110 | 113.1874 | 1.0909 | 1.0488 | 0.96647 | 211.3887 | 51.7994 | 59113.3697 | 59097.1573 | 1177.9431 | 1077.9431 | 1146.5338
8111 | 108.2142 | 1.0256 | 0.88263 | 0.98693 | 192.7952 | 44.9021 | 44290.4619 | 45171.6348 | 899.4327 | 799.4327 | 1146.5403
8112 | 105.974 | 0.97592 | 1.0222 | 1.0982 | 218.652 | 49.3231 | 60689.632 | 59943.4087 | 1194.8682 | 1094.8682 | 1146.7552
8113 | 113.7332 | 0.98118 | 0.74068 | 1.0381 | 188.3589 | 43.0816 | 38122.2618 | 38762.8402 | 771.2568 | 671.2568 | 1146.7894
8114 | 123.3359 | 1.004 | 1.1788 | 1.1354 | 228.5932 | 51.3422 | 75300.4001 | 77882.0219 | 1553.6404 | 1453.6404 | 1146.8621
8115 | 81.6312 | 0.95167 | 1.029 | 0.96958 | 206.3377 | 49.2587 | 53177.2394 | 51926.3088 | 1034.5262 | 934.5262 | 1146.8766
8116 | 105.6142 | 1.0541 | 1.1281 | 1.0236 | 179.6122 | 52.635 | 66821.4616 | 64183.9531 | 1279.6791 | 1179.6791 | 1146.9833
8117 | 81.0797 | 0.96338 | 0.71931 | 0.85587 | 188.1584 | 44.594 | 31125.1251 | 30961.1094 | 615.2222 | 515.2222 | 1147.1707
8118 | 110.137 | 1.0216 | 1.1818 | 0.91571 | 194.8336 | 51.0385 | 61169.5754 | 59945.5792 | 1194.9116 | 1094.9116 | 1147.2941
8119 | 101.4246 | 1.001 | 0.89697 | 0.97238 | 190.0265 | 53.2627 | 52052.8491 | 53001.8578 | 1056.0372 | 956.0372 | 1147.363
8120 | 103.1139 | 0.99519 | 1.011 | 1.0385 | 190.8677 | 45.3494 | 52460.3101 | 53604.5562 | 1068.0911 | 968.0911 | 1147.3746
8121 | 106.0631 | 1.0217 | 1.0385 | 0.90556 | 208.5365 | 51.8062 | 54543.1909 | 54231.2861 | 1080.6257 | 980.6257 | 1147.3799
8122 | 120.1032 | 1.008 | 0.85094 | 0.88866 | 205.5878 | 51.1505 | 45078.1434 | 44686.7954 | 889.7359 | 789.7359 | 1147.4183
8123 | 106.4556 | 1.0465 | 0.83756 | 0.96557 | 211.8913 | 52.5663 | 48579.9964 | 49338.4078 | 982.7682 | 882.7682 | 1147.5191
8124 | 95.4615 | 1.0421 | 1.2243 | 0.97302 | 194.2368 | 52.9882 | 68587.8805 | 68182.8115 | 1359.6562 | 1259.6562 | 1147.5916
8125 | 103.5257 | 0.97964 | 0.94873 | 0.93955 | 198.5991 | 46.2972 | 46162.2335 | 45299.4476 | 901.989 | 801.989 | 1147.7235
8126 | 92.335 | 1.071 | 0.86621 | 0.89023 | 203.6745 | 54.1742 | 47335.8779 | 47001.4603 | 936.0292 | 836.0292 | 1147.7873
8127 | 98.9297 | 1.036 | 1.2428 | 1.0026 | 207.3406 | 50.286 | 68021.1718 | 68754.4062 | 1371.0881 | 1271.0881 | 1147.8217
8128 | 100.5997 | 1.0543 | 1.1238 | 1.0482 | 172.2083 | 48.7569 | 62774.1635 | 61645.0317 | 1228.9006 | 1128.9006 | 1147.8664
8129 | 96.6066 | 1.0493 | 0.95724 | 0.95027 | 188.3386 | 45.3558 | 46043.225 | 45004.2471 | 896.0849 | 796.0849 | 1148.0863
8130 | 98.153 | 0.97808 | 0.98294 | 0.93379 | 204.5242 | 49.2096 | 50096.226 | 49063.0692 | 977.2614 | 877.2614 | 1148.0929
8131 | 102.6987 | 1.0146 | 0.90695 | 0.96966 | 213.1617 | 50.6035 | 49988.2034 | 48441.6557 | 964.8331 | 864.8331 | 1148.176
8132 | 102.5487 | 1.0568 | 0.77053 | 0.97113 | 200.6973 | 49.4097 | 42527.9143 | 43891.2546 | 873.8251 | 773.8251 | 1148.181
8133 | 93.1587 | 0.92942 | 0.84285 | 0.98044 | 191.9781 | 54.6582 | 50092.0375 | 50434.5181 | 1004.6904 | 904.6904 | 1148.1918
8134 | 106.0677 | 1.0091 | 1.1323 | 0.9824 | 222.2615 | 48.6502 | 59546.0638 | 60261.6646 | 1201.2333 | 1101.2333 | 1148.1933
8135 | 107.3927 | 0.9319 | 1.1479 | 1.0492 | 178.5498 | 54.3588 | 71084.7201 | 70838.5481 | 1412.771 | 1312.771 | 1148.2802
8136 | 88.8516 | 0.97233 | 0.97276 | 0.95271 | 186.0912 | 47.4379 | 48248.0147 | 47705.4596 | 950.1092 | 850.1092 | 1148.3995
8137 | 109.3467 | 0.99246 | 1.063 | 1.0446 | 198.6963 | 53.9714 | 65985.0442 | 64881.3711 | 1293.6274 | 1193.6274 | 1148.4198
8138 | 96.4578 | 0.88005 | 1.025 | 1.0192 | 199.833 | 50.7929 | 57575.0148 | 58135.6222 | 1158.7124 | 1058.7124 | 1148.4799
8139 | 126.6347 | 0.86894 | 0.97213 | 0.98845 | 176.9001 | 45.405 | 48802.5524 | 49657.9241 | 989.1585 | 889.1585 | 1148.5163
8140 | 115.2871 | 0.98069 | 1.1893 | 0.95473 | 201.8485 | 53.9293 | 67534.4299 | 65246.1129 | 1300.9223 | 1200.9223 | 1148.5192
8141 | 98.9415 | 0.97554 | 0.70093 | 1.0255 | 196.8067 | 51.0691 | 41835.544 | 41214.3427 | 820.2869 | 720.2869 | 1148.6056
8142 | 97.2405 | 1.0217 | 1.0723 | 1.0521 | 220.6757 | 51.2524 | 63135.0821 | 62478.7423 | 1245.5748 | 1145.5748 | 1148.6513
8143 | 83.785 | 1.0407 | 0.94479 | 0.91894 | 208.2786 | 54.6228 | 52394.6979 | 53147.3806 | 1058.9476 | 958.9476 | 1148.7346
8144 | 122.5223 | 0.93873 | 0.96671 | 0.90911 | 202.498 | 52.26 | 52141.6555 | 53611.0592 | 1068.2212 | 968.2212 | 1148.7463
8145 | 85.0906 | 0.91463 | 1.1324 | 0.99886 | 224.4704 | 51.4578 | 62435.6461 | 63224.5547 | 1260.4911 | 1160.4911 | 1148.8688
8146 | 105.2815 | 0.97826 | 0.76746 | 0.99917 | 214.3158 | 50.039 | 43738.9381 | 43066.3495 | 857.327 | 757.327 | 1149.0723
8147 | 95.5744 | 1.0211 | 0.85487 | 0.94353 | 217.799 | 49.7799 | 45228.2225 | 45131.9566 | 898.6391 | 798.6391 | 1149.1454
8148 | 87.5066 | 0.94383 | 0.72954 | 1.0039 | 200.7324 | 54.3386 | 44485.0763 | 44692.4119 | 889.8482 | 789.8482 | 1149.1677
8149 | 101.4523 | 0.98159 | 1.0492 | 0.98905 | 209.8282 | 49.1964 | 56161.1573 | 55756.0712 | 1111.1214 | 1011.1214 | 1149.2852
8150 | 105.3747 | 1.0327 | 0.98927 | 1.0187 | 209.5176 | 46.0065 | 51579.9769 | 49527.9123 | 986.5582 | 886.5582 | 1149.3327
8151 | 96.1015 | 0.89862 | 1.0776 | 1.0009 | 195.4199 | 53.2307 | 62209.554 | 61100.1768 | 1218.0035 | 1118.0035 | 1149.38
8152 | 108.9852 | 0.9753 | 0.91868 | 1.0769 | 203.1435 | 49.6307 | 54577.3183 | 56178.0447 | 1119.5609 | 1019.5609 | 1149.6741
8153 | 103.8692 | 1.0183 | 0.96969 | 1.0376 | 207.1543 | 51.6011 | 57584.3469 | 58263.8482 | 1161.277 | 1061.277 | 1149.6834
8154 | 113.7743 | 1.013 | 0.57559 | 0.96779 | 190.8835 | 46.7047 | 31590.6041 | 30900.1193 | 614.0024 | 514.0024 | 1149.6892
8155 | 107.498 | 1.0772 | 0.82035 | 1.0264 | 195.0861 | 46.9074 | 45122.0897 | 45692.7773 | 909.8555 | 809.8555 | 1149.699
8156 | 107.5599 | 0.99864 | 0.90125 | 1.0038 | 188.9972 | 49.1798 | 49964.5546 | 49382.3336 | 983.6467 | 883.6467 | 1149.7456
8157 | 95.9933 | 0.962 | 1.2284 | 0.97685 | 202.5987 | 48.9951 | 63520.9846 | 64111.1822 | 1278.2236 | 1178.2236 | 1149.8159
8158 | 107.2577 | 1.0258 | 1.0462 | 1.0145 | 193.5189 | 48.0391 | 56466.8809 | 57776.6085 | 1151.5322 | 1051.5322 | 1149.8204
8159 | 105.2837 | 1.0636 | 1.1926 | 1.0355 | 201.4806 | 48.0112 | 64866.6499 | 64723.1108 | 1290.4622 | 1190.4622 | 1149.8456
8160 | 88.8183 | 0.9597 | 1.1043 | 0.99918 | 201.6031 | 47.0817 | 56164.0998 | 54555.5706 | 1087.1114 | 987.1114 | 1149.8555
8161 | 110.0276 | 1.0613 | 0.7039 | 0.89695 | 206.7844 | 53.5344 | 40257.7696 | 39342.0468 | 782.8409 | 682.8409 | 1149.925
8162 | 107.3767 | 0.97976 | 1.0432 | 0.92596 | 210.8528 | 47.096 | 50660.1436 | 51458.7894 | 1025.1758 | 925.1758 | 1149.9404
8163 | 101.4196 | 0.9984 | 0.98606 | 0.94184 | 200.6756 | 52.2593 | 54026.0383 | 53940.5604 | 1074.8112 | 974.8112 | 1149.9861
8164 | 90.5165 | 0.99925 | 1.2568 | 0.98579 | 209.6241 | 47.2444 | 63014.9494 | 63362.9568 | 1263.2591 | 1163.2591 | 1150.054
8165 | 102.1183 | 0.9948 | 1.0552 | 0.9499 | 202.3858 | 52.158 | 57778.4524 | 57541.4215 | 1146.8284 | 1046.8284 | 1150.1069
8166 | 99.8034 | 0.95482 | 1.1674 | 0.97942 | 224.333 | 46.8893 | 58302.929 | 56752.6576 | 1131.0532 | 1031.0532 | 1150.1649
8167 | 100.2922 | 0.94657 | 0.97502 | 1.1112 | 180.3045 | 50.507 | 59696.4875 | 60636.6447 | 1208.7329 | 1108.7329 | 1150.216
8168 | 102.4645 | 1.0254 | 0.90602 | 0.93562 | 186.8068 | 48.9448 | 46819.5533 | 47752.9872 | 951.0597 | 851.0597 | 1150.2423
8169 | 96.0541 | 0.91366 | 1.1234 | 1.0372 | 209.0681 | 50.5254 | 63514.3758 | 64819.222 | 1292.3844 | 1192.3844 | 1150.2615
8170 | 101.0108 | 0.93926 | 0.95393 | 0.98613 | 187.7355 | 50.7813 | 52775.4772 | 53360.2556 | 1063.2051 | 963.2051 | 1150.2984
8171 | 110.1467 | 0.96385 | 1.0894 | 0.95602 | 200.2124 | 48.6102 | 55986.9637 | 56303.9007 | 1122.078 | 1022.078 | 1150.4069
8172 | 104.7211 | 0.92744 | 0.85116 | 0.98934 | 205.8462 | 49.3563 | 46561.6179 | 46528.4399 | 926.5688 | 826.5688 | 1150.4346
8173 | 93.2964 | 1.019 | 1.0935 | 1.0153 | 202.754 | 51.0514 | 61730.8467 | 58266.7905 | 1161.3358 | 1061.3358 | 1150.5153
8174 | 111.9082 | 1.0418 | 1.0717 | 0.97146 | 211.3029 | 51.5231 | 59862.0543 | 59298.8672 | 1181.9773 | 1081.9773 | 1150.6043
8175 | 101.4706 | 0.99658 | 1.0808 | 1.0143 | 199.5389 | 51.0685 | 61351.7193 | 60716.5049 | 1210.3301 | 1110.3301 | 1150.7288
8176 | 101.1426 | 1.0014 | 1.2046 | 1.0425 | 214.7817 | 50.9412 | 69341.2237 | 68691.7112 | 1369.8342 | 1269.8342 | 1150.7546
8177 | 127.2923 | 0.98558 | 1.1393 | 0.94638 | 200.53 | 46.6522 | 56354.3668 | 56098.6326 | 1117.9727 | 1017.9727 | 1150.8356
8178 | 85.8406 | 1.0867 | 1.4671 | 1.0457 | 204.5619 | 54.2467 | 88488.6323 | 91028.5316 | 1816.5706 | 1716.5706 | 1151.0215
8179 | 101.8237 | 1.048 | 0.83276 | 1.0499 | 188.8741 | 50.686 | 49914.6269 | 49930.6106 | 994.6122 | 894.6122 | 1151.1387
8180 | 96.4312 | 0.98735 | 1.2515 | 1.0404 | 214.6687 | 51.0162 | 71498.2889 | 71118.4485 | 1418.369 | 1318.369 | 1151.2271
8181 | 97.9659 | 0.9918 | 0.88673 | 0.94123 | 206.8931 | 53.1719 | 49751.3276 | 49858.977 | 993.1795 | 893.1795 | 1151.3525
8182 | 106.3267 | 1.0439 | 0.84639 | 0.97536 | 196.2059 | 48.4949 | 45613.0542 | 45723.1211 | 910.4624 | 810.4624 | 1151.4697
8183 | 106.5315 | 0.90599 | 0.89829 | 1.0174 | 199.6994 | 47.96 | 48659.0654 | 49173.4166 | 979.4683 | 879.4683 | 1151.5076
8184 | 94.3182 | 0.99171 | 0.92545 | 0.95063 | 189.913 | 50.626 | 49463.7603 | 47846.7472 | 952.9349 | 852.9349 | 1151.5943
8185 | 104.6316 | 1.0251 | 1.0772 | 0.96155 | 192.0656 | 57.6887 | 66131.9999 | 66010.9508 | 1316.219 | 1216.219 | 1151.6679
8186 | 94.6084 | 0.99339 | 1.0478 | 0.94955 | 202.4833 | 50.2231 | 54890.378 | 55166.6918 | 1099.3338 | 999.3338 | 1151.8209
8187 | 112.9931 | 0.97043 | 1.3354 | 0.98345 | 182.6143 | 47.5555 | 67850.9175 | 67579.6205 | 1347.5924 | 1247.5924 | 1151.8393
8188 | 106.9925 | 1.0077 | 0.95691 | 1.0155 | 191.4367 | 50.6707 | 54893.5395 | 55544.7519 | 1106.895 | 1006.895 | 1151.8449
8189 | 100.6456 | 0.9251 | 1.1147 | 1.0447 | 197.2432 | 44.151 | 55719.6188 | 55704.0048 | 1110.0801 | 1010.0801 | 1151.9098
8190 | 116.7007 | 1.0893 | 0.87712 | 1.081 | 225.269 | 45.8249 | 49501.4942 | 48767.3012 | 971.346 | 871.346 | 1151.9945
8191 | 111.5152 | 1.0167 | 1.334 | 1.055 | 186.1038 | 51.1726 | 78008.611 | 78032.0676 | 1556.6414 | 1456.6414 | 1152.025
8192 | 105.5224 | 0.95134 | 1.1217 | 1.0405 | 183.7728 | 45.6129 | 57997.5109 | 57276.2252 | 1141.5245 | 1041.5245 | 1152.0267
8193 | 95.0502 | 0.95814 | 0.98686 | 1.0458 | 192.6633 | 52.3023 | 58935.4841 | 59085.913 | 1177.7183 | 1077.7183 | 1152.0344
8194 | 97.471 | 0.98042 | 0.64954 | 1.0117 | 214.066 | 45.8177 | 34700.6504 | 33920.7982 | 674.416 | 574.416 | 1152.2033
8195 | 107.9693 | 1.0122 | 1.0411 | 0.99783 | 195.4811 | 50.275 | 57919.4469 | 57348.484 | 1142.9697 | 1042.9697 | 1152.2176
8196 | 98.6575 | 0.99791 | 1.0446 | 0.98403 | 187.4118 | 47.5787 | 53778.1998 | 54639.3281 | 1088.7866 | 988.7866 | 1152.2219
8197 | 95.1408 | 0.99101 | 0.80603 | 1.0094 | 199.2038 | 50.0876 | 45674.8128 | 45574.3443 | 907.4869 | 807.4869 | 1152.2631
8198 | 123.9491 | 0.97621 | 1.0893 | 0.96302 | 207.6341 | 51.7856 | 60797.792 | 61487.1039 | 1225.7421 | 1125.7421 | 1152.3204
8199 | 99.6987 | 0.99052 | 1.1065 | 0.97042 | 186.8826 | 48.982 | 57618.1626 | 54730.3116 | 1090.6062 | 990.6062 | 1152.331
8200 | 92.4532 | 0.92656 | 0.86788 | 1.0101 | 208.6045 | 47.9023 | 46304.2626 | 46285.5959 | 921.7119 | 821.7119 | 1152.3314
8201 | 99.0869 | 1.0533 | 0.74612 | 1.0567 | 193.9527 | 50.9675 | 45697.7557 | 45241.574 | 900.8315 | 800.8315 | 1152.4103
8202 | 123.6969 | 1.0043 | 0.88836 | 1.0148 | 205.6401 | 48.4392 | 49893.3825 | 48524.5705 | 966.4914 | 866.4914 | 1152.5801
8203 | 91.3242 | 0.97852 | 1.0457 | 0.93016 | 196.4901 | 47.4073 | 50545.4055 | 52500.1861 | 1046.0037 | 946.0037 | 1152.5937
8204 | 92.5403 | 0.99126 | 1.1682 | 1.0629 | 216.9074 | 52.5852 | 70332.8932 | 69572.3073 | 1387.4461 | 1287.4461 | 1152.6581
8205 | 109.3993 | 0.95869 | 0.71943 | 0.98769 | 218.242 | 47.6949 | 39111.0841 | 39386.772 | 783.7354 | 683.7354 | 1152.7906
8206 | 98.1692 | 0.96253 | 0.84816 | 0.91354 | 211.9203 | 49.0266 | 42831.772 | 42183.1919 | 839.6638 | 739.6638 | 1152.805
8207 | 107.4855 | 0.98841 | 1.1988 | 1.0427 | 214.6807 | 45.0082 | 61254.6642 | 61661.4782 | 1229.2296 | 1129.2296 | 1152.8807
8208 | 94.8458 | 1.0507 | 1.0214 | 1.0058 | 183.7253 | 50.7546 | 57386.941 | 57013.9238 | 1136.2785 | 1036.2785 | 1153.1495
8209 | 104.4236 | 0.96669 | 1.0863 | 0.94597 | 203.412 | 47.6724 | 54006.1124 | 53103.6401 | 1058.0728 | 958.0728 | 1153.1816
8210 | 103.75 | 0.93463 | 0.81923 | 1.0241 | 193.2833 | 51.553 | 48442.0745 | 49314.078 | 982.2816 | 882.2816 | 1153.1859
8211 | 95.2293 | 1.0762 | 0.92377 | 0.96364 | 202.2188 | 52.9898 | 52803.3686 | 53767.103 | 1071.3421 | 971.3421 | 1153.1896
8212 | 94.8467 | 0.99441 | 1.1403 | 1.0098 | 217.9826 | 51.2613 | 64078.0299 | 63312.5638 | 1262.2513 | 1162.2513 | 1153.407
8213 | 80.9854 | 1.0116 | 0.80945 | 0.97363 | 208.3365 | 48.9298 | 42778.4931 | 41593.7885 | 827.8758 | 727.8758 | 1153.4132
8214 | 109.8026 | 0.96326 | 1.2326 | 0.92585 | 225.0552 | 49.3863 | 61807.9348 | 61422.6771 | 1224.4535 | 1124.4535 | 1153.4486
8215 | 88.4212 | 0.98538 | 0.9547 | 0.95998 | 195.6279 | 54.2524 | 54644.6179 | 54908.1996 | 1094.164 | 994.164 | 1153.523
8216 | 87.0496 | 1.0168 | 0.72231 | 0.97792 | 188.9111 | 49.7325 | 39719.9274 | 39926.226 | 794.5245 | 694.5245 | 1153.5372
8217 | 109.2352 | 0.99391 | 0.93379 | 0.99175 | 201.8159 | 54.2594 | 56341.6301 | 57980.0692 | 1155.6014 | 1055.6014 | 1153.5575
8218 | 99.4292 | 1.0138 | 0.99278 | 0.8742 | 214.6794 | 47.5941 | 46318.7743 | 47252.6477 | 941.053 | 841.053 | 1153.5775
8219 | 112.0583 | 0.94913 | 1.0389 | 0.94553 | 172.3564 | 51.5181 | 56258.3332 | 57163.0609 | 1139.2612 | 1039.2612 | 1153.5875
8220 | 109.7358 | 0.96697 | 0.80548 | 1.0761 | 206.5897 | 49.1684 | 48043.0966 | 47305.6336 | 942.1127 | 842.1127 | 1153.7656
8221 | 92.2966 | 0.93018 | 1.1287 | 0.9483 | 193.5933 | 45.5945 | 52909.7184 | 53191.6495 | 1059.833 | 959.833 | 1153.8279
8222 | 103.8211 | 0.91468 | 1.0818 | 1.0479 | 199.9267 | 49.3418 | 60817.5088 | 59152.5285 | 1179.0506 | 1079.0506 | 1153.8747
8223 | 104.7867 | 0.8992 | 0.91082 | 0.98614 | 192.0716 | 54.9295 | 54704.8219 | 54798.3535 | 1091.9671 | 991.9671 | 1153.8908
8224 | 82.54 | 1.0715 | 0.94604 | 0.93376 | 196.9097 | 48.6922 | 47516.7724 | 48220.0626 | 960.4013 | 860.4013 | 1153.9333
8225 | 86.7993 | 1.1174 | 1.1275 | 0.93372 | 210.2496 | 49.7838 | 57450.8057 | 59279.3578 | 1181.5872 | 1081.5872 | 1154.124
8226 | 103.3847 | 0.98743 | 0.91443 | 0.99431 | 204.691 | 47.0294 | 47766.2264 | 47701.2835 | 950.0257 | 850.0257 | 1154.1897
8227 | 109.7542 | 1.0231 | 0.9894 | 1.0218 | 187.9156 | 54.5725 | 61484.8491 | 59958.6368 | 1195.1727 | 1095.1727 | 1154.2401
8228 | 102.5145 | 0.97421 | 1.0093 | 1.0109 | 214.3045 | 50.3202 | 56581.5909 | 55242.1792 | 1100.8436 | 1000.8436 | 1154.2488
8229 | 95.9748 | 0.937 | 0.92621 | 1.0122 | 193.6421 | 46.0413 | 47498.7197 | 46886.1253 | 933.7225 | 833.7225 | 1154.2797
8230 | 104.6462 | 0.95295 | 0.92014 | 0.97363 | 204.1624 | 52.9099 | 52881.3958 | 52879.1564 | 1053.5831 | 953.5831 | 1154.7363
8231 | 85.0889 | 1.1185 | 1.1901 | 0.96237 | 187.5292 | 49.6124 | 61730.323 | 61975.0691 | 1235.5014 | 1135.5014 | 1154.7428
8232 | 94.9065 | 0.99767 | 1.2489 | 0.98068 | 164.169 | 50.2482 | 66462.419 | 65142.8807 | 1298.8576 | 1198.8576 | 1154.751
8233 | 102.2216 | 0.96446 | 0.91949 | 1.0756 | 188.0718 | 50.5838 | 55203.0034 | 55204.1494 | 1100.083 | 1000.083 | 1154.8643
8234 | 94.1708 | 0.91643 | 1.0751 | 0.98038 | 195.8857 | 47.5331 | 54398.2676 | 55370.6342 | 1103.4127 | 1003.4127 | 1154.9127
8235 | 94.7615 | 0.96988 | 0.70798 | 0.91813 | 201.0654 | 50.4898 | 37661.0267 | 38530.4475 | 766.609 | 666.609 | 1154.9748
8236 | 95.0067 | 1.0209 | 0.94135 | 0.99237 | 192.1032 | 45.7679 | 47386.2245 | 49077.7856 | 977.5557 | 877.5557 | 1155.189
8237 | 99.4829 | 1.0344 | 1.0485 | 0.96867 | 201.543 | 49.8031 | 55909.3427 | 57810.7741 | 1152.2155 | 1052.2155 | 1155.2911
8238 | 99.1062 | 1.0499 | 0.90051 | 0.97956 | 205.7303 | 50.0925 | 49604.7186 | 49535.3795 | 986.7076 | 886.7076 | 1155.3675
8239 | 104.1137 | 1.0475 | 1.0468 | 0.89852 | 218.932 | 54.8869 | 57828.9675 | 57760.6538 | 1151.2131 | 1051.2131 | 1155.4224
8240 | 110.0802 | 0.9712 | 1.1496 | 0.96786 | 198.2106 | 50.7409 | 62079.1575 | 62401.1733 | 1244.0235 | 1144.0235 | 1155.4582
8241 | 113.2221 | 1.1028 | 0.98695 | 0.96923 | 212.8369 | 48.3763 | 52529.3051 | 53819.5863 | 1072.3917 | 972.3917 | 1155.5284
8242 | 112.3512 | 1.0582 | 1.0813 | 0.98703 | 192.9988 | 46.7532 | 55650.7284 | 55010.064 | 1096.2013 | 996.2013 | 1155.5967
8243 | 93.4512 | 1.0244 | 0.88421 | 0.93396 | 197.116 | 54.0423 | 50000.0414 | 49631.9062 | 988.6381 | 888.6381 | 1155.6031
8244 | 101.2061 | 0.99369 | 0.9543 | 1.1308 | 200.7139 | 46.5497 | 55112.6507 | 53843.0825 | 1072.8617 | 972.8617 | 1155.6262
8245 | 87.489 | 0.95198 | 1.0153 | 1.0313 | 197.3355 | 51.4854 | 58396.3393 | 60658.6986 | 1209.174 | 1109.174 | 1155.7297
8246 | 108.3214 | 0.94027 | 1.1633 | 0.97895 | 184.0367 | 47.6621 | 59315.9525 | 59489.6244 | 1185.7925 | 1085.7925 | 1155.8877
8247 | 116.9113 | 1.0336 | 1.015 | 0.85876 | 185.6924 | 50.1651 | 49972.2627 | 49151.017 | 979.0203 | 879.0203 | 1155.8883
8248 | 95.1962 | 1.0101 | 0.898 | 1.0306 | 194.9152 | 50.1811 | 51460.0475 | 52595.7552 | 1047.9151 | 947.9151 | 1155.9224
8249 | 84.9184 | 0.95448 | 0.73048 | 1.0486 | 196.9383 | 51.8808 | 44141.2492 | 43384.9972 | 863.6999 | 763.6999 | 1155.9808
8250 | 97.916 | 0.96631 | 0.78168 | 1.0534 | 182.0853 | 49.9687 | 46055.5317 | 47270.0357 | 941.4007 | 841.4007 | 1156.0282
8251 | 96.706 | 0.99644 | 0.89176 | 1.0045 | 193.0324 | 49.7769 | 49577.3549 | 48826.6199 | 972.5324 | 872.5324 | 1156.0447
8252 | 94.4439 | 1.0982 | 0.96203 | 1.0649 | 203.2464 | 45.5305 | 51568.4515 | 50808.0246 | 1012.1605 | 912.1605 | 1156.091
8253 | 99.9436 | 0.91016 | 1.0198 | 0.98697 | 209.6903 | 47.2239 | 52037.696 | 51285.7142 | 1021.7143 | 921.7143 | 1156.1313
8254 | 111.2536 | 1.0606 | 0.87109 | 0.99563 | 182.7606 | 49.1842 | 48642.9525 | 47970.0287 | 955.4006 | 855.4006 | 1156.1736
8255 | 93.5551 | 1.0542 | 0.89528 | 1.0874 | 205.0829 | 45.8383 | 49352.3001 | 48443.6486 | 964.873 | 864.873 | 1156.2473
8256 | 95.5981 | 1.0338 | 1.0334 | 0.93025 | 206.1021 | 51.8722 | 55196.7382 | 56145.3514 | 1118.907 | 1018.907 | 1156.2807
8257 | 81.2974 | 1.1015 | 0.9029 | 1.0191 | 196.6326 | 49.4818 | 50157.1649 | 49529.8692 | 986.5974 | 886.5974 | 1156.4875
8258 | 91.6434 | 1.025 | 1.0166 | 0.90456 | 204.5579 | 50.4156 | 51300.5767 | 51315.3572 | 1022.3071 | 922.3071 | 1156.5843
8259 | 98.4697 | 0.95954 | 1.0614 | 0.93103 | 200.9544 | 51.0662 | 55488.8811 | 54535.534 | 1086.7107 | 986.7107 | 1156.5976
8260 | 112.9134 | 0.95704 | 0.80867 | 1.0314 | 192.8804 | 49.8653 | 47171.8104 | 46953.6249 | 935.0725 | 835.0725 | 1156.6456
8261 | 90.3488 | 1.0727 | 0.74151 | 1.0377 | 211.6422 | 47.0855 | 41005.1138 | 42083.7921 | 837.6758 | 737.6758 | 1156.8266
8262 | 108.5311 | 1.0834 | 0.88375 | 0.97939 | 194.3208 | 53.2002 | 52496.5265 | 52169.1006 | 1039.382 | 939.382 | 1156.8705
8263 | 104.4805 | 0.97025 | 1.0964 | 0.99883 | 204.4416 | 49.9223 | 59937.6215 | 59626.6674 | 1188.5333 | 1088.5333 | 1156.8837
8264 | 107.2327 | 0.99187 | 0.87809 | 0.85092 | 208.459 | 52.9223 | 45379.838 | 43003.7109 | 856.0742 | 756.0742 | 1156.9748
8265 | 102.356 | 1.0431 | 0.70933 | 1.0218 | 231.5257 | 51.8258 | 43327.3942 | 43574.5791 | 867.4916 | 767.4916 | 1157.205
8266 | 110.2152 | 0.99625 | 1.0344 | 1.0088 | 193.734 | 50.6721 | 58632.8324 | 60686.0172 | 1209.7203 | 1109.7203 | 1157.2386
8267 | 95.9423 | 1.0917 | 1.0012 | 1.0316 | 198.7618 | 51.9101 | 59251.2279 | 59251.7929 | 1181.0359 | 1081.0359 | 1157.2786
8268 | 113.6931 | 0.95688 | 0.96006 | 0.97779 | 211.6867 | 50.0391 | 52628.7823 | 49780.411 | 991.6082 | 891.6082 | 1157.3424
8269 | 102.0439 | 0.98608 | 0.80172 | 1.0861 | 194.3277 | 50.6454 | 49390.3817 | 48455.7215 | 965.1144 | 865.1144 | 1157.4299
8270 | 100.2735 | 0.94601 | 1.1083 | 0.91554 | 201.4345 | 50.9794 | 56764.7333 | 58129.4828 | 1158.5897 | 1058.5897 | 1157.4947
8271 | 95.2746 | 1.0469 | 0.95885 | 0.89857 | 202.3734 | 55.6935 | 53742.4349 | 55340.3985 | 1102.808 | 1002.808 | 1157.4989
8272 | 107.7991 | 1.0477 | 1.2639 | 1.0201 | 206.598 | 53.506 | 75236.067 | 78124.6462 | 1558.4929 | 1458.4929 | 1157.5011
8273 | 108.6283 | 1.0192 | 1.0979 | 0.94239 | 202.13 | 47.0066 | 54040.7447 | 53700.034 | 1070.0007 | 970.0007 | 1157.5328
8274 | 105.6261 | 1.0002 | 0.79882 | 0.98702 | 190.1429 | 48.9108 | 43921.356 | 44829.1245 | 892.5825 | 792.5825 | 1157.5509
8275 | 110.2324 | 0.93474 | 1.1164 | 1.0099 | 208.083 | 49.6192 | 61261.2388 | 60062.8448 | 1197.2569 | 1097.2569 | 1157.6513
8276 | 94.9449 | 0.94201 | 1.2299 | 0.99259 | 201.8497 | 49.9579 | 65657.3362 | 65658.4133 | 1309.1683 | 1209.1683 | 1157.6594
8277 | 96.0328 | 0.93241 | 0.76444 | 1.0288 | 207.7775 | 52.0453 | 45800.9093 | 45305.06 | 902.1012 | 802.1012 | 1157.7571
8278 | 95.0194 | 0.93355 | 1.0508 | 0.95824 | 206.8039 | 52.2449 | 57445.4443 | 58447.8679 | 1164.9574 | 1064.9574 | 1157.926
8279 | 102.0638 | 0.96055 | 1.1219 | 1.0254 | 201.2236 | 49.2664 | 61707.1087 | 60388.3311 | 1203.7666 | 1103.7666 | 1158.0099
8280 | 79.3514 | 0.97395 | 1.0428 | 1.0452 | 204.0464 | 47.5756 | 55735.1253 | 55142.4736 | 1098.8495 | 998.8495 | 1158.0367
8281 | 86.9403 | 0.9284 | 1.0745 | 0.96686 | 212.8344 | 50.2355 | 56458.9576 | 55631.3415 | 1108.6268 | 1008.6268 | 1158.1121
8282 | 104.3072 | 1.063 | 1.1139 | 1.0427 | 204.1594 | 52.098 | 66489.1457 | 64926.0867 | 1294.5217 | 1194.5217 | 1158.2113
8283 | 85.5952 | 1.0467 | 1.0987 | 1.0236 | 207.9575 | 50.7878 | 61874.2312 | 61784.3205 | 1231.6864 | 1131.6864 | 1158.2231
8284 | 113.5736 | 1.115 | 1.097 | 0.98901 | 206.8203 | 49.0731 | 59662.3869 | 60729.2345 | 1210.5847 | 1110.5847 | 1158.2745
8285 | 104.2596 | 0.96981 | 0.68001 | 0.98141 | 198.7023 | 49.6612 | 38362.2282 | 38597.8871 | 767.9577 | 667.9577 | 1158.3065
8286 | 114.1355 | 0.9297 | 0.8563 | 0.98885 | 203.7877 | 48.5939 | 46507.3463 | 45750.4587 | 911.0092 | 811.0092 | 1158.5717
8287 | 116.0291 | 1.092 | 1.0549 | 0.88759 | 178.3894 | 49.3305 | 52617.5057 | 51969.4214 | 1035.3884 | 935.3884 | 1158.6261
8288 | 85.729 | 1.0179 | 1.1 | 1.0756 | 198.1304 | 52.9145 | 67423.9049 | 67714.6264 | 1350.2925 | 1250.2925 | 1158.6353
8289 | 98.5648 | 1.0327 | 0.997 | 1.0371 | 204.1612 | 49.2494 | 56142.4513 | 56452.5238 | 1125.0505 | 1025.0505 | 1158.6684
8290 | 106.0289 | 1.0297 | 1.1215 | 0.99006 | 213.0399 | 46.8383 | 57335.8842 | 55921.8956 | 1114.4379 | 1014.4379 | 1158.7478
8291 | 93.1486 | 1.0111 | 0.86395 | 1.0654 | 197.9037 | 48.3773 | 49283.7518 | 48932.0703 | 974.6414 | 874.6414 | 1158.9279
8292 | 99.5792 | 0.99455 | 1.1568 | 1.0232 | 200.6753 | 51.4276 | 66168.031 | 66754.0109 | 1331.0802 | 1231.0802 | 1158.9305
8293 | 111.9045 | 0.97759 | 0.96471 | 0.97866 | 215.635 | 49.0607 | 51902.4328 | 53273.7255 | 1061.4745 | 961.4745 | 1158.9822
8294 | 80.6496 | 1.0112 | 0.86669 | 1.0042 | 210.1446 | 48.4393 | 46317.9504 | 46361.258 | 923.2252 | 823.2252 | 1159.0494
8295 | 96.2546 | 1.0382 | 0.77305 | 1.0199 | 202.389 | 50.4009 | 44975.5215 | 45555.4653 | 907.1093 | 807.1093 | 1159.2415
8296 | 104.6133 | 1.0191 | 0.93754 | 1.0752 | 199.4122 | 49.1508 | 54983.844 | 56502.4314 | 1126.0486 | 1026.0486 | 1159.2706
8297 | 106.7885 | 0.99132 | 1.0318 | 0.91707 | 223.6408 | 46.6168 | 49269.8734 | 48898.6108 | 973.9722 | 873.9722 | 1159.2961
8298 | 98.0903 | 0.93451 | 0.86765 | 1.0192 | 197.3925 | 48.3447 | 47379.7264 | 47682.0506 | 949.641 | 849.641 | 1159.3364
8299 | 98.7466 | 0.995 | 0.77332 | 0.98427 | 197.3181 | 53.5855 | 46249.3405 | 46107.4198 | 918.1484 | 818.1484 | 1159.3552
8300 | 90.0118 | 1.025 | 0.87336 | 0.97771 | 206.7347 | 52.4096 | 49794.7871 | 50343.8192 | 1002.8764 | 902.8764 | 1159.4402
8301 | 105.0089 | 0.92506 | 1.0175 | 0.98853 | 200.5767 | 46.4091 | 51390.5571 | 53300.7954 | 1062.0159 | 962.0159 | 1159.7568
8302 | 92.6688 | 1.029 | 1.235 | 0.87238 | 211.2386 | 47.6733 | 56120.7649 | 56881.2731 | 1133.6255 | 1033.6255 | 1159.766
8303 | 93.7726 | 1.0117 | 0.94262 | 1.082 | 186.6036 | 47.938 | 53628.6005 | 54071.7101 | 1077.4342 | 977.4342 | 1159.8335
8304 | 118.7485 | 1.0432 | 1.0919 | 1.0183 | 198.2808 | 51.4527 | 63786.0941 | 65064.3481 | 1297.287 | 1197.287 | 1160.2146
8305 | 83.0619 | 1.0548 | 0.90287 | 1.0458 | 197.8357 | 52.7792 | 54655.4655 | 54872.9442 | 1093.4589 | 993.4589 | 1160.2546
8306 | 97.5854 | 1.0272 | 1.0949 | 0.99009 | 201.314 | 47.5217 | 56480.1861 | 56756.592 | 1131.1318 | 1031.1318 | 1160.3958
8307 | 107.1742 | 1.0622 | 1.0508 | 0.9758 | 210.8219 | 49.3835 | 56470.6871 | 57011.612 | 1136.2322 | 1036.2322 | 1160.4369
8308 | 116.0718 | 1.0202 | 1.1541 | 0.95433 | 222.9293 | 51.1348 | 62598.8374 | 63794.301 | 1271.886 | 1171.886 | 1160.4502
8309 | 102.409 | 1.0283 | 1.0119 | 0.99767 | 200.0995 | 50.3653 | 56352.0853 | 55880.0063 | 1113.6001 | 1013.6001 | 1160.4551
8310 | 75.3942 | 0.98982 | 0.8098 | 0.98644 | 190.0898 | 52.9538 | 46442.763 | 45748.9455 | 910.9789 | 810.9789 | 1160.4911
8311 | 101.0868 | 1.0558 | 0.93826 | 1.0002 | 194.3055 | 50.806 | 53296.831 | 55063.0148 | 1097.2603 | 997.2603 | 1160.4939
8312 | 109.4605 | 1.0331 | 0.98198 | 0.94736 | 212.3185 | 48.7259 | 51051.4842 | 52960.2224 | 1055.2044 | 955.2044 | 1160.6203
8313 | 89.2629 | 0.95233 | 0.95993 | 0.98465 | 205.9433 | 50.2728 | 51997.3801 | 51949.3128 | 1034.9863 | 934.9863 | 1160.6902
8314 | 95.5416 | 1.0356 | 1.046 | 1.0141 | 185.0564 | 47.3314 | 55074.6593 | 55803.8182 | 1112.0764 | 1012.0764 | 1160.6937
8315 | 104.2502 | 1.0347 | 1.169 | 1.0846 | 198.2718 | 46.8654 | 64674.1403 | 61388.6597 | 1223.7732 | 1123.7732 | 1160.731
8316 | 99.7036 | 1.0582 | 1.0063 | 1.0928 | 202.5175 | 50.796 | 61421.3411 | 61687.2602 | 1229.7452 | 1129.7452 | 1160.82
8317 | 101.633 | 1.0976 | 0.81843 | 1.0322 | 203.3267 | 48.3031 | 46398.589 | 45829.2767 | 912.5855 | 812.5855 | 1160.832
8318 | 96.0041 | 0.95141 | 1.0305 | 0.97141 | 194.2966 | 48.7079 | 53404.0682 | 53342.6886 | 1062.8538 | 962.8538 | 1160.8539
8319 | 78.3883 | 0.99676 | 0.91869 | 1.0654 | 214.6914 | 49.2843 | 52301.4656 | 51809.3895 | 1032.1878 | 932.1878 | 1160.8552
8320 | 103.412 | 0.99462 | 0.99172 | 1.0677 | 203.9176 | 51.5173 | 60050.1994 | 60298.7739 | 1201.9755 | 1101.9755 | 1160.9112
8321 | 81.1569 | 1.0017 | 1.1316 | 1.0699 | 188.8772 | 49.4206 | 64037.5595 | 65772.7782 | 1311.4556 | 1211.4556 | 1160.95
8322 | 112.857 | 0.93559 | 1.2626 | 1.023 | 210.6044 | 47.7863 | 66978.9718 | 63805.6037 | 1272.1121 | 1172.1121 | 1161.0231
8323 | 99.7234 | 1.0009 | 1.1207 | 0.98657 | 192.1803 | 49.7641 | 60179.975 | 60037.7837 | 1196.7557 | 1096.7557 | 1161.1911
8324 | 84.224 | 0.95037 | 1.0174 | 1.068 | 213.5814 | 49.2898 | 57716.4989 | 57459.0994 | 1145.182 | 1045.182 | 1161.2109
8325 | 101.8093 | 0.97219 | 0.73529 | 1.0103 | 189.6938 | 43.7417 | 37012.6847 | 37835.1214 | 752.7024 | 652.7024 | 1161.307
8326 | 92.2374 | 0.95182 | 1.1169 | 1.0229 | 205.9067 | 48.5681 | 59960.9755 | 58241.0702 | 1160.8214 | 1060.8214 | 1161.3883
8327 | 103.2882 | 1.0507 | 0.91156 | 0.98042 | 194.6025 | 44.6565 | 44951.2245 | 43579.9671 | 867.5993 | 767.5993 | 1161.463
8328 | 91.2183 | 1.0222 | 0.91094 | 0.9834 | 197.6546 | 50.5222 | 50166.849 | 50159.5674 | 999.1913 | 899.1913 | 1161.5916
8329 | 79.4855 | 1.0057 | 0.91812 | 0.9611 | 204.5688 | 48.3137 | 46698.6351 | 46800.9436 | 932.0189 | 832.0189 | 1161.7204
8330 | 101.8557 | 1.0015 | 1.0449 | 0.96065 | 205.3591 | 47.1391 | 52329.6279 | 51667.0769 | 1029.3415 | 929.3415 | 1161.773
8331 | 116.4289 | 0.96029 | 0.89599 | 0.96707 | 185.0088 | 49.6008 | 48709.1781 | 48215.9604 | 960.3192 | 860.3192 | 1161.8152
8332 | 99.9971 | 1.08 | 0.93713 | 0.95187 | 212.8533 | 51.1212 | 51335.4349 | 50563.7429 | 1007.2749 | 907.2749 | 1161.9679
8333 | 95.1591 | 1.0059 | 1.0533 | 1.0447 | 209.7914 | 49.0118 | 58832.0819 | 59072.7682 | 1177.4554 | 1077.4554 | 1162.0285
8334 | 106.9193 | 1.0353 | 0.99884 | 0.93755 | 194.6353 | 54.588 | 57357.1588 | 56667.1206 | 1129.3424 | 1029.3424 | 1162.0362
8335 | 107.6811 | 0.91178 | 0.87907 | 1.0032 | 194.1591 | 45.3555 | 44644.3646 | 44703.7784 | 890.0756 | 790.0756 | 1162.0781
8336 | 95.8782 | 0.94842 | 1.0742 | 0.97512 | 199.2982 | 52.877 | 60393.9272 | 60179.4484 | 1199.589 | 1099.589 | 1162.1011
8337 | 108.929 | 1.1208 | 0.96931 | 1.0088 | 203.2123 | 49.6783 | 54844.6538 | 53708.5547 | 1070.1711 | 970.1711 | 1162.2192
8338 | 99.7137 | 0.92698 | 0.99883 | 1.0169 | 213.8682 | 51.0067 | 56734.5565 | 58168.9016 | 1159.378 | 1059.378 | 1162.2319
8339 | 92.0696 | 0.94627 | 1.246 | 1.0507 | 187.5158 | 48.785 | 68306.6606 | 69144.392 | 1378.8878 | 1278.8878 | 1162.2513
8340 | 90.6327 | 0.96474 | 0.88557 | 1.1274 | 201.3496 | 44.5111 | 48531.2068 | 47989.9227 | 955.7985 | 855.7985 | 1162.2856
8341 | 115.4814 | 0.99035 | 0.72884 | 0.93352 | 205.4779 | 51.4721 | 41113.136 | 41255.002 | 821.1 | 721.1 | 1162.2924
8342 | 89.2137 | 1.0336 | 1.0153 | 1.033 | 189.2597 | 46.9614 | 53770.4333 | 54315.3789 | 1082.3076 | 982.3076 | 1162.3239
8343 | 89.3023 | 1.0921 | 0.96365 | 0.97752 | 188.6444 | 49.3717 | 51511.1106 | 52245.7643 | 1040.9153 | 940.9153 | 1162.4467
8344 | 105.4798 | 1.0326 | 1.1351 | 1.0878 | 199.2767 | 51.4602 | 69346.4283 | 68893.6425 | 1373.8729 | 1273.8729 | 1162.4971
8345 | 102.4357 | 0.91658 | 0.79635 | 1.0441 | 199.4686 | 52.6305 | 48900.1373 | 49149.2983 | 978.986 | 878.986 | 1162.5705
8346 | 115.9015 | 1.0284 | 1.2791 | 1.05 | 199.6605 | 48.7745 | 71518.1316 | 73173.3911 | 1459.4678 | 1359.4678 | 1162.6944
8347 | 104.4806 | 0.99945 | 0.89976 | 0.97127 | 209.0717 | 50.5697 | 49683.0739 | 51316.4516 | 1022.329 | 922.329 | 1162.7157
8348 | 96.5958 | 1.0787 | 1.0535 | 1.1146 | 210.5198 | 50.3023 | 64517.0364 | 63645.1854 | 1268.9037 | 1168.9037 | 1162.8684
8349 | 99.8129 | 0.95734 | 1.2565 | 1.0563 | 196.9399 | 47.5232 | 67808.5901 | 69760.4066 | 1391.2081 | 1291.2081 | 1162.8819
8350 | 105.6259 | 0.98784 | 0.88222 | 1.0519 | 213.7186 | 52.3834 | 54290.9358 | 54467.6338 | 1085.3527 | 985.3527 | 1162.9111
8351 | 105.7981 | 0.96558 | 1.0807 | 1.026 | 202.5326 | 53.1454 | 64558.3949 | 64551.9652 | 1287.0393 | 1187.0393 | 1162.9413
8352 | 85.3144 | 0.97961 | 0.91373 | 0.93369 | 190.4764 | 50.7863 | 47762.7194 | 49461.7007 | 985.234 | 885.234 | 1162.9562
8353 | 115.1444 | 0.93941 | 1.0136 | 0.99966 | 202.9098 | 46.1037 | 51905.3687 | 50813.4833 | 1012.2697 | 912.2697 | 1162.9778
8354 | 94.2876 | 0.89657 | 0.88518 | 0.92806 | 198.0853 | 46.5732 | 42394.9218 | 41310.2485 | 822.205 | 722.205 | 1163.0029
8355 | 128.0069 | 0.92297 | 1.0905 | 0.93735 | 208.5945 | 50.1744 | 57424.237 | 57522.451 | 1146.449 | 1046.449 | 1163.0178
8356 | 105.214 | 0.96868 | 1.023 | 1.0625 | 204.1111 | 49.4787 | 59027.0827 | 58068.7668 | 1157.3753 | 1057.3753 | 1163.0898
8357 | 99.4074 | 0.94812 | 0.97472 | 1.0057 | 210.5856 | 51.0264 | 55037.8574 | 55846.169 | 1112.9234 | 1012.9234 | 1163.1012
8358 | 91.6707 | 0.91098 | 1.0535 | 1.0221 | 189.0649 | 50.5674 | 58862.2973 | 60564.742 | 1207.2948 | 1107.2948 | 1163.1118
8359 | 98.8562 | 0.98131 | 1.0457 | 0.95097 | 184.6518 | 50.2418 | 55021.3254 | 56516.4333 | 1126.3287 | 1026.3287 | 1163.2136
8360 | 78.404 | 1.0366 | 1.226 | 1.0373 | 223.2399 | 49.335 | 66970.6335 | 68462.6467 | 1365.2529 | 1265.2529 | 1163.2591
8361 | 90.0476 | 1.0498 | 0.8336 | 0.97143 | 188.5151 | 53.2808 | 48370.9166 | 50208.2045 | 1000.1641 | 900.1641 | 1163.408
8362 | 103.4508 | 0.99112 | 0.9666 | 0.97479 | 200.5382 | 49.3105 | 51718.7607 | 50947.6094 | 1014.9522 | 914.9522 | 1163.4896
8363 | 95.2818 | 0.97716 | 0.93809 | 0.92461 | 210.0549 | 46.4456 | 44819.6908 | 45693.1564 | 909.8631 | 809.8631 | 1163.5431
8364 | 100.3423 | 0.9294 | 0.6822 | 1.0452 | 186.2953 | 47.2174 | 38258.9089 | 37578.302 | 747.566 | 647.566 | 1163.6062
8365 | 101.2979 | 0.98823 | 1.1243 | 0.96038 | 195.9365 | 51.5041 | 60963.9699 | 63241.5987 | 1260.832 | 1160.832 | 1163.6592
8366 | 100.8513 | 0.98074 | 1.0011 | 0.86942 | 204.3597 | 46.5454 | 45319.9834 | 46065.3565 | 917.3071 | 817.3071 | 1163.663
8367 | 97.956 | 1.0864 | 0.93441 | 1.0042 | 197.3901 | 52.9857 | 55555.492 | 57232.801 | 1140.656 | 1040.656 | 1164.0685
8368 | 112.8208 | 0.98793 | 0.8826 | 0.97491 | 206.871 | 49.1474 | 47973.6988 | 49487.3864 | 985.7477 | 885.7477 | 1164.2541
8369 | 91.4733 | 0.98675 | 0.84936 | 1.0691 | 199.4909 | 47.0783 | 47199.9138 | 47512.5943 | 946.2519 | 846.2519 | 1164.3354
8370 | 112.3749 | 0.93729 | 0.89856 | 1.0397 | 194.7956 | 50.6084 | 52807.3687 | 53601.0859 | 1068.0217 | 968.0217 | 1164.3555
8371 | 93.3484 | 1.0115 | 1.0866 | 0.98598 | 197.8074 | 51.8503 | 60644.8188 | 61265.3432 | 1221.3069 | 1121.3069 | 1164.3979
8372 | 99.5793 | 0.92994 | 0.65666 | 0.9741 | 200.0373 | 46.3424 | 34134.3096 | 34348.9886 | 682.9798 | 582.9798 | 1164.4248
8373 | 96.962 | 1.0136 | 1.0322 | 1.0302 | 208.7615 | 51.1111 | 59580.8013 | 58324.7919 | 1162.4958 | 1062.4958 | 1164.5404
8374 | 87.1332 | 0.97941 | 0.94643 | 0.9581 | 212.6051 | 47.5223 | 47360.2761 | 46830.5639 | 932.6113 | 832.6113 | 1164.7903
8375 | 85.9951 | 1.0085 | 1.1791 | 1.0238 | 199.649 | 54.7679 | 71068.2973 | 69848.9505 | 1392.979 | 1292.979 | 1164.8672
8376 | 81.1706 | 0.91116 | 0.88386 | 0.92783 | 178.686 | 48.4987 | 43538.2001 | 43030.1963 | 856.6039 | 756.6039 | 1164.9194
8377 | 106.2311 | 1.027 | 0.86873 | 0.96596 | 198.986 | 47.4825 | 45224.4011 | 43717.4721 | 870.3494 | 770.3494 | 1164.9511
8378 | 121.1778 | 1.0113 | 1.1122 | 1.0185 | 206.0289 | 51.932 | 65393.3229 | 66381.6049 | 1323.6321 | 1223.6321 | 1164.9822
8379 | 115.2175 | 1.0205 | 1.0234 | 0.97065 | 191.6525 | 50.8236 | 56652.7487 | 55216.7716 | 1100.3354 | 1000.3354 | 1164.9971
8380 | 74.079 | 0.98651 | 1.0639 | 1.0183 | 217.9429 | 46.3794 | 53852.6864 | 55295.3809 | 1101.9076 | 1001.9076 | 1165.1257
8381 | 94.5199 | 1.0048 | 0.90551 | 0.95141 | 198.7411 | 52.4492 | 50365.0723 | 48848.2231 | 972.9645 | 872.9645 | 1165.1692
8382 | 106.4518 | 0.96335 | 1.0523 | 1.0355 | 201.3824 | 45.0944 | 53960.5752 | 52946.6973 | 1054.9339 | 954.9339 | 1165.3396
8383 | 98.9781 | 1.0814 | 0.94046 | 1.0506 | 193.5817 | 50.0578 | 55012.5854 | 55482.4973 | 1105.6499 | 1005.6499 | 1165.3948
8384 | 97.5497 | 0.89067 | 1.0415 | 1.1226 | 217.7649 | 53.2842 | 67147.3554 | 66054.832 | 1317.0966 | 1217.0966 | 1165.4324
8385 | 97.4923 | 0.99954 | 0.85018 | 1.0658 | 213.2732 | 48.1382 | 48525.2162 | 48954.7469 | 975.0949 | 875.0949 | 1165.4616
8386 | 94.0057 | 0.996 | 1.1018 | 1.0169 | 208.0875 | 48.3591 | 58919.1504 | 59366.9725 | 1183.3395 | 1083.3395 | 1165.5404
8387 | 94.9783 | 1.0358 | 0.92955 | 1.0199 | 204.9746 | 49.2357 | 51728.1947 | 49678 | 989.56 | 889.56 | 1165.56
8388 | 122.4445 | 1.0956 | 0.81701 | 0.9564 | 200.3932 | 48.166 | 44298.5566 | 46396.0639 | 923.9213 | 823.9213 | 1165.6221
8389 | 95.794 | 0.87119 | 1.0087 | 1.0005 | 192.2074 | 46.5392 | 51044.9381 | 51973.1638 | 1035.4633 | 935.4633 | 1165.7481
8390 | 94.7112 | 0.90676 | 0.88093 | 1.0581 | 215.9411 | 49.9437 | 51057.634 | 50652.5252 | 1009.0505 | 909.0505 | 1165.9774
8391 | 97.794 | 1.0003 | 0.98529 | 0.95459 | 200.4426 | 52.3712 | 54580.9271 | 53787.1936 | 1071.7439 | 971.7439 | 1166.0496
8392 | 102.11 | 1.0397 | 0.89383 | 0.98323 | 198.1453 | 51.0019 | 50435.2518 | 50655.5632 | 1009.1113 | 909.1113 | 1166.0847
8393 | 112.0471 | 0.98257 | 1.0269 | 1.0329 | 202.1097 | 50.7266 | 59591.0509 | 60435.6729 | 1204.7135 | 1104.7135 | 1166.1457
8394 | 107.4404 | 1.0128 | 0.8632 | 1.0293 | 202.2053 | 50.1791 | 50244.0473 | 50448.0239 | 1004.9605 | 904.9605 | 1166.2061
8395 | 115.6229 | 1.0854 | 0.91165 | 0.94393 | 208.7452 | 53.8513 | 53307.7797 | 54816.3122 | 1092.3262 | 992.3262 | 1166.2435
8396 | 102.6429 | 1.0661 | 0.91287 | 0.96624 | 206.2261 | 49.9158 | 49696.4525 | 50608.7239 | 1008.1745 | 908.1745 | 1166.2527
8397 | 97.1225 | 1.0409 | 1.0242 | 1.0467 | 184.0494 | 48.6816 | 57292.616 | 58173.1419 | 1159.4628 | 1059.4628 | 1166.3483
8398 | 100.0833 | 1.0238 | 1.088 | 1.0396 | 212.5617 | 50.6767 | 62724.0011 | 60762.9327 | 1211.2587 | 1111.2587 | 1166.3557
8399 | 96.3652 | 0.91342 | 1.0102 | 0.95884 | 195.0086 | 53.4872 | 56710.7083 | 56414.559 | 1124.2912 | 1024.2912 | 1166.4665
8400 | 94.5585 | 1.0097 | 0.98371 | 0.94242 | 201.6406 | 48.2041 | 49492.4038 | 49970.8052 | 995.4161 | 895.4161 | 1166.4831
8401 | 94.6804 | 1.0272 | 0.84613 | 1.0522 | 205.6741 | 46.2382 | 45868.7965 | 48865.5363 | 973.3107 | 873.3107 | 1166.6882
8402 | 94.6103 | 1.0789 | 0.69269 | 1.0114 | 193.9362 | 53.3519 | 43016.9964 | 42564.412 | 847.2882 | 747.2882 | 1166.6999
8403 | 100.3377 | 0.98845 | 1.1177 | 1.0259 | 195.9838 | 43.8441 | 54818.5288 | 54369.5422 | 1083.3908 | 983.3908 | 1166.7207
8404 | 108.5804 | 0.98767 | 1.0952 | 1.0208 | 187.1764 | 50.5752 | 62152.4086 | 60684.6917 | 1209.6938 | 1109.6938 | 1166.8102
8405 | 93.0059 | 1.0774 | 0.78091 | 0.9807 | 179.5248 | 50.4208 | 43846.0931 | 43198.0158 | 859.9603 | 759.9603 | 1166.8916
8406 | 91.5593 | 0.9464 | 0.81164 | 1.034 | 197.1776 | 50.0061 | 46498.6088 | 46774.9335 | 931.4987 | 831.4987 | 1167.074
8407 | 101.3176 | 0.98084 | 1.1752 | 0.96379 | 198.2763 | 47.7059 | 58971.6602 | 58942.2594 | 1174.8452 | 1074.8452 | 1167.1446
8408 | 118.4138 | 1.0837 | 1.0476 | 1.1165 | 183.9273 | 49.6223 | 64590.7678 | 62619.9741 | 1248.3995 | 1148.3995 | 1167.1947
8409 | 89.562 | 1.0181 | 1.1422 | 1.0888 | 206.9085 | 45.3508 | 60740.9613 | 59472.0687 | 1185.4414 | 1085.4414 | 1167.3042
8410 | 92.3557 | 0.97702 | 1.0234 | 0.95508 | 200.2404 | 48.2664 | 51733.6434 | 52708.8406 | 1050.1768 | 950.1768 | 1167.4294
8411 | 91.1089 | 1.0462 | 0.93897 | 1.0131 | 200.8618 | 47.1578 | 49556.5811 | 48408.4272 | 964.1685 | 864.1685 | 1167.4592
8412 | 94.9126 | 1.0504 | 1.1279 | 0.9301 | 220.5281 | 49.4971 | 57081.4266 | 56851.5581 | 1133.0312 | 1033.0312 | 1167.64
8413 | 96.8735 | 1.0122 | 0.81926 | 1.0111 | 173.4091 | 48.8284 | 45408.0872 | 45983.1003 | 915.662 | 815.662 | 1167.7765
8414 | 93.041 | 1.0674 | 1.0527 | 1.0265 | 207.3495 | 47.7297 | 56525.014 | 56157.8001 | 1119.156 | 1019.156 | 1167.7781
8415 | 97.9971 | 1.018 | 0.95134 | 1.047 | 191.4 | 53.5163 | 58835.5628 | 59178.8023 | 1179.576 | 1079.576 | 1167.836
8416 | 100.0895 | 0.97701 | 1.145 | 1.0064 | 205.2835 | 54.146 | 67892.2011 | 67820.252 | 1352.405 | 1252.405 | 1167.9007
8417 | 98.0681 | 0.92584 | 0.81482 | 1.0164 | 191.6693 | 53.4203 | 49282.8405 | 50787.6715 | 1011.7534 | 911.7534 | 1167.9837
8418 | 108.5683 | 1.0013 | 0.77458 | 0.98117 | 197.8779 | 46.81 | 40862.1204 | 40620.7318 | 808.4146 | 708.4146 | 1168.0155
8419 | 97.4577 | 1.0141 | 0.88045 | 1.0783 | 195.6724 | 51.7853 | 54477.003 | 51892.8431 | 1033.8569 | 933.8569 | 1168.1047
8420 | 99.6842 | 0.9811 | 1.047 | 0.96042 | 207.7911 | 49.6127 | 54946.8574 | 53919.6638 | 1074.3933 | 974.3933 | 1168.3677
8421 | 105.3041 | 1.013 | 1.0616 | 0.97513 | 192.5424 | 52.2879 | 59896.5245 | 59856.4823 | 1193.1296 | 1093.1296 | 1168.5023
8422 | 103.6372 | 0.96399 | 0.95211 | 0.90875 | 202.6199 | 53.0136 | 51368.0698 | 50546.1012 | 1006.922 | 906.922 | 1168.5276
8423 | 103.1034 | 1.0441 | 0.92819 | 0.95459 | 212.0847 | 50.693 | 50585.0998 | 50573.1977 | 1007.464 | 907.464 | 1168.7616
8424 | 92.544 | 0.94023 | 0.92535 | 0.91282 | 190.8831 | 50.3417 | 47093.9336 | 46643.1288 | 928.8626 | 828.8626 | 1168.7783
8425 | 93.3658 | 0.99957 | 1.0702 | 0.97562 | 203.4348 | 46.2917 | 52855.5323 | 52810.4857 | 1052.2097 | 952.2097 | 1168.8668
8426 | 111.2753 | 1.059 | 1.1572 | 1.0625 | 180.762 | 50.1514 | 67754.3601 | 66653.6874 | 1329.0737 | 1229.0737 | 1168.9037
8427 | 81.4486 | 0.99384 | 1.1693 | 1.0234 | 197.2664 | 52.7045 | 67531.5732 | 69605.2458 | 1388.1049 | 1288.1049 | 1168.9473
8428 | 99.4691 | 1.0807 | 1.0422 | 0.98435 | 186.0504 | 48.6232 | 55295.0405 | 57145.1381 | 1138.9028 | 1038.9028 | 1168.9518
8429 | 101.4155 | 0.93171 | 1.0432 | 1.0421 | 207.7664 | 52.6829 | 62459.2047 | 60953.7735 | 1215.0755 | 1115.0755 | 1169.0957
8430 | 86.6207 | 1.0714 | 0.90445 | 1.0137 | 190.2362 | 49.1952 | 49857.9889 | 51142.7521 | 1018.855 | 918.855 | 1169.1304
8431 | 87.882 | 0.92389 | 1.0793 | 1.0304 | 191.9041 | 49.0541 | 58731.6776 | 60236.0907 | 1200.7218 | 1100.7218 | 1169.2563
8432 | 88.5008 | 0.97112 | 1.3335 | 0.95303 | 192.4796 | 52.4821 | 71400.0847 | 69190.544 | 1379.8109 | 1279.8109 | 1169.3237
8433 | 93.8212 | 0.97969 | 1.1996 | 1.0406 | 206.7343 | 55.1347 | 74096.9738 | 76547.3552 | 1526.9471 | 1426.9471 | 1169.4312
8434 | 87.7981 | 0.90379 | 0.77446 | 1.1041 | 201.1417 | 49.5292 | 46481.9183 | 45968.5788 | 915.3716 | 815.3716 | 1169.449
8435 | 102.377 | 0.96477 | 0.94732 | 0.96106 | 214.7708 | 50.1236 | 50799.8855 | 51128.77 | 1018.5754 | 918.5754 | 1169.5542
8436 | 75.2639 | 1.0006 | 1.2968 | 0.91343 | 187.2895 | 52.587 | 66438.0461 | 66796.2134 | 1331.9243 | 1231.9243 | 1169.6266
8437 | 102.9105 | 0.94806 | 1.0652 | 0.95368 | 207.0033 | 51.4552 | 57499.9084 | 56708.5347 | 1130.1707 | 1030.1707 | 1169.629
8438 | 107.8954 | 0.89662 | 0.91333 | 1.0061 | 205.1743 | 45.9315 | 46855.361 | 47049.8782 | 936.9976 | 836.9976 | 1169.6462
8439 | 96.7398 | 0.98987 | 1.0908 | 0.91584 | 205.5115 | 48.04 | 52796.304 | 55136.8366 | 1098.7367 | 998.7367 | 1169.7111
8440 | 87.3172 | 1.0453 | 0.85102 | 1.0076 | 219.7579 | 50.9664 | 48574.6578 | 48977.7324 | 975.5546 | 875.5546 | 1169.7575
8441 | 108.0291 | 0.96823 | 1.1425 | 0.99506 | 208.5276 | 47.0094 | 58566.5887 | 60254.3935 | 1201.0879 | 1101.0879 | 1169.8157
8442 | 88.1687 | 1.0016 | 0.91245 | 1.0518 | 214.1403 | 47.6529 | 50153.6186 | 49764.6148 | 991.2923 | 891.2923 | 1169.8462
8443 | 91.7611 | 1.0152 | 1.0142 | 0.99014 | 194.7738 | 48.0269 | 52898.4473 | 53365.0431 | 1063.3009 | 963.3009 | 1169.868
8444 | 98.1023 | 1.0023 | 1.2216 | 1.0097 | 200.5953 | 55.2991 | 73845.1769 | 74983.782 | 1495.6756 | 1395.6756 | 1169.9513
8445 | 107.1028 | 1.0001 | 1.0496 | 1.05 | 198.3088 | 49.5412 | 60100.7621 | 59344.3743 | 1182.8875 | 1082.8875 | 1169.9658
8446 | 111.3809 | 0.90946 | 0.88956 | 1.0338 | 187.721 | 52.5526 | 53838.8841 | 54268.6882 | 1081.3738 | 981.3738 | 1170.135
8447 | 116.5006 | 0.99054 | 0.97105 | 0.97494 | 217.5898 | 48.2023 | 51414.0626 | 51706.0654 | 1030.1213 | 930.1213 | 1170.2217
8448 | 104.1712 | 0.98499 | 0.98679 | 1.0048 | 198.8451 | 52.5827 | 57730.1273 | 57136.3824 | 1138.7276 | 1038.7276 | 1170.2462
8449 | 104.6507 | 0.97577 | 0.99445 | 0.97874 | 218.9761 | 46.8032 | 50552.1402 | 50767.7894 | 1011.3558 | 911.3558 | 1170.2696
8450 | 92.6991 | 0.91132 | 1.0179 | 1.0793 | 203.3307 | 48.4004 | 57463.3457 | 57001.5165 | 1136.0303 | 1036.0303 | 1170.2768
8451 | 94.2104 | 0.99084 | 0.9145 | 1.0444 | 201.3539 | 48.6411 | 51196.9947 | 50022.3368 | 996.4467 | 896.4467 | 1170.4753
8452 | 96.068 | 1.0134 | 0.87551 | 0.96829 | 223.6861 | 50.8615 | 48292.9604 | 46773.0921 | 931.4618 | 831.4618 | 1170.4817
8453 | 99.5722 | 1.0603 | 0.87006 | 0.97406 | 201.732 | 48.262 | 46198.274 | 45313.269 | 902.2654 | 802.2654 | 1170.4974
8454 | 104.6785 | 0.98499 | 1.0032 | 1.0248 | 204.7609 | 47.2979 | 53705.9094 | 52964.3387 | 1055.2868 | 955.2868 | 1170.6282
8455 | 95.6872 | 0.98037 | 0.91067 | 1.0822 | 202.2443 | 48.8479 | 52923.7133 | 52076.741 | 1037.5348 | 937.5348 | 1170.7454
8456 | 105.8788 | 0.92214 | 1.0737 | 1.0242 | 197.2845 | 51.6392 | 62026.9685 | 61107.4342 | 1218.1487 | 1118.1487 | 1170.7759
8457 | 86.2511 | 1.0484 | 1.0781 | 0.91345 | 193.9826 | 46.1592 | 49826.1827 | 51719.3962 | 1030.3879 | 930.3879 | 1170.8373
8458 | 111.792 | 1.0338 | 0.96356 | 1.0997 | 203.4209 | 46.6559 | 55035.1242 | 55898.6558 | 1113.9731 | 1013.9731 | 1171.0754
8459 | 100.7698 | 0.94916 | 1.0298 | 1.0597 | 186.1286 | 47.0323 | 56009.0577 | 55028.5417 | 1096.5708 | 996.5708 | 1171.1584
8460 | 96.2771 | 0.93865 | 0.95544 | 1.0168 | 202.3002 | 50.9914 | 54348.7938 | 54259.35 | 1081.187 | 981.187 | 1171.2949
8461 | 86.1702 | 0.94949 | 0.95742 | 1.0508 | 196.654 | 47.1937 | 51538.6519 | 51487.1513 | 1025.743 | 925.743 | 1171.4584
8462 | 105.2981 | 0.98807 | 0.99234 | 1.0137 | 186.115 | 48.6474 | 54184.5714 | 55365.3219 | 1103.3064 | 1003.3064 | 1171.5387
8463 | 86.1765 | 0.94653 | 1.0939 | 0.97902 | 196.6628 | 48.5912 | 56200.7188 | 57500.549 | 1146.011 | 1046.011 | 1171.6257
8464 | 84.0426 | 1.0876 | 1.1472 | 0.95238 | 197.0755 | 54.8022 | 65083.658 | 65915.2223 | 1314.3044 | 1214.3044 | 1171.886
8465 | 104.3109 | 1.0743 | 0.96377 | 1.0267 | 192.9908 | 43.0222 | 47584.5962 | 47878.3882 | 953.5678 | 853.5678 | 1171.9058
8466 | 80.2236 | 1.04 | 0.94328 | 1.0539 | 215.2965 | 50.4212 | 54544.9984 | 53726.9337 | 1070.5387 | 970.5387 | 1172.0031
8467 | 96.8055 | 0.99706 | 1.1284 | 0.96827 | 187.6664 | 45.1944 | 53927.4829 | 53120.4856 | 1058.4097 | 958.4097 | 1172.005
8468 | 118.9424 | 0.9984 | 0.99172 | 0.98291 | 197.1423 | 52.9166 | 58062.9148 | 59251.0779 | 1181.0216 | 1081.0216 | 1172.0462
8469 | 90.0739 | 1.0061 | 1.1921 | 1.0382 | 199.6516 | 45.6739 | 60867.8783 | 59959.2446 | 1195.1849 | 1095.1849 | 1172.1062
8470 | 89.9404 | 0.90447 | 0.91554 | 0.99677 | 199.3551 | 51.7976 | 51682.6505 | 51391.4022 | 1023.828 | 923.828 | 1172.1121
8471 | 109.2382 | 0.94109 | 0.89798 | 0.95381 | 181.8479 | 49.5197 | 47686.1062 | 47372.007 | 943.4401 | 843.4401 | 1172.1219
8472 | 99.0427 | 1.0167 | 0.99766 | 1.0091 | 208.4412 | 46.346 | 51535.1203 | 50936.7744 | 1014.7355 | 914.7355 | 1172.1807
8473 | 104.5229 | 1.0314 | 0.77562 | 1.0502 | 206.8353 | 51.986 | 48157.5194 | 46242.4721 | 920.8494 | 820.8494 | 1172.1894
8474 | 102.4318 | 1.0342 | 1.0118 | 1.0642 | 206.4975 | 56.5683 | 67106.7709 | 65371.9316 | 1303.4386 | 1203.4386 | 1172.243
8475 | 105.2217 | 1.0276 | 0.96588 | 1.003 | 202.3766 | 52.2645 | 56486.4835 | 57333.1092 | 1142.6622 | 1042.6622 | 1172.279
8476 | 101.6699 | 0.96091 | 1.1199 | 0.98512 | 195.9113 | 48.2103 | 58092.4075 | 57264.8036 | 1141.2961 | 1041.2961 | 1172.3002
8477 | 94.6206 | 0.97327 | 0.77865 | 0.99107 | 195.3525 | 49.9966 | 43382.1214 | 40634.1583 | 808.6832 | 708.6832 | 1172.4542
8478 | 79.3806 | 0.96693 | 1.0692 | 1.0656 | 205.6411 | 49.26 | 60110.8632 | 58436.6614 | 1164.7332 | 1064.7332 | 1172.4893
8479 | 105.804 | 1.0933 | 1.0864 | 1.0048 | 208.1551 | 53.9944 | 65393.9074 | 65185.4185 | 1299.7084 | 1199.7084 | 1172.5129
8480 | 99.9083 | 1.0266 | 0.96335 | 1.0119 | 210.0679 | 48.232 | 52174.3917 | 52471.0718 | 1045.4214 | 945.4214 | 1172.5247
8481 | 103.3325 | 1.0177 | 1.1419 | 0.88159 | 187.5235 | 46.5879 | 51986.1672 | 53280.5695 | 1061.6114 | 961.6114 | 1172.5384
8482 | 102.7606 | 1.0207 | 1.0851 | 0.89374 | 182.1187 | 50.7572 | 54730.3168 | 56608.6189 | 1128.1724 | 1028.1724 | 1172.8955
8483 | 101.2642 | 0.95606 | 1.2573 | 0.99908 | 205.2823 | 49.2688 | 66863.3265 | 66092.2607 | 1317.8452 | 1217.8452 | 1173.0389
8484 | 97.6248 | 1.0376 | 0.98619 | 0.99298 | 215.5387 | 45.0289 | 48871.8193 | 48361.9495 | 963.239 | 863.239 | 1173.0812
8485 | 94.2873 | 1.002 | 1.229 | 0.96441 | 185.3959 | 46.4462 | 59624.6182 | 58925.5385 | 1174.5108 | 1074.5108 | 1173.2598
8486 | 95.8686 | 0.97156 | 1.2271 | 0.99287 | 197.4581 | 50.9327 | 66993.55 | 64589.463 | 1287.7893 | 1187.7893 | 1173.3123
8487 | 101.0936 | 0.98111 | 1.1 | 0.91494 | 201.9628 | 52.7464 | 58521.4666 | 59349.4632 | 1182.9893 | 1082.9893 | 1173.3884
8488 | 109.8124 | 0.95256 | 1.1681 | 0.98997 | 189.5149 | 49.8361 | 63031.4345 | 63456.2826 | 1265.1257 | 1165.1257 | 1173.4328
8489 | 96.2904 | 1.0578 | 1.1114 | 1.0576 | 197.8728 | 49.7515 | 63747.1786 | 64950.2334 | 1295.0047 | 1195.0047 | 1173.4851
8490 | 90.1181 | 0.95583 | 0.93336 | 0.99269 | 201.017 | 51.7823 | 52639.6332 | 52225.4002 | 1040.508 | 940.508 | 1173.5835
8491 | 91.9349 | 0.98915 | 0.99646 | 1.0302 | 201.6956 | 49.4648 | 55478.0349 | 56164.5625 | 1119.2912 | 1019.2912 | 1173.7111
8492 | 96.4058 | 0.98344 | 1.0107 | 1.037 | 189.0847 | 46.1337 | 52917.2209 | 53188.6311 | 1059.7726 | 959.7726 | 1174.0018
8493 | 114.1497 | 1.0067 | 0.92035 | 1.0361 | 194.201 | 43.63 | 46813.2386 | 46354.4081 | 923.0882 | 823.0882 | 1174.0527
8494 | 99.6606 | 0.98617 | 0.85836 | 0.98898 | 201.2779 | 51.4874 | 48969.0039 | 48882.3658 | 973.6473 | 873.6473 | 1174.0596
8495 | 97.6 | 0.95277 | 1.1299 | 1.0481 | 212.5857 | 52.8972 | 67776.2041 | 68650.9623 | 1369.0192 | 1269.0192 | 1174.0698
8496 | 88.8337 | 0.99291 | 1.0616 | 0.96866 | 195.8381 | 44.4856 | 49865.9939 | 49101.0303 | 978.0206 | 878.0206 | 1174.3293
8497 | 95.2145 | 1.0531 | 0.92744 | 0.9107 | 188.3649 | 54.1128 | 51318.7498 | 51429.0086 | 1024.5802 | 924.5802 | 1174.3814
8498 | 99.9216 | 0.97991 | 1.2115 | 1.0483 | 190.046 | 52.3941 | 71858.8894 | 75529.6958 | 1506.5939 | 1406.5939 | 1174.4052
8499 | 99.7789 | 0.97319 | 0.95829 | 0.93207 | 188.9003 | 51.2592 | 50950.6638 | 51399.7003 | 1023.994 | 923.994 | 1174.4106
8500 | 96.8783 | 1.0326 | 1.1377 | 0.95713 | 197.2356 | 51.8419 | 61838.0714 | 59577.7452 | 1187.5549 | 1087.5549 | 1174.4498
8501 | 89.467 | 0.98301 | 1.0617 | 0.90041 | 187.3484 | 53.2281 | 55755.055 | 54233.3703 | 1080.6674 | 980.6674 | 1174.4593
8502 | 104.138 | 1.006 | 0.90726 | 0.95872 | 221.6085 | 50.8703 | 49798.7909 | 49921.8512 | 994.437 | 894.437 | 1174.4678
8503 | 99.0352 | 1.0192 | 1.0503 | 1.0105 | 200.4836 | 50.4612 | 58852.5937 | 59389.0172 | 1183.7803 | 1083.7803 | 1174.4706
8504 | 99.1331 | 0.97167 | 1.1574 | 1.0586 | 183.8081 | 50.4276 | 66826.3624 | 67575.8465 | 1347.5169 | 1247.5169 | 1174.497
8505 | 91.7945 | 0.96108 | 1.1866 | 1.0549 | 206.5979 | 50.0072 | 67215.3816 | 66328.6116 | 1322.5722 | 1222.5722 | 1174.5562
8506 | 104.5551 | 1.0798 | 0.85828 | 1.0235 | 208.317 | 48.2213 | 48010.6686 | 48654.2825 | 969.0856 | 869.0856 | 1174.596
8507 | 93.8562 | 1.0065 | 1.026 | 0.97829 | 209.8681 | 46.0955 | 50830.6964 | 51234.1828 | 1020.6837 | 920.6837 | 1174.63
8508 | 80.881 | 1.0476 | 1.212 | 0.99811 | 201.9642 | 49.7008 | 64535.9819 | 66094.151 | 1317.883 | 1217.883 | 1174.6978
8509 | 100.5689 | 0.99091 | 0.8275 | 1.0394 | 196.5789 | 51.0857 | 49224.9714 | 49453.7302 | 985.0746 | 885.0746 | 1174.9004
8510 | 92.7255 | 1.0473 | 1.0594 | 1.0248 | 214.3909 | 51.4727 | 61099.7001 | 61053.9327 | 1217.0787 | 1117.0787 | 1174.9599
8511 | 109.0433 | 1.017 | 0.86777 | 1.061 | 187.4291 | 48.6982 | 50426.4396 | 51799.9799 | 1031.9996 | 931.9996 | 1175.0197
8512 | 112.6099 | 1.0613 | 1.0215 | 0.97668 | 209.6526 | 49.0879 | 55050.3717 | 54655.6063 | 1089.1121 | 989.1121 | 1175.3583
8513 | 97.1662 | 1.0611 | 0.68514 | 1.0727 | 206.4408 | 48.5439 | 40889.2935 | 42068.3936 | 837.3679 | 737.3679 | 1175.5454
8514 | 110.5378 | 0.97395 | 0.7801 | 0.9859 | 192.7541 | 49.3689 | 43477.211 | 44550.6449 | 887.0129 | 787.0129 | 1175.5547
8515 | 98.1122 | 0.96453 | 0.9948 | 1.0274 | 202.1525 | 48.0962 | 53910.1947 | 53699.812 | 1069.9962 | 969.9962 | 1175.7662
8516 | 111.7594 | 0.99689 | 0.78359 | 0.96605 | 205.2889 | 51.8332 | 45217.21 | 46073.7938 | 917.4759 | 817.4759 | 1175.7665
8517 | 115.2044 | 1.027 | 1.013 | 1.0705 | 207.8123 | 53.5479 | 64614.1297 | 64189.2678 | 1279.7854 | 1179.7854 | 1175.8123
8518 | 100.3087 | 0.86501 | 1.197 | 1.0366 | 195.6831 | 49.4562 | 65849.2261 | 67759.3835 | 1351.1877 | 1251.1877 | 1175.8434
8519 | 116.2816 | 0.97843 | 1.1064 | 0.94097 | 208.5947 | 48.1 | 55757.8423 | 56473.6166 | 1125.4723 | 1025.4723 | 1175.8799
8520 | 100.1358 | 1.0103 | 0.91643 | 0.9107 | 202.4762 | 51.9572 | 48822.1607 | 48160.3734 | 959.2075 | 859.2075 | 1175.9044
8521 | 105.9791 | 0.96961 | 1.1295 | 1.0024 | 199.675 | 50.2533 | 62264.9454 | 62620.9891 | 1248.4198 | 1148.4198 | 1175.9493
8522 | 98.7562 | 1.0212 | 0.86831 | 1.0563 | 212.8917 | 46.6465 | 47701.5301 | 48699.0178 | 969.9804 | 869.9804 | 1176.0236
8523 | 90.0544 | 0.90093 | 1.0776 | 0.97848 | 211.1727 | 47.8754 | 54576.2369 | 54194.9366 | 1079.8987 | 979.8987 | 1176.0556
8524 | 111.3374 | 1.0274 | 1.0656 | 0.97774 | 191.2461 | 49.6875 | 57644.9212 | 58364.2002 | 1163.284 | 1063.284 | 1176.1188
8525 | 87.2026 | 0.94075 | 0.96389 | 0.99083 | 208.5374 | 49.4305 | 51472.4536 | 51879.2431 | 1033.5849 | 933.5849 | 1176.1547
8526 | 112.5584 | 1.0127 | 0.76321 | 1.0005 | 187.4516 | 47.4059 | 41788.0583 | 41642.7972 | 828.8559 | 728.8559 | 1176.2356
8527 | 102.7564 | 1.0867 | 1.0594 | 1.0977 | 208.1933 | 53.025 | 67788.8175 | 66442.7484 | 1324.855 | 1224.855 | 1176.4724
8528 | 99.2232 | 0.9497 | 1.1079 | 1.0065 | 197.0587 | 46.7155 | 56693.6062 | 57824.8488 | 1152.497 | 1052.497 | 1176.4835
8529 | 95.4213 | 0.91482 | 0.61054 | 0.99693 | 201.0506 | 50.0256 | 35016.6677 | 34961.0238 | 695.2205 | 595.2205 | 1176.5119
8530 | 101.9409 | 0.90244 | 0.9768 | 0.98718 | 182.3716 | 50.6068 | 53636.981 | 53591.3931 | 1067.8279 | 967.8279 | 1176.571
8531 | 87.5405 | 0.87287 | 0.9927 | 0.9687 | 198.4252 | 50.254 | 52363.6949 | 53150.9774 | 1059.0195 | 959.0195 | 1176.6088
8532 | 95.4547 | 1.0343 | 0.87961 | 1.0326 | 214.0044 | 51.7843 | 52363.2951 | 52810.8939 | 1052.2179 | 952.2179 | 1176.7007
8533 | 86.5083 | 1.0552 | 1.2126 | 0.97356 | 202.621 | 50.0939 | 63914.9291 | 62976.4178 | 1255.5284 | 1155.5284 | 1176.8323
8534 | 108.6155 | 0.96257 | 0.81049 | 0.9528 | 195.2023 | 48.0734 | 42345.6076 | 42690.515 | 849.8103 | 749.8103 | 1176.8444
8535 | 86.9656 | 1.024 | 1.1642 | 1.0052 | 200.6945 | 48.305 | 61033.7934 | 59399.2548 | 1183.9851 | 1083.9851 | 1176.9351
8536 | 92.506 | 1.0246 | 1.1258 | 0.96649 | 206.7489 | 48.5301 | 57609.9382 | 55355.4581 | 1103.1092 | 1003.1092 | 1176.9396
8537 | 88.9054 | 1.0103 | 0.96257 | 0.96019 | 207.5477 | 51.2607 | 52189.6622 | 52009.2886 | 1036.1858 | 936.1858 | 1176.9564
8538 | 106.4664 | 1.0355 | 1.0662 | 1.0209 | 205.7109 | 48.6832 | 58564.3193 | 58725.8993 | 1170.518 | 1070.518 | 1176.9878
8539 | 101.3607 | 0.99015 | 0.8587 | 1.017 | 204.9012 | 52.9649 | 51772.6803 | 51208.8193 | 1020.1764 | 920.1764 | 1177.0049
8540 | 101.8743 | 0.89152 | 1.0376 | 1.0685 | 205.6814 | 49.6261 | 59734.0967 | 59501.7124 | 1186.0342 | 1086.0342 | 1177.0381
8541 | 101.0386 | 1.028 | 0.97909 | 1.0456 | 191.3781 | 48.6661 | 55067.7898 | 56576.2892 | 1127.5258 | 1027.5258 | 1177.1005
8542 | 91.4503 | 0.96671 | 0.98314 | 0.96734 | 188.0919 | 51.4605 | 53677.9814 | 52854.936 | 1053.0987 | 953.0987 | 1177.1072
8543 | 89.585 | 0.97341 | 0.92615 | 0.89821 | 216.4901 | 50.1816 | 46337.3398 | 46435.4791 | 924.7096 | 824.7096 | 1177.117
8544 | 110.7963 | 0.93757 | 1.0232 | 1.0177 | 187.4433 | 47.1152 | 54139.4658 | 54556.4654 | 1087.1293 | 987.1293 | 1177.1834
8545 | 95.7184 | 0.97971 | 0.95248 | 0.96826 | 192.213 | 53.7041 | 54756.8749 | 53556.9779 | 1067.1396 | 967.1396 | 1177.2607
8546 | 83.337 | 1.0279 | 0.9464 | 1.0549 | 197.7658 | 46.0954 | 50166.0723 | 49053.2531 | 977.0651 | 877.0651 | 1177.2737
8547 | 98.5668 | 1.05 | 1.0469 | 1.0088 | 182.1564 | 53.338 | 62029.2238 | 59745.2178 | 1190.9044 | 1090.9044 | 1177.4966
8548 | 101.1699 | 0.94227 | 1.0503 | 0.93947 | 190.2645 | 48.5279 | 52699.9558 | 53142.5458 | 1058.8509 | 958.8509 | 1177.6094
8549 | 86.319 | 0.96295 | 0.96814 | 0.94894 | 200.3879 | 48.4018 | 48690.5048 | 47530.3299 | 946.6066 | 846.6066 | 1177.6126
8550 | 95.1097 | 1.0054 | 0.92664 | 0.97008 | 211.7775 | 48.2645 | 48212.9442 | 48432.4276 | 964.6486 | 864.6486 | 1177.8871
8551 | 86.5341 | 0.95435 | 0.79545 | 0.97226 | 193.3739 | 48.3528 | 41581.7593 | 42461.8015 | 845.236 | 745.236 | 1177.9117
8552 | 96.0391 | 1.0253 | 1.1951 | 1 | 201.7191 | 50.1735 | 65106.8558 | 65987.7702 | 1315.7554 | 1215.7554 | 1177.9539
8553 | 98.9106 | 0.99663 | 0.85654 | 0.93255 | 195.8425 | 53.9075 | 48569.6138 | 47306.2049 | 942.1241 | 842.1241 | 1178.0704
8554 | 96.0234 | 0.95863 | 0.8519 | 0.9859 | 186.8943 | 48.7606 | 45628.7784 | 45570.4103 | 907.4082 | 807.4082 | 1178.2032
8555 | 96.1923 | 0.98869 | 1.0433 | 1.0794 | 203.9632 | 51.092 | 62597.9302 | 64103.5189 | 1278.0704 | 1178.0704 | 1178.2236
8556 | 107.3814 | 0.92923 | 1.0759 | 0.93575 | 197.2176 | 47.945 | 53248.6724 | 53041.0506 | 1056.821 | 956.821 | 1178.2609
8557 | 91.7205 | 1.0344 | 1.1262 | 1.0882 | 212.7794 | 50.6604 | 67106.1185 | 68791.3265 | 1371.8265 | 1271.8265 | 1178.3124
8558 | 87.5639 | 1.0352 | 1.0089 | 0.98463 | 182.3562 | 48.4207 | 52670.5426 | 51448.9509 | 1024.979 | 924.979 | 1178.3193
8559 | 113.6787 | 1.0224 | 0.92558 | 0.95263 | 188.0886 | 48.0437 | 48134.1528 | 47907.1436 | 954.1429 | 854.1429 | 1178.4003
8560 | 92.5558 | 0.96779 | 0.87177 | 0.95573 | 215.3357 | 51.0646 | 47335.4009 | 45087.2075 | 897.7441 | 797.7441 | 1178.5329
8561 | 108.9402 | 0.98961 | 0.74941 | 1.1294 | 188.068 | 49.2681 | 47197.3783 | 46097.5656 | 917.9513 | 817.9513 | 1178.5561
8562 | 85.4922 | 0.88661 | 1.1204 | 1.0241 | 197.6967 | 52.5119 | 64427.8973 | 64963.4293 | 1295.2686 | 1195.2686 | 1178.5857
8563 | 94.665 | 0.98697 | 1.1425 | 0.99873 | 184.7609 | 46.611 | 57724.5552 | 58485.8179 | 1165.7164 | 1065.7164 | 1178.6248
8564 | 100.4186 | 1.025 | 1.0941 | 1.0409 | 196.0377 | 51.2103 | 63784.7475 | 63279.5794 | 1261.5916 | 1161.5916 | 1178.6318
8565 | 112.4776 | 0.97337 | 0.99293 | 0.92936 | 201.8389 | 48.1094 | 49863.5455 | 48613.8296 | 968.2766 | 868.2766 | 1178.6586
8566 | 102.8735 | 1.0718 | 1.2396 | 1.0157 | 199.4222 | 43.7023 | 60041.0702 | 59704.1989 | 1190.084 | 1090.084 | 1178.7135
8567 | 101.7356 | 0.92661 | 0.96336 | 1.0499 | 209.3719 | 47.6247 | 52867.9757 | 52659.4888 | 1049.1898 | 949.1898 | 1178.7483
8568 | 104.9231 | 0.94401 | 0.88984 | 0.94618 | 212.4424 | 52.0616 | 49202.2087 | 47318.1337 | 942.3627 | 842.3627 | 1178.98
8569 | 109.2399 | 0.96874 | 1.1184 | 0.97497 | 199.8955 | 45.4221 | 54537.4136 | 55363.1054 | 1103.2621 | 1003.2621 | 1179.1564
8570 | 93.0725 | 1.0328 | 1.1373 | 0.94216 | 196.0662 | 48.0035 | 56247.6915 | 54953.9421 | 1095.0788 | 995.0788 | 1179.1992
8571 | 99.3728 | 1.0125 | 1.0272 | 1.0351 | 200.5072 | 45.1723 | 52773.8267 | 52028.3519 | 1036.567 | 936.567 | 1179.312
8572 | 123.5489 | 1.02 | 1.0367 | 1.0635 | 201.7736 | 49.3621 | 60847.4059 | 62388.3001 | 1243.766 | 1143.766 | 1179.3378
8573 | 99.4966 | 1.0775 | 0.77738 | 0.9988 | 192.4187 | 50.1413 | 44499.9115 | 44842.2799 | 892.8456 | 792.8456 | 1179.4533
8574 | 87.5704 | 0.89052 | 0.93909 | 1.0203 | 212.4994 | 49.276 | 51270.0326 | 51219.0168 | 1020.3803 | 920.3803 | 1179.6261
8575 | 123.0164 | 1.0536 | 0.76038 | 1.0433 | 222.4637 | 53.6332 | 49722.2574 | 48935.2513 | 974.705 | 874.705 | 1179.6448
8576 | 94.1582 | 0.98939 | 1.0691 | 1.0068 | 212.1835 | 47.5828 | 55862.992 | 57660.7906 | 1149.2158 | 1049.2158 | 1179.668
8577 | 110.5804 | 0.90101 | 0.75809 | 1.0615 | 211.0833 | 51.718 | 46982.4517 | 47095.2547 | 937.9051 | 837.9051 | 1179.6753
8578 | 89.7752 | 0.94784 | 0.89841 | 1.0523 | 208.2715 | 52.6197 | 54429.936 | 53627.525 | 1068.5505 | 968.5505 | 1179.6791
8579 | 90.4839 | 1.0544 | 0.97108 | 1.0465 | 214.6968 | 49.9754 | 55769.1447 | 55780.3303 | 1111.6066 | 1011.6066 | 1179.7067
8580 | 108.6683 | 0.88147 | 0.86904 | 1.0346 | 205.5619 | 49.5637 | 49517.5799 | 50481.3011 | 1005.626 | 905.626 | 1179.7854
8581 | 92.9366 | 1.0599 | 1.0768 | 1.0768 | 199.0306 | 50.4634 | 63680.695 | 63523.3263 | 1266.4665 | 1166.4665 | 1179.8222
8582 | 101.0081 | 1.0281 | 0.91981 | 0.90454 | 191.0896 | 48.796 | 45857.2119 | 48112.0894 | 958.2418 | 858.2418 | 1179.8555
8583 | 100.153 | 0.94594 | 1.0838 | 1.0575 | 195.9852 | 51.4687 | 64060.7844 | 64218.0559 | 1280.3611 | 1180.3611 | 1179.9285
8584 | 79.7602 | 1.0137 | 1.2428 | 0.95181 | 196.7352 | 51.9669 | 65870.1657 | 63724.0873 | 1270.4817 | 1170.4817 | 1180.1513
8585 | 106.9247 | 0.99914 | 1.1392 | 0.98961 | 202.2959 | 52.4803 | 64970.9088 | 64848.307 | 1292.9661 | 1192.9661 | 1180.1586
8586 | 97.066 | 1.0515 | 0.87357 | 0.94409 | 183.8627 | 47.186 | 43915.7483 | 42740.6087 | 850.8122 | 750.8122 | 1180.1913
8587 | 106.0791 | 0.97094 | 1.0549 | 0.97857 | 201.0417 | 49.0558 | 55893.4752 | 57240.388 | 1140.8078 | 1040.8078 | 1180.3278
8588 | 100.4671 | 0.94559 | 1.0598 | 1.0491 | 195.0281 | 49.6767 | 60144.9862 | 58600.5395 | 1168.0108 | 1068.0108 | 1180.3611
8589 | 92.8421 | 0.99674 | 1.0889 | 1.0332 | 213.2824 | 48.9937 | 59871.3183 | 61621.7795 | 1228.4356 | 1128.4356 | 1180.4483
8590 | 114.7759 | 0.99502 | 0.92042 | 0.91015 | 205.4567 | 45.9021 | 43900.6889 | 44208.6024 | 880.172 | 780.172 | 1180.4924
8591 | 101.5662 | 1.0369 | 1.123 | 1.0572 | 208.9902 | 48.3141 | 62658.9479 | 64505.2012 | 1286.104 | 1186.104 | 1180.4933
8592 | 90.6829 | 0.93296 | 0.98551 | 0.98487 | 196.0175 | 53.5227 | 56673.462 | 56604.9216 | 1128.0984 | 1028.0984 | 1180.4983
8593 | 119.8376 | 0.99795 | 1.0739 | 0.96141 | 192.9812 | 51.7097 | 59762.9558 | 55343.7098 | 1102.8742 | 1002.8742 | 1180.5473
8594 | 100.2798 | 0.9034 | 0.82283 | 1.0316 | 190.8986 | 46.5906 | 43959.7357 | 44578.2455 | 887.5649 | 787.5649 | 1180.5823
8595 | 104.4253 | 0.98446 | 1.0228 | 0.99008 | 193.7298 | 50.33 | 56334.8099 | 56698.1742 | 1129.9635 | 1029.9635 | 1180.742
8596 | 103.2541 | 1.0454 | 0.80887 | 0.96627 | 192.3982 | 51.7952 | 46265.5544 | 46072.627 | 917.4525 | 817.4525 | 1180.9406
8597 | 106.0876 | 0.93992 | 1.147 | 0.90005 | 198.4419 | 49.0335 | 55707.2252 | 55555.6504 | 1107.113 | 1007.113 | 1181.0074
8598 | 109.1964 | 1.0112 | 0.52411 | 1.0743 | 205.9601 | 52.9823 | 35888.5016 | 35909.5729 | 714.1915 | 614.1915 | 1181.0087
8599 | 102.5536 | 1.0131 | 0.94038 | 0.94482 | 201.9593 | 50.4638 | 50281.3338 | 49762.219 | 991.2444 | 891.2444 | 1181.0137
8600 | 83.8869 | 1.0031 | 1.191 | 0.92897 | 195.4301 | 52.8061 | 63062.0724 | 62022.8315 | 1236.4566 | 1136.4566 | 1181.1508
8601 | 100.6752 | 0.99071 | 1.272 | 0.91858 | 197.9854 | 46.6563 | 59366.133 | 59603.8969 | 1188.0779 | 1088.0779 | 1181.159
8602 | 99.2092 | 0.9508 | 1.0364 | 1.0631 | 193.2836 | 48.6642 | 58400.8879 | 58918.3227 | 1174.3665 | 1074.3665 | 1181.1866
8603 | 100.695 | 1.0035 | 1.1064 | 0.94427 | 178.8762 | 49.7663 | 57199.3444 | 54736.6834 | 1090.7337 | 990.7337 | 1181.418
8604 | 89.3977 | 0.96201 | 0.71919 | 0.92427 | 198.2618 | 49.9075 | 37665.2052 | 38392.0439 | 763.8409 | 663.8409 | 1181.4272
8605 | 97.7511 | 0.8803 | 0.98873 | 0.99717 | 216.418 | 53.9898 | 58092.1723 | 57947.9964 | 1154.9599 | 1054.9599 | 1181.4985
8606 | 104.0292 | 1.0285 | 0.81626 | 1.0395 | 180.688 | 54.5531 | 52304.5627 | 52271.9328 | 1041.4387 | 941.4387 | 1181.5031
8607 | 103.978 | 1.0459 | 1.0205 | 1.0679 | 215.1691 | 44.9355 | 54069.621 | 54050.5564 | 1077.0111 | 977.0111 | 1181.7129
8608 | 103.3974 | 1.008 | 0.94417 | 0.94534 | 191.2637 | 41.6047 | 41662.0378 | 41203.0975 | 820.062 | 720.062 | 1181.7249
8609 | 110.2723 | 0.93836 | 0.9038 | 0.95594 | 216.6877 | 46.6323 | 45330.9807 | 45297.2639 | 901.9453 | 801.9453 | 1181.7811
8610 | 106.6245 | 1.0203 | 0.63947 | 0.97856 | 202.1546 | 48.8907 | 36114.84 | 37031.9247 | 736.6385 | 636.6385 | 1182.0036
8611 | 95.1634 | 1.0389 | 1.0036 | 0.94482 | 187.2256 | 53.5809 | 56290.0032 | 56049.4963 | 1116.9899 | 1016.9899 | 1182.0171
8612 | 112.5176 | 1.0714 | 0.76888 | 0.98557 | 199.0645 | 51.5281 | 45458.1577 | 46719.3801 | 930.3876 | 830.3876 | 1182.046
8613 | 99.6583 | 1.058 | 1.1679 | 0.95516 | 205.232 | 49.7655 | 60969.3347 | 60988.9822 | 1215.7796 | 1115.7796 | 1182.1183
8614 | 100.6105 | 1.0622 | 0.97296 | 1.016 | 191.9263 | 52.7276 | 57948.4126 | 57302.153 | 1142.0431 | 1042.0431 | 1182.1431
8615 | 86.8571 | 1.0002 | 1.0495 | 1.1038 | 212.1898 | 50.135 | 62642.8503 | 62632.5668 | 1248.6513 | 1148.6513 | 1182.2393
8616 | 98.2005 | 1.0049 | 1.2028 | 1.0757 | 189.9534 | 44.0729 | 61559.3763 | 61639.0164 | 1228.7803 | 1128.7803 | 1182.2503
8617 | 89.0034 | 1.0012 | 0.99954 | 1.0056 | 189.7682 | 45.1066 | 49549.9244 | 49391.2976 | 983.826 | 883.826 | 1182.3336
8618 | 85.0694 | 1.1441 | 0.93617 | 0.94298 | 191.6202 | 47.0866 | 46341.7593 | 48284.0263 | 961.6805 | 861.6805 | 1182.3721
8619 | 101.8487 | 1.0046 | 1.0355 | 0.91246 | 203.7335 | 49.1382 | 51657.9002 | 53190.8104 | 1059.8162 | 959.8162 | 1182.5001
8620 | 91.4877 | 1.0404 | 0.92724 | 0.95329 | 207.3464 | 46.343 | 45582.0968 | 46136.2399 | 918.7248 | 818.7248 | 1182.5842
8621 | 94.942 | 0.96129 | 1.0252 | 0.97724 | 192.7345 | 49.5565 | 54365.3539 | 55044.4834 | 1096.8897 | 996.8897 | 1182.7252
8622 | 97.6265 | 0.91862 | 0.92803 | 1.0246 | 181.7012 | 50.8404 | 53084.0642 | 56284.1491 | 1121.683 | 1021.683 | 1182.7418
8623 | 107.5997 | 0.98169 | 1.0652 | 1.0082 | 210.1946 | 54.2729 | 64227.9366 | 64880.004 | 1293.6001 | 1193.6001 | 1182.8199
8624 | 107.2429 | 1.0165 | 0.92488 | 1.0863 | 198.0614 | 50.5365 | 56479.4907 | 55680.0586 | 1109.6012 | 1009.6012 | 1182.8556
8625 | 110.9933 | 1.0632 | 0.63263 | 0.95529 | 204.4571 | 50.3894 | 36603.342 | 35843.8955 | 712.8779 | 612.8779 | 1182.8981
8626 | 107.6339 | 1.0973 | 0.84836 | 1.0156 | 187.0099 | 48.6051 | 47806.3232 | 47376.8742 | 943.5375 | 843.5375 | 1183.1623
8627 | 116.56 | 0.99697 | 0.93934 | 0.99653 | 180.9732 | 45.3716 | 47924.8671 | 48776.0529 | 971.5211 | 871.5211 | 1183.18
8628 | 111.923 | 1.0491 | 1.0128 | 1.0688 | 201.0745 | 48.0686 | 57877.1127 | 56483.9989 | 1125.68 | 1025.68 | 1183.232
8629 | 102.2388 | 0.98036 | 1.0278 | 1.0361 | 206.2554 | 50.8528 | 59455.6809 | 59600.4743 | 1188.0095 | 1088.0095 | 1183.2455
8630 | 96.4849 | 1.112 | 1.0322 | 1.0316 | 192.9371 | 54.1839 | 63704.4642 | 65199.0036 | 1299.9801 | 1199.9801 | 1183.2593
8631 | 98.137 | 1.032 | 1.052 | 1.0147 | 196.4943 | 45.5711 | 53456.321 | 53808.0991 | 1072.162 | 972.162 | 1183.3262
8632 | 100.8847 | 1.0453 | 1.0759 | 0.97955 | 182.3661 | 48.0449 | 55882.6852 | 57534.9574 | 1146.6991 | 1046.6991 | 1183.3878
8633 | 111.3008 | 1.0195 | 0.7316 | 0.92205 | 192.609 | 49.0612 | 38855.3193 | 39129.28 | 778.5856 | 678.5856 | 1183.4947
8634 | 100.2539 | 1.0093 | 1.2446 | 0.97324 | 198.5258 | 54.528 | 71765.9747 | 72243.4305 | 1440.8686 | 1340.8686 | 1183.5097
8635 | 110.3803 | 0.99343 | 0.88791 | 0.88545 | 195.305 | 52.3545 | 47097.2946 | 46968.0587 | 935.3612 | 835.3612 | 1183.5517
8636 | 102.0835 | 0.98385 | 0.85619 | 1.0347 | 211.1045 | 52.4154 | 51909.6251 | 51475.6398 | 1025.5128 | 925.5128 | 1183.5534
8637 | 98.0999 | 0.97836 | 1.101 | 1.0181 | 186.0315 | 51.194 | 62480.7384 | 63348.8905 | 1262.9778 | 1162.9778 | 1183.6829
8638 | 89.2028 | 0.95461 | 1.1609 | 1.0008 | 207.9608 | 49.8165 | 62325.4867 | 61038.9175 | 1216.7784 | 1116.7784 | 1183.7029
8639 | 81.641 | 0.96516 | 0.93816 | 1.0516 | 216.2941 | 52.2104 | 55837.3765 | 57073.7896 | 1137.4758 | 1037.4758 | 1183.7239
8640 | 103.5385 | 0.97404 | 1.0745 | 0.97987 | 205.4708 | 49.6137 | 57448.1391 | 57490.0408 | 1145.8008 | 1045.8008 | 1183.9738
8641 | 77.1827 | 0.93322 | 1.2929 | 1.0038 | 193.5485 | 50.0906 | 68813.917 | 69380.4981 | 1383.61 | 1283.61 | 1184.0895
8642 | 101.1454 | 1.0178 | 0.77914 | 0.9678 | 211.8276 | 45.1616 | 38915.3843 | 39257.8467 | 781.1569 | 681.1569 | 1184.2752
8643 | 91.5309 | 0.99933 | 0.85186 | 0.99881 | 204.1409 | 52.5638 | 49735.8585 | 50274.933 | 1001.4987 | 901.4987 | 1184.2847
8644 | 76.4291 | 1.0107 | 1.0328 | 0.95677 | 192.6965 | 48.6771 | 52054.4771 | 52479.3276 | 1045.5866 | 945.5866 | 1184.4164
8645 | 99.9847 | 1.0345 | 0.89697 | 1.0707 | 207.982 | 49.543 | 52914.0576 | 53686.2889 | 1069.7258 | 969.7258 | 1184.4266
8646 | 125.9096 | 1.1109 | 1.0342 | 1.0694 | 199.3649 | 48.9207 | 61146.0714 | 59815.5994 | 1192.312 | 1092.312 | 1184.5015
8647 | 90.2445 | 1.0222 | 0.91637 | 1.0313 | 181.5263 | 51.6053 | 53712.2214 | 53624.3811 | 1068.4876 | 968.4876 | 1184.5964
8648 | 82.4558 | 0.99484 | 0.92364 | 0.92973 | 193.2352 | 54.3885 | 51360.2851 | 49812.154 | 992.2431 | 892.2431 | 1184.8233
8649 | 105.5241 | 1.0347 | 0.81944 | 0.99037 | 204.9309 | 53.5473 | 49507.2884 | 50306.085 | 1002.1217 | 902.1217 | 1184.8394
8650 | 115.5282 | 0.97056 | 1.1719 | 1.0016 | 213.6205 | 48.8941 | 63088.7269 | 62894.5397 | 1253.8908 | 1153.8908 | 1184.8528
8651 | 110.0903 | 0.98849 | 1.0902 | 1.0172 | 192.0774 | 50.8196 | 62080.2456 | 62609.6629 | 1248.1933 | 1148.1933 | 1184.9522
8652 | 93.4129 | 1.0122 | 0.97786 | 0.98339 | 203.8226 | 48.2105 | 51122.7765 | 52864.2012 | 1053.284 | 953.284 | 1185.0884
8653 | 107.673 | 0.95203 | 0.99646 | 0.91761 | 195.9979 | 51.4392 | 52503.0713 | 54686.3568 | 1089.7271 | 989.7271 | 1185.2046
8654 | 100.8724 | 1.0367 | 1.2107 | 0.95107 | 208.5141 | 57.2449 | 72108.5074 | 73108.1306 | 1458.1626 | 1358.1626 | 1185.2224
8655 | 94.038 | 0.98855 | 1.0453 | 0.99215 | 198.1366 | 48.79 | 55331.1866 | 56573.0777 | 1127.4616 | 1027.4616 | 1185.293
8656 | 110.6785 | 1.0519 | 1.1186 | 1.1005 | 204.8848 | 48.929 | 66130.7753 | 68501.1154 | 1366.0223 | 1266.0223 | 1185.3083
8657 | 93.9896 | 0.99495 | 0.72146 | 0.99139 | 198.4114 | 51.0776 | 41508.0012 | 41282.304 | 821.6461 | 721.6461 | 1185.4496
8658 | 110.4866 | 1.1631 | 0.80636 | 0.98201 | 198.7552 | 47.9654 | 44344.3194 | 42666.2209 | 849.3244 | 749.3244 | 1185.4986
8659 | 104.7541 | 1.004 | 0.95058 | 0.9805 | 190.9427 | 52.628 | 54777.5361 | 55380.7251 | 1103.6145 | 1003.6145 | 1185.6448
8660 | 97.3117 | 1.0266 | 1.0677 | 1.0254 | 212.5949 | 52.1216 | 62483.6647 | 61530.3208 | 1226.6064 | 1126.6064 | 1185.658
8661 | 107.3387 | 1.085 | 0.7892 | 1.0559 | 185.5622 | 51.7637 | 49350.6238 | 49686.5125 | 989.7303 | 889.7303 | 1185.7424
8662 | 93.3403 | 1.0386 | 1.0114 | 0.95561 | 212.068 | 53.8125 | 57439.3436 | 56251.6874 | 1121.0337 | 1021.0337 | 1185.9068
8663 | 105.7575 | 0.9878 | 0.91172 | 0.99008 | 188.4809 | 52.5809 | 53144.9728 | 52810.4297 | 1052.2086 | 952.2086 | 1185.9755
8664 | 102.7743 | 0.98191 | 0.70905 | 0.96069 | 195.62 | 49.8892 | 39213.9218 | 39117.1422 | 778.3428 | 678.3428 | 1186.0469
8665 | 92.3528 | 0.9766 | 1.1591 | 1.0499 | 190.7943 | 53.2319 | 69775.8329 | 70698.3731 | 1409.9675 | 1309.9675 | 1186.1036
8666 | 117.6727 | 1.0195 | 0.82866 | 0.97087 | 204.7239 | 44.2084 | 41075.0237 | 40331.6006 | 802.632 | 702.632 | 1186.104
8667 | 99.7831 | 0.95038 | 1.0418 | 1.069 | 203.0592 | 53.7122 | 65112.1615 | 65280.1011 | 1301.602 | 1201.602 | 1186.2174
8668 | 109.1841 | 1.0421 | 0.81707 | 1.0302 | 206.1057 | 51.5558 | 49467.8369 | 49055.8007 | 977.116 | 877.116 | 1186.2535
8669 | 89.0857 | 1.0522 | 1.1764 | 1.054 | 193.6096 | 48.0801 | 64314.9804 | 64454.4204 | 1285.0884 | 1185.0884 | 1186.331
8670 | 98.5545 | 0.93584 | 1.1058 | 0.97364 | 202.1718 | 50.6596 | 59414.8359 | 60370.4538 | 1203.4091 | 1103.4091 | 1186.5565
8671 | 108.8979 | 0.95631 | 0.57623 | 1.0116 | 184.5366 | 48.8961 | 33779.7699 | 34287.8837 | 681.7577 | 581.7577 | 1186.6878
8672 | 95.1311 | 1.0166 | 1.2159 | 0.99757 | 209.5336 | 53.8133 | 70684.9176 | 69903.0227 | 1394.0605 | 1294.0605 | 1186.845
8673 | 97.83 | 0.9994 | 1.0791 | 0.98354 | 195.3113 | 51.4926 | 59881.09 | 58788.5884 | 1171.7718 | 1071.7718 | 1186.9392
8674 | 82.4182 | 0.96877 | 0.99026 | 1.0282 | 201.6545 | 51.2465 | 56470.1896 | 57276.2555 | 1141.5251 | 1041.5251 | 1186.9582
8675 | 107.0429 | 1.0237 | 0.8106 | 1.0413 | 182.5228 | 50.8321 | 48660.31 | 47879.6791 | 953.5936 | 853.5936 | 1186.969
8676 | 102.3523 | 1.0491 | 0.92043 | 0.96935 | 205.747 | 50.7741 | 50959.5491 | 50127.5587 | 998.5512 | 898.5512 | 1187.0393
8677 | 111.3806 | 1.0015 | 1.0703 | 0.95143 | 179.9909 | 54.7312 | 62019.7964 | 60306.8041 | 1202.1361 | 1102.1361 | 1187.0798
8678 | 99.8192 | 1.0159 | 1.1033 | 0.96166 | 208.2038 | 51.4032 | 59962.1632 | 58964.1302 | 1175.2826 | 1075.2826 | 1187.1313
8679 | 105.6129 | 0.93891 | 0.59216 | 1.0041 | 194.0267 | 50.3314 | 35112.5356 | 34867.8371 | 693.3567 | 593.3567 | 1187.179
8680 | 99.9273 | 1.0658 | 1.3458 | 1.0996 | 206.1776 | 49.0475 | 78013.5641 | 77126.2227 | 1538.5245 | 1438.5245 | 1187.2479
8681 | 90.019 | 0.96193 | 1.0783 | 0.89327 | 184.0238 | 48.6882 | 51298.3161 | 51914.7028 | 1034.2941 | 934.2941 | 1187.3618
8682 | 90.839 | 1.0236 | 1.1713 | 1.0369 | 189.9468 | 52.7721 | 69192.1061 | 68226.367 | 1360.5273 | 1260.5273 | 1187.407
8683 | 85.8823 | 1.1048 | 0.91713 | 0.98111 | 202.878 | 49.5248 | 49464.8292 | 48646.4627 | 968.9293 | 868.9293 | 1187.5163
8684 | 91.8156 | 1.0462 | 1.0595 | 0.90253 | 199.7306 | 52.0469 | 54966.128 | 54942.4016 | 1094.848 | 994.848 | 1187.5701
8685 | 113.0706 | 1.0484 | 1.1784 | 1.0041 | 192.5274 | 56.7791 | 74109.6114 | 72216.7968 | 1440.3359 | 1340.3359 | 1187.5845
8686 | 122.2523 | 1.0729 | 0.87774 | 0.93299 | 197.685 | 45.7822 | 43694.4354 | 43693.9267 | 869.8785 | 769.8785 | 1187.7893
8687 | 108.0798 | 0.94688 | 0.82334 | 0.96747 | 211.099 | 48.4999 | 43807.6726 | 43213.7335 | 860.2747 | 760.2747 | 1187.7987
8688 | 104.3158 | 1.0439 | 0.9008 | 1.0001 | 194.9559 | 50.4352 | 51124.1846 | 49909.497 | 994.1899 | 894.1899 | 1188.0656
8689 | 103.7635 | 0.94597 | 0.81104 | 0.98766 | 189.5785 | 56.4646 | 50962.0205 | 50208.5644 | 1000.1713 | 900.1713 | 1188.1111
8690 | 90.0171 | 1.0714 | 1.067 | 1.086 | 202.5585 | 50.7154 | 63865.3308 | 64461.1196 | 1285.2224 | 1185.2224 | 1188.1112
8691 | 98.0808 | 0.98552 | 1.1107 | 0.96979 | 188.4816 | 46.0127 | 54197.5953 | 54254.5352 | 1081.0907 | 981.0907 | 1188.2508
8692 | 84.6409 | 1.0301 | 1.1405 | 1.0408 | 207.9982 | 52.0408 | 66518.5091 | 66378.8688 | 1323.5774 | 1223.5774 | 1188.2852
8693 | 113.1092 | 1.0127 | 1.1295 | 0.98052 | 195.9334 | 51.7063 | 63382.0115 | 64041.615 | 1276.8323 | 1176.8323 | 1188.3489
8694 | 104.7366 | 0.87177 | 0.96181 | 1.0662 | 194.0966 | 49.0081 | 54924.8092 | 53443.0492 | 1064.861 | 964.861 | 1188.3626
8695 | 99.5902 | 0.99456 | 0.61668 | 0.98584 | 201.5868 | 53.0954 | 37739.9097 | 37334.418 | 742.6884 | 642.6884 | 1188.533
8696 | 108.5729 | 1.0684 | 1.0025 | 0.98777 | 192.7886 | 50.8457 | 56438.5688 | 57246.7225 | 1140.9344 | 1040.9344 | 1188.5879
8697 | 91.8326 | 1.041 | 1.147 | 1.0767 | 202.4344 | 50.6447 | 67589.2636 | 67676.4804 | 1349.5296 | 1249.5296 | 1188.6305
8698 | 91.7044 | 1.0202 | 0.98865 | 1.0033 | 195.1471 | 49.3379 | 53750.3995 | 52541.7337 | 1046.8347 | 946.8347 | 1188.6609
8699 | 105.1717 | 1.0115 | 1.1234 | 0.99988 | 183.5745 | 46.2357 | 57037.92 | 57445.4161 | 1144.9083 | 1044.9083 | 1188.6771
8700 | 95.3789 | 0.97205 | 0.95379 | 0.89734 | 198.6993 | 50.0012 | 47629.2385 | 47854.0472 | 953.0809 | 853.0809 | 1188.7582
8701 | 118.162 | 0.98814 | 0.89329 | 0.9608 | 190.6406 | 45.1022 | 44166.6478 | 43829.3964 | 872.5879 | 772.5879 | 1188.8035
8702 | 112.1569 | 0.96852 | 1.0218 | 0.97806 | 198.8634 | 49.2435 | 54761.8839 | 53911.1142 | 1074.2223 | 974.2223 | 1188.8881
8703 | 102.4855 | 1.0203 | 0.99525 | 1.0867 | 209.9817 | 55.3677 | 65884.0077 | 66540.7991 | 1326.816 | 1226.816 | 1188.8891
8704 | 85.005 | 0.96527 | 1.0669 | 1.0996 | 200.772 | 46.0017 | 57944.0309 | 57512.9616 | 1146.2592 | 1046.2592 | 1188.8999
8705 | 101.503 | 1.0356 | 0.9211 | 1.0676 | 180.5333 | 50.1354 | 54752.2942 | 56884.6016 | 1133.692 | 1033.692 | 1188.9166
8706 | 99.541 | 1.0066 | 0.81719 | 0.96769 | 197.8205 | 53.0122 | 47430.4237 | 48493.2315 | 965.8646 | 865.8646 | 1188.9383
8707 | 118.3137 | 0.98659 | 1.0221 | 0.97349 | 192.6492 | 48.4672 | 54072.9309 | 55113.5791 | 1098.2716 | 998.2716 | 1188.9984
8708 | 104.2029 | 1.0315 | 1.1029 | 1.0383 | 204.6195 | 44.0231 | 55351.1872 | 56274.7536 | 1121.4951 | 1021.4951 | 1189.0961
8709 | 84.5544 | 1.0646 | 1.0635 | 1.0348 | 189.6817 | 50.5799 | 60407.3566 | 59830.0744 | 1192.6015 | 1092.6015 | 1189.2166
8710 | 79.259 | 0.94768 | 0.9181 | 0.94595 | 200.1564 | 52.3159 | 49565.0999 | 48041.8072 | 956.8361 | 856.8361 | 1189.3143
8711 | 108.9821 | 1.0284 | 1.1375 | 0.98523 | 198.5933 | 49.6089 | 61356.3647 | 63427.0223 | 1264.5404 | 1164.5404 | 1189.4131
8712 | 105.2556 | 0.96138 | 1.0574 | 1.0341 | 192.7574 | 50.536 | 60563.5813 | 60632.9236 | 1208.6585 | 1108.6585 | 1189.4553
8713 | 91.4822 | 1.0397 | 1.2301 | 0.98455 | 179.7579 | 47.5381 | 62274.5309 | 62683.7032 | 1249.6741 | 1149.6741 | 1189.4646
8714 | 98.9116 | 1.0498 | 0.87268 | 1.0403 | 198.7938 | 48.7729 | 49544.0659 | 50175.2283 | 999.5046 | 899.5046 | 1189.4839
8715 | 87.4422 | 0.98487 | 1.1413 | 0.89845 | 201.1851 | 51.0829 | 56979.0912 | 56084.3786 | 1117.6876 | 1017.6876 | 1189.5062
8716 | 88.1363 | 1.0054 | 0.93657 | 1.0822 | 202.4972 | 48.1091 | 53228.5359 | 53582.1828 | 1067.6437 | 967.6437 | 1189.5234
8717 | 115.3709 | 1.0535 | 0.95481 | 0.99069 | 205.0122 | 50.5056 | 54117.8248 | 51860.6331 | 1033.2127 | 933.2127 | 1189.5443
8718 | 99.7878 | 0.93748 | 1.2854 | 1.0285 | 216.7292 | 47.5649 | 67545.4553 | 66763.5762 | 1331.2715 | 1231.2715 | 1189.5595
8719 | 101.6671 | 1.0463 | 1.1926 | 1.0776 | 201.9717 | 53.7692 | 75018.0955 | 77967.7139 | 1555.3543 | 1455.3543 | 1189.769
8720 | 101.9658 | 1.0029 | 1.1217 | 0.96246 | 198.3552 | 48.9238 | 58020.6164 | 58843.2345 | 1172.8647 | 1072.8647 | 1189.8489
8721 | 107.092 | 0.98917 | 1.0064 | 0.96323 | 202.1278 | 43.3283 | 46794.3147 | 46875.1837 | 933.5037 | 833.5037 | 1189.8945
8722 | 99.7808 | 1.0246 | 0.87526 | 1.0118 | 210.2345 | 48.2726 | 47896.2348 | 48715.0269 | 970.3005 | 870.3005 | 1189.8961
8723 | 100.5695 | 1.0454 | 0.891 | 1.0335 | 190.6226 | 48.4468 | 49896.5428 | 48263.189 | 961.2638 | 861.2638 | 1189.9036
8724 | 111.1283 | 0.98478 | 1.0011 | 0.95851 | 202.1299 | 50.9205 | 54635.7429 | 53579.5119 | 1067.5902 | 967.5902 | 1190.0267
8725 | 96.6356 | 1.0422 | 1.3872 | 1.0772 | 203.8553 | 46.2582 | 73983.7145 | 72858.6336 | 1453.1727 | 1353.1727 | 1190.1464
8726 | 87.1704 | 1.0113 | 0.91511 | 1.0493 | 208.8661 | 47.5704 | 50082.6431 | 49704.5329 | 990.0907 | 890.0907 | 1190.2373
8727 | 107.147 | 1.0669 | 1.1756 | 1.0537 | 202.8256 | 47.5146 | 64494.7034 | 63822.7117 | 1272.4542 | 1172.4542 | 1190.2518
8728 | 91.0044 | 0.96571 | 0.77266 | 1.0225 | 191.0309 | 49.9557 | 44049.5722 | 43020.2753 | 856.4055 | 756.4055 | 1190.3154
8729 | 92.0301 | 1.0473 | 1.1354 | 0.99793 | 197.1805 | 52.3523 | 64562.4478 | 64633.8554 | 1288.6771 | 1188.6771 | 1190.46
8730 | 99.9918 | 1.0134 | 1.3255 | 0.94059 | 201.0208 | 45.9345 | 62124.9058 | 60717.0022 | 1210.34 | 1110.34 | 1190.4622
8731 | 89.8842 | 0.99682 | 0.91264 | 1.0378 | 209.0137 | 52.866 | 55019.6102 | 56344.0535 | 1122.8811 | 1022.8811 | 1190.7459
8732 | 116.2928 | 0.9229 | 0.86084 | 0.97726 | 189.8899 | 49.5499 | 47192.1898 | 48287.9307 | 961.7586 | 861.7586 | 1190.8764
8733 | 91.211 | 1.0011 | 0.92108 | 1.0274 | 200.2846 | 48.3674 | 50387.0983 | 50453.2181 | 1005.0644 | 905.0644 | 1190.9194
8734 | 100.8206 | 1.1084 | 0.81075 | 0.98477 | 198.9403 | 48.9614 | 44761.2318 | 43568.0472 | 867.3609 | 767.3609 | 1190.9798
8735 | 96.1392 | 1.0186 | 1.1632 | 0.99991 | 203.5483 | 51.3064 | 64902.477 | 65832.7842 | 1312.6557 | 1212.6557 | 1190.9872
8736 | 121.5173 | 1.0367 | 1.1299 | 1.1087 | 192.5231 | 50.2764 | 69511.6481 | 71748.0246 | 1430.9605 | 1330.9605 | 1191.0819
8737 | 106.5583 | 0.89586 | 0.95384 | 0.96021 | 197.1958 | 47.1799 | 47912.8176 | 48102.6693 | 958.0534 | 858.0534 | 1191.1948
8738 | 86.197 | 0.955 | 1.022 | 1.0019 | 198.3412 | 53.021 | 58855.0069 | 58141.0982 | 1158.822 | 1058.822 | 1191.5054
8739 | 103.751 | 1.0853 | 0.97948 | 0.93028 | 200.3758 | 53.4184 | 54889.5633 | 55245.2026 | 1100.9041 | 1000.9041 | 1191.5117
8740 | 106.546 | 0.95538 | 1.0395 | 0.94862 | 207.8705 | 53.344 | 58239.5355 | 58773.5704 | 1171.4714 | 1071.4714 | 1191.5377
8741 | 87.2525 | 0.9344 | 1.0329 | 1.021 | 205.1728 | 48.3793 | 55169.4426 | 53320.4492 | 1062.409 | 962.409 | 1191.611
8742 | 97.6812 | 1.0629 | 1.062 | 1.0142 | 187.14 | 45.9166 | 54408.6466 | 53324.4478 | 1062.489 | 962.489 | 1191.7191
8743 | 92.571 | 1.068 | 0.80872 | 0.97924 | 201.8257 | 45.978 | 41158.6053 | 39887.4896 | 793.7498 | 693.7498 | 1191.8563
8744 | 110.0529 | 0.94038 | 0.88781 | 1.0158 | 205.2143 | 49.6594 | 50127.6051 | 50234.5158 | 1000.6903 | 900.6903 | 1191.8953
8745 | 105.312 | 0.89514 | 1.082 | 0.98576 | 187.3852 | 53.1489 | 61886.6738 | 60881.3787 | 1213.6276 | 1113.6276 | 1191.9029
8746 | 100.5793 | 1.0955 | 0.92121 | 0.98958 | 183.5078 | 51.6529 | 52962.5636 | 54566.0839 | 1087.3217 | 987.3217 | 1191.9754
8747 | 94.6783 | 0.96831 | 1.1053 | 1.0015 | 176.9568 | 50.5761 | 60801.4243 | 61352.9506 | 1223.059 | 1123.059 | 1192.0444
8748 | 109.273 | 0.94756 | 0.73472 | 0.99269 | 177.6547 | 47.2499 | 39532.0029 | 40405.0765 | 804.1015 | 704.1015 | 1192.0513
8749 | 91.8099 | 0.93731 | 0.70813 | 1.0192 | 205.6265 | 49.9898 | 40585.1396 | 41462.4849 | 825.2497 | 725.2497 | 1192.075
8750 | 101.8859 | 0.97334 | 1.1932 | 1.014 | 212.8907 | 51.7276 | 67923.2256 | 69593.1365 | 1387.8627 | 1287.8627 | 1192.2543
8751 | 106.3573 | 0.99725 | 1.053 | 0.94019 | 210.8027 | 52.6866 | 57962.0478 | 57450.1759 | 1145.0035 | 1045.0035 | 1192.2685
8752 | 100.7606 | 0.92244 | 1.0983 | 1.0123 | 173.3833 | 51.6043 | 62346.1419 | 62073.6917 | 1237.4738 | 1137.4738 | 1192.3784
8753 | 101.8248 | 0.95208 | 1.1849 | 0.93009 | 223.4375 | 49.1878 | 59198.616 | 59307.3167 | 1182.1463 | 1082.1463 | 1192.3844
8754 | 88.843 | 0.93495 | 1.103 | 0.94984 | 201.4908 | 49.7324 | 56437.8988 | 56581.0447 | 1127.6209 | 1027.6209 | 1192.3913
8755 | 127.9294 | 1.005 | 1.1461 | 1.0642 | 200.412 | 53.0847 | 71767.3461 | 69851.5548 | 1393.0311 | 1293.0311 | 1192.5322
8756 | 105.2247 | 1.0802 | 1.0302 | 0.94903 | 205.0679 | 51.1685 | 56048.5285 | 54387.9977 | 1083.76 | 983.76 | 1192.551
8757 | 109.9123 | 1.0256 | 0.73219 | 1.0747 | 202.8671 | 53.3473 | 48195.7602 | 47021.4147 | 936.4283 | 836.4283 | 1192.6275
8758 | 106.7571 | 1.0001 | 1.0181 | 0.97132 | 198.807 | 48.5791 | 53424.1426 | 53795.3718 | 1071.9074 | 971.9074 | 1192.6347
8759 | 94.4054 | 1.0567 | 1.176 | 0.9331 | 216.4817 | 48.6962 | 58509.1812 | 59213.6337 | 1180.2727 | 1080.2727 | 1192.9661
8760 | 104.8755 | 0.88715 | 0.98765 | 0.9914 | 190.5586 | 47.7821 | 51422.6074 | 51314.7359 | 1022.2947 | 922.2947 | 1192.9763
8761 | 78.6599 | 1.0362 | 0.68951 | 1.0229 | 204.2604 | 48.8484 | 38637.6558 | 38360.3613 | 763.2072 | 663.2072 | 1193.0916
8762 | 123.7157 | 0.94724 | 0.8315 | 0.88859 | 190.9509 | 49.8182 | 42837.5606 | 42364.7805 | 843.2956 | 743.2956 | 1193.0925
8763 | 119.2823 | 0.91197 | 1.1142 | 0.97921 | 208.5941 | 49.3515 | 59422.4443 | 60707.2205 | 1210.1444 | 1110.1444 | 1193.1178
8764 | 88.0169 | 0.9605 | 0.90264 | 1.021 | 205.447 | 49.4375 | 49945.5214 | 50315.7248 | 1002.3145 | 902.3145 | 1193.1733
8765 | 122.0167 | 1.0387 | 1.2859 | 0.98571 | 204.6408 | 49.1811 | 68773.7696 | 68811.4675 | 1372.2294 | 1272.2294 | 1193.2016
8766 | 95.3539 | 1.0074 | 1.0068 | 0.98698 | 198.8641 | 52.6492 | 57574.9269 | 56832.2072 | 1132.6441 | 1032.6441 | 1193.2836
8767 | 113.0194 | 1.0084 | 1.0359 | 1.0371 | 196.1955 | 47.8373 | 57042.9333 | 56991.0141 | 1135.8203 | 1035.8203 | 1193.5234
8768 | 88.58 | 0.95498 | 0.7552 | 1.0063 | 198.7337 | 52.5665 | 44594.4179 | 45161.0455 | 899.2209 | 799.2209 | 1193.5336
8769 | 87.2389 | 1.0234 | 1.2148 | 1.0004 | 176.4818 | 47.9596 | 62742.8663 | 60750.0864 | 1211.0017 | 1111.0017 | 1193.5768
8770 | 103.4593 | 1.0071 | 0.83428 | 1.0223 | 200.066 | 49.7762 | 47839.3069 | 49048.5913 | 976.9718 | 876.9718 | 1193.5872
8771 | 100.8053 | 1.0203 | 1.1713 | 0.94838 | 206.8387 | 47.2895 | 57601.1003 | 59852.3716 | 1193.0474 | 1093.0474 | 1193.6001
8772 | 79.2108 | 1.0449 | 1.1114 | 1.0666 | 199.3046 | 49.1563 | 62538.9564 | 61091.8454 | 1217.8369 | 1117.8369 | 1193.6274
8773 | 85.3214 | 0.92656 | 1.2062 | 0.96542 | 184.2089 | 45.8126 | 57156.2008 | 56590.4584 | 1127.8092 | 1027.8092 | 1193.7036
8774 | 95.6696 | 1.0345 | 0.90158 | 1.0336 | 194.4544 | 52.2646 | 54071.4017 | 54490.7348 | 1085.8147 | 985.8147 | 1193.7164
8775 | 112.7328 | 1.0489 | 1.0197 | 1.0019 | 185.2707 | 48.7983 | 55807.2954 | 57575.1884 | 1147.5038 | 1047.5038 | 1193.9063
8776 | 107.422 | 0.94556 | 1.0488 | 0.96176 | 209.1659 | 51.0923 | 56933.899 | 54146.3711 | 1078.9274 | 978.9274 | 1194.3027
8777 | 100.0529 | 1.0221 | 1.0597 | 1.0336 | 220.6995 | 48.3494 | 58122.0591 | 59385.1047 | 1183.7021 | 1083.7021 | 1194.3108
8778 | 100.9487 | 1.0797 | 1.2893 | 0.96094 | 195.3772 | 45.3108 | 61271.2421 | 63314.6196 | 1262.2924 | 1162.2924 | 1194.3982
8779 | 93.4846 | 1.0379 | 1.098 | 0.91276 | 196.7035 | 50.1615 | 55335.8211 | 55372.318 | 1103.4464 | 1003.4464 | 1194.5148
8780 | 97.1865 | 1.0408 | 1.0824 | 1.0826 | 210.8641 | 52.4664 | 66995.7354 | 66681.7924 | 1329.6358 | 1229.6358 | 1194.5217
8781 | 89.8117 | 0.99557 | 0.92563 | 0.96726 | 209.8511 | 51.7442 | 51164.1618 | 51204.6695 | 1020.0934 | 920.0934 | 1194.5246
8782 | 104.6413 | 0.94486 | 0.97447 | 0.97246 | 206.0806 | 48.9835 | 51467.4986 | 50828.4081 | 1012.5682 | 912.5682 | 1194.5286
8783 | 92.3631 | 0.95426 | 0.87114 | 0.99691 | 214.6774 | 45.7109 | 43941.1507 | 44281.5137 | 881.6303 | 781.6303 | 1194.6161
8784 | 88.8385 | 1.0092 | 1.0515 | 0.94133 | 214.2596 | 50.8022 | 55053.8706 | 54944.164 | 1094.8833 | 994.8833 | 1194.7801
8785 | 73.4551 | 0.98357 | 0.9748 | 0.98588 | 192.1135 | 51.4544 | 53359.0654 | 52746.5631 | 1050.9313 | 950.9313 | 1194.7829
8786 | 96.567 | 0.92424 | 0.77273 | 1.0503 | 200.3465 | 48.9422 | 44289.8714 | 44426.7632 | 884.5353 | 784.5353 | 1194.8004
8787 | 93.7643 | 1.0396 | 0.86433 | 0.93093 | 213.9564 | 51.6472 | 46805.3045 | 46724.7688 | 930.4954 | 830.4954 | 1194.8281
8788 | 94.3427 | 1.0445 | 1.0715 | 1.0823 | 207.2891 | 46.3305 | 58498.6471 | 60291.9713 | 1201.8394 | 1101.8394 | 1195.0047
8789 | 100.0018 | 0.90282 | 1.2231 | 0.99808 | 207.9946 | 48.8836 | 64298.5537 | 63067.1192 | 1257.3424 | 1157.3424 | 1195.0563
8790 | 103.6023 | 0.98448 | 0.93221 | 1.0607 | 203.0674 | 50.1664 | 54926.3244 | 54279.1758 | 1081.5835 | 981.5835 | 1195.2037
8791 | 105.7368 | 1.0481 | 1.1117 | 0.97895 | 209.7171 | 52.7361 | 63444.5712 | 64879.3617 | 1293.5872 | 1193.5872 | 1195.2686
8792 | 107.5673 | 1.0325 | 0.87534 | 1.0271 | 201.2149 | 47.3924 | 48073.5177 | 47775.7449 | 951.5149 | 851.5149 | 1195.3655
8793 | 107.1047 | 0.97515 | 1.119 | 1.0517 | 193.3287 | 44.5646 | 57295.1362 | 56371.9515 | 1123.439 | 1023.439 | 1195.4843
8794 | 99.3794 | 0.98756 | 0.89188 | 0.9559 | 206.4076 | 45.9899 | 43928.6551 | 44809.0114 | 892.1802 | 792.1802 | 1195.5739
8795 | 115.477 | 0.88535 | 1.2161 | 1.0351 | 204.272 | 52.7851 | 72046.1038 | 72160.9664 | 1439.2193 | 1339.2193 | 1195.6936
8796 | 106.6089 | 1.0388 | 0.90745 | 0.98106 | 200.9738 | 50.5826 | 50834.5049 | 49255.9439 | 981.1189 | 881.1189 | 1195.7744
8797 | 104.663 | 0.95234 | 0.96169 | 0.98239 | 193.0714 | 51.5927 | 54078.0237 | 53680.7607 | 1069.6152 | 969.6152 | 1195.7895
8798 | 103.2819 | 0.97651 | 1.2405 | 1.0258 | 189.6764 | 50.8202 | 69982.7224 | 69825.0001 | 1392.5 | 1292.5 | 1195.8697
8799 | 94.0658 | 1.0762 | 0.78114 | 0.94593 | 179.7965 | 47.8443 | 40375.9661 | 40649.8528 | 808.9971 | 708.9971 | 1195.8904
8800 | 106.0279 | 1.0424 | 0.972 | 1.0514 | 200.2289 | 49.2986 | 56030.7488 | 56104.4037 | 1118.0881 | 1018.0881 | 1195.9129
8801 | 85.0612 | 1.0151 | 1.0947 | 0.99636 | 198.2717 | 52.0975 | 61520.5829 | 62083.4388 | 1237.6688 | 1137.6688 | 1195.9599
8802 | 116.5332 | 1.0741 | 0.885 | 0.92506 | 197.7878 | 50.9266 | 48264.8555 | 47707.3717 | 950.1474 | 850.1474 | 1195.9674
8803 | 97.8959 | 0.93227 | 0.86474 | 1.0235 | 203.561 | 48.4714 | 47526.778 | 48123.9654 | 958.4793 | 858.4793 | 1196.1223
8804 | 100.3865 | 0.96912 | 0.92999 | 1.011 | 206.1073 | 45.3028 | 47209.8242 | 47079.1325 | 937.5826 | 837.5826 | 1196.1807
8805 | 119.7934 | 0.9506 | 1.1048 | 0.91513 | 191.2604 | 52.3573 | 59087.5562 | 59236.673 | 1180.7335 | 1080.7335 | 1196.2681
8806 | 110.4332 | 1.0787 | 0.99113 | 0.98079 | 205.4848 | 46.5986 | 51054.4156 | 48447.3272 | 964.9465 | 864.9465 | 1196.2829
8807 | 106.4092 | 1.0108 | 1.1705 | 1.0033 | 198.8974 | 44.1884 | 56843.9862 | 57869.0525 | 1153.381 | 1053.381 | 1196.4261
8808 | 94.9015 | 1.0225 | 1.1574 | 1.0076 | 214.5402 | 48.7406 | 61786.9369 | 62197.3593 | 1239.9472 | 1139.9472 | 1196.4532
8809 | 93.2661 | 0.97327 | 0.84458 | 1.0994 | 221.0045 | 50.2963 | 51488.825 | 50839.7974 | 1012.7959 | 912.7959 | 1196.8676
8810 | 93.9507 | 0.95342 | 0.85505 | 0.98695 | 193.3686 | 50.1766 | 47031.8397 | 47081.3756 | 937.6275 | 837.6275 | 1196.9005
8811 | 101.7365 | 0.93202 | 0.92595 | 1.0119 | 217.8865 | 55.7442 | 57733.874 | 59748.0926 | 1190.9619 | 1090.9619 | 1196.9097
8812 | 119.4688 | 0.91115 | 1.1186 | 0.94939 | 207.9786 | 52.7988 | 62029.472 | 61944.5125 | 1234.8902 | 1134.8902 | 1197.0833
8813 | 76.1059 | 1.0233 | 0.96355 | 1.1297 | 207.2967 | 48.721 | 57033.6262 | 57631.142 | 1148.6228 | 1048.6228 | 1197.0932
8814 | 106.3578 | 0.97652 | 1.2385 | 0.99857 | 205.2492 | 52.4294 | 70492.519 | 70626.457 | 1408.5291 | 1308.5291 | 1197.21
8815 | 106.0163 | 0.9527 | 1.0642 | 0.97807 | 193.3387 | 52.3349 | 59953.408 | 59429.6761 | 1184.5935 | 1084.5935 | 1197.2742
8816 | 66.7304 | 0.91759 | 1.019 | 0.98452 | 202.3935 | 50.416 | 53866.9842 | 54432.2779 | 1084.6456 | 984.6456 | 1197.287
8817 | 121.6034 | 0.95957 | 0.96168 | 1.0334 | 173.1005 | 50.9422 | 56745.7378 | 58556.9291 | 1167.1386 | 1067.1386 | 1197.4312
8818 | 85.2839 | 0.91541 | 0.90837 | 0.95845 | 197.9423 | 46.3504 | 44170.1317 | 44327.1714 | 882.5434 | 782.5434 | 1197.5093
8819 | 130.1285 | 1.077 | 0.87786 | 0.97392 | 170.3019 | 46.268 | 46212.5812 | 46621.0677 | 928.4214 | 828.4214 | 1197.513
8820 | 81.1038 | 1.0941 | 1.1926 | 0.98466 | 186.0221 | 51.203 | 64855.4338 | 64442.6381 | 1284.8528 | 1184.8528 | 1197.5409
8821 | 98.5163 | 0.98847 | 0.92775 | 0.99547 | 193.6506 | 55.0111 | 56356.3653 | 54884.4064 | 1093.6881 | 993.6881 | 1197.633
8822 | 87.5022 | 0.98158 | 0.81036 | 1.0629 | 213.3981 | 53.2924 | 50694.0981 | 48897.513 | 973.9503 | 873.9503 | 1197.725
8823 | 92.9451 | 0.98122 | 0.79756 | 0.99811 | 210.6819 | 53.0126 | 47246.5461 | 48864.5223 | 973.2904 | 873.2904 | 1197.74
8824 | 110.3257 | 0.9049 | 0.93843 | 1.001 | 205.6021 | 48.8334 | 50955.4198 | 51199.582 | 1019.9916 | 919.9916 | 1197.9834
8825 | 99.4721 | 1.0506 | 0.96306 | 0.99558 | 211.7099 | 50.3 | 53695.8866 | 54411.2896 | 1084.2258 | 984.2258 | 1198.0793
8826 | 111.0652 | 0.98973 | 0.87528 | 1.0756 | 194.12 | 47.4016 | 50032.8389 | 49852.1373 | 993.0427 | 893.0427 | 1198.133
8827 | 109.3602 | 1.0644 | 0.83493 | 0.99649 | 222.4453 | 51.5106 | 49075.1612 | 49295.1574 | 981.9031 | 881.9031 | 1198.2694
8828 | 87.7972 | 0.9777 | 1.0032 | 1.0019 | 207.6469 | 47.7157 | 52264.1011 | 51857.6571 | 1033.1531 | 933.1531 | 1198.4103
8829 | 101.3598 | 0.96114 | 0.93717 | 0.98674 | 181.021 | 51.3633 | 52682.4984 | 52969.7084 | 1055.3942 | 955.3942 | 1198.5032
8830 | 115.1327 | 0.96443 | 0.63391 | 0.9708 | 204.0417 | 46.1268 | 33712.0023 | 34010.7554 | 676.2151 | 576.2151 | 1198.6309
8831 | 98.8462 | 1.0634 | 0.93053 | 1.0058 | 211.1097 | 50.3275 | 52601.8159 | 50198.0016 | 999.96 | 899.96 | 1198.6393
8832 | 98.7882 | 0.99408 | 1.0349 | 1.0135 | 200.687 | 51.7387 | 59552.8829 | 61307.0077 | 1222.1402 | 1122.1402 | 1198.8576
8833 | 93.908 | 1.0667 | 0.6738 | 1.041 | 192.5348 | 51.9028 | 41797.1471 | 42982.5365 | 855.6507 | 755.6507 | 1198.8898
8834 | 85.0556 | 1.0162 | 0.93464 | 0.94734 | 205.9859 | 51.0486 | 49817.8848 | 50733.3293 | 1010.6666 | 910.6666 | 1198.9231
8835 | 103.4154 | 1.1289 | 1.0596 | 0.95452 | 208.7747 | 48.8043 | 55266.0006 | 55928.3051 | 1114.5661 | 1014.5661 | 1198.9712
8836 | 88.8672 | 1.1648 | 1.3089 | 0.99281 | 182.6321 | 49.9212 | 70219.7437 | 72446.5568 | 1444.9311 | 1344.9311 | 1199.0861
8837 | 111.4691 | 1.0315 | 1.1196 | 0.93801 | 206.0794 | 52.2925 | 61135.3139 | 60248.9527 | 1200.9791 | 1100.9791 | 1199.2046
8838 | 106.8982 | 1.0353 | 0.84744 | 1.0424 | 195.9779 | 52.0217 | 51908.7502 | 51422.7403 | 1024.4548 | 924.4548 | 1199.2188
8839 | 97.0304 | 0.99343 | 0.97145 | 0.9647 | 220.5679 | 56.0749 | 58176.7768 | 60135.0058 | 1198.7001 | 1098.7001 | 1199.2538
8840 | 107.6863 | 0.94666 | 1.0317 | 0.97475 | 188.2238 | 49.5117 | 55028.3578 | 56636.7735 | 1128.7355 | 1028.7355 | 1199.5864
8841 | 104.8211 | 0.98061 | 0.98422 | 1.1006 | 191.7685 | 50.8194 | 60465.7585 | 61243.8555 | 1220.8771 | 1120.8771 | 1199.7084
8842 | 89.0197 | 1.0687 | 0.98218 | 0.96836 | 183.7028 | 51.3039 | 53859.6612 | 53084.8391 | 1057.6968 | 957.6968 | 1199.7962
8843 | 88.5275 | 0.98439 | 1.1571 | 1.0238 | 203.8706 | 47.3101 | 60368.8573 | 59882.6966 | 1193.6539 | 1093.6539 | 1199.9508
8844 | 101.4704 | 0.99973 | 0.88006 | 1.057 | 217.6779 | 54.6637 | 56613.4037 | 56548.4893 | 1126.9698 | 1026.9698 | 1199.9801
8845 | 89.224 | 1.0296 | 0.95867 | 1.0203 | 182.3253 | 50.8101 | 54550.231 | 53202.4093 | 1060.0482 | 960.0482 | 1200.0571
8846 | 99.4569 | 1.0135 | 1.015 | 1.0072 | 201.2207 | 47.8083 | 53894.9508 | 53728.9089 | 1070.5782 | 970.5782 | 1200.2138
8847 | 94.3993 | 0.94578 | 0.91048 | 0.91103 | 215.5574 | 54.1616 | 49976.6871 | 49034.6655 | 976.6933 | 876.6933 | 1200.4141
8848 | 101.6001 | 1.075 | 1.152 | 1.0248 | 206.9495 | 49.5908 | 64169.8907 | 62254.2733 | 1241.0855 | 1141.0855 | 1200.5378
8849 | 87.4347 | 0.97094 | 0.89329 | 0.99488 | 181.3615 | 52.3808 | 51179.8038 | 51134.2733 | 1018.6855 | 918.6855 | 1200.6102
8850 | 90.908 | 0.96022 | 1.1505 | 1.0206 | 206.2976 | 55.5932 | 70336.8368 | 71012.4896 | 1416.2498 | 1316.2498 | 1200.6626
8851 | 93.329 | 0.9324 | 1.2688 | 0.8972 | 211.1285 | 52.156 | 64121.67 | 62906.1975 | 1254.124 | 1154.124 | 1200.9223
8852 | 116.3578 | 1.0841 | 0.97597 | 0.88507 | 212.9023 | 44.4081 | 44174.445 | 44410.9932 | 884.2199 | 784.2199 | 1200.9886
8853 | 97.4675 | 0.90326 | 1.2399 | 0.98908 | 186.6237 | 48.3286 | 63709.2162 | 61093.4483 | 1217.869 | 1117.869 | 1201.0317
8854 | 110.7963 | 1.077 | 0.72164 | 1.0063 | 188.7278 | 51.3224 | 43580.8457 | 43321.536 | 862.4307 | 762.4307 | 1201.1555
8855 | 82.2906 | 0.93401 | 0.88487 | 0.94323 | 190.6595 | 51.2703 | 46923.6432 | 46216.6192 | 920.3324 | 820.3324 | 1201.1684
8856 | 116.5746 | 1.0163 | 1.0874 | 1.0962 | 203.7092 | 51.1421 | 67220.3674 | 67663.0345 | 1349.2607 | 1249.2607 | 1201.602
8857 | 85.5279 | 0.92326 | 0.9698 | 0.96098 | 200.9544 | 45.417 | 46114.071 | 44778.7167 | 891.5743 | 791.5743 | 1201.6705
8858 | 95.7985 | 1.0031 | 1.0456 | 0.924 | 189.8442 | 48.0667 | 51249.6722 | 50706.9405 | 1010.1388 | 910.1388 | 1201.9164
8859 | 91.0823 | 0.94702 | 0.99512 | 1.109 | 194.8692 | 50.4802 | 60258.9013 | 61146.5424 | 1218.9308 | 1118.9308 | 1201.9728
8860 | 101.6315 | 1.07 | 1.1554 | 0.97994 | 218.8793 | 46.6149 | 58065.6442 | 60192.9209 | 1199.8584 | 1099.8584 | 1202.0071
8861 | 96.3884 | 1.0342 | 1.0964 | 1.0329 | 196.2953 | 51.4963 | 63642.8974 | 63557.231 | 1267.1446 | 1167.1446 | 1202.0239
8862 | 85.1248 | 1.0236 | 1.2023 | 0.97313 | 187.8674 | 47.9466 | 60461.6103 | 61623.8613 | 1228.4772 | 1128.4772 | 1202.1747
8863 | 82.9294 | 1.0138 | 0.95714 | 1.0091 | 187.1125 | 55.0476 | 57982.8644 | 58156.2529 | 1159.1251 | 1059.1251 | 1202.5703
8864 | 98.7098 | 1.1079 | 0.86518 | 0.95745 | 208.2549 | 47.6606 | 44900.8405 | 43228.7146 | 860.5743 | 760.5743 | 1202.6426
8865 | 87.6463 | 1.0193 | 1.0466 | 0.99731 | 186.355 | 47.5442 | 54059.6455 | 55036.6451 | 1096.7329 | 996.7329 | 1202.6991
8866 | 111.1886 | 0.95402 | 0.94481 | 1.0537 | 198.538 | 53.9136 | 59591.0298 | 59547.2901 | 1186.9458 | 1086.9458 | 1202.8196
8867 | 102.2096 | 1.0678 | 1.0729 | 1.0061 | 214.4109 | 47.9999 | 57264.3032 | 56289.4364 | 1121.7887 | 1021.7887 | 1202.8636
8868 | 101.2461 | 0.94789 | 1.1124 | 0.99215 | 194.2719 | 49.5391 | 59621.2001 | 60171.3014 | 1199.426 | 1099.426 | 1202.8872
8869 | 105.4096 | 1.0521 | 1.1511 | 0.93431 | 198.7426 | 52.8813 | 62935.1992 | 62705.345 | 1250.1069 | 1150.1069 | 1203.0006
8870 | 93.9207 | 0.97641 | 1.2488 | 0.98672 | 205.9542 | 47.8964 | 63617.6913 | 64644.454 | 1288.8891 | 1188.8891 | 1203.1727
8871 | 111.1291 | 0.98456 | 0.98836 | 1.0187 | 199.7478 | 53.7804 | 60230.0779 | 57866.2697 | 1153.3254 | 1053.3254 | 1203.2564
8872 | 98.8594 | 0.96969 | 1.0585 | 1.0513 | 189.4324 | 48.1315 | 58364.2108 | 59399.829 | 1183.9966 | 1083.9966 | 1203.4386
8873 | 104.4333 | 1.0191 | 1.1902 | 0.96307 | 209.3393 | 57.1474 | 71797.6051 | 73891.7828 | 1473.8357 | 1373.8357 | 1203.6075
8874 | 115.8563 | 1.0149 | 1.0056 | 1.0251 | 223.7503 | 49.0216 | 56524.5604 | 57744.6046 | 1150.8921 | 1050.8921 | 1203.6084
8875 | 104.7338 | 0.9914 | 1.1358 | 0.97081 | 190.6798 | 50.4527 | 61060.9261 | 60880.6982 | 1213.614 | 1113.614 | 1203.756
8876 | 110.9266 | 0.98434 | 0.84182 | 1.0157 | 199.0373 | 49.0163 | 47460.1759 | 50615.2455 | 1008.3049 | 908.3049 | 1203.8532
8877 | 87.6174 | 0.97756 | 0.90467 | 1.0016 | 198.3812 | 49.6746 | 49466.0757 | 49282.8493 | 981.657 | 881.657 | 1203.8729
8878 | 126.2533 | 1.0092 | 0.802 | 0.88826 | 207.0878 | 50.6068 | 42706.7946 | 43661.0049 | 869.2201 | 769.2201 | 1203.9116
8879 | 86.9463 | 0.88619 | 0.81562 | 0.94147 | 199.4392 | 52.7821 | 44796.5101 | 44176.7518 | 879.535 | 779.535 | 1203.9509
8880 | 82.3434 | 1.0054 | 0.70409 | 0.99776 | 198.1102 | 51.442 | 40595.7417 | 41231.8965 | 820.6379 | 720.6379 | 1204.4282
8881 | 97.1962 | 0.96849 | 0.95776 | 1.017 | 198.2674 | 51.1466 | 54831.9362 | 55664.1084 | 1109.2822 | 1009.2822 | 1204.4622
8882 | 119.0952 | 0.94695 | 0.8618 | 0.96659 | 203.6803 | 48.1974 | 45787.7698 | 46082.2207 | 917.6444 | 817.6444 | 1204.5104
8883 | 97.3513 | 0.90069 | 1.118 | 0.90468 | 204.2472 | 51.7824 | 57119.8964 | 59748.0567 | 1190.9611 | 1090.9611 | 1204.6145
8884 | 110.2445 | 0.99241 | 0.94981 | 1.0615 | 196.2101 | 49.6021 | 55634.2325 | 55436.7355 | 1104.7347 | 1004.7347 | 1204.6506
8885 | 110.916 | 0.98753 | 0.7032 | 1.0348 | 213.7577 | 50.613 | 42587.2392 | 42618.29 | 848.3658 | 748.3658 | 1204.6583
8886 | 109.7661 | 0.98105 | 1.0802 | 1.038 | 188.5132 | 52.7138 | 64973.1058 | 63328.526 | 1262.5705 | 1162.5705 | 1204.9841
8887 | 106.4605 | 1.0354 | 0.82598 | 0.97903 | 200.3395 | 50.6614 | 46752.6262 | 45382.2357 | 903.6447 | 803.6447 | 1205.0718
8888 | 117.7573 | 0.94219 | 0.91469 | 0.97351 | 209.0359 | 54.0408 | 54325.9907 | 55654.1836 | 1109.0837 | 1009.0837 | 1205.1508
8889 | 108.9094 | 0.97247 | 1.2565 | 1.049 | 199.0216 | 46.7593 | 66784.6123 | 66056.789 | 1317.1358 | 1217.1358 | 1205.1551
8890 | 72.2841 | 1.0455 | 0.93975 | 1.0769 | 199.9087 | 54.7263 | 59720.399 | 59058.8729 | 1177.1775 | 1077.1775 | 1205.1867
8891 | 107.5293 | 1.0119 | 1.061 | 1.1042 | 201.6069 | 51.6215 | 66293.6669 | 66101.7699 | 1318.0354 | 1218.0354 | 1205.1868
8892 | 100.338 | 1.0588 | 1.1724 | 0.98663 | 184.3814 | 50.6268 | 64124.9366 | 64250.4341 | 1281.0087 | 1181.0087 | 1205.1998
8893 | 87.017 | 0.94599 | 0.83614 | 0.98869 | 198.9998 | 47.1291 | 43039.0615 | 42225.518 | 840.5104 | 740.5104 | 1205.2024
8894 | 93.7975 | 1.0284 | 0.81464 | 0.99362 | 228.5641 | 51.7103 | 47073.2063 | 46176.9516 | 919.539 | 819.539 | 1205.2348
8895 | 97.8686 | 0.93951 | 0.93096 | 1.0178 | 198.0413 | 46.9957 | 49048.1367 | 47428.1231 | 944.5625 | 844.5625 | 1205.2787
8896 | 101.8964 | 0.96924 | 0.82071 | 0.99859 | 195.8455 | 53.7783 | 49580.8944 | 49505.7916 | 986.1158 | 886.1158 | 1205.2915
8897 | 98.095 | 1.0437 | 0.89714 | 1.0166 | 187.0488 | 50.1558 | 51065.1419 | 51433.482 | 1024.6696 | 924.6696 | 1205.3685
8898 | 99.4056 | 0.93252 | 0.94515 | 1.0603 | 182.6216 | 47.0082 | 51649.1439 | 51401.5593 | 1024.0312 | 924.0312 | 1205.4601
8899 | 96.9529 | 1.0588 | 0.8594 | 1.022 | 222.316 | 50.2684 | 49531.888 | 46649.4092 | 928.9882 | 828.9882 | 1205.5105
8900 | 94.4243 | 0.97282 | 1.0134 | 1.0126 | 195.2038 | 54.3641 | 60975.3627 | 60956.5814 | 1215.1316 | 1115.1316 | 1205.7501
8901 | 92.5653 | 1.0017 | 1.0668 | 0.93714 | 194.1721 | 47.8074 | 52423.6773 | 52876.7854 | 1053.5357 | 953.5357 | 1205.829
8902 | 96.1034 | 1.0041 | 1.0183 | 0.99352 | 219.0027 | 55.3273 | 61534.1976 | 61906.2246 | 1234.1245 | 1134.1245 | 1206.0616
8903 | 84.8495 | 0.93178 | 0.99678 | 0.97158 | 208.4268 | 52.4412 | 55140.8609 | 56003.8519 | 1116.077 | 1016.077 | 1206.2527
8904 | 101.453 | 1.0275 | 1.3862 | 0.93611 | 214.517 | 51.0833 | 71828.4833 | 70845.069 | 1412.9014 | 1312.9014 | 1206.2566
8905 | 101.1274 | 1.0115 | 1.0576 | 0.98614 | 213.8114 | 52.5302 | 60371.3558 | 61023.5154 | 1216.4703 | 1116.4703 | 1206.271
8906 | 101.6792 | 0.92176 | 1.1401 | 0.95154 | 185.8737 | 53.3975 | 63120.9495 | 62741.7793 | 1250.8356 | 1150.8356 | 1206.4123
8907 | 91.1855 | 1.0794 | 0.87279 | 1.0019 | 207.3808 | 45.6038 | 44575.5605 | 46075.3161 | 917.5063 | 817.5063 | 1206.4156
8908 | 116.1751 | 1.0779 | 1.2707 | 0.99471 | 208.2464 | 52.4773 | 73111.2801 | 71313.545 | 1422.2709 | 1322.2709 | 1206.4695
8909 | 85.3514 | 0.94709 | 1.4498 | 0.99321 | 204.933 | 50.5558 | 77087.1676 | 75975.9609 | 1515.5192 | 1415.5192 | 1206.5644
8910 | 108.6686 | 0.9513 | 0.88467 | 1.0201 | 196.6259 | 50.9925 | 51485.8478 | 50790.0734 | 1011.8015 | 911.8015 | 1206.617
8911 | 91.6261 | 1.0662 | 0.64776 | 0.98803 | 203.215 | 49.5543 | 36759.4392 | 36634.5379 | 728.6908 | 628.6908 | 1206.7784
8912 | 113.7484 | 1.0328 | 1.0816 | 1.0128 | 183.5734 | 53.6566 | 65265.5647 | 64510.8708 | 1286.2174 | 1186.2174 | 1206.8478
8913 | 86.8825 | 0.94602 | 0.96497 | 1.0011 | 198.0198 | 55.0411 | 57892.1714 | 57404.3235 | 1144.0865 | 1044.0865 | 1206.9267
8914 | 105.3627 | 1.0297 | 1.0916 | 1.0821 | 198.8487 | 48.3548 | 62564.2257 | 62026.5532 | 1236.5311 | 1136.5311 | 1206.9416
8915 | 96.6516 | 1.0034 | 0.89734 | 1.0526 | 207.2533 | 51.6689 | 54023.0823 | 52929.9287 | 1054.5986 | 954.5986 | 1207.1613
8916 | 96.5091 | 0.97781 | 1.3259 | 1.0025 | 178.8011 | 55.7752 | 79583.6955 | 77694.552 | 1549.891 | 1449.891 | 1207.222
8917 | 95.0158 | 1.0732 | 1.0708 | 0.98872 | 201.2322 | 52.69 | 61359.9238 | 62156.8919 | 1239.1378 | 1139.1378 | 1207.3486
8918 | 110.2909 | 1.1355 | 1.2038 | 0.95773 | 215.0389 | 52.2962 | 67057.5786 | 65657.1405 | 1309.1428 | 1209.1428 | 1207.4343
8919 | 96.0208 | 1.0683 | 0.80605 | 1.1126 | 197.1571 | 49.7751 | 49942.1824 | 49332.3887 | 982.6478 | 882.6478 | 1207.4863
8920 | 93.2181 | 1.034 | 1.066 | 0.96574 | 199.594 | 52.0998 | 58855.7194 | 59321.9104 | 1182.4382 | 1082.4382 | 1207.521
8921 | 94.1538 | 1.0687 | 1.1075 | 1.04 | 199.2704 | 44.9237 | 56462.1445 | 55218.0514 | 1100.361 | 1000.361 | 1207.6057
8922 | 129.108 | 0.97047 | 1.0629 | 0.97727 | 177.5548 | 50.7386 | 59240.0191 | 58455.9195 | 1165.1184 | 1065.1184 | 1207.6219
8923 | 85.0116 | 1.0956 | 1.3068 | 0.9273 | 195.3471 | 50.5519 | 66163.894 | 66245.0142 | 1320.9003 | 1220.9003 | 1207.6871
8924 | 105.8288 | 1.0179 | 0.91109 | 0.9261 | 210.7996 | 49.658 | 47459.7528 | 46428.0489 | 924.561 | 824.561 | 1207.712
8925 | 107.5805 | 1.0517 | 0.70126 | 0.96893 | 212.9203 | 53.1831 | 42366.4956 | 41121.7741 | 818.4355 | 718.4355 | 1207.8369
8926 | 109.2648 | 1.0138 | 0.93392 | 0.96607 | 207.3019 | 46.9805 | 47798.5198 | 47554.4707 | 947.0894 | 847.0894 | 1207.8375
8927 | 106.8362 | 1.0052 | 1.1163 | 1.0407 | 219.2129 | 50.1005 | 63802.6254 | 66494.1239 | 1325.8825 | 1225.8825 | 1207.8459
8928 | 93.4033 | 1.0655 | 0.94281 | 0.93265 | 214.5738 | 52.4525 | 51557.041 | 51717.6594 | 1030.3532 | 930.3532 | 1208.1158
8929 | 98.4578 | 1.0578 | 0.95263 | 0.95514 | 197.5373 | 49.1201 | 50007.3383 | 49110.5874 | 978.2117 | 878.2117 | 1208.16
8930 | 92.1137 | 0.97026 | 0.71636 | 1.0125 | 191.8102 | 48.077 | 39361.0196 | 39191.772 | 779.8354 | 679.8354 | 1208.1839
8931 | 104.1081 | 0.98849 | 1.0665 | 1.0098 | 220.2644 | 48.5712 | 57525.1169 | 58886.3973 | 1173.7279 | 1073.7279 | 1208.1908
8932 | 86.8308 | 0.96353 | 1.1106 | 0.95303 | 200.4912 | 47.8771 | 54882.5797 | 54522.9046 | 1086.4581 | 986.4581 | 1208.1935
8933 | 98.2821 | 1.0578 | 0.9287 | 1.0041 | 207.3269 | 47.3204 | 49253.9015 | 49333.6038 | 982.6721 | 882.6721 | 1208.2189
8934 | 115.5304 | 0.95629 | 0.84451 | 1.0162 | 198.8154 | 47.7955 | 46495.5704 | 46802.6814 | 932.0536 | 832.0536 | 1208.2341
8935 | 92.3713 | 1.1179 | 0.86384 | 1.0776 | 205.7259 | 50.8567 | 52798.3766 | 53883.1571 | 1073.6631 | 973.6631 | 1208.3019
8936 | 99.6921 | 0.87474 | 0.89544 | 0.9673 | 189.5737 | 49.9034 | 47765.6597 | 46012.0034 | 916.2401 | 816.2401 | 1208.4871
8937 | 89.6815 | 0.97715 | 0.989 | 0.92993 | 200.9401 | 51.5155 | 52094.493 | 51627.7723 | 1028.5554 | 928.5554 | 1208.5339
8938 | 116.1193 | 0.99072 | 1.0206 | 0.99096 | 184.8906 | 48.7589 | 55108.4895 | 53766.3982 | 1071.328 | 971.328 | 1208.5952
8939 | 97.853 | 1.0503 | 1.0215 | 0.99721 | 208.9632 | 47.3693 | 53329.9358 | 53466.9031 | 1065.3381 | 965.3381 | 1208.6232
8940 | 98.277 | 0.92059 | 0.94484 | 1.003 | 198.45 | 48.9245 | 50989.8402 | 52367.061 | 1043.3412 | 943.3412 | 1208.7253
8941 | 95.6399 | 1.0142 | 1.0134 | 0.98216 | 204.3125 | 47.7635 | 52375.3704 | 52120.3134 | 1038.4063 | 938.4063 | 1208.7381
8942 | 100.1438 | 1.0549 | 1.0354 | 1.0283 | 188.2452 | 47.8997 | 56248.6959 | 55345.2935 | 1102.9059 | 1002.9059 | 1208.8614
8943 | 123.135 | 1.0214 | 0.79704 | 0.93598 | 202.9618 | 53.3685 | 46728.4598 | 45744.5904 | 910.8918 | 810.8918 | 1209.1428
8944 | 108.2629 | 0.99774 | 0.80029 | 0.93298 | 199.0387 | 47.7494 | 41009.1003 | 40422.2893 | 804.4458 | 704.4458 | 1209.1683
8945 | 80.7603 | 0.98096 | 0.93259 | 0.92908 | 191.795 | 50.8869 | 48313.9053 | 49219.002 | 980.38 | 880.38 | 1209.2215
8946 | 93.9114 | 1.0814 | 0.83064 | 1.0094 | 198.0765 | 50.2134 | 47399.4853 | 47371.0844 | 943.4217 | 843.4217 | 1209.6399
8947 | 108.5516 | 0.95467 | 0.8999 | 0.95999 | 191.5375 | 50.3159 | 48873.6015 | 48153.1544 | 959.0631 | 859.0631 | 1209.9026
8948 | 112.7946 | 1.0134 | 0.81558 | 1.0059 | 211.0193 | 52.2795 | 49075.5373 | 50168.5899 | 999.3718 | 899.3718 | 1209.9662
8949 | 106.2706 | 1.072 | 0.97829 | 1.0343 | 194.2862 | 50.1011 | 56595.4225 | 55003.0916 | 1096.0618 | 996.0618 | 1210.0282
8950 | 100.2522 | 1.0826 | 1.1432 | 0.91334 | 185.9404 | 50.9602 | 58924.1593 | 59876.402 | 1193.528 | 1093.528 | 1210.0387
8951 | 101.7881 | 0.92929 | 1.1813 | 1.0665 | 203.3567 | 49.1946 | 66834.5961 | 65764.2705 | 1311.2854 | 1211.2854 | 1210.0546
8952 | 116.02 | 0.9919 | 1.0753 | 1.0116 | 203.4422 | 45.234 | 54611.6336 | 51533.6503 | 1026.673 | 926.673 | 1210.0722
8953 | 88.7225 | 1.0183 | 1.0744 | 0.99704 | 179.7685 | 52.1881 | 60800.3677 | 60344.0427 | 1202.8809 | 1102.8809 | 1210.0945
8954 | 101.279 | 0.97048 | 0.77122 | 1.1023 | 182.7427 | 49.0612 | 46710.8108 | 46561.5004 | 927.23 | 827.23 | 1210.2478
8955 | 90.7335 | 0.9567 | 1.0781 | 1.0322 | 181.8745 | 49.8497 | 59980.7475 | 58596.0264 | 1167.9205 | 1067.9205 | 1210.7789
8956 | 100.115 | 0.91291 | 0.88811 | 0.93998 | 211.5996 | 53.4686 | 49734.6013 | 49734.0999 | 990.682 | 890.682 | 1210.8304
8957 | 104.4852 | 0.96724 | 1.002 | 0.94248 | 181.4929 | 51.4007 | 53917.6267 | 54764.277 | 1091.2855 | 991.2855 | 1210.9296
8958 | 99.1408 | 0.88676 | 0.81797 | 0.93844 | 203.0801 | 53.8265 | 46252.8762 | 46214.1342 | 920.2827 | 820.2827 | 1211.0115
8959 | 96.8379 | 0.96755 | 1.2153 | 0.96301 | 191.108 | 51.1747 | 64875.9722 | 63374.4783 | 1263.4896 | 1163.4896 | 1211.0955
8960 | 124.4484 | 0.98975 | 1.0267 | 1.0918 | 205.3698 | 50.5107 | 63042.7945 | 63014.0363 | 1256.2807 | 1156.2807 | 1211.1742
8961 | 100.7515 | 0.93299 | 0.97618 | 0.96713 | 192.5656 | 50.391 | 52502.8466 | 52543.0914 | 1046.8618 | 946.8618 | 1211.1771
8962 | 103.1552 | 0.99163 | 1.0738 | 1.0599 | 205.018 | 45.9205 | 57163.2465 | 57535.5169 | 1146.7103 | 1046.7103 | 1211.2489
8963 | 93.1986 | 1.057 | 1.2297 | 1.0276 | 211.2959 | 53.1371 | 72592.1282 | 75124.4261 | 1498.4885 | 1398.4885 | 1211.2854
8964 | 106.4811 | 1.0523 | 1.0594 | 0.95783 | 219.0544 | 53.4911 | 60493.3459 | 61810.2176 | 1232.2044 | 1132.2044 | 1211.4202
8965 | 119.2818 | 1.0259 | 1.1395 | 1.071 | 218.4488 | 49.5294 | 66727.3444 | 65380.377 | 1303.6075 | 1203.6075 | 1211.4556
8966 | 103.3476 | 1.0701 | 1.083 | 0.98765 | 175.3964 | 50.3354 | 59580.9336 | 61903.4761 | 1234.0695 | 1134.0695 | 1211.5184
8967 | 104.2597 | 0.90505 | 0.83989 | 0.97719 | 219.7274 | 48.2432 | 44366.757 | 43677.2406 | 869.5448 | 769.5448 | 1211.5973
8968 | 91.7138 | 1.0567 | 0.96654 | 0.91402 | 200.0251 | 47.8846 | 47144.1832 | 46886.1206 | 933.7224 | 833.7224 | 1211.6235
8969 | 104.2376 | 1.0101 | 1.0809 | 0.95202 | 194.1773 | 48.3335 | 55020.5881 | 55859.5688 | 1113.1914 | 1013.1914 | 1211.6487
8970 | 98.7243 | 0.99079 | 0.83819 | 1.0245 | 215.8818 | 50.8691 | 48873.3351 | 49289.7412 | 981.7948 | 881.7948 | 1212.0518
8971 | 118.5564 | 0.93473 | 0.89454 | 1.0003 | 214.9416 | 50.7494 | 51248.6519 | 52043.6843 | 1036.8737 | 936.8737 | 1212.0652
8972 | 104.8767 | 0.96099 | 1.1167 | 1.0703 | 192.0215 | 51.4442 | 66862.2222 | 69434.0518 | 1384.681 | 1284.681 | 1212.0759
8973 | 97.12 | 0.94909 | 1.0588 | 0.87878 | 214.0392 | 50.2646 | 51614.42 | 51667.2261 | 1029.3445 | 929.3445 | 1212.2512
8974 | 84.4636 | 1.0028 | 1.1469 | 0.94885 | 200.7294 | 49.2752 | 57995.6659 | 57864.6462 | 1153.2929 | 1053.2929 | 1212.3511
8975 | 116.8923 | 0.93876 | 1.042 | 0.99888 | 202.6223 | 54.9787 | 63458.9947 | 65045.4845 | 1296.9097 | 1196.9097 | 1212.3517
8976 | 98.3225 | 0.98033 | 1.0283 | 0.98849 | 192.4279 | 51.1958 | 57165.3411 | 55927.6947 | 1114.5539 | 1014.5539 | 1212.6557
8977 | 107.0686 | 0.94953 | 1.016 | 0.97656 | 201.4605 | 51.7688 | 56826.5236 | 56664.2774 | 1129.2855 | 1029.2855 | 1212.7195
8978 | 78.9371 | 1.0633 | 1.0712 | 0.93875 | 213.7852 | 51.8128 | 56663.3401 | 56435.3838 | 1124.7077 | 1024.7077 | 1212.7948
8979 | 104.3541 | 1.0106 | 1.0395 | 0.93622 | 214.9289 | 52.3379 | 56669.0508 | 57504.6328 | 1146.0927 | 1046.0927 | 1212.8195
8980 | 98.3249 | 0.99211 | 0.63387 | 0.94546 | 204.6184 | 51.2042 | 35886.2705 | 35538.1188 | 706.7624 | 606.7624 | 1212.8618
8981 | 102.7476 | 0.97741 | 0.97033 | 0.88158 | 202.3767 | 48.5877 | 46644.8259 | 46107.6552 | 918.1531 | 818.1531 | 1213.0086
8982 | 81.8321 | 0.9853 | 1.0158 | 1.1005 | 211.5232 | 49.1136 | 59076.1296 | 58966.9741 | 1175.3395 | 1075.3395 | 1213.0712
8983 | 100.321 | 0.94575 | 1.0281 | 1.0445 | 197.1509 | 49.025 | 57494.9227 | 54876.7728 | 1093.5355 | 993.5355 | 1213.0861
8984 | 98.9918 | 1.031 | 1.0383 | 1.0064 | 176.0606 | 48.9927 | 56372.8851 | 54399.1957 | 1083.9839 | 983.9839 | 1213.1579
8985 | 92.2566 | 0.96485 | 0.99379 | 1.0146 | 196.2102 | 51.973 | 57225.7484 | 56177.3483 | 1119.547 | 1019.547 | 1213.178
8986 | 89.0679 | 0.96575 | 0.89636 | 1.0102 | 211.336 | 49.203 | 48996.2272 | 47739.1114 | 950.7822 | 850.7822 | 1213.2357
8987 | 85.625 | 0.95489 | 0.98901 | 1.004 | 199.9364 | 47.8886 | 51665.8588 | 51666.5646 | 1029.3313 | 929.3313 | 1213.7038
8988 | 102.2333 | 0.95999 | 0.96642 | 0.99945 | 201.465 | 50.2174 | 53633.9858 | 53177.6331 | 1059.5527 | 959.5527 | 1213.7615
8989 | 100.0935 | 0.87662 | 1.198 | 1.0078 | 196.8339 | 51.8231 | 67313.8445 | 66531.5672 | 1326.6313 | 1226.6313 | 1213.7939
8990 | 112.1061 | 1.0391 | 1.0174 | 1.0464 | 205.8304 | 54.2154 | 64235.5133 | 65461.7381 | 1305.2348 | 1205.2348 | 1213.9395
8991 | 102.4036 | 0.94953 | 1.1693 | 1.0024 | 193.0575 | 49.4728 | 62994.7763 | 63682.3111 | 1269.6462 | 1169.6462 | 1214.0604
8992 | 102.5688 | 0.95051 | 0.91693 | 0.9508 | 193.2708 | 49.4443 | 48119.9476 | 46477.5218 | 925.5504 | 825.5504 | 1214.0624
8993 | 118.0216 | 0.97234 | 0.73342 | 0.97258 | 207.4808 | 47.5501 | 39582.2733 | 41020.1699 | 816.4034 | 716.4034 | 1214.1909
8994 | 89.9091 | 1.0815 | 1.0116 | 1.0608 | 199.6862 | 52.2144 | 61306.9854 | 61551.028 | 1227.0206 | 1127.0206 | 1214.2331
8995 | 78.1374 | 0.95188 | 1.1111 | 1.0242 | 208.1036 | 51.3176 | 62424.822 | 63245.5595 | 1260.9112 | 1160.9112 | 1214.2469
8996 | 107.8907 | 1.0056 | 0.88946 | 0.9935 | 190.2457 | 47.8765 | 47692.1336 | 49222.4229 | 980.4485 | 880.4485 | 1214.3044
8997 | 90.4682 | 1.0495 | 1.1025 | 1.0063 | 198.6713 | 48.0587 | 58079.3761 | 56020.2846 | 1116.4057 | 1016.4057 | 1214.334
8998 | 98.3837 | 1.0457 | 1.3146 | 0.97225 | 203.4976 | 53.9782 | 74747.8375 | 77046.1567 | 1536.9231 | 1436.9231 | 1214.5764
8999 | 90.1849 | 1.0462 | 1.0577 | 1.0023 | 180.6308 | 50.9783 | 59034.431 | 56135.3015 | 1118.706 | 1018.706 | 1214.7775
9000 | 105.9649 | 1.0347 | 1.1997 | 0.93893 | 190.3377 | 46.9778 | 58256.5056 | 58037.3424 | 1156.7468 | 1056.7468 | 1214.8619
9001 | 95.2466 | 0.98945 | 0.98357 | 0.98177 | 199.5236 | 50.2189 | 53425.5602 | 51841.9419 | 1032.8388 | 932.8388 | 1214.871
9002 | 105.863 | 0.98297 | 1.0029 | 1.0011 | 203.2947 | 49.0131 | 54513.9717 | 54209.704 | 1080.1941 | 980.1941 | 1214.898
9003 | 94.5312 | 0.9842 | 1.0815 | 1.0948 | 200.0008 | 52.1739 | 66831.7941 | 64885.8176 | 1293.7164 | 1193.7164 | 1215.0359
9004 | 105.949 | 0.94749 | 1.255 | 0.97329 | 202.0855 | 48.3802 | 64152.1287 | 63166.8206 | 1259.3364 | 1159.3364 | 1215.5768
9005 | 106.6244 | 1.0218 | 1.0755 | 1.0284 | 202.8136 | 48.6526 | 59311.5182 | 59699.3953 | 1189.9879 | 1089.9879 | 1215.7554
9006 | 92.2039 | 1.0337 | 1.1722 | 1.0164 | 189.0076 | 52.6919 | 67988.8304 | 69839.5034 | 1392.7901 | 1292.7901 | 1216.0634
9007 | 104.9294 | 1.0111 | 1.1267 | 0.92122 | 204.0566 | 48.3374 | 55503.4167 | 56080.646 | 1117.6129 | 1017.6129 | 1216.1209
9008 | 92.676 | 0.97582 | 1.106 | 1.0153 | 212.468 | 48.6328 | 59218.2121 | 57860.6351 | 1153.2127 | 1053.2127 | 1216.2082
9009 | 95.4916 | 0.93307 | 1.0006 | 0.96358 | 180.3242 | 47.9714 | 50708.2035 | 49339.2498 | 982.785 | 882.785 | 1216.219
9010 | 85.4999 | 0.94754 | 1.0666 | 0.98822 | 205.7237 | 53.0172 | 60383.2403 | 58114.8097 | 1158.2962 | 1058.2962 | 1216.2431
9011 | 102.3642 | 1.1365 | 0.87171 | 0.86975 | 190.6967 | 51.5045 | 45231.7794 | 45796.0129 | 911.9203 | 811.9203 | 1216.3366
9012 | 117.3395 | 1.0339 | 1.1798 | 1.006 | 220.328 | 50.0348 | 65676.0016 | 66636.1265 | 1328.7225 | 1228.7225 | 1216.3984
9013 | 110.2553 | 1.1197 | 0.92651 | 0.97584 | 195.8903 | 48.8199 | 50362.4983 | 50427.0802 | 1004.5416 | 904.5416 | 1216.474
9014 | 83.5715 | 1.0095 | 1.2264 | 0.92389 | 189.2579 | 45.9555 | 56136.8521 | 56705.748 | 1130.115 | 1030.115 | 1216.5732
9015 | 113.7139 | 1.0122 | 1.0832 | 1.0192 | 203.5409 | 51.272 | 62709.5573 | 63163.5323 | 1259.2706 | 1159.2706 | 1216.5745
9016 | 94.3717 | 0.98792 | 1.1918 | 1.0208 | 197.615 | 50.2021 | 65953.1512 | 66010.4109 | 1316.2082 | 1216.2082 | 1216.6539
9017 | 90.2768 | 0.99645 | 1.0031 | 0.94263 | 209.0249 | 49.8958 | 51877.6772 | 51902.9848 | 1034.0597 | 934.0597 | 1216.7485
9018 | 102.4933 | 0.97569 | 0.97762 | 1.0172 | 217.579 | 52.9306 | 58147.2009 | 58564.5946 | 1167.2919 | 1067.2919 | 1216.761
9019 | 107.6974 | 0.99317 | 1.0554 | 1.005 | 203.5525 | 50.5858 | 59268.425 | 57829.1453 | 1152.5829 | 1052.5829 | 1216.7961
9020 | 99.2692 | 1.0677 | 0.93336 | 1.0311 | 207.3556 | 55.2868 | 59272.3786 | 59041.581 | 1176.8316 | 1076.8316 | 1216.8288
9021 | 112.1271 | 1.0316 | 1.0049 | 0.97181 | 210.7958 | 47.4862 | 52076.4811 | 52901.3007 | 1054.026 | 954.026 | 1216.9889
9022 | 80.3001 | 0.94138 | 1.0504 | 1.059 | 198.4929 | 47.4685 | 56589.6483 | 56556.1471 | 1127.1229 | 1027.1229 | 1217.0966
9023 | 86.5349 | 0.98102 | 0.8476 | 1.0327 | 170.5432 | 47.7177 | 45988.9365 | 45176.5472 | 899.5309 | 799.5309 | 1217.1358
9024 | 94.0692 | 0.98381 | 1.1121 | 0.99673 | 216.059 | 48.8335 | 58865.9778 | 58560.4283 | 1167.2086 | 1067.2086 | 1217.1719
9025 | 96.3987 | 0.96542 | 0.82207 | 1.0427 | 206.8651 | 49.1224 | 46886.3584 | 45020.5575 | 896.4111 | 796.4111 | 1217.2579
9026 | 101.541 | 1.0641 | 1.0577 | 1.0754 | 207.7088 | 53.6056 | 66976.7335 | 68562.7557 | 1367.2551 | 1267.2551 | 1217.2679
9027 | 88.2567 | 1.0262 | 0.80769 | 1.0653 | 199.806 | 47.6709 | 45535.3008 | 44667.558 | 889.3512 | 789.3512 | 1217.4934
9028 | 91.9042 | 0.9949 | 1.0415 | 1.0485 | 187.5784 | 45.8563 | 54456.6214 | 55632.4118 | 1108.6482 | 1008.6482 | 1217.5574
9029 | 91.5667 | 0.99938 | 1.0138 | 1.0607 | 215.175 | 51.7538 | 60606.3638 | 58774.9396 | 1171.4988 | 1071.4988 | 1217.5972
9030 | 103.7744 | 0.99623 | 0.85421 | 0.97628 | 193.6033 | 49.2144 | 46323.9328 | 46036.4583 | 916.7292 | 816.7292 | 1217.6034
9031 | 105.4162 | 0.98105 | 0.94937 | 1.022 | 170.9281 | 45.8345 | 49381.8474 | 49002.2755 | 976.0455 | 876.0455 | 1217.6168
9032 | 95.2488 | 1.0828 | 0.7068 | 0.98563 | 223.396 | 46.9363 | 37761.9323 | 38350.7224 | 763.0144 | 663.0144 | 1217.8006
9033 | 99.8462 | 0.93005 | 0.92344 | 0.99833 | 197.264 | 51.6136 | 52572.4884 | 52752.644 | 1051.0529 | 951.0529 | 1217.8452
9034 | 87.8401 | 1.0444 | 1.0316 | 1.0374 | 210.4983 | 44.4763 | 51887.544 | 51750.0566 | 1031.0011 | 931.0011 | 1217.883
9035 | 104.8097 | 1.0192 | 1.0777 | 0.9758 | 214.0933 | 52.9376 | 61540.7198 | 60338.351 | 1202.767 | 1102.767 | 1218.0354
9036 | 102.7895 | 0.91728 | 1.121 | 1.1214 | 196.2978 | 47.648 | 64590.1255 | 64858.6641 | 1293.1733 | 1193.1733 | 1218.0683
9037 | 105.1446 | 0.97541 | 0.855 | 0.9918 | 200.1676 | 50.9777 | 48656.863 | 48508.8517 | 966.177 | 866.177 | 1218.1732
9038 | 106.7208 | 0.96253 | 0.96289 | 1.0278 | 200.7909 | 53.2119 | 58326.2892 | 59244.5546 | 1180.8911 | 1080.8911 | 1218.2816
9039 | 105.8446 | 1.0452 | 1.1026 | 1.1344 | 230.0007 | 49.3738 | 67449.4003 | 69201.4567 | 1380.0291 | 1280.0291 | 1218.3929
9040 | 99.8088 | 0.98634 | 0.93567 | 1.0467 | 194.7142 | 52.668 | 56963.0327 | 58246.1305 | 1160.9226 | 1060.9226 | 1218.6412
9041 | 108.9135 | 1.0435 | 0.92171 | 0.92138 | 214.5558 | 51.3153 | 49625.9514 | 49083.814 | 977.6763 | 877.6763 | 1218.8489
9042 | 110.8624 | 1.0067 | 0.94693 | 1.021 | 210.8195 | 56.9595 | 61637.4923 | 61750.2036 | 1231.0041 | 1131.0041 | 1218.8695
9043 | 100.8786 | 1.0224 | 1.1845 | 0.98737 | 222.6638 | 50.8795 | 64973.31 | 67188.1208 | 1339.7624 | 1239.7624 | 1218.9008
9044 | 109.6629 | 0.97248 | 0.96161 | 0.91172 | 196.5822 | 50.2167 | 49578.1007 | 51912.2177 | 1034.2444 | 934.2444 | 1219.0354
9045 | 99.9716 | 0.99268 | 1.2208 | 0.95727 | 194.3729 | 50.3038 | 63973.2823 | 64487.1208 | 1285.7424 | 1185.7424 | 1219.0808
9046 | 101.434 | 1.087 | 0.95074 | 1.0093 | 204.9578 | 51.2733 | 55056.7696 | 53448.1427 | 1064.9629 | 964.9629 | 1219.0904
9047 | 122.7061 | 1.0655 | 0.76022 | 0.95719 | 181.8555 | 51.987 | 44808.3214 | 43778.943 | 871.5789 | 771.5789 | 1219.2221
9048 | 93.3647 | 1.0398 | 1.1772 | 0.93588 | 187.1846 | 49.9255 | 60036.9505 | 59066.4997 | 1177.33 | 1077.33 | 1219.4511
9049 | 100.6233 | 0.98774 | 1.1574 | 0.97266 | 180.0315 | 47.7868 | 58727.3712 | 57320.313 | 1142.4063 | 1042.4063 | 1219.5277
9050 | 94.2056 | 0.89587 | 1.0796 | 1.0581 | 197.0135 | 51.8344 | 63780.1939 | 62321.1259 | 1242.4225 | 1142.4225 | 1219.55
9051 | 89.1575 | 0.93965 | 1.0887 | 1.0358 | 201.5686 | 46.2447 | 56226.102 | 55284.1129 | 1101.6823 | 1001.6823 | 1219.5514
9052 | 102.9014 | 1.0043 | 0.80306 | 1.0381 | 195.2255 | 51.6595 | 48600.2031 | 48092.5525 | 957.851 | 857.851 | 1219.7883
9053 | 124.0738 | 1.0446 | 1.3217 | 1.0879 | 226.3852 | 46.0733 | 72448.223 | 70620.7469 | 1408.4149 | 1308.4149 | 1219.7968
9054 | 84.4217 | 0.94048 | 1.0962 | 0.9413 | 207.625 | 52.2989 | 58323.4939 | 58228.7352 | 1160.5747 | 1060.5747 | 1220.3068
9055 | 103.7064 | 0.95524 | 0.83733 | 0.97221 | 200.5792 | 48.1622 | 44178.4274 | 43126.632 | 858.5326 | 758.5326 | 1220.5958
9056 | 108.7485 | 0.9729 | 0.83247 | 1.0395 | 196.613 | 51.9729 | 50671.1782 | 50956.4605 | 1015.1292 | 915.1292 | 1220.6693
9057 | 98.574 | 1.0044 | 1.0352 | 0.9084 | 197.3137 | 49.4484 | 51591.7525 | 52357.185 | 1043.1437 | 943.1437 | 1220.8286
9058 | 104.2512 | 0.98803 | 1.0489 | 1.0444 | 187.5379 | 52.6189 | 63251.9589 | 64095.5864 | 1277.9117 | 1177.9117 | 1220.9003
9059 | 92.1269 | 0.94012 | 0.87427 | 1.0353 | 180.7148 | 44.4617 | 44275.2277 | 43146.9854 | 858.9397 | 758.9397 | 1220.9313
9060 | 104.7863 | 0.99139 | 0.76625 | 1.0171 | 198.9757 | 45.8624 | 40705.3012 | 41316.1458 | 822.3229 | 722.3229 | 1221.1662
9061 | 96.8486 | 1.0707 | 0.97421 | 1.0214 | 188.2446 | 50.14 | 55281.5791 | 56214.1705 | 1120.2834 | 1020.2834 | 1221.2065
9062 | 99.6265 | 0.97773 | 0.95226 | 1.0329 | 209.25 | 51.9404 | 56357.2746 | 53933.2571 | 1074.6651 | 974.6651 | 1221.2649
9063 | 112.0109 | 0.89555 | 1.0779 | 0.98273 | 204.1488 | 53.0777 | 61750.4111 | 61887.9832 | 1233.7597 | 1133.7597 | 1221.3928
9064 | 95.4699 | 0.94708 | 0.9611 | 0.96661 | 181.6687 | 51.1135 | 52288.2991 | 51452.1324 | 1025.0426 | 925.0426 | 1221.5226
9065 | 93.5045 | 0.88008 | 1.1546 | 0.98427 | 215.3199 | 48.575 | 59415.7258 | 58402.5461 | 1164.0509 | 1064.0509 | 1221.685
9066 | 106.9479 | 0.98089 | 0.97026 | 0.98994 | 208.933 | 47.2272 | 50524.9918 | 50973.6572 | 1015.4731 | 915.4731 | 1221.8654
9067 | 88.8872 | 0.85412 | 0.88411 | 0.99712 | 199.2077 | 50.9823 | 49013.9743 | 50018.5361 | 996.3707 | 896.3707 | 1222.0956
9068 | 93.9957 | 0.93132 | 1.1967 | 1.0433 | 181.536 | 50.8967 | 68181.1861 | 69195.8121 | 1379.9162 | 1279.9162 | 1222.3155
9069 | 110.9022 | 1.0138 | 1.0195 | 0.97776 | 210.7544 | 46.1457 | 51397.7511 | 51743.3658 | 1030.8673 | 930.8673 | 1222.4463
9070 | 106.7473 | 0.94947 | 0.86076 | 0.92159 | 187.7701 | 52.5002 | 47155.4585 | 47067.9392 | 937.3588 | 837.3588 | 1222.5295
9071 | 116.1591 | 1.0838 | 0.82622 | 0.96948 | 202.6301 | 50.2701 | 46797.9341 | 46268.6931 | 921.3739 | 821.3739 | 1222.5683
9072 | 99.0791 | 0.92314 | 1.2492 | 1.0649 | 208.5229 | 50.8206 | 72463.1359 | 73187.5338 | 1459.7507 | 1359.7507 | 1222.5711
9073 | 97.4634 | 0.97371 | 0.86295 | 0.99486 | 210.5217 | 54.3698 | 52047.6394 | 52859.0714 | 1053.1814 | 953.1814 | 1222.5722
9074 | 93.84 | 0.95323 | 1.1326 | 0.92563 | 191.4484 | 50.5635 | 57722.4501 | 56227.927 | 1120.5585 | 1020.5585 | 1222.6977
9075 | 95.4853 | 1.0119 | 1.0119 | 0.97333 | 186.8756 | 52.364 | 56821.2067 | 56344.0401 | 1122.8808 | 1022.8808 | 1222.703
9076 | 84.2589 | 1.0299 | 0.85068 | 1.0958 | 211.1445 | 52.7185 | 53929.4686 | 55364.4013 | 1103.288 | 1003.288 | 1222.7858
9077 | 108.8159 | 0.9439 | 0.81001 | 1.0182 | 195.8697 | 48.9781 | 45619.7705 | 44276 | 881.52 | 781.52 | 1222.9012
9078 | 97.8098 | 0.93596 | 1.0525 | 0.98606 | 190.6516 | 52.1309 | 59066.6707 | 57275.2808 | 1141.5056 | 1041.5056 | 1223.0006
9079 | 112.0577 | 1.0461 | 1.0756 | 1.0836 | 196.6744 | 47.9661 | 61725.0816 | 61563.6182 | 1227.2724 | 1127.2724 | 1223.0459
9080 | 91.3498 | 0.97497 | 0.98435 | 1.0039 | 222.2427 | 50.317 | 54424.6378 | 55915.3257 | 1114.3065 | 1014.3065 | 1223.2238
9081 | 113.4046 | 1.0046 | 0.98666 | 1.0347 | 201.0855 | 50.2948 | 57274.0129 | 58005.0247 | 1156.1005 | 1056.1005 | 1223.2936
9082 | 104.9594 | 0.96516 | 0.9232 | 1.0129 | 209.5042 | 46.4175 | 48318.3272 | 49099.431 | 977.9886 | 877.9886 | 1223.4357
9083 | 92.3263 | 1.1288 | 0.98093 | 1.0055 | 193.4738 | 46.3129 | 50697.9068 | 51268.8177 | 1021.3764 | 921.3764 | 1223.4909
9084 | 104.0555 | 0.99616 | 1.2253 | 0.97421 | 198.24 | 45.3337 | 59012.5794 | 59369.2331 | 1183.3847 | 1083.3847 | 1223.5641
9085 | 113.9112 | 1.0433 | 0.76081 | 0.92921 | 184.7949 | 49.1674 | 40787.1142 | 40529.4461 | 806.5889 | 706.5889 | 1223.5774
9086 | 95.1771 | 0.93468 | 0.86264 | 0.99125 | 208.1896 | 48.4908 | 45986.3107 | 45644.1471 | 908.8829 | 808.8829 | 1223.5928
9087 | 93.9788 | 1.0184 | 1.0036 | 1.0243 | 209.3392 | 52.7614 | 59500.9746 | 58553.3464 | 1167.0669 | 1067.0669 | 1223.6268
9088 | 109.4145 | 1.0834 | 0.89184 | 1.0766 | 202.4096 | 50.8762 | 55080.0894 | 55577.8847 | 1107.5577 | 1007.5577 | 1223.6321
9089 | 90.722 | 1.0442 | 1.1174 | 0.99488 | 195.1674 | 46.5274 | 56324.9763 | 56470.6316 | 1125.4126 | 1025.4126 | 1223.682
9090 | 95.1805 | 0.99178 | 1.1896 | 1.0181 | 185.6376 | 51.3912 | 67274.0189 | 67959.4268 | 1355.1885 | 1255.1885 | 1223.8252
9091 | 111.475 | 1.0194 | 0.94942 | 1.0361 | 192.0495 | 48.0893 | 52961.8829 | 54597.2369 | 1087.9447 | 987.9447 | 1223.9263
9092 | 109.8962 | 0.98076 | 0.88733 | 0.92197 | 190.5386 | 48.6023 | 45190.1943 | 46616.725 | 928.3345 | 828.3345 | 1223.9392
9093 | 109.135 | 1.025 | 0.91039 | 0.93486 | 192.8838 | 46.088 | 44573.5086 | 45088.7433 | 897.7749 | 797.7749 | 1224.0428
9094 | 103.856 | 1.0177 | 1.1244 | 1.0771 | 184.0348 | 52.9258 | 69871.3651 | 71296.4426 | 1421.9289 | 1321.9289 | 1224.369
9095 | 98.2603 | 1.0474 | 1.1965 | 0.9327 | 184.8922 | 56.3958 | 68924.5102 | 69148.4831 | 1378.9697 | 1278.9697 | 1224.5211
9096 | 121.3432 | 0.99529 | 1.1203 | 0.99254 | 189.738 | 50.0871 | 61934.2815 | 61862.6287 | 1233.2526 | 1133.2526 | 1224.5698
9097 | 109.0906 | 1.0043 | 0.93616 | 1.0774 | 196.6696 | 49.8383 | 55926.0185 | 55242.5339 | 1100.8507 | 1000.8507 | 1224.7944
9098 | 108.9706 | 0.88453 | 1.356 | 0.95617 | 204.4526 | 51.3573 | 71744.8235 | 71533.9189 | 1426.6784 | 1326.6784 | 1224.855
9099 | 88.6615 | 0.93285 | 0.99919 | 1.0205 | 207.4062 | 45.1577 | 49987.2821 | 50798.8459 | 1011.9769 | 911.9769 | 1224.9557
9100 | 113.6427 | 0.9818 | 1.1897 | 0.97645 | 208.9361 | 48.9736 | 62566.5354 | 60650.2196 | 1209.0044 | 1109.0044 | 1224.9908
9101 | 92.8357 | 1.0064 | 1.1017 | 0.98109 | 196.5644 | 49.3 | 58091.9412 | 58502.9905 | 1166.0598 | 1066.0598 | 1225.0031
9102 | 110.38 | 0.99474 | 0.87466 | 1.0286 | 204.0997 | 50.3847 | 51068.1984 | 53208.4668 | 1060.1693 | 960.1693 | 1225.0792
9103 | 94.5621 | 0.95902 | 0.75972 | 0.96921 | 178.8089 | 52.205 | 43353.2626 | 44246.6561 | 880.9331 | 780.9331 | 1225.121
9104 | 82.1483 | 1.0323 | 1.1757 | 0.99408 | 208.4735 | 52.8621 | 66474.3009 | 68035.9057 | 1356.7181 | 1256.7181 | 1225.1815
9105 | 111.9689 | 0.92454 | 1.2126 | 1.0415 | 179.0313 | 48.5811 | 66562.9246 | 64993.483 | 1295.8697 | 1195.8697 | 1225.3629
9106 | 96.2701 | 0.98005 | 1.0017 | 1.1163 | 204.6533 | 46.8417 | 56999.6246 | 59670.2306 | 1189.4046 | 1089.4046 | 1225.4648
9107 | 88.8686 | 1.0485 | 0.8525 | 1.0441 | 205.5156 | 49.7407 | 49114.3544 | 48473.5033 | 965.4701 | 865.4701 | 1225.4712
9108 | 97.0682 | 1.037 | 0.8583 | 0.90089 | 210.6748 | 51.8292 | 45503.9475 | 45110.2445 | 898.2049 | 798.2049 | 1225.6437
9109 | 108.2797 | 0.97479 | 1.3679 | 0.97787 | 189.2776 | 47.681 | 69001.2046 | 69196.2496 | 1379.925 | 1279.925 | 1225.7868
9110 | 106.4689 | 1.0614 | 1.2105 | 1.0376 | 202.8575 | 50.8817 | 69860.6143 | 71708.8179 | 1430.1764 | 1330.1764 | 1225.8825
9111 | 88.2447 | 0.96 | 1.0514 | 1.0119 | 196.9733 | 49.4619 | 57008.4812 | 55941.4211 | 1114.8284 | 1014.8284 | 1226.1675
9112 | 128.2629 | 1.0718 | 0.94021 | 0.93985 | 192.6702 | 44.5146 | 45647.6245 | 44161.3748 | 879.2275 | 779.2275 | 1226.2583
9113 | 90.7541 | 0.97615 | 0.86383 | 0.99499 | 198.1879 | 53.713 | 51122.8052 | 49690.3842 | 989.8077 | 889.8077 | 1226.4341
9114 | 103.6522 | 0.90849 | 1.0795 | 0.98439 | 188.7836 | 46.6267 | 54125.2678 | 52946.856 | 1054.9371 | 954.9371 | 1226.485
9115 | 99.0729 | 1.0552 | 0.72448 | 1.0154 | 194.1463 | 51.0488 | 43082.5164 | 41999.2456 | 835.9849 | 735.9849 | 1226.5356
9116 | 101.9334 | 0.99145 | 0.97787 | 0.99324 | 192.6152 | 50.377 | 54213.0011 | 54097.6056 | 1077.9521 | 977.9521 | 1226.6313
9117 | 88.2973 | 0.89092 | 1.0399 | 0.93283 | 209.2743 | 47.785 | 50320.6547 | 50798.5684 | 1011.9714 | 911.9714 | 1226.6897
9118 | 98.6457 | 0.93615 | 1.0382 | 0.86229 | 194.921 | 51.9211 | 51471.1482 | 52041.4248 | 1036.8285 | 936.8285 | 1226.7257
9119 | 103.7335 | 0.9957 | 0.77995 | 0.95483 | 204.6778 | 47.4393 | 40433.7915 | 40244.8434 | 800.8969 | 700.8969 | 1226.816
9120 | 92.8365 | 0.97858 | 0.98943 | 0.95845 | 179.5164 | 51.9574 | 54171.9912 | 53570.5371 | 1067.4107 | 967.4107 | 1226.8539
9121 | 104.1302 | 0.9923 | 1.0586 | 1.0158 | 203.2921 | 47.8875 | 56643.8899 | 55659.6032 | 1109.1921 | 1009.1921 | 1226.9271
9122 | 104.3787 | 0.98149 | 1.3146 | 0.98507 | 199.9717 | 49.2565 | 69032.2314 | 69696.0902 | 1389.9218 | 1289.9218 | 1227.0427
9123 | 99.6499 | 1.042 | 1.0377 | 0.98645 | 196.8283 | 50.8646 | 57544.5206 | 56507.5357 | 1126.1507 | 1026.1507 | 1227.1112
9124 | 107.3482 | 1.0463 | 1.1873 | 1.0738 | 214.0215 | 48.9362 | 68101.6428 | 67942.995 | 1354.8599 | 1254.8599 | 1227.113
9125 | 92.3708 | 0.98857 | 0.88608 | 0.9804 | 199.0102 | 44.8191 | 43226.8344 | 42155.4914 | 839.1098 | 739.1098 | 1227.2526
9126 | 87.2635 | 0.9949 | 0.97519 | 1.0363 | 187.807 | 52.1366 | 57401.9894 | 55702.0734 | 1110.0415 | 1010.0415 | 1227.2594
9127 | 112.7544 | 0.98275 | 1.1386 | 0.95923 | 208.1413 | 43.0461 | 51992.9958 | 53536.9862 | 1066.7397 | 966.7397 | 1227.4678
9128 | 104.3975 | 0.93188 | 1.0549 | 0.96567 | 193.1078 | 48.5136 | 54332.8139 | 54716.2782 | 1090.3256 | 990.3256 | 1227.507
9129 | 108.0513 | 0.99328 | 1.0536 | 1.051 | 183.1461 | 50.9128 | 62023.6154 | 64030.4388 | 1276.6088 | 1176.6088 | 1227.6047
9130 | 112.8703 | 1.0051 | 1.3632 | 0.99977 | 198.4572 | 55.1726 | 81652.6765 | 83914.4255 | 1674.2885 | 1574.2885 | 1227.6491
9131 | 104.9431 | 1.0092 | 0.98657 | 0.94435 | 188.3118 | 53.0951 | 55278.1243 | 55910.976 | 1114.2195 | 1014.2195 | 1227.997
9132 | 97.8732 | 0.9897 | 0.90867 | 1.0778 | 203.4928 | 51.7222 | 55869.9489 | 56751.2024 | 1131.024 | 1031.024 | 1228.482
9133 | 92.5682 | 1.0655 | 1.0673 | 0.96076 | 209.7102 | 49.534 | 55889.7991 | 56891.1738 | 1133.8235 | 1033.8235 | 1228.4823
9134 | 93.7952 | 0.95659 | 0.8075 | 0.96499 | 191.1633 | 51.954 | 45336.7429 | 44873.3926 | 893.4679 | 793.4679 | 1228.536
9135 | 103.6247 | 1.0424 | 0.83343 | 0.9989 | 199.2679 | 49.5627 | 46814.5407 | 46832.2959 | 932.6459 | 832.6459 | 1228.6125
9136 | 115.3553 | 0.98627 | 0.74336 | 0.99119 | 201.2855 | 50.483 | 43141.4904 | 43087.938 | 857.7588 | 757.7588 | 1228.6308
9137 | 114.9004 | 1.0219 | 1.0657 | 1.0289 | 196.8121 | 50.9944 | 62101.1281 | 62872.4279 | 1253.4486 | 1153.4486 | 1228.664
9138 | 89.7195 | 0.97784 | 0.95336 | 1.0221 | 199.2099 | 50.9437 | 54310.1582 | 55116.9656 | 1098.3393 | 998.3393 | 1228.7225
9139 | 106.304 | 1.0367 | 0.85893 | 1.0628 | 187.9803 | 49.6349 | 50966.443 | 51587.0666 | 1027.7413 | 927.7413 | 1228.8166
9140 | 94.1705 | 0.92984 | 1.2933 | 0.97867 | 200.6731 | 44.1705 | 59977.414 | 59163.4071 | 1179.2681 | 1079.2681 | 1229.0083
9141 | 100.6425 | 1.0468 | 0.97872 | 1.0546 | 198.2685 | 46.2972 | 52859.7202 | 52480.2335 | 1045.6047 | 945.6047 | 1229.0737
9142 | 97.2827 | 1.0861 | 0.84825 | 0.98924 | 203.0977 | 49.7528 | 47208.3967 | 47067.728 | 937.3546 | 837.3546 | 1229.216
9143 | 118.3017 | 0.96267 | 0.95142 | 0.9208 | 186.0794 | 50.6943 | 50371.339 | 49192.5193 | 979.8504 | 879.8504 | 1229.2882
9144 | 106.9569 | 1.0146 | 0.84297 | 1.0581 | 208.2597 | 49.7466 | 49977.6897 | 50470.7295 | 1005.4146 | 905.4146 | 1229.3372
9145 | 92.5946 | 1.0375 | 0.88259 | 1.0707 | 207.8518 | 49.9359 | 52195.1469 | 52273.1135 | 1041.4623 | 941.4623 | 1229.4185
9146 | 100.2172 | 0.97911 | 0.98601 | 1.0163 | 188.1912 | 48.6651 | 53730.6965 | 54048.653 | 1076.9731 | 976.9731 | 1229.5361
9147 | 122.7588 | 0.98506 | 1.154 | 0.95073 | 199.5242 | 53.059 | 64827.817 | 64802.2182 | 1292.0444 | 1192.0444 | 1229.5744
9148 | 117.4329 | 1.0084 | 0.81975 | 1.0606 | 194.5182 | 50.7602 | 50335.575 | 50014.8651 | 996.2973 | 896.2973 | 1229.6358
9149 | 104.8614 | 1.0849 | 0.91075 | 1.0756 | 193.3094 | 46.7825 | 51341.656 | 52620.1309 | 1048.4026 | 948.4026 | 1229.9438
9150 | 125.7057 | 1.1278 | 0.95357 | 0.97753 | 193.8991 | 53.5605 | 57712.9745 | 57840.249 | 1152.805 | 1052.805 | 1230.1201
9151 | 98.2559 | 1.0303 | 0.8475 | 0.97653 | 201.4844 | 53.6705 | 50052.6298 | 51026.0333 | 1016.5207 | 916.5207 | 1230.2426
9152 | 113.2069 | 0.96445 | 0.96344 | 1.0559 | 189.8486 | 52.7407 | 59601.854 | 57340.0337 | 1142.8007 | 1042.8007 | 1230.332
9153 | 85.4993 | 0.97482 | 1.0576 | 1.1298 | 188.3241 | 51.1812 | 65608.6606 | 63581.9992 | 1267.64 | 1167.64 | 1230.4431
9154 | 88.7068 | 1.0004 | 1.01 | 1.0308 | 193.6082 | 48.5283 | 55026.9268 | 55155.9639 | 1099.1193 | 999.1193 | 1230.5622
9155 | 94.6489 | 0.98486 | 0.96421 | 1.0375 | 198.6376 | 50.9842 | 55951.8515 | 55552.8914 | 1107.0578 | 1007.0578 | 1230.8161
9156 | 104.7067 | 1.0809 | 0.84838 | 1.0808 | 201.3599 | 47.3601 | 48986.1535 | 50044.174 | 996.8835 | 896.8835 | 1230.9459
9157 | 98.4989 | 0.96361 | 1.2434 | 0.92642 | 195.2369 | 48.8299 | 61077.5097 | 59434.8332 | 1184.6967 | 1084.6967 | 1231.0802
9158 | 88.7669 | 0.91337 | 0.8107 | 0.98704 | 213.3739 | 48.202 | 42692.4042 | 42398.4566 | 843.9691 | 743.9691 | 1231.2715
9159 | 104.197 | 0.98612 | 0.87039 | 1.0072 | 195.7433 | 51.5625 | 50698.7368 | 52015.442 | 1036.3088 | 936.3088 | 1231.4967
9160 | 100.9558 | 0.99736 | 1.1915 | 0.9162 | 201.1902 | 49.0728 | 58712.9436 | 58760.1475 | 1171.203 | 1071.203 | 1231.8105
9161 | 106.2653 | 0.92098 | 1.145 | 0.98827 | 206.2739 | 53.2632 | 65688.2345 | 65636.266 | 1308.7253 | 1208.7253 | 1231.8131
9162 | 92.6424 | 1.0498 | 0.97113 | 0.98114 | 208.9051 | 51.394 | 54175.7308 | 54150.7714 | 1079.0154 | 979.0154 | 1231.8389
9163 | 92.1941 | 0.94035 | 1.1818 | 1.0119 | 191.0527 | 49.5171 | 63697.8278 | 61863.355 | 1233.2671 | 1133.2671 | 1231.8693
9164 | 94.5382 | 0.95542 | 1.1277 | 0.9973 | 197.6429 | 48.443 | 59054.6034 | 59947.6171 | 1194.9523 | 1094.9523 | 1231.8796
9165 | 99.5004 | 1.045 | 0.79627 | 1.0059 | 200.0077 | 50.1076 | 45542.8393 | 45784.682 | 911.6936 | 811.6936 | 1231.9118
9166 | 108.2027 | 0.89381 | 1.1252 | 0.98441 | 209.4783 | 48.9102 | 59116.6066 | 57228.3884 | 1140.5678 | 1040.5678 | 1231.9243
9167 | 97.8918 | 0.97093 | 0.8551 | 1.038 | 180.0381 | 52.5987 | 51865.3425 | 51493.8821 | 1025.8776 | 925.8776 | 1231.9259
9168 | 89.2216 | 0.93261 | 0.89118 | 1.0204 | 183.1037 | 47.2451 | 47078.0786 | 46516.2367 | 926.3247 | 826.3247 | 1232.2358
9169 | 99.4272 | 0.92725 | 0.86208 | 1.1066 | 204.5796 | 52.1287 | 54741.7759 | 55832.7595 | 1112.6552 | 1012.6552 | 1232.4885
9170 | 104.0652 | 0.9613 | 0.80699 | 0.9919 | 193.3748 | 50.7773 | 45917.6623 | 45679.6808 | 909.5936 | 809.5936 | 1232.5805
9171 | 88.7332 | 0.99303 | 0.82519 | 1.0125 | 190.1193 | 47.3988 | 43969.5937 | 44591.6886 | 887.8338 | 787.8338 | 1232.5815
9172 | 119.7336 | 1.0221 | 1.2523 | 1.0014 | 179.3262 | 53.9509 | 74434.0844 | 76836.2932 | 1532.7259 | 1432.7259 | 1232.5899
9173 | 92.8474 | 0.92636 | 1.0016 | 0.9715 | 195.449 | 51.4289 | 54664.3015 | 54342.3268 | 1082.8465 | 982.8465 | 1232.7587
9174 | 102.0557 | 1.0761 | 0.96432 | 1.0181 | 180.0265 | 49.1949 | 53883.3121 | 54146.7188 | 1078.9344 | 978.9344 | 1232.9282
9175 | 105.2054 | 1.0024 | 1.0283 | 0.93691 | 190.6664 | 48.3595 | 51880.0267 | 51762.578 | 1031.2516 | 931.2516 | 1233.1961
9176 | 96.2959 | 0.98896 | 1.1948 | 1.0378 | 195.6993 | 51.5947 | 69082.5314 | 70789.1134 | 1411.7823 | 1311.7823 | 1233.3315
9177 | 119.6007 | 1.1083 | 1.0685 | 0.99253 | 196.6692 | 53.6416 | 64194.5377 | 64802.5639 | 1292.0513 | 1192.0513 | 1233.3644
9178 | 93.7602 | 1.0261 | 0.86317 | 1.0188 | 224.1558 | 48.5706 | 47608.6887 | 47165.8838 | 939.3177 | 839.3177 | 1233.4215
9179 | 103.0091 | 1.0249 | 0.81645 | 0.88633 | 216.4689 | 51.9959 | 43332.3299 | 43778.682 | 871.5736 | 771.5736 | 1233.4346
9180 | 123.1533 | 0.91504 | 1.0784 | 0.9758 | 208.384 | 46.1883 | 54016.1347 | 53373.4404 | 1063.4688 | 963.4688 | 1233.6405
9181 | 101.9829 | 1.0163 | 0.83938 | 1.0351 | 209.171 | 48.0484 | 46933.6825 | 48719.106 | 970.3821 | 870.3821 | 1233.7463
9182 | 90.0051 | 1.06 | 1.194 | 0.85245 | 199.2071 | 47.6637 | 53260.1788 | 53908.2604 | 1074.1652 | 974.1652 | 1233.8066
9183 | 100.2001 | 1.046 | 1.0821 | 0.93345 | 210.323 | 45.0224 | 50404.9599 | 49641.4427 | 988.8289 | 888.8289 | 1233.8367
9184 | 96.5286 | 0.98761 | 0.88857 | 1.0822 | 213.726 | 51.6416 | 54797.6708 | 53688.2341 | 1069.7647 | 969.7647 | 1233.9089
9185 | 81.2591 | 1.091 | 0.83266 | 1.0581 | 183.2789 | 47.4078 | 46154.6555 | 45118.1066 | 898.3621 | 798.3621 | 1234.0239
9186 | 92.5088 | 1.0504 | 0.71922 | 0.93388 | 198.1961 | 49.4791 | 38239.3811 | 39004.546 | 776.0909 | 676.0909 | 1234.1308
9187 | 101.6805 | 1.03 | 0.89391 | 0.96145 | 200.7302 | 51.7088 | 50057.6021 | 49834.5189 | 992.6904 | 892.6904 | 1234.3286
9188 | 107.6962 | 1.006 | 1.0602 | 0.91105 | 193.5616 | 52.609 | 56709.2668 | 56631.598 | 1128.632 | 1028.632 | 1234.5083
9189 | 105.6176 | 1.0733 | 1.0733 | 0.99754 | 204.7053 | 44.6673 | 53090.1545 | 52436.7512 | 1044.735 | 944.735 | 1234.516
9190 | 110.9696 | 1.0319 | 1.2585 | 0.94898 | 182.2154 | 51.7363 | 67893.9958 | 68086.8553 | 1357.7371 | 1257.7371 | 1234.5729
9191 | 87.0801 | 1.0354 | 0.88097 | 0.96462 | 206.1462 | 51.5332 | 48645.3876 | 48556.1198 | 967.1224 | 867.1224 | 1234.589
9192 | 96.9662 | 0.94625 | 0.86362 | 0.98726 | 194.7105 | 46.3047 | 43923.6589 | 43042.2236 | 856.8445 | 756.8445 | 1234.7031
9193 | 100.3203 | 1.0812 | 1.0649 | 1.0905 | 219.5238 | 46.5176 | 59284.1951 | 58364.2637 | 1163.2853 | 1063.2853 | 1234.9303
9194 | 100.66 | 1.0148 | 0.93156 | 0.98533 | 200.8596 | 51.8962 | 53137.3432 | 52818.4365 | 1052.3687 | 952.3687 | 1235.1567
9195 | 117.9894 | 0.97449 | 0.97536 | 0.92146 | 189.4248 | 49.2925 | 50159.0179 | 49143.1584 | 978.8632 | 878.8632 | 1235.2095
9196 | 110.8659 | 0.98958 | 0.99696 | 0.89583 | 206.6555 | 50.9025 | 51252.523 | 51811.1489 | 1032.223 | 932.223 | 1235.2281
9197 | 99.8919 | 1.0073 | 0.95584 | 1.0649 | 186.1893 | 52.6344 | 59058.4797 | 58789.3858 | 1171.7877 | 1071.7877 | 1235.2482
9198 | 87.2824 | 1.0037 | 0.85784 | 1.0143 | 180.0642 | 48.5304 | 46659.2134 | 46693.818 | 929.8764 | 829.8764 | 1235.2933
9199 | 100.8616 | 1.0144 | 1.1508 | 0.99149 | 191.256 | 48.2436 | 60175.6777 | 60548.6577 | 1206.9732 | 1106.9732 | 1235.3568
9200 | 93.3344 | 1.0015 | 1.1053 | 0.98959 | 185.5606 | 52.3633 | 62355.1062 | 60958.9559 | 1215.1791 | 1115.1791 | 1235.4762
9201 | 112.5848 | 1.0425 | 1.0955 | 0.94344 | 179.3869 | 51.0863 | 58976.7427 | 57879.2789 | 1153.5856 | 1053.5856 | 1235.6412
9202 | 101.7698 | 1.0043 | 0.77928 | 1.0833 | 188.884 | 49.0588 | 46617.4715 | 46411.0346 | 924.2207 | 824.2207 | 1235.6738
9203 | 108.4814 | 1.0025 | 1.1854 | 1.0386 | 212.0198 | 48.5019 | 65199.6456 | 65009.035 | 1296.1807 | 1196.1807 | 1235.6813
9204 | 96.115 | 1.1022 | 0.81685 | 1.0154 | 216.9631 | 51.21 | 48118.0484 | 48268.9183 | 961.3784 | 861.3784 | 1235.7472
9205 | 96.5297 | 1.1104 | 1.2791 | 0.97988 | 204.8255 | 50.0741 | 68330.9137 | 68979.2842 | 1375.5857 | 1275.5857 | 1235.7662
9206 | 117.7833 | 1.0102 | 0.89576 | 1.0542 | 208.3201 | 49.5304 | 52873.4595 | 52615.919 | 1048.3184 | 948.3184 | 1235.894
9207 | 76.6349 | 1.0341 | 0.86787 | 1.0455 | 211.5381 | 46.1317 | 45723.8864 | 45765.1191 | 911.3024 | 811.3024 | 1235.918
9208 | 115.9283 | 0.99029 | 1.1108 | 1.0305 | 209.7094 | 50.4606 | 63766.793 | 60927.5675 | 1214.5513 | 1114.5513 | 1236.1792
9209 | 80.691 | 1.0332 | 1.0535 | 0.99764 | 188.6604 | 45.0859 | 51334.6374 | 51745.543 | 1030.9109 | 930.9109 | 1236.2094
9210 | 121.7858 | 0.98206 | 1.2097 | 1.0116 | 214.1283 | 48.4299 | 65272.5763 | 65779.8643 | 1311.5973 | 1211.5973 | 1236.2849
9211 | 103.9113 | 1.0439 | 0.98374 | 1.0143 | 200.1448 | 50.8756 | 56483.841 | 55987.1747 | 1115.7435 | 1015.7435 | 1236.9568
9212 | 116.3236 | 0.93667 | 1.0156 | 1.0186 | 201.0846 | 45.6638 | 52417.1934 | 51044.6459 | 1016.8929 | 916.8929 | 1236.9628
9213 | 99.7667 | 0.95603 | 1.1718 | 0.98791 | 198.4749 | 49.9948 | 62840.274 | 63517.4127 | 1266.3483 | 1166.3483 | 1237.3502
9214 | 112.0725 | 1.0429 | 1.0416 | 0.91696 | 190.306 | 46.4324 | 49964.2915 | 50416.6436 | 1004.3329 | 904.3329 | 1237.3526
9215 | 109.1791 | 1.0176 | 1.056 | 1.0075 | 206.023 | 47.0044 | 55435.903 | 58078.0857 | 1157.5617 | 1057.5617 | 1237.4657
9216 | 81.6036 | 0.97656 | 0.97077 | 1.102 | 196.627 | 48.3654 | 55790.8447 | 55450.073 | 1105.0015 | 1005.0015 | 1237.4785
9217 | 100.5656 | 1.0327 | 1.2801 | 0.89966 | 206.0085 | 45.7123 | 57599.055 | 60460.7134 | 1205.2143 | 1105.2143 | 1237.6891
9218 | 107.3817 | 1.0159 | 0.84752 | 1.0035 | 199.8119 | 50.5552 | 48710.6389 | 47800.1655 | 952.0033 | 852.0033 | 1237.7903
9219 | 96.8304 | 0.93107 | 1.0249 | 1.0156 | 200.5247 | 50.7852 | 57642.0927 | 58821.7906 | 1172.4358 | 1072.4358 | 1237.9274
9220 | 96.0044 | 0.92348 | 0.9181 | 0.97302 | 194.6285 | 48.4195 | 47741.7889 | 46828.1138 | 932.5623 | 832.5623 | 1237.9511
9221 | 101.1811 | 1.0041 | 0.7958 | 0.98847 | 200.4068 | 46.98 | 41929.0139 | 41856.6181 | 833.1324 | 733.1324 | 1238.2221
9222 | 89.3032 | 0.94054 | 0.81186 | 1.0047 | 193.6438 | 52.7736 | 47670.6221 | 48242.4977 | 960.85 | 860.85 | 1238.3082
9223 | 103.2552 | 0.98306 | 0.9527 | 0.97428 | 193.4319 | 50.7917 | 52493.8267 | 52699.1904 | 1049.9838 | 949.9838 | 1238.863
9224 | 107.3858 | 1.034 | 1.1221 | 1.0311 | 187.2852 | 49.8093 | 63349.9166 | 62397.8941 | 1243.9579 | 1143.9579 | 1239.0112
9225 | 110.0515 | 0.97562 | 1.2062 | 1.0241 | 202.3519 | 45.6687 | 61520.2127 | 61988.0018 | 1235.76 | 1135.76 | 1239.0288
9226 | 90.1237 | 1.008 | 1.0789 | 1.0568 | 205.5401 | 50.2444 | 62061.0111 | 60594.3942 | 1207.8879 | 1107.8879 | 1239.2825
9227 | 90.2238 | 0.97415 | 0.93501 | 1.013 | 199.4017 | 49.8889 | 51836.7466 | 53795.7824 | 1071.9156 | 971.9156 | 1239.441
9228 | 99.3555 | 1.0193 | 1.0321 | 1.012 | 184.8322 | 52.4443 | 60275.7286 | 58982.2039 | 1175.6441 | 1075.6441 | 1239.5367
9229 | 86.6979 | 1.057 | 0.91543 | 0.99745 | 176.449 | 50.8773 | 51294.9264 | 52063.2713 | 1037.2654 | 937.2654 | 1239.7624
9230 | 95.4542 | 1.0041 | 1.1515 | 1.0011 | 202.0576 | 49.3936 | 61873.155 | 62266.6268 | 1241.3325 | 1141.3325 | 1239.7714
9231 | 96.2031 | 1.0169 | 1.0058 | 0.97784 | 194.4135 | 51.448 | 55824.8916 | 55323.2092 | 1102.4642 | 1002.4642 | 1239.7958
9232 | 114.9964 | 0.95527 | 1.2608 | 0.97351 | 190.603 | 50.9497 | 68322.3172 | 66526.7807 | 1326.5356 | 1226.5356 | 1239.9063
9233 | 101.4657 | 0.9907 | 0.94119 | 0.99487 | 184.5821 | 51.601 | 53689.0577 | 54572.4015 | 1087.448 | 987.448 | 1239.9184
9234 | 89.4106 | 1.0166 | 0.94121 | 0.96876 | 195.3825 | 45.1606 | 45477.7034 | 46232.3932 | 920.6479 | 820.6479 | 1240.1534
9235 | 94.7929 | 0.91646 | 0.73465 | 1.0061 | 188.7438 | 44.5663 | 37001.1556 | 35420.5831 | 704.4117 | 604.4117 | 1240.1781
9236 | 102.0185 | 1.07 | 0.89091 | 1.0258 | 216.4143 | 49.4469 | 50802.2254 | 49073.3948 | 977.4679 | 877.4679 | 1240.2878
9237 | 108.197 | 1.0037 | 1.1475 | 1.02 | 194.041 | 51.5854 | 66179.221 | 64342.7799 | 1282.8556 | 1182.8556 | 1240.398
9238 | 95.0511 | 0.98914 | 0.88552 | 1.0267 | 185.5669 | 48.5831 | 48925.4354 | 47943.1704 | 954.8634 | 854.8634 | 1240.5139
9239 | 105.1402 | 1.0241 | 1.0687 | 0.95753 | 190.1252 | 46.9618 | 53305.0224 | 51223.7352 | 1020.4747 | 920.4747 | 1240.5786
9240 | 112.3025 | 1.0426 | 0.98478 | 1.0459 | 195.4134 | 48.8777 | 56261.3559 | 57269.4314 | 1141.3886 | 1041.3886 | 1240.5834
9241 | 82.1544 | 1.1395 | 0.83137 | 0.93424 | 190.2878 | 45.8217 | 40069.8591 | 42038.3936 | 836.7679 | 736.7679 | 1240.7102
9242 | 96.9044 | 0.8904 | 1.2425 | 1.1351 | 197.7179 | 49.0158 | 73556.401 | 70989.6852 | 1415.7937 | 1315.7937 | 1240.8069
9243 | 90.3939 | 1.0784 | 0.97583 | 0.93717 | 205.2054 | 50.158 | 50964.9821 | 50318.9865 | 1002.3797 | 902.3797 | 1240.8728
9244 | 98.6612 | 0.97727 | 0.97905 | 0.97972 | 178.4077 | 49.1543 | 52066.4733 | 50823.0603 | 1012.4612 | 912.4612 | 1241.0413
9245 | 99.3126 | 0.90074 | 1.2504 | 1.0837 | 204.1324 | 47.6362 | 69016.9531 | 67485.1297 | 1345.7026 | 1245.7026 | 1241.0908
9246 | 89.1283 | 1.0424 | 0.88606 | 0.91371 | 192.9999 | 48.478 | 43944.9672 | 45462.7661 | 905.2553 | 805.2553 | 1241.1139
9247 | 105.7576 | 1.0384 | 0.90593 | 0.98313 | 210.4004 | 49.4487 | 49682.1722 | 49770.3383 | 991.4068 | 891.4068 | 1241.1521
9248 | 87.1573 | 1.0593 | 0.82745 | 1.014 | 205.4303 | 51.3344 | 48015.7619 | 47198.3563 | 939.9671 | 839.9671 | 1241.1525
9249 | 88.1308 | 1.0039 | 0.9492 | 0.98435 | 204.7643 | 46.4404 | 47704.9334 | 47898.0883 | 953.9618 | 853.9618 | 1241.6088
9250 | 93.4369 | 0.96924 | 0.8837 | 1.0406 | 174.8229 | 49.3563 | 50029.6279 | 50587.5708 | 1007.7514 | 907.7514 | 1241.9273
9251 | 93.675 | 0.96269 | 1.0013 | 0.92961 | 204.3355 | 51.9003 | 53192.053 | 51683.0332 | 1029.6607 | 929.6607 | 1242.049
9252 | 85.3547 | 0.94104 | 0.88034 | 0.95994 | 175.7137 | 48.9785 | 45500.3662 | 45442.6946 | 904.8539 | 804.8539 | 1242.2194
9253 | 100.84 | 1.0129 | 1.0278 | 1.0375 | 191.1694 | 49.4411 | 57962.3176 | 57885.5266 | 1153.7105 | 1053.7105 | 1242.3395
9254 | 110.6881 | 0.98097 | 0.95073 | 1.1042 | 204.6861 | 45.8232 | 53285.5035 | 54322.2893 | 1082.4458 | 982.4458 | 1242.5292
9255 | 95.966 | 0.93762 | 1.1142 | 0.9903 | 196.2235 | 52.778 | 63180.3058 | 61577.9587 | 1227.5592 | 1127.5592 | 1242.8686
9256 | 100.2141 | 0.97573 | 0.95798 | 0.93454 | 212.2489 | 49.4466 | 49314.8908 | 49457.3344 | 985.1467 | 885.1467 | 1242.9504
9257 | 100.2042 | 1.1355 | 1.187 | 1.0622 | 212.4908 | 47.5639 | 65592.5222 | 65043.3812 | 1296.8676 | 1196.8676 | 1242.9746
9258 | 95.0499 | 0.9942 | 0.83391 | 1.0473 | 201.734 | 54.5633 | 53011.0782 | 53541.938 | 1066.8388 | 966.8388 | 1243.5165
9259 | 100.2462 | 0.98418 | 0.87885 | 0.9679 | 196.9875 | 47.2812 | 45081.2038 | 45693.2891 | 909.8658 | 809.8658 | 1243.6685
9260 | 110.8222 | 1.0218 | 0.82877 | 0.93955 | 195.0094 | 50.7979 | 45501.7527 | 46157.7292 | 919.1546 | 819.1546 | 1243.7481
9261 | 116.8729 | 1.0019 | 1.1503 | 0.96825 | 224.9448 | 49.0271 | 60569.5108 | 61484.9874 | 1225.6997 | 1125.6997 | 1243.9601
9262 | 103.4294 | 1.0302 | 1.0429 | 1.0359 | 211.093 | 51.835 | 61735.3015 | 58869.2904 | 1173.3858 | 1073.3858 | 1243.994
9263 | 95.0315 | 1.0798 | 0.73803 | 1.0228 | 196.2962 | 51.418 | 44285.7118 | 45560.0403 | 907.2008 | 807.2008 | 1244.0131
9264 | 90.4085 | 0.95003 | 1.0827 | 1.0322 | 191.9031 | 54.6134 | 65912.4564 | 66364.6802 | 1323.2936 | 1223.2936 | 1244.053
9265 | 102.5661 | 1.0181 | 0.9939 | 0.99466 | 206.2669 | 50.0767 | 54940.8097 | 53225.8245 | 1060.5165 | 960.5165 | 1244.3648
9266 | 106.1462 | 1.0203 | 0.96971 | 0.9272 | 199.0564 | 53.8763 | 54475.2465 | 54551.7146 | 1087.0343 | 987.0343 | 1245.2222
9267 | 88.3174 | 0.90565 | 0.97745 | 1.0646 | 213.558 | 47.774 | 53748.1618 | 53398.6174 | 1063.9723 | 963.9723 | 1245.5976
9268 | 96.5808 | 1.0233 | 1.1549 | 1.0189 | 209.3168 | 48.7637 | 62411.3479 | 62666.6366 | 1249.3327 | 1149.3327 | 1245.6015
9269 | 96.3504 | 1.0257 | 1.0498 | 0.94694 | 208.5085 | 48.8717 | 53623.2192 | 54217.7945 | 1080.3559 | 980.3559 | 1245.6096
9270 | 113.3601 | 0.92795 | 0.87182 | 1.0114 | 179.2963 | 46.6662 | 46237.052 | 45671.6888 | 909.4338 | 809.4338 | 1245.7026
9271 | 98.4562 | 0.98059 | 1.0053 | 1.0405 | 182.7761 | 50.5279 | 57915.2828 | 59626.0067 | 1188.5201 | 1088.5201 | 1245.8894
9272 | 97.2945 | 0.99378 | 1.0045 | 0.99308 | 200.2252 | 53.7409 | 59007.8778 | 58119.3636 | 1158.3873 | 1058.3873 | 1246.1518
9273 | 107.6715 | 1.0314 | 1.0108 | 1.0516 | 204.1054 | 54.4562 | 64137.693 | 63383.1502 | 1263.663 | 1163.663 | 1246.1756
9274 | 100.9886 | 1.098 | 1.1323 | 0.92175 | 176.1688 | 51.7158 | 59885.1093 | 60669.2399 | 1209.3848 | 1109.3848 | 1246.2889
9275 | 88.7558 | 1.0207 | 1.0923 | 1.0106 | 208.665 | 51.4394 | 61650.0269 | 61754.4432 | 1231.0889 | 1131.0889 | 1246.3243
9276 | 104.3931 | 1.0358 | 1.0541 | 1.0063 | 203.6581 | 49.5492 | 58119.8701 | 56750.551 | 1131.011 | 1031.011 | 1246.3609
9277 | 97.0266 | 0.98607 | 0.80112 | 1.026 | 196.6111 | 46.8187 | 43157.6793 | 42528.3077 | 846.5662 | 746.5662 | 1246.525
9278 | 99.8942 | 0.94614 | 0.84891 | 0.94571 | 188.8386 | 49.6603 | 44750.982 | 44733.4452 | 890.6689 | 790.6689 | 1246.5313
9279 | 91.7695 | 1.0685 | 0.94739 | 0.90747 | 193.1298 | 45.2329 | 43516.3446 | 41681.5783 | 829.6316 | 729.6316 | 1246.6373
9280 | 91.0661 | 0.94233 | 0.89034 | 0.96573 | 208.7466 | 48.2022 | 45791.0649 | 46104.7301 | 918.0946 | 818.0946 | 1246.8849
9281 | 106.4205 | 0.95144 | 1.0962 | 1.0235 | 189.3561 | 45.6028 | 55970.047 | 54272.2182 | 1081.4444 | 981.4444 | 1246.9804
9282 | 93.231 | 0.99389 | 0.91067 | 1.0153 | 213.6447 | 50.789 | 51878.7524 | 52680.3761 | 1049.6075 | 949.6075 | 1247.0178
9283 | 85.394 | 1.0012 | 0.97624 | 0.93788 | 201.8767 | 47.3014 | 47554.8886 | 47138.8595 | 938.7772 | 838.7772 | 1247.0773
9284 | 107.5955 | 0.96777 | 0.79547 | 1.0043 | 226.6948 | 53.6859 | 48706.8438 | 48066.0491 | 957.321 | 857.321 | 1247.1968
9285 | 107.7345 | 1.092 | 1.1822 | 1.0335 | 193.4447 | 46.3972 | 62337.7576 | 63137.3884 | 1258.7478 | 1158.7478 | 1247.1984
9286 | 98.3724 | 1.0025 | 0.85421 | 1.091 | 205.6002 | 47.5536 | 49212.0222 | 50351.7749 | 1003.0355 | 903.0355 | 1247.366
9287 | 87.3918 | 1.0071 | 0.8696 | 1.0641 | 191.2202 | 51.0406 | 51915.5168 | 51325.2429 | 1022.5049 | 922.5049 | 1247.3687
9288 | 122.5693 | 1.0172 | 0.76863 | 0.91503 | 215.2854 | 53.7427 | 44714.1639 | 45793.744 | 911.8749 | 811.8749 | 1247.382
9289 | 109.3394 | 1.0496 | 1.3457 | 0.91692 | 209.0282 | 50.4665 | 68271.3543 | 67010.47 | 1336.2094 | 1236.2094 | 1247.5169
9290 | 98.223 | 1.0442 | 0.81658 | 1.0354 | 187.7793 | 51.4244 | 48941.5492 | 48567.9559 | 967.3591 | 867.3591 | 1247.5888
9291 | 99.4354 | 0.98782 | 1.0663 | 1.048 | 201.8586 | 46.2733 | 56456.2649 | 56089.1064 | 1117.7821 | 1017.7821 | 1247.5924
9292 | 97.5694 | 0.97249 | 1.2721 | 0.98321 | 205.6699 | 51.2166 | 69124.3043 | 68351.632 | 1363.0326 | 1263.0326 | 1247.831
9293 | 99.6301 | 0.97714 | 0.94914 | 1.0442 | 207.905 | 56.0732 | 61242.1641 | 60597.0328 | 1207.9407 | 1107.9407 | 1248.1932
9294 | 92.8274 | 0.97854 | 0.98533 | 0.9074 | 188.3912 | 49.7339 | 49172.3497 | 49662.3188 | 989.2464 | 889.2464 | 1248.2902
9295 | 122.7915 | 1.0283 | 1.2844 | 1.0151 | 199.4571 | 49.1502 | 70489.8957 | 71682.4569 | 1429.6491 | 1329.6491 | 1248.3537
9296 | 104.2836 | 1.0251 | 1.2071 | 0.95218 | 205.8954 | 47.4614 | 59832.1254 | 59124.2487 | 1178.485 | 1078.485 | 1248.7725
9297 | 95.8463 | 1.0616 | 1.2939 | 0.9845 | 193.7216 | 49.9305 | 68878.877 | 70416.7641 | 1404.3353 | 1304.3353 | 1248.804
9298 | 116.6731 | 0.96731 | 1.2136 | 0.93156 | 187.5314 | 51.3791 | 64073.5182 | 64191.1097 | 1279.8222 | 1179.8222 | 1248.9828
9299 | 83.1047 | 1.0155 | 1.168 | 0.97764 | 201.3708 | 52.9353 | 65115.2425 | 66716.6021 | 1330.332 | 1230.332 | 1249.1773
9300 | 100.6572 | 1.0969 | 0.95753 | 1.0048 | 182.1174 | 52.868 | 56885.7466 | 57290.4471 | 1141.8089 | 1041.8089 | 1249.2607
9301 | 114.5926 | 1.0599 | 1.0301 | 1.1013 | 199.619 | 50.713 | 63889.7754 | 65081.6502 | 1297.633 | 1197.633 | 1249.2618
9302 | 107.3677 | 0.94002 | 0.97143 | 1.1076 | 195.0536 | 51.7224 | 61066.0635 | 60000.4544 | 1196.0091 | 1096.0091 | 1249.4702
9303 | 97.3281 | 0.9922 | 0.83689 | 0.87739 | 212.071 | 52.3849 | 43736.1278 | 44095.8564 | 877.9171 | 777.9171 | 1249.4881
9304 | 101.7608 | 1.0082 | 1.0869 | 1.0107 | 200.3295 | 47.287 | 56997.666 | 57107.8974 | 1138.1579 | 1038.1579 | 1249.5296
9305 | 83.3615 | 1.0253 | 1.1255 | 1.0668 | 213.065 | 51.5283 | 66488.764 | 68508.1845 | 1366.1637 | 1266.1637 | 1249.6099
9306 | 98.1184 | 0.96781 | 0.92898 | 1.0553 | 204.5196 | 45.7151 | 49364.3277 | 49179.7997 | 979.596 | 879.596 | 1249.7356
9307 | 98.1232 | 1.0702 | 1.0843 | 1.0575 | 185.4852 | 54.0779 | 67875.2956 | 68082.6634 | 1357.6533 | 1257.6533 | 1250.0227
9308 | 84.1206 | 1.0354 | 1.07 | 1.0383 | 173.8706 | 47.5287 | 57118.5647 | 57572.678 | 1147.4536 | 1047.4536 | 1250.2925
9309 | 98.3007 | 1.091 | 1.1431 | 1.0469 | 204.4078 | 53.9817 | 70597.165 | 69928.3193 | 1394.5664 | 1294.5664 | 1250.3531
9310 | 100.348 | 0.86771 | 0.85373 | 1.0156 | 201.3279 | 53.7609 | 51496.6603 | 51571.3418 | 1027.4268 | 927.4268 | 1250.4352
9311 | 105.987 | 0.97262 | 1.248 | 0.95995 | 200.4852 | 48.3975 | 63168.6217 | 61256.6468 | 1221.1329 | 1121.1329 | 1250.6085
9312 | 94.892 | 0.89009 | 1.161 | 0.90817 | 204.2008 | 50.691 | 57931.2154 | 58055.9095 | 1157.1182 | 1057.1182 | 1250.6839
9313 | 92.6827 | 1.0298 | 1.0224 | 1.019 | 192.6911 | 51.4955 | 58758.1628 | 58550.1354 | 1167.0027 | 1067.0027 | 1250.6895
9314 | 104.5729 | 0.97586 | 1.1058 | 1.0008 | 192.75 | 52.2763 | 63378.1671 | 64421.3295 | 1284.4266 | 1184.4266 | 1250.8288
9315 | 96.3462 | 1.0152 | 1.1535 | 0.98312 | 201.9715 | 51.7556 | 63958.7754 | 62691.0187 | 1249.8204 | 1149.8204 | 1251.1877
9316 | 89.6036 | 0.98787 | 0.93794 | 0.99102 | 208.7964 | 48.9402 | 50031.7124 | 51706.5212 | 1030.1304 | 930.1304 | 1251.3354
9317 | 101.439 | 0.96488 | 0.92876 | 0.98781 | 188.529 | 46.7944 | 47699.7729 | 48936.8318 | 974.7366 | 874.7366 | 1251.3965
9318 | 84.9126 | 0.98659 | 1.2061 | 0.9856 | 201.9744 | 49.4468 | 63123.6064 | 63082.5641 | 1257.6513 | 1157.6513 | 1251.4176
9319 | 110.5506 | 0.9746 | 1.1522 | 1.0226 | 219.1053 | 50.6867 | 65400.7937 | 65380.4183 | 1303.6084 | 1203.6084 | 1251.486
9320 | 98.5928 | 1.0311 | 1.0272 | 0.98083 | 210.2839 | 46.4758 | 51759.2375 | 51411.9889 | 1024.2398 | 924.2398 | 1251.6196
9321 | 89.6292 | 1.0309 | 1.0108 | 0.99488 | 183.9084 | 50.8685 | 56037.2232 | 52492.2609 | 1045.8452 | 945.8452 | 1251.8996
9322 | 104.5846 | 0.98722 | 1.2546 | 1.0509 | 194.2306 | 52.7193 | 75145.9068 | 75389.558 | 1503.7912 | 1403.7912 | 1251.9467
9323 | 102.8623 | 1.0081 | 0.95589 | 1.0414 | 199.3897 | 48.0581 | 53021.0712 | 54355.9864 | 1083.1197 | 983.1197 | 1252.1454
9324 | 108.6689 | 1.0046 | 1.0527 | 1.0198 | 204.175 | 44.7824 | 53169.004 | 53094.1204 | 1057.8824 | 957.8824 | 1252.2413
9325 | 123.2801 | 1.0263 | 0.98143 | 0.99011 | 198.1897 | 50.9592 | 56163.8574 | 57400.0482 | 1144.001 | 1044.001 | 1252.405
9326 | 94.9152 | 0.98383 | 1.0953 | 0.86037 | 210.5003 | 47.1763 | 49072.3114 | 49062.1852 | 977.2437 | 877.2437 | 1254.022
9327 | 99.7631 | 0.99284 | 1.1847 | 1.0532 | 183.1569 | 50.4134 | 68076.418 | 68588.9438 | 1367.7789 | 1267.7789 | 1254.0897
9328 | 99.6698 | 1.0201 | 1.2249 | 1.1407 | 197.1882 | 52.3086 | 78601.902 | 77796.5677 | 1551.9314 | 1451.9314 | 1254.4003
9329 | 106.8459 | 0.98596 | 0.77074 | 0.97469 | 201.8684 | 48.7086 | 41924.7105 | 43717.322 | 870.3464 | 770.3464 | 1254.4171
9330 | 94.3591 | 1.0354 | 1.0546 | 0.94566 | 203.8603 | 53.0909 | 58337.9221 | 59368.8148 | 1183.3763 | 1083.3763 | 1254.8599
9331 | 110.2102 | 1.0401 | 1.1962 | 1.0825 | 196.1588 | 48.9854 | 69243.5712 | 70753.3483 | 1411.067 | 1311.067 | 1255.183
9332 | 94.5329 | 1.0548 | 0.87909 | 1.0272 | 193.3854 | 50.1361 | 50464.9152 | 50672.9996 | 1009.46 | 909.46 | 1255.1885
9333 | 97.0789 | 0.92333 | 1.1671 | 0.88718 | 197.9509 | 50.7834 | 57332.2362 | 57094.9138 | 1137.8983 | 1037.8983 | 1255.3982
9334 | 91.6318 | 1.0101 | 1.0529 | 0.94369 | 197.2425 | 53.2677 | 58056.9835 | 57426.1974 | 1144.5239 | 1044.5239 | 1255.4467
9335 | 93.8871 | 0.96292 | 1.1269 | 1.0496 | 208.924 | 43.8558 | 56043.7107 | 57856.6026 | 1153.1321 | 1053.1321 | 1255.4781
9336 | 117.1439 | 0.99472 | 0.7297 | 1.0294 | 218.8109 | 50.7216 | 44229.1749 | 44176.087 | 879.5217 | 779.5217 | 1255.8927
9337 | 103.2049 | 0.99304 | 0.86785 | 1.0681 | 195.2049 | 50.9336 | 52626.0994 | 51995.7159 | 1035.9143 | 935.9143 | 1256.0744
9338 | 113.1711 | 1.0195 | 0.88954 | 0.96406 | 206.9053 | 49.1983 | 48074.3488 | 47059.9502 | 937.199 | 837.199 | 1256.6437
9339 | 104.6657 | 0.99547 | 1.196 | 0.98004 | 193.8726 | 46.0649 | 58986.6912 | 57237.2782 | 1140.7456 | 1040.7456 | 1256.7181
9340 | 85.8834 | 1.0411 | 0.93904 | 1.0272 | 206.3689 | 47.3151 | 50077.0382 | 50999.4772 | 1015.9895 | 915.9895 | 1256.9116
9341 | 100.2279 | 0.98794 | 1.0079 | 1.1376 | 186.1581 | 48.4211 | 60499.2459 | 59507.4254 | 1186.1485 | 1086.1485 | 1256.9203
9342 | 82.0134 | 0.96723 | 1.0538 | 0.94037 | 214.6757 | 50.1554 | 53894.9881 | 55068.9429 | 1097.3789 | 997.3789 | 1256.9954
9343 | 122.7382 | 1.0118 | 1.2553 | 0.97865 | 205.7334 | 52.4638 | 71174.3483 | 69957.3346 | 1395.1467 | 1295.1467 | 1257.2805
9344 | 86.4889 | 1.1027 | 0.96608 | 0.96283 | 180.2643 | 46.1624 | 47521.6384 | 48463.0676 | 965.2614 | 865.2614 | 1257.4151
9345 | 77.3209 | 1.0574 | 0.79758 | 1.0851 | 210.6312 | 51.3655 | 48862.9358 | 48445.0249 | 964.9005 | 864.9005 | 1257.6299
9346 | 91.0909 | 1.0636 | 0.98375 | 1.0477 | 195.8787 | 50.9831 | 57681.7284 | 58353.9476 | 1163.079 | 1063.079 | 1257.6533
9347 | 90.7398 | 0.99422 | 0.81023 | 0.96328 | 193.3658 | 48.5775 | 42489.9155 | 42197.9445 | 839.9589 | 739.9589 | 1257.7371
9348 | 87.7475 | 1.0043 | 1.0238 | 0.965 | 192.5782 | 50.9064 | 54970.3998 | 55090.9028 | 1097.8181 | 997.8181 | 1258.0139
9349 | 107.4662 | 0.95701 | 1.0897 | 0.99388 | 211.5082 | 52.485 | 62451.2905 | 61631.5302 | 1228.6306 | 1128.6306 | 1258.1197
9350 | 92.3289 | 0.93403 | 0.94258 | 1.1139 | 193.6182 | 48.9586 | 55818.3906 | 55813.6937 | 1112.2739 | 1012.2739 | 1258.2912
9351 | 96.9924 | 1.0629 | 1.0843 | 0.94007 | 187.6389 | 53.5208 | 60259.0765 | 60271.8791 | 1201.4376 | 1101.4376 | 1258.3029
9352 | 124.5241 | 0.86826 | 1.0167 | 1.06 | 208.8743 | 47.6959 | 56767.1348 | 55928.396 | 1114.5679 | 1014.5679 | 1258.3067
9353 | 108.8827 | 0.98973 | 0.99966 | 0.99147 | 205.8471 | 46.9995 | 51853.5198 | 51982.235 | 1035.6447 | 935.6447 | 1258.336
9354 | 99.755 | 1.0186 | 1.0337 | 1.016 | 194.9306 | 49.7126 | 57455.8695 | 58643.2466 | 1168.8649 | 1068.8649 | 1258.4981
9355 | 88.2833 | 1.009 | 1.1079 | 0.9301 | 234.4018 | 50.2869 | 56532.998 | 58305.0624 | 1162.1012 | 1062.1012 | 1258.7272
9356 | 86.4974 | 0.93264 | 1.0742 | 0.9754 | 191.5826 | 46.9056 | 53124.1077 | 54152.0958 | 1079.0419 | 979.0419 | 1258.7897
9357 | 93.3728 | 1.0474 | 0.65709 | 1.024 | 191.3969 | 49.7786 | 38553.317 | 40675.7715 | 809.5154 | 709.5154 | 1258.9172
9358 | 80.8465 | 1.0322 | 1.0528 | 1.1008 | 181.5953 | 52.236 | 65076.9378 | 65487.5059 | 1305.7501 | 1205.7501 | 1258.9917
9359 | 103.0264 | 0.99753 | 0.84719 | 0.97369 | 214.098 | 48.9567 | 45630.0217 | 44901.3564 | 894.0271 | 794.0271 | 1259.4846
9360 | 94.734 | 0.97389 | 0.94844 | 1.0131 | 193.719 | 55.0979 | 58218.2821 | 58471.273 | 1165.4255 | 1065.4255 | 1259.5189
9361 | 98.5498 | 1.0088 | 1.0718 | 1.1264 | 201.8428 | 51.6561 | 67699.5553 | 69558.3318 | 1387.1666 | 1287.1666 | 1259.5388
9362 | 116.6552 | 0.98286 | 0.93342 | 0.92011 | 210.0556 | 50.6637 | 49531.4185 | 48904.9082 | 974.0982 | 874.0982 | 1259.5466
9363 | 94.2262 | 1.0459 | 1.1217 | 1.1013 | 186.3193 | 50.311 | 67299.4072 | 67096.3705 | 1337.9274 | 1237.9274 | 1259.5972
9364 | 93.0952 | 1.0515 | 0.92468 | 1.0774 | 195.2909 | 51.8225 | 56894.5408 | 56589.8105 | 1127.7962 | 1027.7962 | 1259.6562
9365 | 103.8755 | 0.99355 | 0.92452 | 1.0264 | 199.0086 | 49.86 | 52660.6366 | 53192.3242 | 1059.8465 | 959.8465 | 1259.9802
9366 | 111.8039 | 1.0146 | 1.1391 | 0.92639 | 205.1145 | 48.021 | 56327.8561 | 55846.7579 | 1112.9352 | 1012.9352 | 1260.1045
9367 | 96.073 | 0.95833 | 1.1062 | 0.91326 | 206.4674 | 50.0949 | 55428.7232 | 56644.3676 | 1128.8874 | 1028.8874 | 1260.4153
9368 | 97.7967 | 0.9884 | 0.80931 | 0.98252 | 206.0964 | 48.8815 | 43799.9179 | 43666.5976 | 869.332 | 769.332 | 1260.5155
9369 | 103.0225 | 0.95062 | 0.98079 | 0.99965 | 220.8188 | 46.1732 | 50013.1304 | 51077.5773 | 1017.5515 | 917.5515 | 1260.5273
9370 | 92.8801 | 1.0338 | 0.81355 | 0.9812 | 172.6474 | 48.7087 | 43731.7402 | 43806.5558 | 872.1311 | 772.1311 | 1260.6018
9371 | 111.0266 | 1.0224 | 1.0148 | 0.91698 | 196.9086 | 49.6079 | 51989.3276 | 53319.1757 | 1062.3835 | 962.3835 | 1260.8352
9372 | 102.4104 | 1.0439 | 0.97953 | 1.0188 | 206.4122 | 52.353 | 58049.6064 | 57820.8146 | 1152.4163 | 1052.4163 | 1260.9664
9373 | 86.0902 | 1.0366 | 0.75166 | 0.99776 | 204.7967 | 50.0556 | 42211.951 | 43025.5594 | 856.5112 | 756.5112 | 1261.3361
9374 | 112.1066 | 0.96155 | 0.87796 | 1.0116 | 216.705 | 50.4518 | 50461.922 | 50634.0545 | 1008.6811 | 908.6811 | 1261.3935
9375 | 94.1356 | 0.97637 | 1.3142 | 0.92367 | 197.9066 | 49.1194 | 64336.514 | 61767.6581 | 1231.3532 | 1131.3532 | 1261.6629
9376 | 96.9734 | 0.94139 | 0.85819 | 1.0145 | 196.3049 | 49.4297 | 47741.6063 | 47464.9464 | 945.2989 | 845.2989 | 1261.6727
9377 | 115.5403 | 0.98621 | 0.7767 | 1.0313 | 208.8219 | 51.5059 | 47334.5427 | 47045.4895 | 936.9098 | 836.9098 | 1261.9843
9378 | 94.2164 | 1.003 | 0.95513 | 0.97457 | 214.9333 | 49.263 | 50726.22 | 51618.1228 | 1028.3625 | 928.3625 | 1262.1714
9379 | 114.4504 | 0.97138 | 1.0531 | 0.94136 | 204.963 | 46.9482 | 51968.2913 | 53101.6499 | 1058.033 | 958.033 | 1262.1795
9380 | 81.3023 | 0.94536 | 0.92752 | 0.93095 | 188.6455 | 52.8036 | 49841.6931 | 49314.2853 | 982.2857 | 882.2857 | 1262.2945
9381 | 102.875 | 0.87661 | 0.82778 | 1.0138 | 189.6411 | 49.7664 | 46440.211 | 46084.0469 | 917.6809 | 817.6809 | 1262.8553
9382 | 96.5906 | 0.98421 | 1.0816 | 1.0855 | 190.5425 | 47.3946 | 60338.2076 | 57347.4772 | 1142.9495 | 1042.9495 | 1262.9581
9383 | 81.2365 | 0.99043 | 1.0703 | 0.90523 | 211.4328 | 47.7098 | 50276.4276 | 50297.8504 | 1001.957 | 901.957 | 1263.0326
9384 | 93.9451 | 0.91005 | 0.87981 | 1.0625 | 209.2096 | 47.188 | 48356.4008 | 46910.0903 | 934.2018 | 834.2018 | 1263.1069
9385 | 107.0206 | 1.021 | 0.76462 | 1.0196 | 188.2291 | 51.0559 | 45570.3266 | 45475.8256 | 905.5165 | 805.5165 | 1263.9556
9386 | 104.3538 | 1.0339 | 1.0438 | 0.97166 | 197.4211 | 49.2426 | 55455.5028 | 56286.963 | 1121.7393 | 1021.7393 | 1264.0479
9387 | 118.5647 | 0.92424 | 0.97958 | 1.0239 | 198.9552 | 51.5145 | 57512.3348 | 58857.2208 | 1173.1444 | 1073.1444 | 1264.4138
9388 | 109.0302 | 0.90745 | 1.1879 | 0.89682 | 203.3488 | 49.4606 | 57790.6035 | 57945.4784 | 1154.9096 | 1054.9096 | 1265.2529
9389 | 112.9998 | 1.079 | 1.308 | 0.94312 | 201.1225 | 46.9036 | 63780.1001 | 64135.6743 | 1278.7135 | 1178.7135 | 1265.2841
9390 | 87.8128 | 0.98999 | 1.2154 | 0.93963 | 205.5044 | 47.0605 | 58041.2425 | 58715.4021 | 1170.308 | 1070.308 | 1265.352
9391 | 91.5777 | 1.1384 | 1.2466 | 0.99347 | 196.7654 | 48.8093 | 65733.0306 | 68003.7222 | 1356.0744 | 1256.0744 | 1265.527
9392 | 97.338 | 1.0298 | 0.88145 | 0.96976 | 213.7318 | 51.8512 | 49733.6553 | 48342.7007 | 962.854 | 862.854 | 1265.7042
9393 | 89.2656 | 0.99966 | 0.79913 | 0.95054 | 200.5514 | 47.4319 | 40462.9832 | 43084.1818 | 857.6836 | 757.6836 | 1265.7959
9394 | 81.0557 | 1.0281 | 1.016 | 1.0644 | 183.9766 | 48.4905 | 56664.5654 | 55311.6775 | 1102.2336 | 1002.2336 | 1265.9701
9395 | 97.4476 | 1.0122 | 1.0287 | 1.0758 | 188.9531 | 52.4416 | 63398.4737 | 63347.8106 | 1262.9562 | 1162.9562 | 1266.0223
9396 | 99.8482 | 0.98279 | 1.0099 | 0.91766 | 210.4925 | 47.8259 | 49226.6536 | 49239.4243 | 980.7885 | 880.7885 | 1266.1637
9397 | 116.5101 | 1.0179 | 1.1224 | 0.93212 | 186.7054 | 48.8272 | 57058.6378 | 56638.3519 | 1128.767 | 1028.767 | 1266.1955
9398 | 86.1527 | 0.98412 | 0.89813 | 0.95998 | 215.5646 | 48.6768 | 46311.3292 | 45758.2264 | 911.1645 | 811.1645 | 1266.2531
9399 | 82.3266 | 0.99185 | 1.006 | 1.083 | 191.2434 | 52.0243 | 61122.0078 | 59471.7902 | 1185.4358 | 1085.4358 | 1266.2684
9400 | 98.2149 | 0.99206 | 0.96322 | 1.0419 | 193.367 | 48.1056 | 53160.0909 | 53960.758 | 1075.2152 | 975.2152 | 1266.9162
9401 | 104.4321 | 1.0175 | 0.82483 | 0.98934 | 211.9499 | 48.6509 | 45082.2238 | 44284.5486 | 881.691 | 781.691 | 1267.1846
9402 | 94.8023 | 1.0367 | 0.8222 | 1.0024 | 195.2382 | 50.4273 | 46711.1434 | 46578.7912 | 927.5758 | 827.5758 | 1267.196
9403 | 92.6607 | 0.97185 | 1.0139 | 0.99104 | 192.6943 | 48.7973 | 53620.3555 | 54923.6579 | 1094.4732 | 994.4732 | 1267.2046
9404 | 99.427 | 0.93451 | 1.2541 | 0.95348 | 207.3271 | 47.8003 | 61805.7259 | 62039.6903 | 1236.7938 | 1136.7938 | 1267.2551
9405 | 112.8314 | 0.92364 | 0.84134 | 1.0183 | 204.8182 | 51.7459 | 49931.7531 | 49162.1253 | 979.2425 | 879.2425 | 1267.3177
9406 | 100.7803 | 0.95948 | 1.0121 | 0.96933 | 210.4202 | 47.1616 | 51041.1288 | 52355.422 | 1043.1084 | 943.1084 | 1267.3936
9407 | 97.9237 | 0.99764 | 0.89438 | 0.91404 | 194.5247 | 52.5814 | 48316.6905 | 49507.5003 | 986.15 | 886.15 | 1267.461
9408 | 103.7392 | 1.0427 | 0.89833 | 1.1711 | 197.5179 | 55.4634 | 64545.561 | 64578.5065 | 1287.5701 | 1187.5701 | 1267.62
9409 | 108.0319 | 0.98605 | 1.1505 | 1.0161 | 199.9806 | 47.3702 | 60622.4613 | 59275.5173 | 1181.5103 | 1081.5103 | 1267.7516
9410 | 90.3363 | 1.0409 | 1.1402 | 1.0695 | 194.7525 | 47.349 | 62387.8364 | 66650.4164 | 1329.0083 | 1229.0083 | 1267.759
9411 | 104.7286 | 0.94403 | 1.0015 | 1.1022 | 210.1659 | 50.3269 | 60740.4719 | 61456.4512 | 1225.129 | 1125.129 | 1267.7789
9412 | 98.8751 | 1.0101 | 0.96455 | 0.9715 | 188.3177 | 49.1939 | 51199.483 | 51812.7505 | 1032.255 | 932.255 | 1268.0216
9413 | 98.5579 | 1.0439 | 0.90142 | 0.95808 | 201.5896 | 46.4242 | 45071.8749 | 44249.9441 | 880.9989 | 780.9989 | 1268.0383
9414 | 110.7446 | 1.0053 | 1.1461 | 0.94532 | 203.7411 | 50.2675 | 60263.7193 | 59974.5528 | 1195.4911 | 1095.4911 | 1268.0828
9415 | 85.7304 | 1.0002 | 1.1045 | 1.0058 | 190.9595 | 51.0869 | 61325.3763 | 60618.5527 | 1208.3711 | 1108.3711 | 1268.2159
9416 | 87.4072 | 1.0556 | 1.0157 | 1.106 | 203.4129 | 51.7638 | 63127.3231 | 62221.2062 | 1240.4241 | 1140.4241 | 1268.3817
9417 | 90.5248 | 0.90175 | 1.1156 | 0.993 | 205.0002 | 49.0483 | 58545.0195 | 57649.2719 | 1148.9854 | 1048.9854 | 1268.6429
9418 | 84.7664 | 0.94469 | 1.127 | 1.0561 | 211.1396 | 48.9874 | 62437.7077 | 62135.3505 | 1238.707 | 1138.707 | 1268.8141
9419 | 111.4097 | 0.93791 | 0.68778 | 1.0785 | 201.6458 | 49.851 | 42388.0787 | 42381.9083 | 843.6382 | 743.6382 | 1268.8285
9420 | 101.8849 | 1.0442 | 0.93541 | 1.0071 | 189.3617 | 53.0575 | 55818.1416 | 55974.5391 | 1115.4908 | 1015.4908 | 1268.8601
9421 | 88.2005 | 0.98641 | 0.94088 | 0.96883 | 190.0835 | 46.951 | 47073.2061 | 48840.2906 | 972.8058 | 872.8058 | 1269.0192
9422 | 104.8232 | 0.95774 | 0.7972 | 1.0298 | 208.2034 | 51.7625 | 47898.4155 | 47865.4675 | 953.3094 | 853.3094 | 1269.1132
9423 | 96.6332 | 1.021 | 0.95563 | 1.0377 | 198.7689 | 44.9989 | 49262.7746 | 50555.5469 | 1007.1109 | 907.1109 | 1269.2813
9424 | 108.7938 | 1.0409 | 1.1117 | 1.1615 | 197.3387 | 49.1453 | 69224.0415 | 70184.6957 | 1399.6939 | 1299.6939 | 1269.3601
9425 | 98.4758 | 1.01 | 1.1058 | 1.0587 | 188.2655 | 48.6447 | 61974.2032 | 59151.5496 | 1179.031 | 1079.031 | 1269.489
9426 | 101.5163 | 0.9702 | 1.0014 | 0.99225 | 209.2888 | 50.3969 | 55251.2831 | 55445.0325 | 1104.9007 | 1004.9007 | 1269.8342
9427 | 82.3884 | 0.97694 | 0.8486 | 0.97285 | 195.3166 | 50.077 | 45567.8091 | 45142.9694 | 898.8594 | 798.8594 | 1269.9188
9428 | 87.8933 | 1.0502 | 1.0877 | 1.037 | 189.753 | 53.6925 | 65709.9325 | 65460.1209 | 1305.2024 | 1205.2024 | 1270.0214
9429 | 92.2316 | 0.93232 | 0.98524 | 1.047 | 183.1797 | 49.0876 | 55041.4921 | 55270.9914 | 1101.4198 | 1001.4198 | 1270.0408
9430 | 111.6772 | 1.1014 | 1.1876 | 1.0176 | 186.4454 | 50.3529 | 67231.7682 | 65857.8964 | 1313.1579 | 1213.1579 | 1270.2778
9431 | 98.0158 | 1.0793 | 1.2965 | 0.99213 | 197.1645 | 54.6731 | 76304.1219 | 76381.0083 | 1523.6202 | 1423.6202 | 1270.4212
9432 | 115.5287 | 0.9624 | 1.1289 | 1.0344 | 213.6349 | 47.8832 | 61455.1512 | 60454.305 | 1205.0861 | 1105.0861 | 1270.5567
9433 | 94.0943 | 1.0185 | 0.87551 | 1.0033 | 184.2348 | 48.7664 | 47692.0931 | 47937.9614 | 954.7592 | 854.7592 | 1270.6882
9434 | 94.9792 | 0.9958 | 1.0303 | 1.0185 | 189.3828 | 47.5018 | 54529.8707 | 54709.9023 | 1090.198 | 990.198 | 1270.7243
9435 | 103.3079 | 1.0637 | 0.76033 | 1.0984 | 205.169 | 51.849 | 49204.254 | 49595.4014 | 987.908 | 887.908 | 1271.0881
9436 | 101.9408 | 0.94792 | 0.82004 | 0.99955 | 207.5256 | 52.8754 | 48656.9415 | 49095.6438 | 977.9129 | 877.9129 | 1271.2556
9437 | 115.6436 | 1.017 | 1.0302 | 0.99791 | 205.7132 | 53.2637 | 61228.3723 | 64110.1613 | 1278.2032 | 1178.2032 | 1271.4426
9438 | 85.6824 | 0.99757 | 1.005 | 0.97483 | 214.986 | 48.7807 | 52173.2583 | 51864.2247 | 1033.2845 | 933.2845 | 1271.565
9439 | 84.9628 | 1.1183 | 1.0404 | 1.0921 | 175.0838 | 43.9611 | 54302.6229 | 53466.97 | 1065.3394 | 965.3394 | 1271.6269
9440 | 95.3815 | 0.9084 | 1.0971 | 0.94834 | 205.0825 | 51.492 | 58241.2183 | 58409.4176 | 1164.1884 | 1064.1884 | 1271.6989
9441 | 92.4629 | 0.97109 | 0.85658 | 1.0182 | 199.8509 | 46.2006 | 44643.9912 | 43753.7725 | 871.0754 | 771.0754 | 1271.8265
9442 | 95.7027 | 0.94654 | 1.2832 | 0.89018 | 212.398 | 48.0321 | 59430.5121 | 58187.8506 | 1159.757 | 1059.757 | 1271.9855
9443 | 92.9545 | 1.0094 | 0.93411 | 0.96417 | 195.4223 | 48.5663 | 48492.8813 | 49443.5804 | 984.8716 | 884.8716 | 1271.9914
9444 | 104.8329 | 0.99669 | 0.87855 | 0.95058 | 200.6162 | 48.1876 | 45478.4436 | 46148.4744 | 918.9695 | 818.9695 | 1272.2294
9445 | 108.3122 | 1.0708 | 0.868 | 1.0966 | 209.049 | 50.2525 | 53871.6555 | 55021.7365 | 1096.4347 | 996.4347 | 1272.2688
9446 | 100.983 | 0.99766 | 0.90778 | 0.96805 | 195.9689 | 53.7262 | 52821.5686 | 53705.8115 | 1070.1162 | 970.1162 | 1272.4283
9447 | 95.872 | 1.0323 | 1.0871 | 0.95193 | 205.2991 | 49.5537 | 56388.2062 | 56723.9749 | 1130.4795 | 1030.4795 | 1272.7412
9448 | 73.5666 | 0.96338 | 0.69357 | 1.0185 | 202.9676 | 49.0804 | 38353.6033 | 38616.1624 | 768.3232 | 668.3232 | 1273.0426
9449 | 77.1728 | 1.1228 | 1.156 | 1.0093 | 202.5964 | 52.4468 | 65936.6959 | 65861.7848 | 1313.2357 | 1213.2357 | 1273.0772
9450 | 79.8507 | 1.0016 | 1.2045 | 1.0054 | 204.0608 | 48.6418 | 62998.3879 | 62791.0425 | 1251.8209 | 1151.8209 | 1273.2339
9451 | 104.6133 | 0.9672 | 0.96419 | 1.0078 | 216.5021 | 51.7507 | 55740.2155 | 55323.076 | 1102.4615 | 1002.4615 | 1273.3917
9452 | 120.2256 | 0.99338 | 0.78721 | 0.93777 | 205.6431 | 49.5814 | 42729.3353 | 42806.114 | 852.1223 | 752.1223 | 1273.5293
9453 | 105.882 | 1.0194 | 1.0838 | 1.0753 | 194.7201 | 52.5382 | 67089.673 | 67240.3463 | 1340.8069 | 1240.8069 | 1273.5467
9454 | 99.0682 | 0.97232 | 0.88678 | 1.0159 | 201.9558 | 49.4148 | 49480.3777 | 50003.1668 | 996.0633 | 896.0633 | 1273.7379
9455 | 110.363 | 1.0128 | 1.1468 | 0.97217 | 220.6651 | 53.7761 | 66183.8326 | 66524.2519 | 1326.485 | 1226.485 | 1273.8198
9456 | 106.7537 | 1.0031 | 0.95713 | 0.95059 | 196.0405 | 53.8619 | 54968.8352 | 53848.0977 | 1072.962 | 972.962 | 1273.8499
9457 | 82.7396 | 1.0123 | 1.0915 | 0.94054 | 207.3961 | 54.0573 | 60230.206 | 59056.7008 | 1177.134 | 1077.134 | 1273.8729
9458 | 85.6863 | 1.0712 | 1.1338 | 1.0074 | 202.6963 | 49.2096 | 60925.4758 | 62948.7423 | 1254.9748 | 1154.9748 | 1274.0808
9459 | 102.3608 | 0.98428 | 1.023 | 1.0796 | 207.9411 | 53.069 | 64165.9563 | 65258.4223 | 1301.1684 | 1201.1684 | 1274.5884
9460 | 91.8566 | 1.0341 | 0.67441 | 1.007 | 181.364 | 50.6177 | 39365.2368 | 39048.8347 | 776.9767 | 676.9767 | 1274.8157
9461 | 95.1919 | 1.0283 | 1.2269 | 1.0278 | 222.8208 | 50.4422 | 68769.8634 | 67959.1496 | 1355.183 | 1255.183 | 1274.8716
9462 | 101.3843 | 0.92059 | 0.97986 | 0.98598 | 207.7 | 49.6538 | 52813.8526 | 54273.3922 | 1081.4678 | 981.4678 | 1275.1384
9463 | 110.5188 | 1.0401 | 1.1491 | 1.0876 | 189.7317 | 50.261 | 68779.6765 | 70183.7831 | 1399.6757 | 1299.6757 | 1275.3657
9464 | 90.8862 | 1.0364 | 1.0416 | 0.94363 | 182.6512 | 47.8823 | 51754.7516 | 51561.1438 | 1027.2229 | 927.2229 | 1275.3786
9465 | 93.5038 | 1.0241 | 0.97879 | 1.006 | 206.9879 | 53.3351 | 57830.35 | 56943.2258 | 1134.8645 | 1034.8645 | 1275.3999
9466 | 98.0702 | 0.98977 | 1.0585 | 0.96659 | 196.3122 | 50.6257 | 56909.6979 | 56853.0438 | 1133.0609 | 1033.0609 | 1275.5857
9467 | 90.2759 | 0.90815 | 1.0601 | 0.99465 | 197.4098 | 50.88 | 58015.4018 | 58034.9863 | 1156.6997 | 1056.6997 | 1275.8372
9468 | 114.4963 | 0.94683 | 0.97697 | 1.0097 | 210.6282 | 48.6566 | 53482.5364 | 56245.3509 | 1120.907 | 1020.907 | 1276.1953
9469 | 105.8297 | 0.9119 | 0.84526 | 0.96434 | 192.1568 | 51.3605 | 47013.5074 | 47900.0275 | 954.0005 | 854.0005 | 1276.3191
9470 | 108.9187 | 0.94323 | 1.2246 | 0.99942 | 202.989 | 51.4021 | 68393.6768 | 67609.6588 | 1348.1932 | 1248.1932 | 1276.5086
9471 | 102.0547 | 1.0221 | 0.93958 | 0.97489 | 215.7454 | 47.1101 | 48281.8632 | 47371.9842 | 943.4397 | 843.4397 | 1276.5604
9472 | 105.1475 | 0.98624 | 0.87763 | 0.99503 | 206.4087 | 51.118 | 50147.0682 | 50329.2363 | 1002.5847 | 902.5847 | 1277.1695
9473 | 111.8815 | 1.0333 | 0.87968 | 1.0258 | 199.6232 | 51.6495 | 52777.373 | 50979.4442 | 1015.5889 | 915.5889 | 1277.1943
9474 | 92.6963 | 1.071 | 1.0086 | 1.114 | 191.3161 | 44.9946 | 55215.9315 | 55046.4099 | 1096.9282 | 996.9282 | 1277.2365
9475 | 82.1086 | 1.0759 | 0.84878 | 1.0356 | 196.4238 | 49.9176 | 48482.8491 | 45901.903 | 914.0381 | 814.0381 | 1277.3458
9476 | 93.421 | 0.98452 | 0.97842 | 0.88322 | 212.4546 | 48.5323 | 46615.9343 | 47030.8369 | 936.6167 | 836.6167 | 1277.3922
9477 | 95.2478 | 0.9802 | 0.87416 | 0.95092 | 199.3544 | 48.3078 | 44865.8058 | 46495.1928 | 925.9039 | 825.9039 | 1277.4852
9478 | 95.2694 | 1.022 | 0.90698 | 1.0425 | 194.735 | 52.112 | 54541.232 | 52868.7009 | 1053.374 | 953.374 | 1277.4858
9479 | 115.3861 | 1.057 | 0.98245 | 0.96339 | 193.3888 | 51.9141 | 55660.2849 | 56639.2291 | 1128.7846 | 1028.7846 | 1277.6223
9480 | 99.8438 | 1.0591 | 1.0135 | 0.94067 | 191.9886 | 49.0179 | 52108.7423 | 49839.9725 | 992.7995 | 892.7995 | 1277.7479
9481 | 93.0555 | 0.99446 | 1.0434 | 1.0104 | 203.3727 | 53.1817 | 61188.7666 | 61237.1932 | 1220.7439 | 1120.7439 | 1277.7613
9482 | 108.3298 | 1.0235 | 0.77669 | 1.0159 | 203.7812 | 48.9024 | 44211.0315 | 44590.7284 | 887.8146 | 787.8146 | 1277.9754
9483 | 100.2662 | 1.0372 | 1.142 | 0.97669 | 197.1117 | 49.8848 | 61026.0604 | 61723.3658 | 1230.4673 | 1130.4673 | 1278.0861
9484 | 100.7171 | 0.93329 | 0.93516 | 0.95916 | 200.5899 | 52.431 | 52158.0163 | 50244.6038 | 1000.8921 | 900.8921 | 1278.0993
9485 | 90.5063 | 0.94553 | 1.0641 | 0.95724 | 188.7174 | 48.9108 | 54196.2513 | 54544.3951 | 1086.8879 | 986.8879 | 1278.0994
9486 | 109.2252 | 1.0905 | 0.99199 | 0.82645 | 183.0847 | 49.897 | 47033.647 | 45978.0885 | 915.5618 | 815.5618 | 1278.4317
9487 | 96.8667 | 0.98133 | 1.2337 | 0.98347 | 197.3403 | 49.3886 | 64813.5794 | 62712.1163 | 1250.2423 | 1150.2423 | 1278.5173
9488 | 101.8098 | 0.96823 | 1.1437 | 0.97412 | 204.539 | 47.7092 | 58059.3187 | 57290.5298 | 1141.8106 | 1041.8106 | 1278.6778
9489 | 88.3583 | 1.0491 | 1.0714 | 1.0071 | 186.2474 | 54.4188 | 63951.7448 | 61333.77 | 1222.6754 | 1122.6754 | 1278.7897
9490 | 83.8688 | 0.93884 | 0.83344 | 0.99867 | 192.0932 | 46.5065 | 42562.9951 | 42172.347 | 839.4469 | 739.4469 | 1278.8631
9491 | 98.8838 | 0.96051 | 0.77979 | 1.0033 | 194.9709 | 49.0378 | 43219.696 | 44394.0194 | 883.8804 | 783.8804 | 1278.8878
9492 | 102.7624 | 0.98524 | 0.96683 | 0.99377 | 197.3847 | 50.869 | 54222.9251 | 54738.7409 | 1090.7748 | 990.7748 | 1278.9697
9493 | 95.9699 | 1.0021 | 1.061 | 1.0148 | 212.1888 | 51.7619 | 60922.0964 | 60992.5623 | 1215.8512 | 1115.8512 | 1279.0345
9494 | 98.8165 | 0.94774 | 0.86207 | 1.0913 | 215.1949 | 49.9717 | 51909.3979 | 52382.312 | 1043.6462 | 943.6462 | 1279.1336
9495 | 107.0542 | 0.91716 | 0.97822 | 0.96769 | 208.1552 | 50.7134 | 53193.6126 | 53743.4947 | 1070.8699 | 970.8699 | 1279.6136
9496 | 83.1933 | 1.0213 | 1.0842 | 0.97376 | 194.8129 | 47.8923 | 54827.6822 | 53372.5927 | 1063.4519 | 963.4519 | 1279.7327
9497 | 115.3142 | 1.0075 | 0.90312 | 1.0106 | 196.6829 | 49.1114 | 50724.0222 | 50157.736 | 999.1547 | 899.1547 | 1279.748
9498 | 92.4217 | 0.99921 | 0.89005 | 1.0153 | 209.1982 | 51.8952 | 51900.0142 | 52967.4484 | 1055.349 | 955.349 | 1279.8109
9499 | 105.2886 | 1.0229 | 0.81409 | 1.0791 | 201.2526 | 52.8897 | 52358.6637 | 51523.2651 | 1026.4653 | 926.4653 | 1279.9162
9500 | 93.5835 | 0.93039 | 0.95518 | 0.98323 | 191.5802 | 57.2792 | 58973.2575 | 56970.5442 | 1135.4109 | 1035.4109 | 1279.925
9501 | 93.987 | 1.0218 | 1.2108 | 0.96668 | 195.9834 | 47.0863 | 59831.0077 | 59839.9514 | 1192.799 | 1092.799 | 1280.0291
9502 | 97.6695 | 0.94398 | 0.93299 | 0.99403 | 200.9728 | 50.6115 | 51805.5641 | 51469.6253 | 1025.3925 | 925.3925 | 1280.1583
9503 | 105.6552 | 0.96577 | 1.1819 | 0.98466 | 195.5183 | 47.6581 | 60521.175 | 63565.2091 | 1267.3042 | 1167.3042 | 1280.1745
9504 | 72.6101 | 0.93229 | 0.88508 | 1.056 | 195.9188 | 49.3674 | 49680.6293 | 49153.008 | 979.0602 | 879.0602 | 1280.2844
9505 | 89.2565 | 1.0416 | 0.88275 | 0.9961 | 196.3286 | 48.6114 | 47460.1078 | 46247.9227 | 920.9585 | 820.9585 | 1280.3209
9506 | 94.5051 | 0.9925 | 0.98375 | 0.99865 | 193.5247 | 50.6097 | 54660.9292 | 54943.2879 | 1094.8658 | 994.8658 | 1280.3857
9507 | 99.6846 | 1.0084 | 1.0896 | 1.0658 | 206.9083 | 47.2187 | 59787.9465 | 60170.0504 | 1199.401 | 1099.401 | 1280.4843
9508 | 77.3002 | 1.0235 | 0.76642 | 0.97108 | 201.6814 | 48.8869 | 40453.9747 | 39675.9501 | 789.519 | 689.519 | 1280.5067
9509 | 104.3826 | 1.0084 | 0.88275 | 0.94233 | 226.107 | 53.3338 | 50205.0484 | 51096.931 | 1017.9386 | 917.9386 | 1280.6465
9510 | 103.3301 | 0.95967 | 1.0769 | 1.0371 | 199.873 | 52.0217 | 63454.2999 | 61959.8997 | 1235.198 | 1135.198 | 1280.7394
9511 | 93.9135 | 0.96837 | 1.3941 | 1.1391 | 209.7504 | 47.129 | 79339.0604 | 78921.8086 | 1574.4362 | 1474.4362 | 1280.8109
9512 | 94.0485 | 1.0077 | 0.74871 | 1.0301 | 187.4491 | 53.4754 | 46498.8878 | 47061.0848 | 937.2217 | 837.2217 | 1281.4577
9513 | 78.2042 | 1.0334 | 0.89627 | 0.95821 | 173.5934 | 48.0121 | 45287.0667 | 46563.2329 | 927.2647 | 827.2647 | 1281.5772
9514 | 89.5476 | 1.0419 | 0.96795 | 0.97883 | 210.737 | 43.8042 | 45800.9199 | 46573.1109 | 927.4622 | 827.4622 | 1282.2592
9515 | 115.3798 | 0.95252 | 1.2651 | 1.0757 | 183.4334 | 48.3164 | 71246.5882 | 73785.9688 | 1471.7194 | 1371.7194 | 1282.5022
9516 | 110.9371 | 0.9979 | 0.7902 | 0.97628 | 206.4724 | 48.5755 | 43058.0915 | 41433.4816 | 824.6696 | 724.6696 | 1282.5257
9517 | 92.9655 | 0.97386 | 1.0932 | 1.0305 | 200.1681 | 49.3794 | 60296.6035 | 61092.0244 | 1217.8405 | 1117.8405 | 1282.7298
9518 | 104.7574 | 0.98073 | 1.0262 | 0.93681 | 186.2899 | 51.6224 | 55115.6548 | 53477.8789 | 1065.5576 | 965.5576 | 1282.8201
9519 | 103.2261 | 1.0041 | 1.2275 | 0.9887 | 185.6095 | 51.0671 | 67453.0043 | 67973.9058 | 1355.4781 | 1255.4781 | 1283.0318
9520 | 92.1418 | 1.0132 | 0.87815 | 0.98794 | 212.8522 | 46.9404 | 45318.6998 | 47361.5051 | 943.2301 | 843.2301 | 1283.1117
9521 | 84.1272 | 1.0243 | 1.1949 | 0.9974 | 189.1134 | 52.8828 | 67769.3576 | 69380.1824 | 1383.6036 | 1283.6036 | 1283.6036
9522 | 110.8834 | 1.0329 | 0.91445 | 0.97921 | 215.1812 | 51.4921 | 52220.5555 | 52753.3135 | 1051.0663 | 951.0663 | 1283.61
9523 | 98.6623 | 0.97583 | 0.80339 | 0.86597 | 214.6502 | 50.4231 | 40148.8786 | 39324.6303 | 782.4926 | 682.4926 | 1283.6168
9524 | 101.8585 | 1.0377 | 1.1814 | 1.0314 | 199.2737 | 48.5243 | 64455.4988 | 65300.3541 | 1302.0071 | 1202.0071 | 1283.6862
9525 | 94.0633 | 1.0342 | 1.0936 | 1.1189 | 193.3032 | 52.589 | 69657.4052 | 69224.2155 | 1380.4843 | 1280.4843 | 1283.7703
9526 | 88.1585 | 0.99231 | 1.0623 | 0.99157 | 209.5775 | 50.6434 | 57982.983 | 56414.5693 | 1124.2914 | 1024.2914 | 1283.9124
9527 | 93.4591 | 0.95803 | 0.77404 | 0.99973 | 215.2702 | 50.8088 | 44081.7359 | 42832.8676 | 852.6574 | 752.6574 | 1283.917
9528 | 89.2918 | 0.95863 | 1.0158 | 0.98439 | 211.205 | 48.4197 | 52774.4707 | 51854.1371 | 1033.0827 | 933.0827 | 1284.0021
9529 | 91.1344 | 1.0759 | 1.0258 | 0.96348 | 219.4058 | 46.9418 | 51215.6019 | 49779.9683 | 991.5994 | 891.5994 | 1284.0789
9530 | 77.6207 | 0.97275 | 0.87524 | 0.98232 | 194.8679 | 52.2885 | 49098.8819 | 47993.3654 | 955.8673 | 855.8673 | 1284.1608
9531 | 98.7105 | 1.0084 | 1.1371 | 1.0076 | 222.3743 | 50.2721 | 62824.3669 | 60927.7758 | 1214.5555 | 1114.5555 | 1284.5906
9532 | 85.2493 | 0.96613 | 0.96213 | 1.0066 | 213.4954 | 52.1389 | 55005.6781 | 56097.258 | 1117.9452 | 1017.9452 | 1284.681
9533 | 94.0551 | 0.96148 | 0.91492 | 0.95762 | 201.1221 | 49.6329 | 48174.8549 | 46346.6583 | 922.9332 | 822.9332 | 1285.2052
9534 | 116.5756 | 0.91346 | 1.1542 | 0.99623 | 201.9156 | 47.5354 | 59920.0616 | 59712.2108 | 1190.2442 | 1090.2442 | 1285.3728
9535 | 106.7594 | 0.96971 | 1.024 | 1.0265 | 221.8062 | 53.8604 | 62411.8498 | 63060.2501 | 1257.205 | 1157.205 | 1285.3909
9536 | 95.1804 | 1.0131 | 1.2205 | 0.99835 | 197.5489 | 49.1207 | 64785.1798 | 66154.5204 | 1319.0904 | 1219.0904 | 1285.769
9537 | 96.7223 | 0.87443 | 1.0898 | 1.1069 | 190.2876 | 52.0621 | 67398.1861 | 69848.8222 | 1392.9764 | 1292.9764 | 1285.9294
9538 | 114.5008 | 0.97626 | 0.81995 | 0.89896 | 192.6796 | 53.8861 | 45935.904 | 46564.6447 | 927.2929 | 827.2929 | 1286.0568
9539 | 82.8648 | 1.099 | 0.9334 | 1.0603 | 202.8961 | 46.1443 | 50072.7907 | 50075.9316 | 997.5186 | 897.5186 | 1286.1928
9540 | 105.5226 | 0.99205 | 0.9505 | 0.94511 | 200.1775 | 48.3323 | 48677.7638 | 45955.0859 | 915.1017 | 815.1017 | 1286.2768
9541 | 94.7511 | 1.0063 | 1.1062 | 1.0917 | 190.7786 | 48.775 | 63744.6139 | 62293.1103 | 1241.8622 | 1141.8622 | 1286.8281
9542 | 104.6481 | 0.99011 | 1.1993 | 0.94962 | 205.9223 | 54.1037 | 67427.4527 | 68695.9398 | 1369.9188 | 1269.9188 | 1286.852
9543 | 87.8493 | 1.0422 | 0.90269 | 1.0292 | 222.0566 | 52.6349 | 53940.4907 | 53225.2614 | 1060.5052 | 960.5052 | 1287.1083
9544 | 85.3824 | 0.98471 | 0.71516 | 1.0149 | 187.9365 | 50.3693 | 40982.3245 | 40585.0725 | 807.7015 | 707.7015 | 1287.1488
9545 | 98.4395 | 1.0155 | 1.0881 | 1.0605 | 187.4186 | 47.9733 | 60337.1803 | 59976.9518 | 1195.539 | 1095.539 | 1287.1666
9546 | 92.7559 | 1.0581 | 1.0791 | 1.0636 | 190.6774 | 48.1538 | 60180.4729 | 60328.7731 | 1202.5755 | 1102.5755 | 1287.4461
9547 | 82.0248 | 0.96427 | 0.84118 | 1.0407 | 201.9454 | 47.5855 | 45623.0382 | 45584.8107 | 907.6962 | 807.6962 | 1287.6211
9548 | 116.9718 | 0.95537 | 0.81379 | 1.0067 | 193.1387 | 48.8357 | 45658.9809 | 45283.8647 | 901.6773 | 801.6773 | 1287.8627
9549 | 107.1777 | 1.0634 | 1.0745 | 1.0418 | 197.5751 | 48.8912 | 60497.6347 | 59646.4447 | 1188.9289 | 1088.9289 | 1287.8681
9550 | 68.5097 | 0.98733 | 1.0605 | 0.98815 | 196.5308 | 54.4223 | 60905.8253 | 59633.7768 | 1188.6755 | 1088.6755 | 1288.1049
9551 | 110.0441 | 0.94794 | 0.86084 | 0.93817 | 195.5737 | 53.0336 | 48558.2552 | 48230.3543 | 960.6071 | 860.6071 | 1288.3525
9552 | 88.1025 | 1.0209 | 1.1809 | 0.95758 | 192.9202 | 47.3879 | 58039.9912 | 56485.1342 | 1125.7027 | 1025.7027 | 1288.6855
9553 | 98.2527 | 1.0113 | 1.2691 | 1.0283 | 209.2769 | 47.9619 | 67560.8499 | 68314.7248 | 1362.2945 | 1262.2945 | 1288.712
9554 | 98.1014 | 0.98054 | 0.65439 | 0.86234 | 196.509 | 49.3514 | 32793.0192 | 30492.9546 | 605.8591 | 505.8591 | 1288.916
9555 | 98.7132 | 0.94639 | 0.8858 | 0.91675 | 197.4792 | 46.8326 | 42603.5455 | 42015.0517 | 836.301 | 736.301 | 1288.9352
9556 | 90.3596 | 0.98826 | 0.8922 | 1.0189 | 181.3153 | 48.1737 | 48274.9235 | 47526.1345 | 946.5227 | 846.5227 | 1289.2232
9557 | 101.8355 | 0.96982 | 0.86531 | 1.042 | 201.4045 | 47.9336 | 48156.0879 | 48113.8711 | 958.2774 | 858.2774 | 1289.5317
9558 | 112.2283 | 1.0032 | 1.0123 | 0.98147 | 206.4357 | 52.292 | 58049.928 | 57783.6398 | 1151.6728 | 1051.6728 | 1289.7009
9559 | 97.6098 | 0.98988 | 1.026 | 1.0322 | 208.842 | 52.6876 | 61102.0563 | 61078.0643 | 1217.5613 | 1117.5613 | 1289.7151
9560 | 84.9094 | 1.0514 | 1.2546 | 1.0754 | 190.667 | 50.0099 | 72129.1852 | 72104.9781 | 1438.0996 | 1338.0996 | 1289.9218
9561 | 81.3592 | 0.99749 | 0.82964 | 0.96717 | 212.6273 | 46.7982 | 41561.4799 | 42495.0603 | 845.9012 | 745.9012 | 1290.5771
9562 | 114.5963 | 1.0442 | 0.78815 | 1.0043 | 199.6106 | 52.5304 | 48067.1339 | 46554.3731 | 927.0875 | 827.0875 | 1290.7157
9563 | 89.3718 | 1.0429 | 1.1502 | 1.0341 | 192.9436 | 49.2686 | 63389.5399 | 63536.0338 | 1266.7207 | 1166.7207 | 1290.9349
9564 | 87.5691 | 0.94278 | 0.95623 | 0.99801 | 191.8429 | 49.6599 | 51683.049 | 52303.8582 | 1042.0772 | 942.0772 | 1290.9526
9565 | 96.9523 | 0.90681 | 1.1179 | 0.97055 | 185.6643 | 51.7123 | 60838.3563 | 63447.5573 | 1264.9511 | 1164.9511 | 1291.0147
9566 | 108.4093 | 1.0131 | 0.87485 | 0.97499 | 195.7622 | 49.233 | 47597.7425 | 45728.9796 | 910.5796 | 810.5796 | 1291.0601
9567 | 90.7971 | 0.98945 | 1.0011 | 1.0262 | 182.4427 | 48.434 | 54291.5916 | 52584.8852 | 1047.6977 | 947.6977 | 1291.2081
9568 | 81.5644 | 1.0034 | 0.9977 | 0.99258 | 214.9207 | 49.169 | 52930.8384 | 53163.4593 | 1059.2692 | 959.2692 | 1292.1381
9569 | 95.7644 | 1.0693 | 1.1409 | 1.0003 | 179.3945 | 49.8172 | 62137.0427 | 61762.6149 | 1231.2523 | 1131.2523 | 1292.1427
9570 | 88.8375 | 0.98749 | 1.1802 | 0.90801 | 224.0912 | 49.2034 | 57270.6792 | 57756.2728 | 1151.1255 | 1051.1255 | 1292.4028
9571 | 98.1925 | 0.99258 | 1.2104 | 0.97413 | 211.6985 | 48.9547 | 62703.7191 | 60521.7494 | 1206.435 | 1106.435 | 1292.5
9572 | 112.694 | 0.96066 | 0.90956 | 0.95815 | 204.7556 | 50.5711 | 49751.6277 | 50056.7819 | 997.1356 | 897.1356 | 1292.7901
9573 | 108.4213 | 1.0194 | 1.2722 | 1.0116 | 199.2107 | 55.0637 | 77147.3456 | 75345.0227 | 1502.9005 | 1402.9005 | 1292.8349
9574 | 102.7736 | 1.0335 | 1.0466 | 0.98325 | 225.1445 | 52.6918 | 60043.6571 | 61262.1631 | 1221.2433 | 1121.2433 | 1292.8635
9575 | 102.3432 | 1.0454 | 0.91025 | 0.96718 | 209.8549 | 49.254 | 48841.254 | 47819.3732 | 952.3875 | 852.3875 | 1292.9764
9576 | 96.7117 | 0.94158 | 0.96055 | 0.90297 | 194.8274 | 55.1827 | 53082.7771 | 51892.5111 | 1033.8502 | 933.8502 | 1292.979
9577 | 94.5144 | 0.9281 | 0.85719 | 0.97641 | 209.3479 | 53.1694 | 49374.7101 | 48180.4882 | 959.6098 | 859.6098 | 1293.0311
9578 | 104.4457 | 0.94131 | 1.056 | 1.0489 | 201.7138 | 51.3601 | 62140.683 | 62110.6559 | 1238.2131 | 1138.2131 | 1293.3582
9579 | 103.2816 | 0.95984 | 1.0465 | 1.0026 | 181.6518 | 51.8199 | 59690.1337 | 59431.0253 | 1184.6205 | 1084.6205 | 1293.3605
9580 | 101.2675 | 1.0839 | 0.8779 | 1.0377 | 205.205 | 49.7682 | 51006.1656 | 52160.3809 | 1039.2076 | 939.2076 | 1293.5789
9581 | 120.1012 | 0.98792 | 0.83721 | 1.0184 | 196.6726 | 49.7693 | 48537.767 | 49101.586 | 978.0317 | 878.0317 | 1293.5794
9582 | 83.8278 | 0.92126 | 1.0088 | 1.0077 | 195.0788 | 46.5029 | 51058.0587 | 53658.66 | 1069.1732 | 969.1732 | 1293.8459
9583 | 97.9601 | 1.0858 | 0.93045 | 0.90583 | 188.4012 | 49.1966 | 46885.8894 | 46524.938 | 926.4988 | 826.4988 | 1294.0605
9584 | 107.81 | 1.0405 | 1.2308 | 0.98896 | 213.2475 | 51.4914 | 68664.7513 | 67111.1068 | 1338.2221 | 1238.2221 | 1294.1157
9585 | 80.0278 | 1.0531 | 1.142 | 1.0683 | 192.3581 | 48.9554 | 64042.2888 | 64677.2141 | 1289.5443 | 1189.5443 | 1294.2107
9586 | 84.3048 | 1.1202 | 0.97484 | 1.05 | 180.9568 | 48.6133 | 54532.9436 | 55973.1801 | 1115.4636 | 1015.4636 | 1294.3268
9587 | 95.1875 | 1.0454 | 1.0736 | 1.0108 | 210.9308 | 47.5392 | 56530.65 | 55099.096 | 1097.9819 | 997.9819 | 1294.3727
9588 | 92.9675 | 0.94675 | 0.78607 | 1.0123 | 193.8023 | 47.2513 | 41952.7431 | 41103.9166 | 818.0783 | 718.0783 | 1294.5664
9589 | 99.3636 | 1.0088 | 1.0434 | 0.907 | 193.9255 | 45.5693 | 47885.948 | 47835.6011 | 952.712 | 852.712 | 1294.5685
9590 | 99.5009 | 1.0225 | 1.2339 | 1.002 | 198.784 | 52.6208 | 70609.3264 | 70444.9512 | 1404.899 | 1304.899 | 1294.606
9591 | 100.5758 | 0.95766 | 1.1494 | 1.018 | 198.9113 | 49.2544 | 62573.8601 | 63149.1079 | 1258.9822 | 1158.9822 | 1294.6125
9592 | 97.4158 | 0.9764 | 0.89611 | 1.0033 | 189.968 | 51.8395 | 51726.1489 | 52022.2053 | 1036.4441 | 936.4441 | 1294.7651
9593 | 101.8998 | 1.0176 | 0.94438 | 1.0018 | 218.2243 | 47.8124 | 50412.9421 | 51313.3351 | 1022.2667 | 922.2667 | 1294.8333
9594 | 109.3988 | 0.94273 | 0.94762 | 1.0397 | 199.0567 | 50.18 | 54813.392 | 53913.3088 | 1074.2662 | 974.2662 | 1295.1467
9595 | 101.4783 | 1.0233 | 0.9664 | 1.0794 | 188.8768 | 53.9025 | 62011.457 | 60821.2673 | 1212.4253 | 1112.4253 | 1295.309
9596 | 96.2453 | 0.96175 | 0.9362 | 1.0303 | 197.3795 | 48.5919 | 51567.4884 | 50909.7035 | 1014.1941 | 914.1941 | 1295.5038
9597 | 101.0439 | 0.99967 | 1.3525 | 1.0458 | 214.0539 | 53.2953 | 80978.5746 | 82802.9865 | 1652.0597 | 1552.0597 | 1295.5172
9598 | 112.5437 | 1.013 | 1.1078 | 0.96148 | 223.1885 | 51.5923 | 61060.023 | 63316.197 | 1262.3239 | 1162.3239 | 1295.5491
9599 | 111.9175 | 1.0259 | 0.97272 | 1.0188 | 216.1219 | 50.4668 | 56024.9335 | 55563.395 | 1107.2679 | 1007.2679 | 1295.763
9600 | 112.7805 | 0.97224 | 0.79295 | 1.1001 | 189.9353 | 50.7601 | 50035.6268 | 49999.6513 | 995.993 | 895.993 | 1296.4359
9601 | 109.408 | 0.92606 | 0.76435 | 0.91832 | 203.1866 | 48.7607 | 39369.737 | 39626.4663 | 788.5293 | 688.5293 | 1296.6993
9602 | 84.4197 | 0.99091 | 0.93434 | 0.91061 | 198.8454 | 49.6079 | 46556.0724 | 48547.246 | 966.9449 | 866.9449 | 1297.9856
9603 | 102.316 | 1.1037 | 1.2408 | 1.0166 | 199.6654 | 53.3575 | 73529.0787 | 73736.9791 | 1470.7396 | 1370.7396 | 1298.0073
9604 | 100.4822 | 1.0307 | 1.0702 | 0.99212 | 197.8385 | 48.5743 | 56805.3664 | 58022.3642 | 1156.4473 | 1056.4473 | 1298.4466
9605 | 101.3144 | 1.0454 | 0.77638 | 1.0094 | 193.9962 | 47.1923 | 42176.6264 | 42979.4715 | 855.5894 | 755.5894 | 1298.4908
9606 | 98.4797 | 0.99938 | 1.3648 | 0.98369 | 190.9024 | 48.8807 | 70623.9171 | 70576.5672 | 1407.5313 | 1307.5313 | 1298.5106
9607 | 87.1306 | 1.0211 | 1.099 | 0.97699 | 216.5332 | 50.5768 | 59020.0805 | 58321.7381 | 1162.4348 | 1062.4348 | 1298.7352
9608 | 94.9112 | 1.1086 | 1.1765 | 0.97702 | 203.0075 | 48.1822 | 60655.3209 | 59884.0858 | 1193.6817 | 1093.6817 | 1298.8239
9609 | 87.1467 | 1.006 | 1.0168 | 1.0181 | 197.9288 | 52.0961 | 58692.3355 | 60751.9294 | 1211.0386 | 1111.0386 | 1298.9809
9610 | 106.6594 | 0.99397 | 1.0093 | 0.97407 | 210.5794 | 48.5059 | 53038.7799 | 53636.7626 | 1068.7353 | 968.7353 | 1299.6757
9611 | 100.3599 | 0.95552 | 0.81432 | 1.1162 | 203.6392 | 49.4479 | 49891.8235 | 48554.9915 | 967.0998 | 867.0998 | 1299.6939
9612 | 94.3762 | 1.0237 | 0.97762 | 0.89775 | 188.6699 | 53.3979 | 52212.6068 | 52347.383 | 1042.9477 | 942.9477 | 1300.1128
9613 | 108.3777 | 1.0444 | 0.73654 | 0.938 | 202.6936 | 51.3614 | 41500.4967 | 42134.6074 | 838.6921 | 738.6921 | 1300.9413
9614 | 94.9773 | 0.93716 | 0.96969 | 0.98017 | 202.8034 | 54.2361 | 56579.6739 | 56029.6719 | 1116.5934 | 1016.5934 | 1301.376
9615 | 104.8308 | 1.0374 | 1.005 | 1.0085 | 207.1964 | 53.7808 | 60564.2485 | 60263.2776 | 1201.2656 | 1101.2656 | 1301.4963
9616 | 98.4435 | 1.0023 | 1.1197 | 1.0165 | 205.9675 | 50.3358 | 62462.4303 | 61851.3173 | 1233.0263 | 1133.0263 | 1301.6444
9617 | 102.3252 | 1.0231 | 1.0283 | 1.0085 | 206.0955 | 48.7217 | 55833.7829 | 54885.313 | 1093.7063 | 993.7063 | 1301.9896
9618 | 106.6798 | 0.97072 | 1.1526 | 1.0117 | 211.2159 | 49.1616 | 62628.5119 | 61079.351 | 1217.587 | 1117.587 | 1302.0935
9619 | 92.5441 | 0.96276 | 0.93918 | 1.0242 | 205.7757 | 52.8766 | 55776.7191 | 56240.5736 | 1120.8115 | 1020.8115 | 1302.5439
9620 | 87.078 | 1.1257 | 1.225 | 0.99943 | 215.6368 | 50.2958 | 66722.0792 | 66973.8076 | 1335.4762 | 1235.4762 | 1302.9102
9621 | 98.3585 | 1.1109 | 1.0845 | 0.89624 | 193.9503 | 49.2018 | 53390.7446 | 56321.6437 | 1122.4329 | 1022.4329 | 1303.0788
9622 | 97.9431 | 1.0615 | 1.039 | 1.004 | 200.8246 | 47.1081 | 54241.7675 | 56042.7287 | 1116.8546 | 1016.8546 | 1303.6594
9623 | 95.1885 | 1.0399 | 0.86395 | 0.92753 | 182.3697 | 50.0626 | 45254.8706 | 44381.9468 | 883.6389 | 783.6389 | 1304.2019
9624 | 109.1703 | 1.063 | 0.97946 | 0.93579 | 195.172 | 48.5428 | 50321.2815 | 50666.2827 | 1009.3257 | 909.3257 | 1304.2465
9625 | 90.415 | 0.98419 | 0.87499 | 0.9612 | 197.2222 | 46.457 | 43403.0899 | 43691.6299 | 869.8326 | 769.8326 | 1304.3204
9626 | 112.2544 | 0.94036 | 1.0997 | 1.0315 | 185.4179 | 46.4655 | 57796.0226 | 58724.0499 | 1170.481 | 1070.481 | 1304.3353
9627 | 105.0018 | 0.95427 | 1.1174 | 0.90281 | 202.9192 | 46.7611 | 52061.8069 | 52521.6709 | 1046.4334 | 946.4334 | 1304.5234
9628 | 106.1205 | 0.99262 | 1.0932 | 0.9704 | 218.0927 | 47.7682 | 55922.6132 | 58341.4024 | 1162.828 | 1062.828 | 1304.5792
9629 | 99.628 | 1.01 | 0.87444 | 1.0251 | 191.4463 | 48.1381 | 48187.7071 | 46581.5824 | 927.6316 | 827.6316 | 1304.6455
9630 | 95.9606 | 0.96215 | 1.0418 | 0.96294 | 198.9328 | 46.6982 | 51356.0701 | 51246.4061 | 1020.9281 | 920.9281 | 1304.7313
9631 | 96.8913 | 1.0137 | 0.83828 | 0.97508 | 193.8019 | 46.5588 | 42823.5459 | 43326.8776 | 862.5376 | 762.5376 | 1304.8738
9632 | 110.9007 | 1.1005 | 1.0363 | 0.95118 | 198.5749 | 48.8661 | 54329.8081 | 52036.7533 | 1036.7351 | 936.7351 | 1304.899
9633 | 102.5873 | 0.98847 | 1.0632 | 0.99586 | 196.5257 | 51.7773 | 60269.6359 | 60453.3531 | 1205.0671 | 1105.0671 | 1304.9831
9634 | 98.732 | 0.97903 | 0.92275 | 1.0371 | 207.9731 | 49.3796 | 52235.3971 | 53228.8596 | 1060.5772 | 960.5772 | 1305.4316
9635 | 87.1894 | 1.0942 | 1.2771 | 1.0514 | 182.2906 | 52.1041 | 75118.7948 | 74989.3964 | 1495.7879 | 1395.7879 | 1305.5986
9636 | 90.7141 | 1.1196 | 1.0428 | 1.0366 | 198.0452 | 51.4536 | 61042.9215 | 60559.1407 | 1207.1828 | 1107.1828 | 1305.6107
9637 | 105.7411 | 0.96518 | 0.56156 | 1.0158 | 200.2936 | 47.5752 | 32194.4165 | 32465.4987 | 645.31 | 545.31 | 1305.8403
9638 | 103.2489 | 1.0301 | 1.0856 | 0.99551 | 202.7836 | 47.1815 | 56212.4818 | 55275.2662 | 1101.5053 | 1001.5053 | 1306.3163
9639 | 92.4707 | 1.0042 | 1.1762 | 0.9755 | 217.5034 | 52.5071 | 65339.0819 | 64311.9633 | 1282.2393 | 1182.2393 | 1306.7418
9640 | 98.892 | 1.0043 | 1.2371 | 0.99593 | 207.4686 | 44.8368 | 59901.2052 | 59377.603 | 1183.5521 | 1083.5521 | 1306.8403
9641 | 112.9068 | 0.9542 | 1.0006 | 1.0252 | 170.1604 | 49.2831 | 56037.2581 | 56718.12 | 1130.3624 | 1030.3624 | 1307.4941
9642 | 94.9988 | 1.1307 | 0.99269 | 0.93828 | 182.9529 | 48.3531 | 50413.9392 | 49364.3701 | 983.2874 | 883.2874 | 1307.5313
9643 | 100.5259 | 1.0808 | 1.0713 | 0.99642 | 181.2545 | 50.1093 | 59115.0023 | 57790.895 | 1151.8179 | 1051.8179 | 1307.8312
9644 | 96.0857 | 0.96195 | 0.7641 | 0.95912 | 196.7136 | 50.0355 | 41490.7202 | 41074.8234 | 817.4965 | 717.4965 | 1308.0388
9645 | 100.4056 | 0.98671 | 1.0578 | 0.9816 | 191.1748 | 52.7608 | 60203.5599 | 60036.7595 | 1196.7352 | 1096.7352 | 1308.0738
9646 | 95.9509 | 0.93985 | 0.96253 | 1.0225 | 187.8408 | 47.3301 | 51036.5263 | 51590.2179 | 1027.8044 | 927.8044 | 1308.3214
9647 | 85.1857 | 0.94398 | 1.087 | 0.99837 | 203.9605 | 50.2145 | 58736.2263 | 58291.2896 | 1161.8258 | 1061.8258 | 1308.4149
9648 | 92.1272 | 0.96357 | 1.0908 | 1.1088 | 213.9803 | 49.9707 | 65087.5045 | 65283.5249 | 1301.6705 | 1201.6705 | 1308.4976
9649 | 97.9651 | 1.0991 | 1.0171 | 1.0579 | 185.1667 | 49.9869 | 59355.0838 | 59568.5682 | 1187.3714 | 1087.3714 | 1308.5291
9650 | 94.4193 | 0.9139 | 1.0279 | 0.93533 | 212.7803 | 50.6025 | 53229.9783 | 52553.7292 | 1047.0746 | 947.0746 | 1308.7958
9651 | 111.0668 | 0.9924 | 1.1628 | 1.0427 | 202.7014 | 51.3618 | 68133.4297 | 68081.4962 | 1357.6299 | 1257.6299 | 1308.9011
9652 | 107.9499 | 1.0035 | 1.087 | 1.0264 | 186.8351 | 46.1562 | 56682.1732 | 55292.4688 | 1101.8494 | 1001.8494 | 1309.0347
9653 | 95.8752 | 1.0836 | 1.0021 | 0.98785 | 201.9156 | 48.6351 | 53397.8652 | 53764.3579 | 1071.2872 | 971.2872 | 1309.1384
9654 | 87.1294 | 1.0203 | 0.97948 | 1.0652 | 204.3998 | 48.3661 | 54965.5866 | 54368.8505 | 1083.377 | 983.377 | 1309.3337
9655 | 109.0055 | 1.1289 | 0.96416 | 0.95215 | 208.3052 | 45.3686 | 47440.5511 | 48197.0088 | 959.9402 | 859.9402 | 1309.4719
9656 | 112.9482 | 1.0216 | 0.95138 | 1.0196 | 190.887 | 50.2349 | 54718.0192 | 55600.2548 | 1108.0051 | 1008.0051 | 1309.4829
9657 | 84.0319 | 1.0266 | 0.68926 | 0.99816 | 198.7985 | 51.7696 | 40281.5259 | 40806.7596 | 812.1352 | 712.1352 | 1309.6788
9658 | 106.5246 | 1.011 | 0.98648 | 1.0027 | 188.3009 | 48.3404 | 53211.7626 | 53961.236 | 1075.2247 | 975.2247 | 1309.8815
9659 | 76.3204 | 0.96387 | 0.88892 | 1.0829 | 190.1226 | 48.4135 | 50356.8351 | 48757.2357 | 971.1447 | 871.1447 | 1309.9074
9660 | 92.6232 | 0.9459 | 0.89865 | 1.0735 | 201.1129 | 47.6621 | 50356.9283 | 49972.4693 | 995.4494 | 895.4494 | 1309.9445
9661 | 96.9313 | 0.96841 | 1.1318 | 0.97542 | 202.3095 | 47.3639 | 56937.8902 | 58260.2262 | 1161.2045 | 1061.2045 | 1309.9565
9662 | 106.1212 | 1.0722 | 0.97089 | 1.0468 | 197.7989 | 50.3768 | 57127.4838 | 57216.7071 | 1140.3341 | 1040.3341 | 1309.9675
9663 | 107.0578 | 1.004 | 0.93011 | 0.93426 | 208.2612 | 54.4416 | 53367.7971 | 53336.3428 | 1062.7269 | 962.7269 | 1310.4784
9664 | 109.5258 | 1.0121 | 1.1121 | 0.8885 | 193.0717 | 50.2963 | 55465.1831 | 55150.6994 | 1099.014 | 999.014 | 1310.8656
9665 | 102.801 | 1.0885 | 1.0901 | 1.0172 | 198.5568 | 49.657 | 60816.3573 | 60076.1105 | 1197.5222 | 1097.5222 | 1310.8841
9666 | 87.436 | 1.0243 | 1.0749 | 0.90385 | 199.3062 | 50.4329 | 53716.0921 | 52508.2524 | 1046.165 | 946.165 | 1311.067
9667 | 85.5795 | 1.0332 | 0.96074 | 1.017 | 209.6491 | 50.8251 | 54363.5679 | 55255.7278 | 1101.1146 | 1001.1146 | 1311.507
9668 | 106.6262 | 1.0356 | 1.0005 | 0.99346 | 210.1726 | 46.9252 | 52033.1511 | 50842.2421 | 1012.8448 | 912.8448 | 1311.7823
9669 | 108.4576 | 0.98296 | 1.0547 | 1.0626 | 192.0028 | 51.908 | 63900.3713 | 64482.2412 | 1285.6448 | 1185.6448 | 1311.8469
9670 | 93.31 | 0.96156 | 1.1248 | 0.96539 | 183.2569 | 51.2682 | 60452.9234 | 58640.2804 | 1168.8056 | 1068.8056 | 1312.2229
9671 | 96.8666 | 1.0745 | 0.97 | 1.0916 | 185.9536 | 52.4107 | 61134.0061 | 62012.9455 | 1236.2589 | 1136.2589 | 1312.6277
9672 | 100.2539 | 0.96824 | 1.0417 | 1.0651 | 215.3809 | 47.5236 | 57555.5145 | 56393.9798 | 1123.8796 | 1023.8796 | 1312.771
9673 | 84.7091 | 0.98884 | 1.1918 | 0.9532 | 202.0701 | 52.5565 | 64308.3795 | 65738.9441 | 1310.7789 | 1210.7789 | 1312.9014
9674 | 95.8475 | 0.94894 | 1.0444 | 0.96837 | 206.2823 | 49.0666 | 54294.8999 | 54144.9457 | 1078.8989 | 978.8989 | 1312.9354
9675 | 101.4506 | 1.0223 | 1.0524 | 1.0511 | 214.4522 | 49.0449 | 59553.6439 | 60049.3139 | 1196.9863 | 1096.9863 | 1313.0981
9676 | 116.2296 | 0.96122 | 0.94993 | 0.95994 | 210.2182 | 50.7352 | 52142.7588 | 51626.7325 | 1028.5346 | 928.5346 | 1313.3119
9677 | 77.7636 | 1.0239 | 0.90191 | 1.0271 | 190.9059 | 49.5994 | 50086.5402 | 49122.2684 | 978.4454 | 878.4454 | 1313.7713
9678 | 106.1244 | 1.0138 | 0.82712 | 1.0441 | 204.6595 | 48.8077 | 47607.0077 | 47750.2986 | 951.006 | 851.006 | 1313.813
9679 | 101.9637 | 0.97188 | 0.82682 | 1.0524 | 203.1351 | 52.9088 | 51483.5576 | 51757.3399 | 1031.1468 | 931.1468 | 1313.9264
9680 | 114.2995 | 1.0594 | 1.0331 | 0.94738 | 208.1245 | 53.7847 | 59361.3162 | 58858.6872 | 1173.1737 | 1073.1737 | 1314.4161
9681 | 100.6076 | 0.98344 | 1.0388 | 0.97629 | 189.784 | 54.5256 | 60885.1849 | 60168.5961 | 1199.3719 | 1099.3719 | 1315.0459
9682 | 93.5466 | 1.0139 | 1.0231 | 0.99367 | 197.8635 | 49.81 | 55560.5201 | 56649.335 | 1128.9867 | 1028.9867 | 1315.244
9683 | 112.4914 | 1.0322 | 1.2297 | 0.98979 | 189.7474 | 48.4383 | 64770.3501 | 62579.579 | 1247.5916 | 1147.5916 | 1315.7937
9684 | 97.399 | 1.0736 | 0.9531 | 1.1273 | 199.275 | 54.4316 | 64373.3015 | 64575.8161 | 1287.5163 | 1187.5163 | 1316.0243
9685 | 110.0914 | 0.96165 | 1.0474 | 1.0435 | 199.5941 | 50.0744 | 60231.772 | 60573.4417 | 1207.4688 | 1107.4688 | 1316.2379
9686 | 79.5001 | 0.92266 | 1.2061 | 0.97287 | 195.1922 | 49.5248 | 61936.785 | 61028.4248 | 1216.5685 | 1116.5685 | 1316.2498
9687 | 114.7371 | 0.9809 | 0.94339 | 0.95057 | 201.3764 | 50.1521 | 50820.119 | 50596.8729 | 1007.9375 | 907.9375 | 1317.1242
9688 | 97.1988 | 1.0045 | 1.0629 | 1.0728 | 216.4032 | 49.3256 | 61274.1143 | 61673.0845 | 1229.4617 | 1129.4617 | 1317.7415
9689 | 103.631 | 1.0534 | 1.0207 | 1.0056 | 210.6533 | 43.8033 | 49953.1222 | 49428.0811 | 984.5616 | 884.5616 | 1317.7837
9690 | 103.0408 | 0.98698 | 1.0917 | 1.0609 | 222.1731 | 53.9438 | 68181.2433 | 67195.3168 | 1339.9063 | 1239.9063 | 1318.3461
9691 | 87.6871 | 1.0127 | 0.92514 | 0.93048 | 186.7761 | 49.7431 | 47424.2428 | 47123.3109 | 938.4662 | 838.4662 | 1318.369
9692 | 110.1238 | 0.93618 | 1.0364 | 1.0813 | 188.8671 | 48.9714 | 60122.1407 | 58114.2164 | 1158.2843 | 1058.2843 | 1318.7203
9693 | 92.1565 | 0.96502 | 0.98452 | 1.0663 | 205.8898 | 51.4627 | 58808.5775 | 61490.4454 | 1225.8089 | 1125.8089 | 1318.8536
9694 | 93.9605 | 1.0117 | 0.77945 | 1.0599 | 201.7727 | 50.3014 | 46540.7072 | 45839.8617 | 912.7972 | 812.7972 | 1319.0367
9695 | 100.8571 | 1.028 | 0.82846 | 1.0117 | 209.1638 | 49.0933 | 46448.8672 | 45557.559 | 907.1512 | 807.1512 | 1319.2268
9696 | 96.7466 | 1.0355 | 1.0633 | 0.98316 | 223.7325 | 48.1022 | 55329.504 | 56311.757 | 1122.2351 | 1022.2351 | 1319.5237
9697 | 111.4182 | 1.0724 | 0.92258 | 1.1747 | 210.5647 | 48.5662 | 58648.7187 | 58602.8048 | 1168.0561 | 1068.0561 | 1319.5732
9698 | 96.141 | 0.96656 | 0.96315 | 0.92734 | 196.459 | 52.8208 | 52282.5861 | 51986.2741 | 1035.7255 | 935.7255 | 1319.692
9699 | 106.7836 | 0.91839 | 1.0681 | 1.0356 | 194.6737 | 48.4746 | 58567.7196 | 57142.4731 | 1138.8495 | 1038.8495 | 1320.5013
9700 | 88.7805 | 1.065 | 0.91419 | 1.0927 | 209.0105 | 47.2499 | 51875.45 | 51863.6555 | 1033.2731 | 933.2731 | 1320.629
9701 | 107.0348 | 0.94994 | 0.94718 | 1.0701 | 212.8506 | 51.1868 | 57298.7789 | 56626.0503 | 1128.521 | 1028.521 | 1321.0802
9702 | 107.8994 | 0.9877 | 1.0324 | 0.99593 | 202.5224 | 54.3438 | 61868.4723 | 59686.3601 | 1189.7272 | 1089.7272 | 1321.1402
9703 | 112.6818 | 0.98533 | 0.99247 | 0.99809 | 208.7038 | 49.0088 | 54196.7128 | 53254.4807 | 1061.0896 | 961.0896 | 1321.9289
9704 | 104.377 | 0.98839 | 1.0134 | 1.0254 | 212.8847 | 51.4748 | 59011.6761 | 59585.5286 | 1187.7106 | 1087.7106 | 1322.2709
9705 | 94.7373 | 0.83681 | 1.0718 | 0.9425 | 204.3678 | 49.2674 | 53880.5436 | 53599.324 | 1067.9865 | 967.9865 | 1322.313
9706 | 109.5186 | 1.0484 | 1.2491 | 1.0426 | 196.3289 | 46.0322 | 65425.9519 | 67780.9791 | 1351.6196 | 1251.6196 | 1322.3264
9707 | 91.1405 | 0.95968 | 1.0626 | 0.96776 | 202.6282 | 52.2482 | 58499.6158 | 58323.2533 | 1162.4651 | 1062.4651 | 1323.4171
9708 | 110.8815 | 0.92337 | 0.98163 | 0.99468 | 195.2746 | 50.8253 | 55025.4918 | 54568.2793 | 1087.3656 | 987.3656 | 1323.9014
9709 | 97.485 | 1.0053 | 1.1709 | 1.0396 | 200.4842 | 49.4583 | 65254.3654 | 66304.7806 | 1322.0956 | 1222.0956 | 1324.0413
9710 | 104.3098 | 0.96299 | 0.85132 | 1.0332 | 204.7109 | 46.9039 | 46170.5895 | 45385.1354 | 903.7027 | 803.7027 | 1324.5283
9711 | 95.7659 | 0.99985 | 0.97405 | 0.93836 | 210.218 | 49.1329 | 49822.8216 | 50164.4579 | 999.2892 | 899.2892 | 1324.6888
9712 | 87.0559 | 1.0242 | 1.136 | 1.0293 | 188.3746 | 49.8323 | 62899.9295 | 63544.5802 | 1266.8916 | 1166.8916 | 1324.8762
9713 | 102.837 | 0.94637 | 0.88631 | 1.06 | 204.4573 | 51.7019 | 53807.609 | 52785.2152 | 1051.7043 | 951.7043 | 1324.8929
9714 | 104.5986 | 0.89395 | 0.76398 | 0.98224 | 213.8163 | 46.5646 | 39510.2658 | 41114.115 | 818.2823 | 718.2823 | 1325.2356
9715 | 82.7104 | 1.0553 | 0.96917 | 0.94208 | 191.4464 | 55.0686 | 55277.5308 | 54875.3777 | 1093.5076 | 993.5076 | 1325.7584
9716 | 105.0956 | 0.94125 | 0.95911 | 0.99151 | 187.6983 | 46.0867 | 48573.2308 | 46750.8045 | 931.0161 | 831.0161 | 1326.0328
9717 | 102.3423 | 1.003 | 1.0974 | 1.0679 | 195.914 | 51.9329 | 66388.8396 | 67640.201 | 1348.804 | 1248.804 | 1326.6784
9718 | 111.1817 | 1.018 | 0.8636 | 1.1145 | 202.8398 | 52.3372 | 56501.0352 | 54564.4299 | 1087.2886 | 987.2886 | 1326.6796
9719 | 98.8037 | 0.94005 | 0.923 | 1.0221 | 183.7259 | 48.9671 | 50924.8748 | 48891.3248 | 973.8265 | 873.8265 | 1326.828
9720 | 113.4101 | 1.0137 | 0.88019 | 0.9074 | 210.5845 | 47.1566 | 43295.0824 | 44177.0612 | 879.5412 | 779.5412 | 1326.9197
9721 | 109.6051 | 1.0679 | 0.86935 | 1.0358 | 196.7082 | 49.1688 | 50225.1289 | 51231.7251 | 1020.6345 | 920.6345 | 1327.2887
9722 | 90.7287 | 0.97848 | 0.98801 | 1.0151 | 190.6058 | 46.4544 | 50905.1464 | 52453.3016 | 1045.066 | 945.066 | 1328.1512
9723 | 105.3778 | 0.95463 | 1.2682 | 1.0362 | 200.5433 | 53.418 | 75776.9062 | 76747.1509 | 1530.943 | 1430.943 | 1328.2875
9724 | 106.4657 | 1.0246 | 1.0566 | 1.0951 | 187.8383 | 48.1062 | 61095.8359 | 61019.32 | 1216.3864 | 1116.3864 | 1328.9319
9725 | 93.4081 | 1.0691 | 1.1522 | 0.97487 | 208.2397 | 52.1499 | 63993.7355 | 61361.8576 | 1223.2372 | 1123.2372 | 1329.2012
9726 | 102.9978 | 1.0431 | 0.90763 | 0.93242 | 199.9561 | 47.7978 | 45786.3727 | 45768.3048 | 911.3661 | 811.3661 | 1329.6491
9727 | 103.1844 | 1.06 | 0.91206 | 1.001 | 199.1556 | 53.8659 | 55267.2695 | 54013.9825 | 1076.2797 | 976.2797 | 1329.9052
9728 | 102.4641 | 1.0725 | 1.0554 | 0.92094 | 192.9057 | 52.4501 | 56935.5926 | 57115.4266 | 1138.3085 | 1038.3085 | 1329.9104
9729 | 85.5419 | 0.92665 | 1.1029 | 0.96099 | 201.1181 | 48.222 | 55131.8153 | 56341.4394 | 1122.8288 | 1022.8288 | 1330.1672
9730 | 94.3955 | 1.0281 | 1.3674 | 1.0596 | 208.031 | 47.44 | 73551.3504 | 73367.4052 | 1463.3481 | 1363.3481 | 1330.1764
9731 | 100.4725 | 1.0103 | 1.1541 | 1.0457 | 196.1262 | 50.067 | 65703.639 | 64998.3695 | 1295.9674 | 1195.9674 | 1330.4164
9732 | 87.8439 | 0.98016 | 1.0571 | 1.0543 | 211.7965 | 50.1165 | 60382.905 | 61978.148 | 1235.563 | 1135.563 | 1330.5553
9733 | 90.2615 | 1.017 | 0.98159 | 0.84051 | 193.8525 | 53.0429 | 48824.8733 | 49842.6109 | 992.8522 | 892.8522 | 1330.9605
9734 | 99.8804 | 0.95493 | 1.1084 | 1.0521 | 220.2001 | 51.1944 | 64803.8345 | 64745.9706 | 1290.9194 | 1190.9194 | 1331.169
9735 | 98.9327 | 0.95922 | 0.79673 | 1.0355 | 222.1139 | 47.3354 | 43766.0735 | 43539.1371 | 866.7827 | 766.7827 | 1331.2611
9736 | 101.1019 | 1.0294 | 0.92813 | 0.9766 | 209.3152 | 49.2192 | 49944.6575 | 49350.0842 | 983.0017 | 883.0017 | 1331.8974
9737 | 118.5291 | 0.97714 | 0.92719 | 0.97497 | 225.2263 | 49.2097 | 50409.1977 | 52147.845 | 1038.9569 | 938.9569 | 1332.0122
9738 | 96.1014 | 1.0001 | 0.89146 | 0.95285 | 209.8333 | 48.2264 | 45809.5377 | 45944.0174 | 914.8803 | 814.8803 | 1332.2232
9739 | 109.4835 | 0.95899 | 1.2895 | 0.97089 | 214.8571 | 50.1864 | 68316.813 | 65701.9337 | 1310.0387 | 1210.0387 | 1333.0091
9740 | 91.5902 | 1.0035 | 0.85055 | 0.97864 | 191.3834 | 48.4202 | 44946.4794 | 44649.443 | 888.9889 | 788.9889 | 1333.1098
9741 | 102.9077 | 0.99664 | 1.048 | 1.0359 | 210.9829 | 50.5138 | 60229.0342 | 60946.9425 | 1214.9389 | 1114.9389 | 1333.7977
9742 | 100.2273 | 0.99628 | 0.99671 | 0.93609 | 198.2649 | 47.9913 | 49766.8502 | 49097.5107 | 977.9502 | 877.9502 | 1333.8286
9743 | 94.572 | 0.99331 | 0.94934 | 0.90641 | 197.969 | 49.8178 | 47745.6668 | 46259.2547 | 921.1851 | 821.1851 | 1334.0781
9744 | 96.4092 | 1.0801 | 0.96369 | 0.992 | 194.883 | 47.9721 | 51050.8708 | 51627.1763 | 1028.5435 | 928.5435 | 1334.3336
9745 | 118.1965 | 1.0217 | 1.0441 | 1.0058 | 194.0365 | 51.8657 | 60922.6175 | 62083.3301 | 1237.6666 | 1137.6666 | 1334.5751
9746 | 85.9543 | 0.94257 | 0.99286 | 1.0407 | 208.4438 | 47.8073 | 53478.5726 | 52243.8893 | 1040.8778 | 940.8778 | 1335.0833
9747 | 87.4708 | 1.0352 | 1.0569 | 0.96916 | 211.5263 | 50.9955 | 57065.5831 | 57006.0151 | 1136.1203 | 1036.1203 | 1335.1487
9748 | 103.0864 | 0.94349 | 0.96042 | 0.97027 | 192.2402 | 47.8754 | 49462.1909 | 49870.1216 | 993.4024 | 893.4024 | 1335.3818
9749 | 95.809 | 0.99672 | 0.80628 | 1.0223 | 200.506 | 50.4196 | 46572.9859 | 47783.7883 | 951.6758 | 851.6758 | 1335.7719
9750 | 101.2214 | 1.1382 | 0.96651 | 0.95211 | 190.5343 | 50.0855 | 52050.2403 | 51177.742 | 1019.5548 | 919.5548 | 1336.2755
9751 | 104.0868 | 1.0364 | 1.0189 | 0.98134 | 202.9289 | 50.5507 | 56200.7052 | 55695.4477 | 1109.909 | 1009.909 | 1336.8751
9752 | 98.8692 | 1.0086 | 1.0147 | 0.96358 | 202.6081 | 50.9025 | 55049.1828 | 55126.309 | 1098.5262 | 998.5262 | 1337.3325
9753 | 93.4743 | 0.91827 | 1.1985 | 0.96871 | 201.8596 | 51.1641 | 63995.2232 | 64876.1692 | 1293.5234 | 1193.5234 | 1337.9064
9754 | 100.8528 | 0.9836 | 0.99865 | 0.99991 | 190.1039 | 48.6258 | 53569.2619 | 52442.6876 | 1044.8538 | 944.8538 | 1338.0996
9755 | 110.6446 | 1.014 | 0.9178 | 0.97552 | 191.9719 | 49.3898 | 49953.5594 | 49468.2441 | 985.3649 | 885.3649 | 1338.405
9756 | 96.2719 | 0.96516 | 0.94579 | 1.0294 | 184.3793 | 51.2302 | 54821.5764 | 55038.7859 | 1096.7757 | 996.7757 | 1338.5305
9757 | 105.4519 | 1.0409 | 1.0752 | 1.0087 | 185.4814 | 47.802 | 57276.3768 | 57993.4687 | 1155.8694 | 1055.8694 | 1338.5617
9758 | 104.5322 | 0.94326 | 1.0451 | 0.91222 | 211.1401 | 49.1624 | 51927.5317 | 52146.2094 | 1038.9242 | 938.9242 | 1339.2193
9759 | 94.0716 | 0.96128 | 0.97793 | 0.97525 | 207.5368 | 49.5729 | 51969.3918 | 50613.862 | 1008.2772 | 908.2772 | 1339.3896
9760 | 104.2895 | 1.0717 | 1.1929 | 0.95654 | 191.2353 | 46.7806 | 58796.7314 | 58427.536 | 1164.5507 | 1064.5507 | 1340.3359
9761 | 93.9952 | 0.98822 | 1.0183 | 0.96797 | 207.6682 | 54.3943 | 58877.4815 | 58974.8171 | 1175.4963 | 1075.4963 | 1340.3792
9762 | 90.1411 | 1.0346 | 1.1321 | 1.0064 | 201.7167 | 46.8921 | 58002.4424 | 57253.9148 | 1141.0783 | 1041.0783 | 1340.4706
9763 | 97.0195 | 0.96887 | 0.84529 | 1.0316 | 206.2706 | 44.9701 | 43648.4937 | 43887.9267 | 873.7585 | 773.7585 | 1340.8686
9764 | 100.1002 | 1.0069 | 0.98502 | 0.90396 | 208.11 | 52.3421 | 52090.2398 | 53086.6889 | 1057.7338 | 957.7338 | 1340.9947
9765 | 91.7883 | 0.95449 | 1.1341 | 0.99802 | 191.8358 | 51.6694 | 63203.2115 | 62859.2969 | 1253.1859 | 1153.1859 | 1341.226
9766 | 101.3583 | 0.9388 | 1.1394 | 1.025 | 185.68 | 48.3025 | 61196.1554 | 60805.8542 | 1212.1171 | 1112.1171 | 1341.316
9767 | 98.7574 | 0.99619 | 0.84003 | 1.0207 | 187.8043 | 50.0779 | 48051.1159 | 49475.0479 | 985.501 | 885.501 | 1341.3871
9768 | 95.069 | 0.97966 | 0.82946 | 0.9742 | 191.9843 | 50.5866 | 45780.4432 | 46654.9236 | 929.0985 | 829.0985 | 1341.4155
9769 | 94.8336 | 1.0387 | 0.88843 | 1.0035 | 195.9307 | 46.935 | 46664.1471 | 45433.8521 | 904.677 | 804.677 | 1341.452
9770 | 90.7516 | 1.0206 | 1.0056 | 0.9916 | 186.7987 | 51.0276 | 55792.7431 | 55735.9342 | 1110.7187 | 1010.7187 | 1341.5327
9771 | 91.8274 | 1.0275 | 0.82621 | 0.96552 | 208.447 | 48.8203 | 43759.7257 | 42642.6701 | 848.8534 | 748.8534 | 1341.6551
9772 | 85.6181 | 0.99003 | 0.90432 | 0.97512 | 208.1735 | 50.491 | 49012.1778 | 46845.8997 | 932.918 | 832.918 | 1341.7811
9773 | 97.2221 | 1.024 | 1.0201 | 0.93246 | 198.6856 | 50.7817 | 53560.0632 | 54566.0054 | 1087.3201 | 987.3201 | 1342.5075
9774 | 103.2551 | 0.99438 | 0.85045 | 0.97176 | 192.0257 | 54.6622 | 50979.3843 | 51570.1681 | 1027.4034 | 927.4034 | 1342.9377
9775 | 108.3801 | 0.99559 | 0.93504 | 0.91468 | 189.2808 | 48.3453 | 46754.1491 | 47443.6189 | 944.8724 | 844.8724 | 1343.5149
9776 | 111.178 | 0.93066 | 1.1533 | 1.113 | 191.5598 | 52.9041 | 73578.7466 | 73924.4171 | 1474.4883 | 1374.4883 | 1344.9311
9777 | 115.3686 | 1.0079 | 1.0555 | 1.0674 | 201.6773 | 49.2182 | 61375.5934 | 64312.5136 | 1282.2503 | 1182.2503 | 1345.3951
9778 | 108.9405 | 1.041 | 0.85677 | 1.1006 | 187.9182 | 51.8216 | 54929.5352 | 55629.4882 | 1108.5898 | 1008.5898 | 1345.554
9779 | 114.9552 | 1.0638 | 1.1319 | 1.0335 | 194.4696 | 46.8056 | 60673.0366 | 59422.3813 | 1184.4476 | 1084.4476 | 1345.5632
9780 | 99.4926 | 1.0133 | 0.92773 | 1.0456 | 206.3065 | 48.1976 | 51820.0421 | 50635.2898 | 1008.7058 | 908.7058 | 1347.1989
9781 | 91.9429 | 1.0209 | 1.1601 | 1.004 | 208.8124 | 51.2561 | 64717.7269 | 65343.1789 | 1302.8636 | 1202.8636 | 1347.3003
9782 | 121.3764 | 1.0095 | 1.1395 | 1.0163 | 198.5236 | 50.672 | 65088.1617 | 64336.2593 | 1282.7252 | 1182.7252 | 1348.0203
9783 | 116.5906 | 1.0274 | 1.0705 | 1.0934 | 188.8569 | 49.9541 | 64639.4408 | 65817.5546 | 1312.3511 | 1212.3511 | 1349.21
9784 | 109.433 | 1.0506 | 1.315 | 0.99836 | 200.782 | 50.5033 | 72310.9675 | 73487.6431 | 1465.7529 | 1365.7529 | 1349.6673
9785 | 85.8742 | 1.0108 | 0.64387 | 1.0164 | 204.1312 | 50.1462 | 37375.0094 | 37265.0791 | 741.3016 | 641.3016 | 1350.0609
9786 | 112.7718 | 0.99223 | 1.0645 | 0.93146 | 198.6327 | 51.8698 | 57433.9123 | 55948.1794 | 1114.9636 | 1014.9636 | 1350.2694
9787 | 87.1861 | 1.0315 | 1.1495 | 0.95151 | 206.1384 | 46.2842 | 54992.0157 | 53964.1211 | 1075.2824 | 975.2824 | 1350.5063
9788 | 93.3689 | 0.88313 | 1.0423 | 0.95568 | 200.4015 | 42.193 | 45710.2438 | 47175.5315 | 939.5106 | 839.5106 | 1350.6286
9789 | 100.0808 | 1.0059 | 1.2955 | 1.0804 | 194.2093 | 50.6519 | 76193.7053 | 77300.2386 | 1542.0048 | 1442.0048 | 1350.9431
9790 | 110.8661 | 0.86685 | 0.8927 | 1.0028 | 206.7139 | 53.6822 | 53420.8522 | 55256.1431 | 1101.1229 | 1001.1229 | 1351.3525
9791 | 107.0542 | 1.0363 | 0.9211 | 0.9674 | 207.2798 | 48.0493 | 48353.3881 | 48449.4696 | 964.9894 | 864.9894 | 1351.9998
9792 | 98.9119 | 0.91645 | 1.0534 | 0.99184 | 208.1658 | 55.3971 | 63108.6944 | 63731.4082 | 1270.6282 | 1170.6282 | 1352.2065
9793 | 99.5314 | 0.94097 | 0.95132 | 1.0933 | 184.5362 | 50.6582 | 57617.9855 | 57375.5209 | 1143.5104 | 1043.5104 | 1353.1563
9794 | 101.355 | 1.0038 | 1.2315 | 0.91445 | 200.1353 | 51.1688 | 63030.6372 | 64885.1782 | 1293.7036 | 1193.7036 | 1353.1727
9795 | 102.4082 | 0.93823 | 0.72758 | 1.0296 | 202.8112 | 47.9426 | 40724.8123 | 41582.444 | 827.6489 | 727.6489 | 1353.4651
9796 | 103.8919 | 0.94428 | 1.0336 | 0.9945 | 195.9479 | 47.9274 | 54161.0837 | 53753.1556 | 1071.0631 | 971.0631 | 1353.8985
9797 | 103.2057 | 0.95853 | 0.9131 | 1.0205 | 209.3694 | 50.0471 | 51795.705 | 51583.91 | 1027.6782 | 927.6782 | 1354.3622
9798 | 125.8827 | 1.019 | 1.0924 | 0.99531 | 223.1244 | 50.4025 | 61488.9592 | 62074.0618 | 1237.4812 | 1137.4812 | 1354.7433
9799 | 91.4322 | 1.1495 | 0.85958 | 0.97927 | 193.5434 | 51.4402 | 48900.7447 | 49115.015 | 978.3003 | 878.3003 | 1354.9882
9800 | 119.664 | 0.96689 | 1.2059 | 1.0517 | 195.9497 | 49.6788 | 68946.4714 | 70247.0656 | 1400.9413 | 1300.9413 | 1355.0268
9801 | 112.4288 | 1.0145 | 0.94364 | 1.0378 | 209.2512 | 53.6471 | 58863.6388 | 57218.1053 | 1140.3621 | 1040.3621 | 1355.0882
9802 | 94.7464 | 0.93914 | 1.1406 | 1.1047 | 205.9525 | 45.6447 | 61781.5639 | 62504.9484 | 1246.099 | 1146.099 | 1355.5656
9803 | 103.7979 | 0.98842 | 0.98285 | 0.94093 | 197.9187 | 48.3527 | 49874.617 | 48386.063 | 963.7213 | 863.7213 | 1356.9649
9804 | 116.8832 | 0.9566 | 0.80162 | 0.89379 | 202.6826 | 54.575 | 45406.4465 | 45460.6541 | 905.2131 | 805.2131 | 1357.0047
9805 | 97.3123 | 0.97212 | 1.0807 | 1.0356 | 190.7013 | 49.8118 | 60649.0634 | 60682.7246 | 1209.6545 | 1109.6545 | 1358.1626
9806 | 99.3828 | 1.0126 | 1.0503 | 0.94569 | 212.7034 | 52.0783 | 57180.8336 | 58829.3459 | 1172.5869 | 1072.5869 | 1358.2186
9807 | 108.2917 | 1.0139 | 0.88273 | 0.9738 | 209.3332 | 51.3591 | 49996.6002 | 48702.4184 | 970.0484 | 870.0484 | 1358.4593
9808 | 87.506 | 0.89861 | 1.2408 | 0.90457 | 189.1119 | 46.3618 | 55871.3187 | 55826.2116 | 1112.5242 | 1012.5242 | 1358.7944
9809 | 102.657 | 1.0524 | 0.99353 | 1.0616 | 192.5576 | 46.289 | 54016.1243 | 57017.7625 | 1136.3552 | 1036.3552 | 1359.4678
9810 | 106.6446 | 1.0226 | 0.85689 | 0.96941 | 211.6411 | 45.0436 | 42540.2863 | 41378.8791 | 823.5776 | 723.5776 | 1359.7507
9811 | 103.1559 | 0.96993 | 1.2647 | 1.0732 | 181.1289 | 49.5505 | 72391.6686 | 73861.8488 | 1473.237 | 1373.237 | 1359.8189
9812 | 115.5154 | 0.9599 | 1.0957 | 0.98671 | 190.1251 | 49.6045 | 59322.1373 | 59597.4713 | 1187.9494 | 1087.9494 | 1359.8983
9813 | 105.2908 | 0.94095 | 0.98608 | 0.91351 | 190.0014 | 51.9307 | 52114.0671 | 53634.3272 | 1068.6865 | 968.6865 | 1360.0158
9814 | 97.4916 | 1.0247 | 0.99779 | 0.95937 | 210.6492 | 48.9755 | 51984.9435 | 52475.8268 | 1045.5165 | 945.5165 | 1360.1259
9815 | 93.2319 | 0.98399 | 0.99838 | 1.0115 | 195.9029 | 54.5475 | 60287.9588 | 59652.0511 | 1189.041 | 1089.041 | 1360.5137
9816 | 94.9876 | 1.0227 | 1.0429 | 1.0584 | 202.5532 | 49.6586 | 59842.9169 | 60241.2015 | 1200.824 | 1100.824 | 1362.1045
9817 | 95.646 | 0.94158 | 1.1901 | 1.0129 | 192.2759 | 46.2227 | 60072.4287 | 61103.4012 | 1218.068 | 1118.068 | 1362.6769
9818 | 118.8797 | 1.0076 | 0.7822 | 1.0045 | 196.8206 | 48.0416 | 43699.3468 | 41997.2258 | 835.9445 | 735.9445 | 1363.1646
9819 | 93.222 | 1.0341 | 1.0047 | 1.022 | 206.0699 | 49.3286 | 55611.3686 | 55863.2847 | 1113.2657 | 1013.2657 | 1363.332
9820 | 78.0198 | 0.98233 | 1.1954 | 1.0636 | 201.5229 | 54.2217 | 73297.2836 | 72810.3241 | 1452.2065 | 1352.2065 | 1363.3481
9821 | 100.6292 | 1.0157 | 1.024 | 0.9545 | 204.3095 | 47.1271 | 51083.9796 | 52010.3914 | 1036.2078 | 936.2078 | 1363.9972
9822 | 94.9343 | 0.95985 | 1.1206 | 0.97642 | 204.0063 | 47.5434 | 56555.9283 | 55487.0248 | 1105.7405 | 1005.7405 | 1364.6253
9823 | 116.4973 | 1.0252 | 1.1296 | 0.99527 | 192.2374 | 53.8094 | 67116.127 | 67721.7583 | 1350.4352 | 1250.4352 | 1365.1122
9824 | 107.127 | 1.0425 | 1.0183 | 0.99565 | 220.7743 | 48.4867 | 54797.3396 | 54453.3265 | 1085.0665 | 985.0665 | 1365.1948
9825 | 93.4987 | 1.0139 | 1.1743 | 0.95909 | 237.2356 | 51.1365 | 62676.9122 | 62609.5916 | 1248.1918 | 1148.1918 | 1365.5155
9826 | 108.684 | 1.0568 | 0.93845 | 1.0271 | 179.2356 | 47.9899 | 51948.1922 | 50865.4585 | 1013.3092 | 913.3092 | 1365.7529
9827 | 97.1991 | 0.99524 | 1.1826 | 0.98354 | 210.3504 | 50.3007 | 63584.0179 | 61730.6831 | 1230.6137 | 1130.6137 | 1369.0674
9828 | 86.3627 | 1.002 | 1.1542 | 1.0283 | 197.2573 | 47.0126 | 60062.3787 | 60422.3881 | 1204.4478 | 1104.4478 | 1369.8976
9829 | 112.2333 | 0.94197 | 0.93256 | 0.97072 | 199.2098 | 50.0728 | 50821.523 | 51356.4486 | 1023.129 | 923.129 | 1370.0148
9830 | 99.6229 | 0.97774 | 1.2103 | 1.0146 | 208.2174 | 47.0642 | 62585.4014 | 61618.8512 | 1228.377 | 1128.377 | 1370.1244
9831 | 111.4717 | 0.97933 | 0.9544 | 0.98807 | 203.4704 | 50.9527 | 53814.5502 | 54809.9389 | 1092.1988 | 992.1988 | 1370.7396
9832 | 103.7651 | 0.96694 | 0.7963 | 0.93943 | 180.8561 | 52.5165 | 44736.1397 | 44553.6622 | 887.0732 | 787.0732 | 1371.0108
9833 | 98.0186 | 1.0334 | 0.84267 | 1.0427 | 199.132 | 53.1449 | 52279.4655 | 52426.0705 | 1044.5214 | 944.5214 | 1371.2657
9834 | 99.7148 | 1.0349 | 0.89714 | 1.0012 | 215.3135 | 54.1246 | 54413.721 | 52873.3586 | 1053.4672 | 953.4672 | 1371.5058
9835 | 96.098 | 0.99459 | 0.8484 | 0.97159 | 203.4341 | 53.457 | 49377.251 | 48541.3648 | 966.8273 | 866.8273 | 1371.5289
9836 | 130.4077 | 1.0297 | 0.81169 | 1.0244 | 211.002 | 53.2885 | 51678.0733 | 53563.0908 | 1067.2618 | 967.2618 | 1371.7194
9837 | 99.0402 | 0.97572 | 0.88018 | 0.96262 | 209.0196 | 51.1202 | 48462.0528 | 50023.7204 | 996.4744 | 896.4744 | 1372.3355
9838 | 92.988 | 1.0341 | 0.98738 | 1.0672 | 209.295 | 50.2112 | 57945.3255 | 57657.1955 | 1149.1439 | 1049.1439 | 1372.4547
9839 | 101.3632 | 0.9957 | 1.004 | 0.93513 | 225.436 | 44.7906 | 46796.3966 | 47185.9319 | 939.7186 | 839.7186 | 1373.113
9840 | 94.8156 | 1.0131 | 0.98308 | 1.0488 | 201.0794 | 53.6547 | 60678.2361 | 61325.0609 | 1222.5012 | 1122.5012 | 1373.237
9841 | 94.1861 | 0.97426 | 1.0787 | 1.0179 | 224.5465 | 48.5672 | 58003.4753 | 58068.0714 | 1157.3614 | 1057.3614 | 1373.4054
9842 | 93.7966 | 0.99801 | 0.93748 | 0.87617 | 212.357 | 54.7524 | 50310.752 | 47620.6682 | 948.4134 | 848.4134 | 1373.6545
9843 | 104.6837 | 1.1035 | 1.106 | 1.0307 | 192.6608 | 46.8949 | 59065.4792 | 59961.4409 | 1195.2288 | 1095.2288 | 1373.7301
9844 | 94.5412 | 1.025 | 1.0947 | 1.0515 | 198.7782 | 49.29 | 61713.9638 | 63311.5954 | 1262.2319 | 1162.2319 | 1373.8357
9845 | 100.8322 | 1.0356 | 0.97089 | 0.99338 | 205.9472 | 48.5007 | 52047.8412 | 52768.7477 | 1051.375 | 951.375 | 1374.4883
9846 | 101.3257 | 1.0009 | 0.90534 | 0.83416 | 216.071 | 50.0852 | 43119.3464 | 42721.8293 | 850.4366 | 750.4366 | 1374.6356
9847 | 96.9118 | 0.94877 | 1.0537 | 0.91664 | 192.3809 | 48.5608 | 51558.3263 | 53522.4547 | 1066.4491 | 966.4491 | 1374.7759
9848 | 103.7728 | 0.98747 | 1.0777 | 0.95789 | 206.544 | 53.3001 | 60688.5161 | 58490.6503 | 1165.813 | 1065.813 | 1374.7828
9849 | 116.8834 | 1.0785 | 0.89785 | 0.94938 | 196.8133 | 47.9491 | 47113.4603 | 47569.6487 | 947.393 | 847.393 | 1375.9105
9850 | 97.4632 | 1.0466 | 0.83941 | 0.99502 | 208.2456 | 52.6628 | 49565.503 | 51561.2254 | 1027.2245 | 927.2245 | 1376.1068
9851 | 110.6809 | 1.0416 | 0.92747 | 0.99004 | 199.4997 | 51.2544 | 53171.8747 | 52458.4639 | 1045.1693 | 945.1693 | 1377.0842
9852 | 98.0008 | 0.95273 | 1.0207 | 1.0341 | 208.2244 | 52.2864 | 60281.9612 | 61992.1456 | 1235.8429 | 1135.8429 | 1378.0204
9853 | 109.8521 | 1.0013 | 1.0925 | 1.0052 | 206.8551 | 54.1049 | 65575.42 | 65432.5299 | 1304.6506 | 1204.6506 | 1378.6761
9854 | 102.5037 | 1.0189 | 1.067 | 0.88338 | 178.5915 | 47.4751 | 49885.8758 | 49506.4829 | 986.1297 | 886.1297 | 1379.2936
9855 | 90.288 | 0.95794 | 0.88219 | 1.055 | 215.0608 | 46.0149 | 47022.1731 | 47241.167 | 940.8233 | 840.8233 | 1379.9352
9856 | 116.3866 | 0.95621 | 0.8078 | 0.95499 | 205.1969 | 50.0525 | 44388.1139 | 44467.3625 | 885.3473 | 785.3473 | 1379.9754
9857 | 106.3578 | 1.0777 | 1.0345 | 0.99183 | 168.0667 | 48.7404 | 55763.6698 | 54007.4239 | 1076.1485 | 976.1485 | 1380.3784
9858 | 114.2304 | 0.97235 | 0.83042 | 1.0683 | 191.7276 | 52.9708 | 53068.0505 | 51616.9028 | 1028.3381 | 928.3381 | 1380.4206
9859 | 78.9611 | 1.0332 | 1.2636 | 0.99029 | 203.9726 | 49.2914 | 65905.1254 | 68573.0483 | 1367.461 | 1267.461 | 1380.7321
9860 | 98.5656 | 1.0281 | 0.94563 | 1.0714 | 208.0004 | 50.7106 | 56725.8074 | 55299.2395 | 1101.9848 | 1001.9848 | 1382.7269
9861 | 79.2793 | 0.97087 | 1.1383 | 1.0291 | 222.2012 | 46.9595 | 58843.1793 | 57486.0586 | 1145.7212 | 1045.7212 | 1383.3189
9862 | 102.6257 | 0.989 | 0.96438 | 0.9971 | 216.4042 | 47.4338 | 50642.3149 | 52065.2104 | 1037.3042 | 937.3042 | 1383.9766
9863 | 92.4296 | 1.0036 | 0.92929 | 1.0471 | 187.2969 | 48.087 | 51438.5644 | 51086.0412 | 1017.7208 | 917.7208 | 1385.4993
9864 | 104.8796 | 1.0312 | 0.8157 | 0.9994 | 206.4863 | 50.7071 | 47027.1316 | 46232.2391 | 920.6448 | 820.6448 | 1386.4261
9865 | 90.5323 | 1.0184 | 0.81924 | 1.0198 | 198.9734 | 52.1051 | 48536.2613 | 47881.1857 | 953.6237 | 853.6237 | 1386.477
9866 | 105.2411 | 0.92231 | 0.83667 | 1.0476 | 191.3661 | 50.4817 | 49337.1021 | 48949.3962 | 974.9879 | 874.9879 | 1386.6387
9867 | 105.8799 | 0.98422 | 1.0496 | 1.0271 | 184.1184 | 45.4531 | 53924.3551 | 54376.0626 | 1083.5213 | 983.5213 | 1386.9994
9868 | 98.6134 | 1.0283 | 1.1566 | 1.0299 | 192.7041 | 49.9254 | 64725.6819 | 65567.4291 | 1307.3486 | 1207.3486 | 1387.2419
9869 | 115.0312 | 0.91611 | 0.98448 | 1.0537 | 205.9596 | 48.0797 | 55146.825 | 53299.6499 | 1061.993 | 961.993 | 1387.4267
9870 | 108.5411 | 0.97009 | 0.70971 | 1.0832 | 199.1133 | 50.0759 | 43968.5795 | 43646.9578 | 868.9392 | 768.9392 | 1388.2092
9871 | 102.0287 | 0.97217 | 1.1995 | 1.0612 | 199.9719 | 50.6603 | 69707.8588 | 69509.6389 | 1386.1928 | 1286.1928 | 1388.8469
9872 | 96.1691 | 1.0553 | 0.93739 | 1.0042 | 198.8592 | 53.5417 | 56033.033 | 55325.4844 | 1102.5097 | 1002.5097 | 1389.0827
9873 | 92.3701 | 0.89258 | 0.87265 | 0.96143 | 181.7228 | 51.1196 | 47285.3036 | 48861.0865 | 973.2217 | 873.2217 | 1391.2076
9874 | 96.791 | 0.99434 | 1.0074 | 0.96989 | 204.3807 | 49.3955 | 53221.7127 | 52239.1564 | 1040.7831 | 940.7831 | 1394.1025
9875 | 78.993 | 1.0288 | 0.68529 | 1.0261 | 206.346 | 53.7331 | 42356.4662 | 42050.8544 | 837.0171 | 737.0171 | 1394.2333
9876 | 96.8987 | 0.90477 | 1.0125 | 1.0632 | 176.7708 | 45.5922 | 53254.1891 | 54847.7231 | 1092.9545 | 992.9545 | 1395.0418
9877 | 104.087 | 1.0591 | 1.2073 | 1.0061 | 207.0091 | 49.7359 | 66101.9917 | 66837.9326 | 1332.7587 | 1232.7587 | 1395.1115
9878 | 84.0604 | 1.0107 | 1.0936 | 1.0554 | 193.9432 | 45.9112 | 57084.1747 | 58239.1644 | 1160.7833 | 1060.7833 | 1395.2449
9879 | 99.5012 | 0.96553 | 1.1555 | 1.0453 | 209.5202 | 49.9828 | 65387.7337 | 66108.6618 | 1318.1732 | 1218.1732 | 1395.6756
9880 | 97.2704 | 0.99561 | 1.0975 | 0.91835 | 200.7374 | 44.7152 | 49597.909 | 49688.6391 | 989.7728 | 889.7728 | 1395.7879
9881 | 109.7361 | 1.0655 | 1.0502 | 0.9886 | 205.5439 | 49.3706 | 57234.4972 | 56763.2277 | 1131.2646 | 1031.2646 | 1396.7588
9882 | 90.7288 | 0.95227 | 0.87799 | 0.99827 | 200.3061 | 52.7698 | 51010.8832 | 52354.2793 | 1043.0856 | 943.0856 | 1396.7904
9883 | 124.5385 | 1.0083 | 0.97573 | 1.0069 | 211.4107 | 52.0209 | 57853.1161 | 56388.0386 | 1123.7608 | 1023.7608 | 1396.9622
9884 | 114.8194 | 0.97955 | 0.86105 | 0.96346 | 192.8168 | 49.7532 | 47063.3905 | 45630.037 | 908.6007 | 808.6007 | 1398.0508
9885 | 103.0619 | 0.9347 | 1.325 | 0.98914 | 213.4662 | 48.1024 | 67890.9969 | 67089.5169 | 1337.7903 | 1237.7903 | 1398.4885
9886 | 95.6605 | 0.98934 | 1.0342 | 1.0396 | 198.969 | 51.8608 | 60866.14 | 62251.8654 | 1241.0373 | 1141.0373 | 1399.4119
9887 | 96.8654 | 1.0213 | 1.1456 | 0.95362 | 195.7498 | 52.8566 | 63166.9584 | 64482.9002 | 1285.658 | 1185.658 | 1400.3782
9888 | 95.051 | 0.9762 | 1.0291 | 0.93882 | 217.3607 | 47.9434 | 50986.0756 | 52169.2544 | 1039.3851 | 939.3851 | 1400.4518
9889 | 90.4977 | 1.0873 | 0.87573 | 1.0519 | 194.1549 | 51.9094 | 53118.0691 | 52698.0209 | 1049.9604 | 949.9604 | 1400.6179
9890 | 99.3729 | 1.0053 | 1.0634 | 0.96509 | 207.4032 | 49.6019 | 56068.9599 | 55863.0227 | 1113.2605 | 1013.2605 | 1402.6559
9891 | 114.9356 | 1.0109 | 1.0373 | 0.97756 | 197.2665 | 48.9038 | 55468.6571 | 55323.494 | 1102.4699 | 1002.4699 | 1402.9005
9892 | 100.7914 | 0.98957 | 0.80407 | 1.096 | 182.1731 | 47.5766 | 46856.139 | 46830.2473 | 932.6049 | 832.6049 | 1403.7912
9893 | 95.4713 | 0.9901 | 1.0621 | 0.94613 | 200.3244 | 45.8307 | 50589.1106 | 49070.24 | 977.4048 | 877.4048 | 1404.0664
9894 | 96.5914 | 0.99089 | 0.91852 | 1.0045 | 193.7363 | 44.7137 | 45728.64 | 45311.9879 | 902.2398 | 802.2398 | 1404.4026
9895 | 93.1974 | 1.0436 | 1.1408 | 0.9787 | 212.8731 | 48.4716 | 59045.4446 | 58407.5269 | 1164.1505 | 1064.1505 | 1406.5939
9896 | 117.6822 | 1.0472 | 1.0731 | 1.0054 | 208.9707 | 53.9271 | 65036.9348 | 64660.8281 | 1289.2166 | 1189.2166 | 1407.0608
9897 | 93.2389 | 0.95736 | 1.1951 | 1.0823 | 202.0057 | 49.8413 | 69117.5055 | 68968.2862 | 1375.3657 | 1275.3657 | 1407.4344
9898 | 95.7629 | 1.0246 | 1.1163 | 1.0282 | 212.9603 | 50.6014 | 63256.1431 | 63927.8097 | 1274.5562 | 1174.5562 | 1407.5067
9899 | 107.9289 | 0.96207 | 0.84247 | 0.99095 | 211.8513 | 47.8045 | 45085.0189 | 45614.0196 | 908.2804 | 808.2804 | 1408.6147
9900 | 105.4583 | 1.0385 | 1.0602 | 1.0203 | 197.9632 | 53.9338 | 64446.7305 | 64605.5607 | 1288.1112 | 1188.1112 | 1408.7007
9901 | 103.4416 | 1.0156 | 0.99315 | 1.0048 | 203.8227 | 49.1011 | 54361.8295 | 54198.1119 | 1079.9622 | 979.9622 | 1409.6027
9902 | 91.7019 | 0.89768 | 0.93678 | 0.98981 | 211.8273 | 44.4376 | 45073.7329 | 46062.9283 | 917.2586 | 817.2586 | 1410.1449
9903 | 106.2699 | 1.0346 | 1.001 | 1.0076 | 197.0305 | 48.9379 | 54935.1457 | 54598.5158 | 1087.9703 | 987.9703 | 1411.4089
9904 | 79.4733 | 1.0004 | 1.2078 | 0.97838 | 198.44 | 47.1742 | 59692.4033 | 58851.1194 | 1173.0224 | 1073.0224 | 1412.0836
9905 | 114.0737 | 1.0558 | 1.2611 | 1.0456 | 205.2595 | 47.9088 | 69146.8004 | 70021.7953 | 1396.4359 | 1296.4359 | 1412.7535
9906 | 97.6411 | 0.99467 | 1.1682 | 0.98795 | 195.6488 | 50.4198 | 63281.497 | 61412.9696 | 1224.2594 | 1124.2594 | 1412.9926
9907 | 93.6591 | 0.92726 | 0.83056 | 0.92716 | 189.3613 | 52.3574 | 45054.7106 | 45403.6099 | 904.0722 | 804.0722 | 1413.1118
9908 | 102.7268 | 0.99483 | 1.0526 | 0.88738 | 198.3675 | 51.0093 | 53057.5182 | 54732.3773 | 1090.6475 | 990.6475 | 1413.8697
9909 | 105.0308 | 0.98572 | 1.2392 | 1.02 | 198.0001 | 45.4455 | 62348.2195 | 62795.4896 | 1251.9098 | 1151.9098 | 1414.3186
9910 | 104.2236 | 0.99097 | 0.9741 | 1.0307 | 209.8854 | 49.1831 | 54667.6116 | 56738.0984 | 1130.762 | 1030.762 | 1415.5192
9911 | 85.4319 | 0.96778 | 1.199 | 0.99297 | 173.5019 | 47.6844 | 60885.7848 | 61645.0194 | 1228.9004 | 1128.9004 | 1417.5964
9912 | 107.9402 | 0.97492 | 1.0972 | 0.98555 | 200.5872 | 50.9413 | 60645.0535 | 59929.111 | 1194.5822 | 1094.5822 | 1419.7393
9913 | 95.149 | 0.9966 | 0.86767 | 1.0325 | 195.1172 | 49.4153 | 49149.441 | 48072.1063 | 957.4421 | 857.4421 | 1420.4902
9914 | 81.7034 | 1.0174 | 0.80159 | 0.99101 | 194.415 | 52.7686 | 46499.3661 | 45639.6839 | 908.7937 | 808.7937 | 1421.645
9915 | 100.6541 | 1.0075 | 1.0859 | 0.91633 | 199.4424 | 49.561 | 54541.9323 | 55528.5004 | 1106.57 | 1006.57 | 1423.1144
9916 | 109.0681 | 1.1055 | 1.0263 | 0.99022 | 195.0692 | 50.4513 | 57552.2668 | 58528.9589 | 1166.5792 | 1066.5792 | 1423.6202
9917 | 97.6133 | 1.0379 | 1.2169 | 0.9912 | 186.1294 | 46.8447 | 61435.5715 | 62091.5459 | 1237.8309 | 1137.8309 | 1424.9524
9918 | 105.8424 | 0.94161 | 1.0628 | 1.0205 | 201.3689 | 51.0596 | 60672.4588 | 61285.9642 | 1221.7193 | 1121.7193 | 1426.312
9919 | 85.4216 | 0.96373 | 1.0354 | 0.98746 | 204.6747 | 51.2263 | 56796.4387 | 58400.2963 | 1164.0059 | 1064.0059 | 1426.9471
9920 | 102.0081 | 0.9416 | 1.0121 | 0.96226 | 208.6923 | 47.7228 | 51269.1661 | 51576.1947 | 1027.5239 | 927.5239 | 1427.5806
9921 | 98.5917 | 1.0054 | 1.1652 | 1.1 | 197.5129 | 51.6784 | 71554.9668 | 71758.4509 | 1431.169 | 1331.169 | 1428.4243
9922 | 97.0989 | 0.96547 | 0.89763 | 1.0214 | 200.141 | 50.4468 | 51179.0942 | 48977.2318 | 975.5446 | 875.5446 | 1430.7145
9923 | 104.7231 | 0.89073 | 1.1377 | 1.0612 | 187.856 | 52.8551 | 68927.0003 | 69156.6794 | 1379.1336 | 1279.1336 | 1430.943
9924 | 95.7114 | 1.0004 | 1.0182 | 1.0539 | 204.536 | 48.3738 | 56742.7259 | 57837.597 | 1152.7519 | 1052.7519 | 1431.5046
9925 | 105.7509 | 1.0062 | 0.74242 | 1.0913 | 192.9999 | 51.6918 | 47573.8986 | 48537.5743 | 966.7515 | 866.7515 | 1431.9471
9926 | 110.7771 | 0.868 | 0.95427 | 1.0353 | 203.1903 | 45.8783 | 49938.9525 | 51247.11 | 1020.9422 | 920.9422 | 1432.422
9927 | 91.7352 | 1.0451 | 1.096 | 1.0383 | 196.4707 | 50.6328 | 62672.8787 | 61666.7969 | 1229.3359 | 1129.3359 | 1432.7259
9928 | 118.269 | 1.0001 | 1.1686 | 0.99357 | 202.0389 | 51.4715 | 66054.5565 | 67243.6417 | 1340.8728 | 1240.8728 | 1433.4448
9929 | 99.847 | 0.99715 | 1.33 | 0.96972 | 206.2504 | 48.9333 | 68187.0215 | 68940.7842 | 1374.8157 | 1274.8157 | 1434.9006
9930 | 92.1591 | 0.9074 | 1.0567 | 0.92247 | 188.8685 | 51.8526 | 55069.9315 | 54937.9709 | 1094.7594 | 994.7594 | 1436.1168
9931 | 95.8037 | 1.0543 | 1.0289 | 1.0424 | 185.8114 | 44.9137 | 52895.4627 | 53272.1822 | 1061.4436 | 961.4436 | 1436.9231
9932 | 88.7532 | 0.91607 | 0.89275 | 0.94982 | 203.0953 | 51.7382 | 48281.1618 | 47655.596 | 949.1119 | 849.1119 | 1438.5245
9933 | 84.2254 | 0.97795 | 0.96587 | 1.0481 | 198.4204 | 48.415 | 53196.2774 | 53088.5377 | 1057.7708 | 957.7708 | 1438.9227
9934 | 107.8212 | 1.0171 | 0.82791 | 1.058 | 170.4049 | 53.7813 | 53178.7619 | 54979.4541 | 1095.5891 | 995.5891 | 1439.8179
9935 | 110.2145 | 1.0365 | 1.1543 | 0.96815 | 198.9928 | 53.2185 | 65749.8868 | 67516.2169 | 1346.3243 | 1246.3243 | 1441.2272
9936 | 97.5304 | 1.0633 | 0.99162 | 1.007 | 206.403 | 53.0074 | 58634.3583 | 58470.3779 | 1165.4076 | 1065.4076 | 1441.2304
9937 | 109.1421 | 0.96977 | 0.91527 | 1.0332 | 195.3326 | 46.7617 | 49365.6599 | 49296.2759 | 981.9255 | 881.9255 | 1442.0048
9938 | 105.1515 | 0.99685 | 1.1599 | 0.94279 | 197.7594 | 47.0022 | 56524.603 | 55197.1907 | 1099.9438 | 999.9438 | 1442.7588
9939 | 94.4897 | 1.0224 | 1.2166 | 0.92586 | 215.2145 | 43.9992 | 54026.6979 | 53571.8682 | 1067.4374 | 967.4374 | 1443.2459
9940 | 89.9204 | 0.98096 | 0.96212 | 0.97654 | 189.5368 | 50.75 | 52348.2004 | 52117.4048 | 1038.3481 | 938.3481 | 1444.3883
9941 | 78.6449 | 1.0775 | 0.95941 | 1.0058 | 191.4281 | 48.6047 | 51213.6819 | 50101.6016 | 998.032 | 898.032 | 1444.5697
9942 | 86.8961 | 1.0058 | 1.1293 | 0.97791 | 182.6568 | 47.6887 | 57017.3373 | 57296.7161 | 1141.9343 | 1041.9343 | 1446.6619
9943 | 106.1386 | 0.9324 | 1.1514 | 1.0331 | 195.977 | 52.2134 | 67472.278 | 68100.694 | 1358.0139 | 1258.0139 | 1447.7017
9944 | 93.6676 | 1.0167 | 1.0077 | 1.0497 | 206.3728 | 50.6105 | 58560.9678 | 59980.4772 | 1195.6095 | 1095.6095 | 1449.891
9945 | 112.1907 | 0.96804 | 1.0552 | 1.0189 | 192.7767 | 43.881 | 52136.5313 | 53515.8126 | 1066.3163 | 966.3163 | 1450.2863
9946 | 90.3623 | 1.023 | 0.9385 | 0.97888 | 209.7115 | 52.8828 | 53680.5956 | 53791.659 | 1071.8332 | 971.8332 | 1450.6775
9947 | 117.4076 | 0.98677 | 0.99095 | 0.98017 | 207.8485 | 49.0717 | 53556.6009 | 52007.8271 | 1036.1565 | 936.1565 | 1451.9314
9948 | 102.1359 | 1.0705 | 0.7935 | 1.0281 | 193.2022 | 52.0387 | 48335.5616 | 49132.7329 | 978.6547 | 878.6547 | 1453.6404
9949 | 100.2214 | 0.95221 | 1.1165 | 0.93402 | 198.6525 | 56.0831 | 64035.6294 | 63128.5851 | 1258.5717 | 1158.5717 | 1453.7408
9950 | 93.5914 | 0.97918 | 1.1622 | 0.96405 | 219.6935 | 54.4225 | 66184.9585 | 66334.8845 | 1322.6977 | 1222.6977 | 1454.3616
9951 | 105.9632 | 1.041 | 1.1722 | 0.94941 | 221.9454 | 51.4581 | 63164.8951 | 64129.2826 | 1278.5857 | 1178.5857 | 1455.3543
9952 | 107.7318 | 1.0004 | 0.9773 | 0.99675 | 208.732 | 52.2435 | 56730.9289 | 57666.9701 | 1149.3394 | 1049.3394 | 1455.453
9953 | 96.9261 | 1.0265 | 0.85273 | 1.0118 | 196.1486 | 50.2444 | 48546.7051 | 48981.2699 | 975.6254 | 875.6254 | 1456.6414
9954 | 100.6174 | 1.054 | 1.1644 | 0.96994 | 198.8905 | 48.7364 | 60411.0515 | 60167.5381 | 1199.3508 | 1099.3508 | 1457.0231
9955 | 112.5194 | 1.0026 | 0.97955 | 0.99904 | 191.2006 | 50.1151 | 54888.2617 | 54244.9699 | 1080.8994 | 980.8994 | 1457.4035
9956 | 95.9948 | 0.93335 | 0.86756 | 0.95849 | 193.4637 | 49.4992 | 45789.2624 | 45720.8584 | 910.4172 | 810.4172 | 1458.4929
9957 | 84.2925 | 1.0977 | 0.76566 | 0.96213 | 192.1976 | 50.135 | 41763.8655 | 40876.9841 | 813.5397 | 713.5397 | 1459.1721
9958 | 97.8231 | 0.91374 | 0.87599 | 0.98543 | 199.732 | 49.5503 | 47401.7382 | 49152.0179 | 979.0404 | 879.0404 | 1462.353
9959 | 114.8656 | 1.0659 | 1.2187 | 0.96644 | 176.9028 | 46.8777 | 61126.7643 | 61302.4544 | 1222.0491 | 1122.0491 | 1462.7788
9960 | 86.6457 | 0.99082 | 0.7908 | 0.96932 | 200.3882 | 47.7679 | 40917.0631 | 42198.5145 | 839.9703 | 739.9703 | 1463.9327
9961 | 102.5614 | 0.97378 | 1.0585 | 1.0294 | 200.9297 | 52.2225 | 62322.6285 | 63234.6842 | 1260.6937 | 1160.6937 | 1467.9865
9962 | 100.8173 | 0.98617 | 1.2087 | 1.0637 | 202.5744 | 54.0052 | 75005.5903 | 74962.247 | 1495.2449 | 1395.2449 | 1470.0555
9963 | 90.3411 | 1.0819 | 0.69325 | 1.0202 | 208.754 | 47.7393 | 38639.6264 | 38002.2061 | 756.0441 | 656.0441 | 1471.6748
9964 | 99.7283 | 1.0688 | 0.93986 | 1.0084 | 192.8774 | 50.5694 | 53512.8415 | 52444.9875 | 1044.8998 | 944.8998 | 1472.7034
9965 | 105.8464 | 0.9798 | 1.0024 | 0.98284 | 203.3551 | 48.5393 | 53056.664 | 53028.146 | 1056.5629 | 956.5629 | 1474.4362
9966 | 97.8608 | 1.0204 | 1.0614 | 1.0259 | 189.5701 | 52.1048 | 62129.9215 | 61198.1346 | 1219.9627 | 1119.9627 | 1474.8833
9967 | 99.8023 | 1.0549 | 0.91318 | 1.0785 | 190.8952 | 50.9021 | 55680.7389 | 55265.4662 | 1101.3093 | 1001.3093 | 1477.6365
9968 | 97.8056 | 0.97756 | 1.0168 | 0.92449 | 186.7895 | 47.922 | 49817.8328 | 50207.5564 | 1000.1511 | 900.1511 | 1486.3008
9969 | 105.6734 | 1.0015 | 1.0623 | 1.0233 | 209.9229 | 48.6258 | 58212.5604 | 57657.3479 | 1149.147 | 1049.147 | 1486.6712
9970 | 95.7768 | 0.88981 | 0.93102 | 1.0313 | 212.8439 | 49.0488 | 51488.5493 | 52470.1834 | 1045.4037 | 945.4037 | 1489.8473
9971 | 115.7165 | 1.0552 | 1.1034 | 1.079 | 213.6507 | 48.1567 | 63428.9739 | 62945.6371 | 1254.9127 | 1154.9127 | 1489.9223
9972 | 94.816 | 1.0261 | 1.1563 | 0.90021 | 195.4557 | 52.5889 | 60052.9695 | 57857.104 | 1153.1421 | 1053.1421 | 1497.65
9973 | 91.6977 | 0.94766 | 1.0946 | 0.92533 | 189.6999 | 48.6613 | 53706.7266 | 54560.1047 | 1087.2021 | 987.2021 | 1498.1345
9974 | 109.4801 | 0.9155 | 1.0033 | 1.0929 | 218.502 | 49.9305 | 59971.9164 | 61197.5744 | 1219.9515 | 1119.9515 | 1498.4428
9975 | 97.2148 | 0.90776 | 1.1717 | 0.93756 | 199.6075 | 50.1617 | 59729.2359 | 58677.198 | 1169.544 | 1069.544 | 1500.4634
9976 | 88.7817 | 0.99768 | 0.90852 | 0.90401 | 180.3211 | 47.0097 | 42953.7644 | 43734.0156 | 870.6803 | 770.6803 | 1503.5417
9977 | 125.81 | 1.0091 | 0.88496 | 0.94846 | 189.8483 | 52.9405 | 51346.8546 | 51345.3972 | 1022.9079 | 922.9079 | 1504.5611
9978 | 91.4571 | 1.0807 | 0.8896 | 1.0213 | 221.8769 | 52.554 | 53165.5061 | 53042.6591 | 1056.8532 | 956.8532 | 1506.3602
9979 | 90.5413 | 0.99185 | 0.8628 | 1.077 | 192.4754 | 48.206 | 49317.6578 | 50296.5914 | 1001.9318 | 901.9318 | 1509.9745
9980 | 87.3596 | 0.92465 | 0.95587 | 0.92049 | 210.3025 | 49.1345 | 47411.3164 | 47255.6392 | 941.1128 | 841.1128 | 1510.3401
9981 | 90.256 | 0.97525 | 1.0273 | 1.0112 | 176.4337 | 52.2009 | 58995.5683 | 58408.2695 | 1164.1654 | 1064.1654 | 1513.218
9982 | 94.5439 | 0.89718 | 1.1907 | 1.0492 | 184.6465 | 51.7342 | 69207.1498 | 72725.3168 | 1450.5063 | 1350.5063 | 1521.3245
9983 | 100.1873 | 1.0193 | 0.94397 | 0.97006 | 184.8326 | 48.7364 | 49790.208 | 50789.499 | 1011.79 | 911.79 | 1522.4615
9984 | 104.5528 | 0.94589 | 0.56366 | 0.9991 | 181.544 | 47.8869 | 31885.1205 | 32281.2417 | 641.6248 | 541.6248 | 1524.8436
9985 | 97.9958 | 0.92485 | 0.80158 | 1.0721 | 196.9902 | 53.58 | 51098.2316 | 49209.454 | 980.1891 | 880.1891 | 1529.2927
9986 | 104.1816 | 1.0543 | 1.0604 | 0.95884 | 196.2376 | 52.4252 | 59259.2306 | 61285.4533 | 1221.7091 | 1121.7091 | 1529.5802
9987 | 107.9815 | 1.0387 | 1.1131 | 0.97211 | 196.7055 | 44.8725 | 53785.7284 | 54788.4861 | 1091.7697 | 991.7697 | 1532.3393
9988 | 96.1201 | 0.99126 | 1.248 | 0.9605 | 198.5268 | 47.5585 | 61738.5347 | 60677.3392 | 1209.5468 | 1109.5468 | 1534.204
9989 | 98.064 | 0.98431 | 1.2261 | 0.98468 | 217.209 | 48.1481 | 62996.3359 | 60572.4821 | 1207.4496 | 1107.4496 | 1536.5544
9990 | 96.9018 | 0.95204 | 0.8533 | 1.014 | 188.1437 | 47.6707 | 45834.6169 | 44162.815 | 879.2563 | 779.2563 | 1540.954
9991 | 116.2027 | 0.99425 | 1.1408 | 0.94787 | 200.9842 | 52.948 | 63575.0335 | 62799.7247 | 1251.9945 | 1151.9945 | 1546.2551
9992 | 84.6169 | 1.0049 | 0.85301 | 0.96278 | 182.7021 | 45.5275 | 41443.7873 | 41802.1438 | 832.0429 | 732.0429 | 1552.0597
9993 | 117.0106 | 0.99607 | 1.0731 | 0.94668 | 206.1605 | 49.9265 | 56744.9898 | 57317.6424 | 1142.3528 | 1042.3528 | 1557.3095
9994 | 100.1295 | 1.0236 | 1.0802 | 1.1223 | 192.4956 | 51.4172 | 67795.0983 | 69098.7678 | 1377.9754 | 1277.9754 | 1567.8051
9995 | 102.4704 | 1.042 | 0.78239 | 0.98356 | 191.7769 | 47.1213 | 41484.5721 | 41732.0958 | 830.6419 | 730.6419 | 1568.8743
9996 | 88.3993 | 1.0512 | 0.8651 | 0.9234 | 207.384 | 45.755 | 41009.6675 | 40368.7372 | 803.3747 | 703.3747 | 1574.2885
9997 | 102.9784 | 1.0205 | 0.646 | 0.97864 | 207.8406 | 52.1962 | 38691.896 | 38325.4552 | 762.5091 | 662.5091 | 1659.338
9998 | 105.0324 | 1.0488 | 0.8018 | 1.0611 | 194.7472 | 47.2437 | 45591.8735 | 47063.7782 | 937.2756 | 837.2756 | 1669.3106
9999 | 105.5494 | 1.0521 | 1.2215 | 1.0019 | 209.9363 | 54.092 | 72416.276 | 71434.4406 | 1424.6888 | 1324.6888 | 1716.5706
[10 empty] | mean | 1001.8633
[10 empty] | stdev | 164.941
[10 empty] | meanB | 1001.8633
[10 empty] | stdevB | 164.941
[10 empty] | low95% | 689.5205
[10 empty] | high95% | 1335.7719
[10 empty] | lowB95% | 690.6994
[10 empty] | highB95% | 1336.2755